## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 239b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 3.6076 | 3.5653 | 3.65 | 3.5833 | 3.632
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 3.501 | 3.4553 | 3.5467 | 3.4751 | 3.5269
Cyanide Leach
Au, Gold (ppm) | 3.2359 | 3.116 | 3.3558 | 3.214 | 3.2577
PhotonAssay
Au, Gold (ppm) | 3.71 | 3.6217 | 3.7982 | 3.7033 | 3.7166
4-Acid Digestion
Ag, Silver (ppm) | 0.27276 | 0.23035 | 0.31516 | 0.24632 | 0.29919
Al, Aluminium (wt.%) | 7.3591 | 7.2061 | 7.5122 | 7.2127 | 7.5056
As, Arsenic (ppm) | 816.0782 | 781.0439 | 851.1125 | 800.4107 | 831.7457
Ba, Barium (ppm) | 678.2737 | 662.9133 | 693.6341 | 667.1171 | 689.4304
Be, Beryllium (ppm) | 2.4153 | 2.298 | 2.5326 | 2.3081 | 2.5225
Bi, Bismuth (ppm) | 0.39168 | 0.36349 | 0.41988 | 0.35483 | 0.42854
Ca, Calcium (wt.%) | 0.86738 | 0.8506 | 0.88417 | 0.84806 | 0.8867
Cd, Cadmium (ppm) | 0.092562 | 0.069326 | 0.1158 | 0.085461 | 0.099663
Ce, Cerium (ppm) | 78.3607 | 75.1766 | 81.5448 | 76.0627 | 80.6588
Co, Cobalt (ppm) | 16.8745 | 16.2341 | 17.5149 | 16.4714 | 17.2776
Cr, Chromium (ppm) | 123.2527 | 116.5104 | 129.995 | 119.0222 | 127.4832
Cs, Caesium (ppm) | 9.3728 | 9.0531 | 9.6924 | 9.1047 | 9.6409
Cu, Copper (ppm) | 30.9206 | 29.779 | 32.0622 | 29.9035 | 31.9377
Dy, Dysprosium (ppm) | 3.4742 | 2.9956 | 3.9529 | 3.3376 | 3.6109
Er, Erbium (ppm) | 1.7632 | 1.5221 | 2.0043 | 1.7132 | 1.8131
Eu, Europium (ppm) | 1.1937 | 1.1205 | 1.267 | 1.1546 | 1.2328
Fe, Iron (wt.%) | 4.0747 | 3.9715 | 4.1779 | 4.0034 | 4.146
Ga, Gallium (ppm) | 19.2786 | 18.2193 | 20.3379 | 18.7301 | 19.8271
Gd, Gadolinium (ppm) | 4.9552 | 4.6806 | 5.2298 | 4.7708 | 5.1396
Hf, Hafnium (ppm) | 3.8083 | 3.5841 | 4.0324 | 3.6737 | 3.9428
Ho, Holmium (ppm) | 0.67848 | 0.52842 | 0.82854 | 0.64647 | 0.7105
In, Indium (ppm) | 0.068141 | 0.058419 | 0.077863 | 0.064436 | 0.071846
K, Potassium (wt.%) | 2.6694 | 2.6107 | 2.7281 | 2.6112 | 2.7275
La, Lanthanum (ppm) | 37.918 | 36.2216 | 39.6143 | 36.7091 | 39.1268
Li, Lithium (ppm) | 48.3104 | 46.2898 | 50.3309 | 47.15 | 49.4707
Lu, Lutetium (ppm) | 0.27896 | 0.23521 | 0.32271 | 0.26352 | 0.2944
Mg, Magnesium (wt.%) | 1.7167 | 1.6799 | 1.7536 | 1.686 | 1.7475
Mn, Manganese (wt.%) | 0.039344 | 0.038178 | 0.040509 | 0.038647 | 0.04004
Mo, Molybdenum (ppm) | 0.98576 | 0.89923 | 1.0723 | 0.90805 | 1.0635
Na, Sodium (wt.%) | 0.68217 | 0.66805 | 0.6963 | 0.66893 | 0.69542
Nb, Niobium (ppm) | 14.1432 | 13.433 | 14.8535 | 13.8011 | 14.4853
Nd, Neodymium (ppm) | 33.1793 | 32.0788 | 34.2798 | 31.9316 | 34.4271
Ni, Nickel (ppm) | 64.8498 | 62.8765 | 66.823 | 63.2932 | 66.4064
P, Phosphorus (wt.%) | 0.065429 | 0.063878 | 0.06698 | 0.064144 | 0.066714
Pb, Lead (ppm) | 23.1693 | 22.0802 | 24.2584 | 22.3074 | 24.0312
Pr, Praseodymium (ppm) | 9.0412 | 8.6509 | 9.4316 | 8.6888 | 9.3937
Rb, Rubidium (ppm) | 153.5534 | 146.5636 | 160.5432 | 148.4509 | 158.6559
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.20816 | 0.19978 | 0.21655 | 0.20239 | 0.21394
Sb, Antimony (ppm) | 727.3939 | 702.2862 | 752.5015 | 714.0227 | 740.765
Sc, Scandium (ppm) | 14.0227 | 13.2856 | 14.7597 | 13.6241 | 14.4212
Sm, Samarium (ppm) | 6.3094 | 6.0449 | 6.5739 | 6.1003 | 6.5184
Sn, Tin (ppm) | 3.6443 | 3.4364 | 3.8522 | 3.4762 | 3.8124
Sr, Strontium (ppm) | 120.6895 | 117.759 | 123.62 | 118.492 | 122.8871
Ta, Tantalum (ppm) | 1.0509 | 0.99403 | 1.1078 | 1.0138 | 1.088
Tb, Terbium (ppm) | 0.67045 | 0.62713 | 0.71377 | 0.64119 | 0.69972
Th, Thorium (ppm) | 14.2122 | 13.6398 | 14.7847 | 13.743 | 14.6815
Ti, Titanium (wt.%) | 0.45955 | 0.44593 | 0.47318 | 0.44909 | 0.47001
Tl, Thallium (ppm) | 0.83153 | 0.78426 | 0.87881 | 0.80158 | 0.86148
Tm, Thulium (ppm) | 0.26147 | 0.21295 | 0.30999 | 0.25085 | 0.27209
U, Uranium (ppm) | 2.6641 | 2.5498 | 2.7785 | 2.5776 | 2.7506
V, Vanadium (ppm) | 105.2798 | 102.6391 | 107.9205 | 102.7741 | 107.7855
W, Tungsten (ppm) | 2.9352 | 2.6685 | 3.2019 | 2.7953 | 3.0751
Y, Yttrium (ppm) | 15.6328 | 14.483 | 16.7825 | 15.0836 | 16.182
Yb, Ytterbium (ppm) | 1.7028 | 1.4989 | 1.9066 | 1.6619 | 1.7436
Zn, Zinc (ppm) | 97.8706 | 95.5523 | 100.1889 | 95.7747 | 99.9665
Zr, Zirconium (ppm) | 130.5435 | 124.802 | 136.285 | 127.2978 | 133.7892
Aqua Regia Digestion
Ag, Silver (ppm) | 0.26607 | 0.25036 | 0.28178 | 0.24457 | 0.28756
Al, Aluminium (wt.%) | 2.5807 | 2.499 | 2.6624 | 2.5475 | 2.6138
As, Arsenic (ppm) | 810.5157 | 788.5465 | 832.485 | 795.0837 | 825.9478
Ba, Barium (ppm) | 106.9023 | 101.5476 | 112.2571 | 103.2109 | 110.5937
Be, Beryllium (ppm) | 1.1903 | 1.1121 | 1.2685 | 1.1395 | 1.2411
Bi, Bismuth (ppm) | 0.37393 | 0.34337 | 0.40449 | 0.32892 | 0.41894
Ca, Calcium (wt.%) | 0.33321 | 0.31961 | 0.34681 | 0.32529 | 0.34113
Cd, Cadmium (ppm) | 0.072145 | 0.059342 | 0.084947 | 0.068565 | 0.075724
Ce, Cerium (ppm) | 54.4618 | 51.1948 | 57.7288 | 53.2334 | 55.6902
Co, Cobalt (ppm) | 15.1646 | 14.4188 | 15.9104 | 14.644 | 15.6851
Cr, Chromium (ppm) | 106.5348 | 102.9656 | 110.1041 | 104.1713 | 108.8984
Cs, Caesium (ppm) | 6.7143 | 6.4433 | 6.9852 | 6.5338 | 6.8948
Cu, Copper (ppm) | 29.557 | 28.2991 | 30.8149 | 28.8857 | 30.2283
Fe, Iron (wt.%) | 3.5146 | 3.4447 | 3.5845 | 3.4645 | 3.5647
Ga, Gallium (ppm) | 8.0876 | 7.7676 | 8.4075 | 7.8666 | 8.3086
Ge, Germanium (ppm) | 0.099892 | 0.070377 | 0.12941 | IND | IND
Hf, Hafnium (ppm) | 0.54012 | 0.45612 | 0.62412 | 0.50707 | 0.57318
In, Indium (ppm) | 0.029758 | 0.02848 | 0.031037 | 0.026708 | 0.032809
K, Potassium (wt.%) | 0.86443 | 0.84063 | 0.88823 | 0.85079 | 0.87807
La, Lanthanum (ppm) | 27.1269 | 25.628 | 28.6258 | 26.243 | 28.0107
Li, Lithium (ppm) | 38.1829 | 36.4821 | 39.8837 | 37.2821 | 39.0838
Lu, Lutetium (ppm) | 0.11511 | 0.075278 | 0.15494 | IND | IND
Mg, Magnesium (wt.%) | 1.4078 | 1.3672 | 1.4483 | 1.3839 | 1.4317
Mn, Manganese (wt.%) | 0.028042 | 0.027165 | 0.028918 | 0.02756 | 0.028523
Mo, Molybdenum (ppm) | 0.89612 | 0.83498 | 0.95726 | 0.84277 | 0.94947
Na, Sodium (wt.%) | 0.090925 | 0.082838 | 0.099012 | 0.087785 | 0.094065
Nb, Niobium (ppm) | 0.40136 | 0.31325 | 0.48948 | 0.36578 | 0.43695
Ni, Nickel (ppm) | 59.3099 | 56.9637 | 61.6562 | 58.0068 | 60.6131
P, Phosphorus (wt.%) | 0.058434 | 0.05669 | 0.060179 | 0.057214 | 0.059654
Pb, Lead (ppm) | 14.434 | 13.3628 | 15.5051 | 13.6658 | 15.2021
Rb, Rubidium (ppm) | 83.3284 | 80.9146 | 85.7421 | 81.378 | 85.2787
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.20722 | 0.19969 | 0.21475 | 0.19987 | 0.21457
Sb, Antimony (ppm) | 553.6866 | 506.1754 | 601.1979 | 543.9228 | 563.4505
Sc, Scandium (ppm) | 6.3277 | 6.0161 | 6.6394 | 6.1673 | 6.4882
Sn, Tin (ppm) | 1.4471 | 1.3483 | 1.5458 | 1.4146 | 1.4795
Sr, Strontium (ppm) | 28.8958 | 27.7691 | 30.0224 | 28.297 | 29.4945
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Th, Thorium (ppm) | 11.7951 | 11.1788 | 12.4114 | 11.4128 | 12.1774
Ti, Titanium (wt.%) | 0.14198 | 0.13544 | 0.14852 | 0.13811 | 0.14585
Tl, Thallium (ppm) | 0.51737 | 0.48718 | 0.54756 | 0.50079 | 0.53396
U, Uranium (ppm) | 1.4506 | 1.3505 | 1.5507 | 1.3982 | 1.503
V, Vanadium (ppm) | 67.0747 | 65.1135 | 69.0359 | 65.9357 | 68.2137
W, Tungsten (ppm) | 0.55586 | 0.4792 | 0.63252 | 0.49619 | 0.61553
Y, Yttrium (ppm) | 8.9394 | 8.6479 | 9.231 | 8.7407 | 9.1381
Yb, Ytterbium (ppm) | 0.86918 | 0.72669 | 1.0117 | 0.79614 | 0.94221
Zn, Zinc (ppm) | 85.7743 | 83.0632 | 88.4853 | 83.7541 | 87.7944
Zr, Zirconium (ppm) | 21.8632 | 20.144 | 23.5824 | 21.2838 | 22.4427
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 239b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 0.85492 | Pt | ppb | 0.56374
4-Acid Digestion
Au | ppm | 3.4836 | Ir | ppm | 0.0003 | Ru | ppm | 0.0032333
B | ppm | 29.835 | Pd | ppb | 11.15 | Se | ppm | 0.69457
Ge | ppm | 0.19958 | Pt | ppb | 1.4667 | Te | ppm | 0.049326
Hg | ppm | 0.032417 | Rh | ppm | 0.00098333
Aqua Regia Digestion
B | ppm | 15.7323 | Ir | ppm | < 0.0019 | Se | ppm | 0.32977
Dy | ppm | 2.0209 | Nd | ppm | 23.1274 | Si | wt.% | 0.40938
Er | ppm | 0.91439 | Pd | ppb | 15.3583 | Sm | ppm | 4.2562
Eu | ppm | 0.64426 | Pr | ppm | 6.1983 | Tb | ppm | 0.49301
Gd | ppm | 4.0014 | Pt | ppb | 2.9 | Te | ppm | 0.026504
Hg | ppm | 0.016441 | Rh | ppm | 0.0012 | Tm | ppm | 0.11356
Ho | ppm | 0.33384 | Ru | ppm | < 0.0056
Borate Fusion XRF
Al2O3 | wt.% | 14.43 | MgO | wt.% | 2.88 | SiO2 | wt.% | 67.795
CaO | wt.% | 1.21 | MnO | wt.% | 0.05 | SO3 | wt.% | 0.513
Fe2O3 | wt.% | 5.83 | Na2O | wt.% | 0.92 | TiO2 | wt.% | 0.81
K2O | wt.% | 3.3 | P2O5 | wt.% | 0.151
Thermogravimetry
LOI1000 | wt.% | 2.495
Infrared Combustion
C | wt.% | 0.11 | S | wt.% | 0.19583
Laser Ablation ICP-MS
Ag | ppm | 0.35 | Hf | ppm | 6.62 | Sm | ppm | 7.315
As | ppm | 847.5 | Ho | ppm | 1.18 | Sn | ppm | 3.8
Ba | ppm | 700 | In | ppm | < 0.05 | Sr | ppm | 120
Be | ppm | 2.6 | La | ppm | 40.6 | Ta | ppm | 1.24
Bi | ppm | 0.43 | Lu | ppm | 0.475 | Tb | ppm | 0.965
Cd | ppm | 0.3 | Mn | wt.% | 0.0427 | Te | ppm | < 0.2
Ce | ppm | 80.35 | Mo | ppm | 1.4 | Th | ppm | 15.2
Co | ppm | 18.1 | Nb | ppm | 15.9 | Ti | wt.% | 0.496
Cr | ppm | 150.5 | Nd | ppm | 36.55 | Tl | ppm | 0.4
Cs | ppm | 9.445 | Ni | ppm | 69 | Tm | ppm | 0.51
Cu | ppm | 33 | Pb | ppm | 25.5 | U | ppm | 3.155
Dy | ppm | 5.775 | Pr | ppm | 9.895 | V | ppm | 112.5
Er | ppm | 3.39 | Rb | ppm | 161.5 | W | ppm | 4
Eu | ppm | 1.355 | Re | ppm | < 0.01 | Y | ppm | 30.75
Ga | ppm | 19.3 | Sb | ppm | 779 | Yb | ppm | 3.22
Gd | ppm | 6.065 | Sc | ppm | 15.4 | Zn | ppm | 100
Ge | ppm | 1.65 | Se | ppm | < 5 | Zr | ppm | 234.5
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 239b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 3.6076 | 0.11032 | 3.387 | 3.8283 | 3.2767 | 3.9386 | 3.058% | 6.116% | 9.174% | 3.4273 | 3.788
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 3.501 | 0.10363 | 3.2937 | 3.7083 | 3.1901 | 3.8119 | 2.96% | 5.92% | 8.88% | 3.3259 | 3.676
Cyanide Leach
Au, ppm | 3.2359 | 0.22074 | 2.7944 | 3.6774 | 2.5737 | 3.8981 | 6.822% | 13.64% | 20.47% | 3.0741 | 3.3977
PhotonAssay
Au, ppm | 3.71 | 0.13801 | 3.4339 | 3.986 | 3.2959 | 4.124 | 3.72% | 7.44% | 11.16% | 3.5245 | 3.8955
4-Acid Digestion
Ag, ppm | 0.27276 | 0.043862 | 0.18503 | 0.36048 | 0.14117 | 0.40434 | 16.08% | 32.16% | 48.24% | 0.25912 | 0.2864
Al, wt.% | 7.3591 | 0.18162 | 6.9959 | 7.7224 | 6.8143 | 7.904 | 2.468% | 4.936% | 7.404% | 6.9912 | 7.7271
As, ppm | 816.0782 | 40.6464 | 734.7855 | 897.371 | 694.1391 | 938.0173 | 4.981% | 9.961% | 14.94% | 775.2743 | 856.8821
Ba, ppm | 678.2737 | 17.8069 | 642.6599 | 713.8875 | 624.853 | 731.6945 | 2.625% | 5.251% | 7.876% | 644.36 | 712.1874
Be, ppm | 2.4153 | 0.18162 | 2.0521 | 2.7785 | 1.8705 | 2.9602 | 7.519% | 15.04% | 22.56% | 2.2946 | 2.5361
Bi, ppm | 0.39168 | 0.020008 | 0.35167 | 0.4317 | 0.33166 | 0.45171 | 5.108% | 10.22% | 15.32% | 0.3721 | 0.41127
Ca, wt.% | 0.86738 | 0.020684 | 0.82601 | 0.90875 | 0.80533 | 0.92943 | 2.385% | 4.769% | 7.154% | 0.82401 | 0.91075
Cd, ppm | 0.092562 | 0.031474 | 0.029614 | 0.15551 | 0 | 0.18698 | 34% | 68.01% | 102% | 0.087934 | 0.09719
Ce, ppm | 78.3607 | 3.844 | 70.6727 | 86.0488 | 66.8287 | 89.8928 | 4.906% | 9.811% | 14.72% | 74.4427 | 82.2788
Co, ppm | 16.8745 | 0.85327 | 15.1679 | 18.581 | 14.3147 | 19.4343 | 5.057% | 10.11% | 15.17% | 16.0308 | 17.7182
Cr, ppm | 123.2527 | 12.3858 | 98.4811 | 148.0243 | 86.0953 | 160.41 | 10.05% | 20.1% | 30.15% | 117.0901 | 129.4153
Cs, ppm | 9.3728 | 0.46849 | 8.4358 | 10.3098 | 7.9673 | 10.7782 | 4.998% | 9.997% | 15% | 8.9041 | 9.8414
Cu, ppm | 30.9206 | 1.3887 | 28.1432 | 33.698 | 26.7545 | 35.0868 | 4.491% | 8.982% | 13.47% | 29.3746 | 32.4666
Dy, ppm | 3.4742 | 0.62606 | 2.2221 | 4.7264 | 1.596 | 5.3524 | 18.02% | 36.04% | 54.06% | 3.3005 | 3.6479
Er, ppm | 1.7632 | 0.30058 | 1.162 | 2.3643 | 0.86143 | 2.6649 | 17.05% | 34.1% | 51.14% | 1.675 | 1.8513
Eu, ppm | 1.1937 | 0.081166 | 1.0314 | 1.356 | 0.95022 | 1.4372 | 6.799% | 13.6% | 20.4% | 1.134 | 1.2534
Fe, wt.% | 4.0747 | 0.094531 | 3.8856 | 4.2637 | 3.7911 | 4.3583 | 2.32% | 4.64% | 6.96% | 3.8709 | 4.2784
Ga, ppm | 19.2786 | 1.1405 | 16.9975 | 21.5597 | 15.857 | 22.7003 | 5.916% | 11.83% | 17.75% | 18.3147 | 20.2426
Gd, ppm | 4.9552 | 0.35472 | 4.2458 | 5.6646 | 3.891 | 6.0193 | 7.158% | 14.32% | 21.48% | 4.7074 | 5.203
Hf, ppm | 3.8083 | 0.24923 | 3.3098 | 4.3067 | 3.0606 | 4.556 | 6.544% | 13.09% | 19.63% | 3.6179 | 3.9987
Ho, ppm | 0.67848 | 0.15726 | 0.36397 | 0.993 | 0.20671 | 1.1503 | 23.18% | 46.36% | 69.53% | 0.64456 | 0.71241
In, ppm | 0.068141 | 0.0064994 | 0.055142 | 0.08114 | 0.048643 | 0.087639 | 9.538% | 19.08% | 28.61% | 0.064734 | 0.071548
K, wt.% | 2.6694 | 0.09616 | 2.4771 | 2.8617 | 2.3809 | 2.9579 | 3.602% | 7.205% | 10.81% | 2.5359 | 2.8029
La, ppm | 37.918 | 2.5183 | 32.8813 | 42.9546 | 30.363 | 45.4729 | 6.642% | 13.28% | 19.92% | 36.0221 | 39.8139
Li, ppm | 48.3104 | 3.3873 | 41.5358 | 55.0849 | 38.1485 | 58.4722 | 7.012% | 14.02% | 21.03% | 45.8948 | 50.7259
Lu, ppm | 0.27896 | 0.057618 | 0.16373 | 0.3942 | 0.10611 | 0.45182 | 20.65% | 41.31% | 61.96% | 0.26501 | 0.29291
Mg, wt.% | 1.7167 | 0.049449 | 1.6178 | 1.8156 | 1.5684 | 1.8651 | 2.88% | 5.761% | 8.641% | 1.6309 | 1.8026
Mn, wt.% | 0.039344 | 0.0017645 | 0.035815 | 0.042872 | 0.03405 | 0.044637 | 4.485% | 8.97% | 13.45% | 0.037376 | 0.041311
Mo, ppm | 0.98576 | 0.068365 | 0.84903 | 1.1225 | 0.78066 | 1.1909 | 6.935% | 13.87% | 20.81% | 0.93647 | 1.035
Na, wt.% | 0.68217 | 0.018033 | 0.64611 | 0.71824 | 0.62807 | 0.73627 | 2.643% | 5.287% | 7.93% | 0.64806 | 0.71628
Nb, ppm | 14.1432 | 1.0844 | 11.9745 | 16.3119 | 10.8902 | 17.3963 | 7.667% | 15.33% | 23% | 13.4361 | 14.8504
Nd, ppm | 33.1793 | 1.4876 | 30.204 | 36.1546 | 28.7164 | 37.6423 | 4.484% | 8.967% | 13.45% | 31.5204 | 34.8383
Ni, ppm | 64.8498 | 2.6579 | 59.5341 | 70.1655 | 56.8762 | 72.8234 | 4.098% | 8.197% | 12.3% | 61.6073 | 68.0923
P, wt.% | 0.065429 | 0.0025941 | 0.060241 | 0.070617 | 0.057647 | 0.073211 | 3.965% | 7.929% | 11.89% | 0.062158 | 0.068701
Pb, ppm | 23.1693 | 1.2266 | 20.7161 | 25.6225 | 19.4895 | 26.8491 | 5.294% | 10.59% | 15.88% | 22.0108 | 24.3278
Pr, ppm | 9.0412 | 0.51279 | 8.0157 | 10.0668 | 7.5029 | 10.5796 | 5.672% | 11.34% | 17.02% | 8.5892 | 9.4933
Rb, ppm | 153.5534 | 9.708 | 134.1374 | 172.9694 | 124.4294 | 182.6775 | 6.322% | 12.64% | 18.97% | 145.8758 | 161.2311
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.20816 | 0.0078024 | 0.19256 | 0.22377 | 0.18476 | 0.23157 | 3.748% | 7.496% | 11.24% | 0.19775 | 0.21857
Sb, ppm | 727.3939 | 32.7256 | 661.9428 | 792.845 | 629.2172 | 825.5705 | 4.499% | 8.998% | 13.5% | 691.0242 | 763.7636
Sc, ppm | 14.0227 | 0.93448 | 12.1537 | 15.8916 | 11.2192 | 16.8261 | 6.664% | 13.33% | 19.99% | 13.3215 | 14.7238
Sm, ppm | 6.3094 | 0.3815 | 5.5464 | 7.0724 | 5.1649 | 7.4539 | 6.047% | 12.09% | 18.14% | 5.9939 | 6.6248
Sn, ppm | 3.6443 | 0.23743 | 3.1694 | 4.1191 | 2.932 | 4.3566 | 6.515% | 13.03% | 19.54% | 3.4621 | 3.8265
Sr, ppm | 120.6895 | 3.28 | 114.1296 | 127.2495 | 110.8496 | 130.5295 | 2.718% | 5.435% | 8.153% | 114.6551 | 126.724
Ta, ppm | 1.0509 | 0.08387 | 0.88318 | 1.2187 | 0.79931 | 1.3025 | 7.981% | 15.96% | 23.94% | 0.99837 | 1.1035
Tb, ppm | 0.67045 | 0.033449 | 0.60356 | 0.73735 | 0.57011 | 0.7708 | 4.989% | 9.978% | 14.97% | 0.63693 | 0.70398
Th, ppm | 14.2122 | 0.6335 | 12.9452 | 15.4792 | 12.3117 | 16.1127 | 4.457% | 8.915% | 13.37% | 13.5016 | 14.9228
Ti, wt.% | 0.45955 | 0.01994 | 0.41967 | 0.49943 | 0.39973 | 0.51937 | 4.339% | 8.678% | 13.02% | 0.43657 | 0.48253
Tl, ppm | 0.83153 | 0.046295 | 0.73894 | 0.92412 | 0.69265 | 0.97042 | 5.567% | 11.13% | 16.7% | 0.78996 | 0.87311
Tm, ppm | 0.26147 | 0.065702 | 0.13007 | 0.39287 | 0.064367 | 0.45858 | 25.13% | 50.26% | 75.38% | 0.2484 | 0.27454
U, ppm | 2.6641 | 0.10168 | 2.4608 | 2.8675 | 2.3591 | 2.9692 | 3.817% | 7.633% | 11.45% | 2.5309 | 2.7973
V, ppm | 105.2798 | 3.4113 | 98.4572 | 112.1024 | 95.0458 | 115.5137 | 3.24% | 6.48% | 9.721% | 100.0158 | 110.5438
W, ppm | 2.9352 | 0.35387 | 2.2275 | 3.643 | 1.8736 | 3.9969 | 12.06% | 24.11% | 36.17% | 2.7885 | 3.082
Y, ppm | 15.6328 | 2.3589 | 10.915 | 20.3505 | 8.5561 | 22.7094 | 15.09% | 30.18% | 45.27% | 14.8511 | 16.4144
Yb, ppm | 1.7028 | 0.27654 | 1.1497 | 2.2559 | 0.87316 | 2.5324 | 16.24% | 32.48% | 48.72% | 1.6176 | 1.7879
Zn, ppm | 97.8706 | 3.6917 | 90.4871 | 105.2541 | 86.7953 | 108.9458 | 3.772% | 7.544% | 11.32% | 92.977 | 102.7641
Zr, ppm | 130.5435 | 9.0607 | 112.422 | 148.6649 | 103.3613 | 157.7257 | 6.941% | 13.88% | 20.82% | 124.0163 | 137.0707
Aqua Regia Digestion
Ag, ppm | 0.26607 | 0.012651 | 0.24076 | 0.29137 | 0.22811 | 0.30402 | 4.755% | 9.51% | 14.26% | 0.25276 | 0.27937
Al, wt.% | 2.5807 | 0.14796 | 2.2847 | 2.8766 | 2.1368 | 3.0246 | 5.734% | 11.47% | 17.2% | 2.4516 | 2.7097
As, ppm | 810.5157 | 31.2543 | 748.0072 | 873.0242 | 716.753 | 904.2785 | 3.856% | 7.712% | 11.57% | 769.9899 | 851.0415
Ba, ppm | 106.9023 | 8.8958 | 89.1108 | 124.6939 | 80.215 | 133.5897 | 8.321% | 16.64% | 24.96% | 101.5572 | 112.2475
Be, ppm | 1.1903 | 0.081857 | 1.0266 | 1.354 | 0.94471 | 1.4359 | 6.877% | 13.75% | 20.63% | 1.1308 | 1.2498
Bi, ppm | 0.37393 | 0.022483 | 0.32896 | 0.41889 | 0.30648 | 0.44137 | 6.013% | 12.03% | 18.04% | 0.35523 | 0.39262
Ca, wt.% | 0.33321 | 0.023331 | 0.28655 | 0.37987 | 0.26321 | 0.4032 | 7.002% | 14% | 21.01% | 0.31655 | 0.34987
Cd, ppm | 0.072145 | 0.015817 | 0.040511 | 0.10378 | 0.024693 | 0.1196 | 21.92% | 43.85% | 65.77% | 0.068538 | 0.075752
Ce, ppm | 54.4618 | 3.2379 | 47.9861 | 60.9375 | 44.7482 | 64.1754 | 5.945% | 11.89% | 17.84% | 51.7387 | 57.1849
Co, ppm | 15.1646 | 1.0254 | 13.1138 | 17.2154 | 12.0884 | 18.2408 | 6.762% | 13.52% | 20.29% | 14.4064 | 15.9228
Cr, ppm | 106.5348 | 4.9418 | 96.6513 | 116.4184 | 91.7095 | 121.3602 | 4.639% | 9.277% | 13.92% | 101.2081 | 111.8616
Cs, ppm | 6.7143 | 0.18376 | 6.3467 | 7.0818 | 6.163 | 7.2655 | 2.737% | 5.474% | 8.21% | 6.3786 | 7.05
Cu, ppm | 29.557 | 1.4841 | 26.5887 | 32.5252 | 25.1046 | 34.0094 | 5.021% | 10.04% | 15.06% | 28.0791 | 31.0348
Fe, wt.% | 3.5146 | 0.10816 | 3.2983 | 3.7309 | 3.1901 | 3.839 | 3.077% | 6.155% | 9.232% | 3.3388 | 3.6903
Ga, ppm | 8.0876 | 0.45348 | 7.1806 | 8.9945 | 6.7271 | 9.448 | 5.607% | 11.21% | 16.82% | 7.6832 | 8.4919
Ge, ppm | 0.099892 | 0.02347 | 0.052952 | 0.14683 | 0.029483 | 0.1703 | 23.5% | 46.99% | 70.49% | 0.094897 | 0.10489
Hf, ppm | 0.54012 | 0.12387 | 0.29238 | 0.78786 | 0.16852 | 0.91173 | 22.93% | 45.87% | 68.8% | 0.51312 | 0.56713
In, ppm | 0.029758 | 0.00079312 | 0.028172 | 0.031345 | 0.027379 | 0.032138 | 2.665% | 5.33% | 7.996% | 0.02827 | 0.031246
K, wt.% | 0.86443 | 0.036189 | 0.79205 | 0.93681 | 0.75586 | 0.973 | 4.186% | 8.373% | 12.56% | 0.82121 | 0.90765
La, ppm | 27.1269 | 2.1991 | 22.7286 | 31.5252 | 20.5295 | 33.7243 | 8.107% | 16.21% | 24.32% | 25.7705 | 28.4832
Li, ppm | 38.1829 | 2.2496 | 33.6836 | 42.6822 | 31.434 | 44.9318 | 5.892% | 11.78% | 17.68% | 36.2738 | 40.0921
Lu, ppm | 0.11511 | 0.032002 | 0.051106 | 0.17912 | 0.019103 | 0.21112 | 27.8% | 55.6% | 83.4% | 0.10936 | 0.12087
Mg, wt.% | 1.4078 | 0.072746 | 1.2623 | 1.5533 | 1.1895 | 1.626 | 5.167% | 10.33% | 15.5% | 1.3374 | 1.4782
Mn, wt.% | 0.028042 | 0.0013621 | 0.025318 | 0.030766 | 0.023955 | 0.032128 | 4.857% | 9.715% | 14.57% | 0.02664 | 0.029444
Mo, ppm | 0.89612 | 0.034807 | 0.82651 | 0.96573 | 0.7917 | 1.0005 | 3.884% | 7.768% | 11.65% | 0.85132 | 0.94093
Na, wt.% | 0.090925 | 0.014837 | 0.061251 | 0.1206 | 0.046414 | 0.13544 | 16.32% | 32.64% | 48.95% | 0.086379 | 0.095471
Nb, ppm | 0.40136 | 0.10421 | 0.19295 | 0.60978 | 0.088748 | 0.71398 | 25.96% | 51.93% | 77.89% | 0.3813 | 0.42143
Ni, ppm | 59.3099 | 3.7009 | 51.9081 | 66.7118 | 48.2071 | 70.4127 | 6.24% | 12.48% | 18.72% | 56.3444 | 62.2754
P, wt.% | 0.058434 | 0.0023607 | 0.053713 | 0.063156 | 0.051352 | 0.065516 | 4.04% | 8.08% | 12.12% | 0.055513 | 0.061356
Pb, ppm | 14.434 | 1.4727 | 11.4886 | 17.3794 | 10.0159 | 18.8521 | 10.2% | 20.41% | 30.61% | 13.7123 | 15.1557
Rb, ppm | 83.3284 | 2.361 | 78.6064 | 88.0503 | 76.2454 | 90.4113 | 2.833% | 5.667% | 8.5% | 79.1619 | 87.4948
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.20722 | 0.010106 | 0.18701 | 0.22743 | 0.17691 | 0.23754 | 4.877% | 9.754% | 14.63% | 0.19686 | 0.21758
Sb, ppm | 553.6866 | 91.952 | 369.7826 | 737.5907 | 277.8306 | 829.5427 | 16.61% | 33.21% | 49.82% | 526.0023 | 581.371
Sc, ppm | 6.3277 | 0.48966 | 5.3484 | 7.3071 | 4.8588 | 7.7967 | 7.738% | 15.48% | 23.22% | 6.0114 | 6.6441
Sn, ppm | 1.4471 | 0.090395 | 1.2663 | 1.6279 | 1.1759 | 1.7183 | 6.247% | 12.49% | 18.74% | 1.3747 | 1.5194
Sr, ppm | 28.8958 | 1.3303 | 26.2352 | 31.5564 | 24.9049 | 32.8866 | 4.604% | 9.208% | 13.81% | 27.451 | 30.3406
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 11.7951 | 0.76097 | 10.2732 | 13.3171 | 9.5122 | 14.078 | 6.452% | 12.9% | 19.35% | 11.2054 | 12.3849
Ti, wt.% | 0.14198 | 0.010113 | 0.12175 | 0.1622 | 0.11164 | 0.17232 | 7.123% | 14.25% | 21.37% | 0.13488 | 0.14908
Tl, ppm | 0.51737 | 0.045323 | 0.42673 | 0.60802 | 0.38141 | 0.65334 | 8.76% | 17.52% | 26.28% | 0.49151 | 0.54324
U, ppm | 1.4506 | 0.11471 | 1.2212 | 1.68 | 1.1065 | 1.7947 | 7.908% | 15.82% | 23.72% | 1.3781 | 1.5231
V, ppm | 67.0747 | 3.0682 | 60.9382 | 73.2112 | 57.87 | 76.2794 | 4.574% | 9.149% | 13.72% | 63.721 | 70.4284
W, ppm | 0.55586 | 0.07642 | 0.40302 | 0.7087 | 0.3266 | 0.78512 | 13.75% | 27.5% | 41.24% | 0.52807 | 0.58365
Y, ppm | 8.9394 | 0.45487 | 8.0297 | 9.8491 | 7.5748 | 10.304 | 5.088% | 10.18% | 15.27% | 8.4924 | 9.3864
Yb, ppm | 0.86918 | 0.11692 | 0.63534 | 1.103 | 0.51843 | 1.2199 | 13.45% | 26.9% | 40.35% | 0.82572 | 0.91263
Zn, ppm | 85.7743 | 3.7595 | 78.2553 | 93.2932 | 74.4958 | 97.0527 | 4.383% | 8.766% | 13.15% | 81.4855 | 90.063
Zr, ppm | 21.8632 | 2.6626 | 16.538 | 27.1885 | 13.8753 | 29.8511 | 12.18% | 24.36% | 36.54% | 20.7701 | 22.9564
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 239b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PAAU02 | PhotonAssay (Au) two-cycle analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 239b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Hermosillo, Sonora, Mexico
Carsurin, Cikarang, West Java, Indonesia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Perth, WA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Kalgoorlie, WA, Australia
MinAnalytical Services, Perth, WA, Australia
MinAnalytical Services, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 239b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 4 | 4 | 219554 | 0.083231 | 3.9121 | 3.864
2 | 1 | ANSLu | 1 | 18 | 18 | 219555 | 0.083031 | 3.7941 | 3.8577
2 | 1 | ANSLu | 1 | 7 | 7 | 219556 | 0.082483 | 3.8327 | 3.8597
2 | 1 | ANSLu | 1 | 2 | 2 | 219557 | 0.084359 | 3.8583 | 3.8611
2 | 1 | ANSLu | 1 | 6 | 6 | 219558 | 0.088031 | 3.8314 | 3.8597
2 | 1 | ANSLu | 1 | 13 | 13 | 219559 | 0.085539 | 3.9041 | 3.8635
2 | 1 | ANSLu | 1 | 9 | 9 | 219560 | 0.087549 | 3.8317 | 3.8597
2 | 1 | ANSLu | 1 | 20 | 20 | 219561 | 0.082593 | 3.737 | 3.8546
2 | 1 | ANSLu | 1 | 3 | 3 | 219562 | 0.083931 | 3.947 | 3.8658
2 | 1 | ANSLu | 1 | 10 | 10 | 219563 | 0.084266 | 3.7491 | 3.8553
2 | 1 | ANSLu | 1 | 16 | 16 | 219564 | 0.086483 | 3.8302 | 3.8596
2 | 1 | ANSLu | 1 | 12 | 12 | 219565 | 0.086629 | 3.7249 | 3.854
2 | 1 | ANSLu | 1 | 17 | 17 | 219566 | 0.086687 | 3.8046 | 3.8582
2 | 1 | ANSLu | 1 | 8 | 8 | 219567 | 0.086708 | 4.4557 | 3.8929
2 | 1 | ANSLu | 1 | 14 | 14 | 219568 | 0.082893 | 3.8341 | 3.8598
2 | 1 | ANSLu | 1 | 19 | 19 | 219569 | 0.085677 | 3.6909 | 3.8522
2 | 1 | ANSLu | 1 | 11 | 11 | 219570 | 0.082993 | 3.8595 | 3.8612
2 | 1 | ANSLu | 1 | 5 | 5 | 219571 | 0.082569 | 3.8734 | 3.8619
2 | 1 | ANSLu | 1 | 1 | 1 | 219572 | 0.086698 | 3.903 | 3.8635
2 | 1 | ANSLu | 1 | 15 | 15 | 219573 | 0.084528 | 3.8513 | 3.8607
Mean | [6 empty] | 0.084844 | 3.8613 | 3.8613
Median | [7 empty] | 3.8334 | 3.8598
SD | [7 empty] | 0.15471 | 0.0082274
RSD | [7 empty] | 4.007% | 0.2131%
RSD Upscale Factor | [7 empty] | 0.05318
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 239b (Certified Value 3.61 ppm) | [32 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*MS | 2
|  |  | 0.085g | 25g | 50g | 50g | 40g | 30g | 40g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 25g | 25g | 3
| 1 | 1 | 3.903 | 3.836 | 3.66 | 3.618 | 3.48 | 3.85 | 3.71 | 3.53 | 3.55 | 3.63 | 3.53 | 3.65 | 3.52 | 3.66 | 3.57 | 3.743 | 3.5282 | 3.584 | 3.98 | 3.56 | 3.74 | 3.64 | 3.64 | 3.76 | 3.568 | 3.45 | 3.54 | 3.682 | 3.5928 | 3.65 | 3.5057 | 1
| 1 | 2 | 3.8583 | 3.779 | 3.72 | 3.595 | 3.5 | 3.47 | 3.69 | 3.5 | 3.53 | 3.65 | 3.501 | 3.64 | 3.54 | 3.69 | 3.73 | 3.781 | 3.5358 | 3.656 | 3.75 | 3.608 | 3.71 | 3.56 | 3.72 | 3.81 | 3.553 | 3.39 | 3.48 | 3.698 | 3.6192 | 3.68 | 3.4489 | 3
| 1 | 3 | 3.947 | 3.67 | 3.54 | 3.448 | 3.49 | 3.52 | 3.69 | 3.55 | 3.52 | 3.67 | 3.491 | 3.58 | 3.49 | 3.77 | 3.69 | 3.686 | 3.5345 | 3.65 | 3.85 | 3.472 | 3.69 | 3.62 | 3.636 | 3.71 | 3.523 | 3.39 | 3.54 | 3.609 | 3.5424 | 3.69 | 3.3973 | 16
| 1 | 4 | 3.9121 | 3.739 | 3.63 | 3.489 | 3.46 | 3.57 | 3.73 | 3.53 | 3.52 | 3.68 | 3.568 | 3.65 | 3.5 | 3.7 | 3.67 | 3.72 | 3.5189 | 3.607 | 3.87 | 3.58 | 3.74 | 3.68 | 3.611 | 3.73 | 3.535 | 3.45 | 3.5 | 3.556 | 3.5136 | 3.75 | 3.4423 | 3.6076
| 1 | 5 | 3.8734 | 3.745 | 3.54 | 3.456 | 3.54 | 3.52 | 3.69 | 3.49 | 3.55 | 3.68 | 3.393 | 3.65 | 3.47 | 3.7 | 3.73 | 3.659 | 3.5533 | 3.544 | 3.92 | 3.44 | 3.79 | 3.75 | 3.662 | 3.79 | 3.573 | 3.47 | 3.43 | 3.559 | 3.648 | 3.71 | 3.4938 | 7
| 1 | 6 | 3.8314 | 3.756 | 3.59 | 3.467 | 3.52 | 3.46 | 3.72 | 3.55 | 3.54 | 3.59 | 3.539 | 3.64 | 3.38 | 3.73 | 3.73 | 3.714 | 3.5471 | 3.634 | 3.8 | 3.588 | 3.8 | 3.64 | 3.662 | 3.68 | 3.525 | 3.56 | 3.34 | 3.52 | 3.648 | 3.73 | 3.7348
|  | 7 | 3.8327
|  | 8 | 4.4557
|  | 9 | 3.8317
|  | 10 | 3.7491
|  | 11 | 3.8595
|  | 12 | 3.7249
|  | 13 | 3.9041
|  | 14 | 3.8341
|  | 15 | 3.8513
|  | 16 | 3.8302
|  | 17 | 3.8046
|  | 18 | 3.7941
|  | 19 | 3.6909
|  | 20 | 3.737
| Mean |  | 3.8613 | 3.7542 | 3.6133 | 3.5122 | 3.4983 | 3.565 | 3.705 | 3.525 | 3.535 | 3.65 | 3.5037 | 3.635 | 3.4833 | 3.7083 | 3.6867 | 3.7172 | 3.5363 | 3.6125 | 3.8617 | 3.5413 | 3.745 | 3.6483 | 3.6552 | 3.7467 | 3.5462 | 3.4517 | 3.4717 | 3.604 | 3.594 | 3.7017 | 3.5038
| Median |  | 3.8334 | 3.7505 | 3.61 | 3.478 | 3.495 | 3.52 | 3.7 | 3.53 | 3.535 | 3.66 | 3.5155 | 3.645 | 3.495 | 3.7 | 3.71 | 3.717 | 3.5352 | 3.6205 | 3.86 | 3.57 | 3.74 | 3.64 | 3.651 | 3.745 | 3.544 | 3.45 | 3.49 | 3.584 | 3.606 | 3.7 | 3.4713
| Std Dev. |  | 0.15471 | 0.054249 | 0.070899 | 0.074714 | 0.028577 | 0.14516 | 0.017607 | 0.0251 | 0.013784 | 0.035214 | 0.06082 | 0.027386 | 0.056095 | 0.037639 | 0.062503 | 0.042696 | 0.01246 | 0.043117 | 0.082321 | 0.068617 | 0.043243 | 0.063377 | 0.037 | 0.049261 | 0.021693 | 0.062743 | 0.076529 | 0.072567 | 0.055848 | 0.036009 | 0.11967
| Rel.Std.Dev. |  | 4.007% | 1.445% | 1.962% | 2.127% | 0.8169% | 4.072% | 0.4752% | 0.7121% | 0.3899% | 0.9648% | 1.736% | 0.7534% | 1.61% | 1.015% | 1.695% | 1.149% | 0.3524% | 1.194% | 2.132% | 1.938% | 1.155% | 1.737% | 1.012% | 1.315% | 0.6117% | 1.818% | 2.204% | 2.014% | 1.554% | 0.9728% | 3.416%
| PDM3 |  | 7.03% | 4.062% | 0.1578% | -2.646% | -3.03% | -1.182% | 2.699% | -2.291% | -2.014% | 1.174% | -2.882% | 0.7584% | -3.446% | 2.791% | 2.191% | 3.036% | -1.977% | 0.1347% | 7.041% | -1.838% | 3.807% | 1.128% | 1.317% | 3.854% | -1.704% | -4.323% | -3.769% | -0.1009% | -0.3781% | 2.606% | -2.879%
| Z-Score (Absolute) |  | NA | 1.11 | 0.04 | 0.73 | 0.84 | 0.33 | 0.74 | 0.64 | 0.56 | 0.32 | 0.8 | 0.2 | 0.95 | 0.76 | 0.6 | 0.83 | 0.55 | 0.03 | 1.93 | 0.51 | 1.04 | 0.31 | 0.36 | 1.06 | 0.47 | 1.2 | 1.04 | 0.03 | 0.11 | 0.71 | 0.8
| Analytical results for Pd in OREAS 239b (Indicative Value 0.855 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | [30 empty] | 1
| No | No | 1.34 | [30 empty] | ppb
|  |  | FA*MS | [30 empty] | 3
|  |  | 25g | [30 empty] | 3
| 1 | 1 | 0.77086 | [30 empty] | 1
| 1 | 2 | 0.86589 | [30 empty] | 1
| 1 | 3 | 0.89005 | [30 empty] | 16
| 1 | 4 | 0.82297 | [30 empty] | 0.85492
| 1 | 5 | 0.85875 | [30 empty] | 7
| 1 | 6 | 0.92101
| Mean |  | 0.85492
| Median |  | 0.86232
| Std Dev. |  | 0.052581
| Rel.Std.Dev. |  | 6.15%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 239b (Indicative Value 0.564 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | [30 empty] | 1
| No | No | 1.34 | [30 empty] | ppb
|  |  | FA*MS | [30 empty] | 3
|  |  | 25g | [30 empty] | 3
| 1 | 1 | 0.61318 | [30 empty] | 1
| 1 | 2 | 0.53804 | [30 empty] | 1
| 1 | 3 | 0.53829 | [30 empty] | 16
| 1 | 4 | 0.52339 | [30 empty] | 0.56374
| 1 | 5 | 0.59248 | [30 empty] | 7
| 1 | 6 | 0.57704
| Mean |  | 0.56374
| Median |  | 0.55767
| Std Dev. |  | 0.035731
| Rel.Std.Dev. |  | 6.338%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 239b (Certified Value 3.5 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 0.085g | 25g | 50g | 30g | 40g | 30g | 30g | 25g | 30g | 25g | 25g | 25g | 25g | 15g | 50g | 25g | 50g | 25g | 10g | 25g | 20g | 15g | 15g | 50g | 3
| 1 | 1 | 3.903 | 3.66 | 3.57 | 3.5196 | 3.41 | > 2 | 3.55 | > 0.5 | 3.49 | 3.42 | 3.79 | 3.21 | 3.49 | 3.605 | 3.42 | 3.67 | 3.49 | 3.22 | 3.47 | 3.553 | 3.521 | 3.2738 | 3.3138 | 3.56 | 1
| 1 | 2 | 3.8583 | 3.61 | 3.56 | 3.5329 | 3.31 | > 2 | 3.58 | > 0.5 | 3.5 | 3.41 | 3.88 | 3.2 | 3.51 | 3.5536 | 3.38 | 3.68 | 3.54 | 3.05 | 3.58 | 3.597 | 3.5085 | 3.1773 | 3.293 | 3.5519 | 3
| 1 | 3 | 3.947 | 3.65 | 3.36 | 3.6099 | 3.43 | > 2 | 3.53 | > 0.5 | 3.44 | 3.47 | 3.5 | 3.36 | 3.47 | 3.3965 | 3.57 | 3.66 | 3.52 | 3.23 | 3.6 | 3.553 | 3.5371 | 3.309 | 3.3292 | 3.5792 | 16
| 1 | 4 | 3.9121 | 3.4 | 3.53 | 3.5905 | 3.39 | > 2 | 3.57 | > 0.5 | 3.49 | 3.2 | 3.61 | 3.58 | 3.51 | 3.5273 | 3.44 | 3.71 | 3.47 | 3.12 | 3.55 | 3.56 | 3.5351 | 3.2831 | 3.3214 | 3.59 | 3.501
| 1 | 5 | 3.8734 | 3.54 | 3.52 | 3.5547 | 3.42 | > 2 | 3.56 | > 0.5 | 3.42 | 3.44 | 3.4 | 3.49 | 3.49 | 3.6384 | 3.39 | 3.7 | 3.42 | 3.18 | 3.52 | 3.54 | 3.5014 | 3.3658 | 3.3613 | 3.57 | 9
| 1 | 6 | 3.8314 | 3.44 | 3.42 | 3.5711 | 3.39 | > 2 | 3.54 | > 0.5 | 3.48 | 3.42 | 3.51 | 3.39 | 3.49 | 3.5991 | 3.38 | 3.71 | 3.52 | 3.09 | 3.54 | 3.517 | 3.4717 | 3.1394 | 3.3433 | 3.53
|  | 7 | 3.8327
|  | 8 | 4.4557
|  | 9 | 3.8317
|  | 10 | 3.7491
|  | 11 | 3.8595
|  | 12 | 3.7249
|  | 13 | 3.9041
|  | 14 | 3.8341
|  | 15 | 3.8513
|  | 16 | 3.8302
|  | 17 | 3.8046
|  | 18 | 3.7941
|  | 19 | 3.6909
|  | 20 | 3.737
| Mean |  | 3.8613 | 3.55 | 3.4933 | 3.5631 | 3.3917 |  | 3.555 |  | 3.47 | 3.3933 | 3.615 | 3.3717 | 3.4933 | 3.5533 | 3.43 | 3.6883 | 3.4933 | 3.1483 | 3.5433 | 3.5533 | 3.5125 | 3.2581 | 3.327 | 3.5635
| Median |  | 3.8334 | 3.575 | 3.525 | 3.5629 | 3.4 |  | 3.555 |  | 3.485 | 3.42 | 3.56 | 3.375 | 3.49 | 3.5763 | 3.405 | 3.69 | 3.505 | 3.15 | 3.545 | 3.553 | 3.5148 | 3.2785 | 3.3253 | 3.565
| Std Dev. |  | 0.15471 | 0.10991 | 0.084301 | 0.034292 | 0.043089 |  | 0.018708 |  | 0.032249 | 0.097091 | 0.18512 | 0.15065 | 0.015055 | 0.086342 | 0.072664 | 0.02137 | 0.043665 | 0.073052 | 0.045898 | 0.026265 | 0.024462 | 0.084475 | 0.023695 | 0.021261
| Rel.Std.Dev. |  | 4.007% | 3.096% | 2.413% | 0.9624% | 1.27% |  | 0.5263% |  | 0.9294% | 2.861% | 5.121% | 4.468% | 0.431% | 2.43% | 2.118% | 0.5794% | 1.25% | 2.32% | 1.295% | 0.7392% | 0.6964% | 2.593% | 0.7122% | 0.5966%
| PDM3 |  | 10.29% | 1.4% | -0.2186% | 1.775% | -3.123% |  | 1.543% |  | -0.8851% | -3.075% | 3.257% | -3.694% | -0.2186% | 1.495% | -2.028% | 5.351% | -0.2186% | -10.07% | 1.21% | 1.495% | 0.3279% | -6.939% | -4.97% | 1.786%
| Z-Score (Absolute) |  | NA | 0.6 | 0 | 0.74 | 1.08 | NA | 0.66 | NA | 0.25 | 1.06 | 1.3 | 1.3 | 0 | 0.64 | 0.67 | 2.08 | 0 | 3.67 | 0.53 | 0.64 | 0.2 | 2.5 | 1.77 | 0.75
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 239b (Certified Value 3.24 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.28 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*OES | 2
|  |  | 0.085g | 200g | 200g | 200g | 50g | 50g | 20g | 200g | 05g | 30g | 30g | 30g | 30g | 30g | 200g | 50g | 30g | 15g | 3
| 1 | 1 | 3.903 | 3.2 | 3.24 | 3.51 | 2.88 | 3.26 | 3.05 | 3.33 | 3.6 | 3.04 | 3.7 | 3.12 | 3.18 | 2.86 | 2.9198 | 3.56 | 2.45 | 3.401 | 1
| 1 | 2 | 3.8583 | 3.15 | 3.3 | 3.48 | 2.89 | 3.24 | 3.01 | 3.3 | 3.6 | 3.06 | 3.59 | 3.1 | 3.2 | 3.01 | 2.9932 | 3.42 | 2.22 | 3.344 | 3
| 1 | 3 | 3.947 | 3.15 | 3.27 | 3.54 | 2.88 | 3.26 | 2.98 | 3.32 | 3.4 | 3.02 | 3.52 | 3.17 | 3.19 | 3 | 2.9935 | 3.391 | 2.15 | 3.344 | 16
| 1 | 4 | 3.9121 | 3.14 | 3.27 | 3.54 | 2.93 | 3.29 | 3.1 | 3.32 | 3.5 | 3.02 | 3.63 | 3.16 | 3.24 | 2.99 | 2.9117 | 3.422 | 2.04 | 3.42 | 3.2359
| 1 | 5 | 3.8734 | 3.15 | 3.44 | 3.49 | 2.89 | 3.22 | 2.96 | 3.3 | 3.5 | 3 | 3.64 | 3.16 | 3.23 | 3.01 | 2.8821 | 3.5 | 1.97 | 3.3725 | 11
| 1 | 6 | 3.8314 | 3.18 | 3.39 | 3.54 | 2.9 | 3.24 | 3.1 | 3.32 | 3.4 | 3.01 | 3.61 | 3.15 | 3.27 | 3.01 | 2.8859 | 3.546 | 2.25 | 3.3535
|  | 7 | 3.8327
|  | 8 | 4.4557
|  | 9 | 3.8317
|  | 10 | 3.7491
|  | 11 | 3.8595
|  | 12 | 3.7249
|  | 13 | 3.9041
|  | 14 | 3.8341
|  | 15 | 3.8513
|  | 16 | 3.8302
|  | 17 | 3.8046
|  | 18 | 3.7941
|  | 19 | 3.6909
|  | 20 | 3.737
| Mean |  | 3.8613 | 3.1617 | 3.3183 | 3.5167 | 2.895 | 3.2517 | 3.0333 | 3.315 | 3.5 | 3.025 | 3.615 | 3.1433 | 3.2183 | 2.98 | 2.931 | 3.4732 | 2.18 | 3.3725
| Median |  | 3.8334 | 3.15 | 3.285 | 3.525 | 2.89 | 3.25 | 3.03 | 3.32 | 3.5 | 3.02 | 3.62 | 3.155 | 3.215 | 3.005 | 2.9158 | 3.461 | 2.185 | 3.363
| Std Dev. |  | 0.15471 | 0.023166 | 0.078846 | 0.027325 | 0.018708 | 0.024014 | 0.059889 | 0.012247 | 0.089443 | 0.021679 | 0.059582 | 0.027325 | 0.034303 | 0.05933 | 0.050378 | 0.071779 | 0.16971 | 0.031793
| Rel.Std.Dev. |  | 4.007% | 0.7327% | 2.376% | 0.777% | 0.6462% | 0.7385% | 1.974% | 0.3695% | 2.556% | 0.7167% | 1.648% | 0.8693% | 1.066% | 1.991% | 1.719% | 2.067% | 7.785% | 0.9427%
| PDM3 |  | 19.33% | -2.293% | 2.548% | 8.677% | -10.53% | 0.4879% | -6.259% | 2.445% | 8.162% | -6.517% | 11.72% | -2.86% | -0.5422% | -7.908% | -9.421% | 7.333% | -32.63% | 4.222%
| Z-Score (Absolute) |  | NA | 0.2 | 0.35 | 1.04 | 1.13 | 0.12 | 0.65 | 0.34 | 0.98 | 0.67 | 1.38 | 0.26 | 0 | 0.83 | 1 | 0.89 | 3.62 | 0.54
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 239b (Certified Value 3.71 ppm) | [6 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | ppm
|  |  | INAA | Raw*PA | Raw*PA | PAAU02 | Raw*PA | 2
3
| 1 | 1 | 3.903 | 3.41 | 3.54 | 3.67 | 3.76 | 1
| 1 | 2 | 3.8583 | 3.7 | 3.61 | 3.78 | 3.61 | 3
| 1 | 3 | 3.947 | 3.58 | 3.59 | 3.78 | 3.52 | 16
| 1 | 4 | 3.9121 | 3.49 | 3.73 | 3.74 | 3.69 | 3.71
| 1 | 5 | 3.8734 | 3.49 | 3.75 | 3.82 | 3.59 | 13
| 1 | 6 | 3.8314 | 3.56 | 3.73 | 3.77 | 3.69
| 2 | 7 | 3.8327 | 3.721 | 3.55 | 3.86 | 3.69
| 2 | 8 | 4.4557 | 3.592 | 3.62 | 3.653 | 3.65
| 2 | 9 | 3.8317 | 3.868 | 3.59 | 3.911 | 3.81
| 2 | 10 | 3.7491 | 3.756 | 3.7 | 4.016 | 3.67
| 2 | 11 | 3.8595 | 3.905 | 3.64 | 4.082 | 3.64
| 2 | 12 | 3.7249 | 3.84 | 3.65 | 3.994 | 3.83
| 3 | 13 | 3.9041 |  | 3.6 | 3.625
| 3 | 14 | 3.8341 |  | 3.84 | 3.616
| 3 | 15 | 3.8513 |  | 3.7 | 3.752
| 3 | 16 | 3.8302 |  | 3.69 | 3.867
| 3 | 17 | 3.8046 |  | 3.7 | 3.99
| 3 | 18 | 3.7941 |  | 3.64 | 4.045
| 4 | 19 | 3.6909 |  | 3.662 | 3.584
| 4 | 20 | 3.737 |  | 3.737 | 3.694
| 4 | 21 |  |  | 3.609 | 3.515
| 4 | 22 |  |  | 3.684 | 3.775
| 4 | 23 |  |  | 3.596 | 3.487
| 4 | 24 |  |  | 3.689 | 3.768
| 5 | 25 |  |  | 3.59 | 3.75
| 5 | 26 |  |  | 3.881 | 3.775
| 5 | 27 |  |  | 3.752 | 3.651
| 5 | 28 |  |  | 3.741 | 3.981
| 5 | 29 |  |  | 3.897 | 3.68
| 5 | 30 |  |  | 3.801 | 3.882
| 6 | 31 |  |  | 3.829 | 3.681
| 6 | 32 |  |  | 3.9315 | 3.875
| 6 | 33 |  |  | 3.9297 | 3.625
| 6 | 34 |  |  | 3.9089 | 3.911
| 6 | 35 |  |  | 3.8558 | 3.751
| 6 | 36 |  |  | 3.8288 | 3.899
| Mean |  | 3.8613 | 3.6593 | 3.7165 | 3.7849 | 3.6792
| Median |  | 3.8334 | 3.646 | 3.7 | 3.7725 | 3.68
| Std Dev. |  | 0.15471 | 0.1628 | 0.11323 | 0.1489 | 0.089082
| Rel.Std.Dev. |  | 4.007% | 4.449% | 3.047% | 3.934% | 2.421%
| PDM3 |  | 4.078% | -1.365% | 0.1754% | 2.019% | -0.8299%
| Z-Score (Absolute) |  | NA | 0.91 | 0.44 | 2.05 | 0.44
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 239b (Certified Value 0.273 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.25 | 0.2 | 0.3 | 0.664 | 0.3 | < 0.5 | 0.3 | 0.5 | 0.25 | 0.26 | 0.28 | 0.28 | 0.19 | < 1 | 0.43 | 0.37 | < 0.5 | 0.25 | 0.29 | 0.229 | 0.40573 | 1.3473 |  |  |  |  | 1
| 1 | 2 | 0.28 | 0.4 | 0.31 | 0.6381 | 0.3 | < 0.5 | 0.3 | 0.4 | 0.29 | 0.25 | 0.26 | 0.29 | 0.16 | < 1 | 0.42 | 0.36 | < 0.5 | 0.27 | 0.26 | 0.238 | 0.42321 | 1.1605 |  |  |  |  | 25
| 1 | 3 | 0.29 | 0.4 | 0.32 | 0.65125 | 0.3 | < 0.5 | 0.4 | 0.5 | 0.26 | 0.25 | 0.26 | 0.25 | 0.22 | < 1 | 0.42 | 0.34 | < 0.5 | 0.24 | 0.3 | 0.242 | 0.35551 | 1.0246 |  |  |  |  | 16
| 1 | 4 | 0.28 | 0.4 | 0.29 | 0.63915 | 0.3 | < 0.5 | 0.4 | 0.5 | 0.26 | 0.3 | 0.28 | 0.26 | 0.21 | < 1 | 0.44 | 0.38 | < 0.5 | 0.19 | 0.28 | 0.246 | 0.45275 | 1.1095 |  |  |  |  | 0.27276
| 1 | 5 | 0.29 | 0.4 | 0.33 | 0.6474 | 0.3 | < 0.5 | 0.4 | 0.5 | 0.27 | 0.29 | 0.25 | 0.28 | 0.2 | < 1 | 0.44 | 0.38 | < 0.5 | 0.23 | 0.27 | 0.237 | 0.47349 | 1.5255 |  |  |  |  | 15
| 1 | 6 | 0.29 | 0.4 | 0.49 | 0.66 | 0.3 | < 0.5 | 0.4 | 0.4 | 0.29 | 0.27 | 0.26 | 0.26 | 0.29 | < 1 | 0.42 | 0.36 | < 0.5 | 0.22 | 0.29 | 0.234 | 0.36362 | 1.2793
| Mean |  | 0.28 | 0.36667 | 0.34 | 0.64998 | 0.3 |  | 0.36667 | 0.46667 | 0.27 | 0.27 | 0.265 | 0.27 | 0.21167 |  | 0.42833 | 0.365 |  | 0.23333 | 0.28167 | 0.23767 | 0.41239 | 1.2411
| Median |  | 0.285 | 0.4 | 0.315 | 0.64932 | 0.3 |  | 0.4 | 0.5 | 0.265 | 0.265 | 0.26 | 0.27 | 0.205 |  | 0.425 | 0.365 |  | 0.235 | 0.285 | 0.2375 | 0.41447 | 1.2199
| Std Dev. |  | 0.015492 | 0.08165 | 0.074833 | 0.010617 | 0 |  | 0.05164 | 0.05164 | 0.016733 | 0.020976 | 0.012247 | 0.015492 | 0.043551 |  | 0.0098319 | 0.015166 |  | 0.027325 | 0.01472 | 0.0059554 | 0.047196 | 0.18128
| Rel.Std.Dev. |  | 5.533% | 22.27% | 22.01% | 1.633% | 0% |  | 14.08% | 11.07% | 6.197% | 7.769% | 4.622% | 5.738% | 20.58% |  | 2.295% | 4.155% |  | 11.71% | 5.226% | 2.506% | 11.44% | 14.61%
| PDM3 |  | 2.655% | 34.43% | 24.65% | 138.3% | 9.988% |  | 34.43% | 71.09% | -1.011% | -1.011% | -2.844% | -1.011% | -22.4% |  | 57.04% | 33.82% |  | -14.45% | 3.266% | -12.87% | 51.19% | 355%
| Z-Score (Absolute) |  | 0.09 | NA | 1.11 | 6.4 | NA | 0.43 | NA | NA | 0.09 | 0.09 | 0.17 | 0.09 | 1.08 | 3.84 | 2.62 | 1.54 | 0.43 | 0.71 | 0.11 | 0.64 | 2.35 | 16.49
| Analytical results for Al in OREAS 239b (Certified Value 7.36 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 7.3763 | 7.2449 | 7.59 | 7.48 | 7.67 | 7.347 | 5.67 | 7.47 | 7.25 | 7.45 | 7.31 | 7.56 | 7.38 | 7.4993 | 7.4035 | 7.304 | 7.33 | 7.25 | 7.16 | 7.55 | 6.98 | 7.1682 | 7.8076 | 7.4981 |  |  | 1
| 1 | 2 | 7.3875 | 7.2231 | 7.49 | 7.01 | 7.6172 | 7.545 | 5.54 | 7.52 | 7.27 | 7.56 | 7.12 | 7.23 | 7.54 | 7.3599 | 7.3109 | 7.462 | 7.39 | 7.24 | 7.02 | 7.54 | 6.97 | 7.4144 | 8.0283 | 7.5043 |  |  | #N/A
| 1 | 3 | 7.4868 | 7.2395 | 7.56 | 5.94 | 7.58 | 7.552 | 6.32 | 7.39 | 7.21 | 7.39 | 7.47 | 7.42 | 7.31 | 7.5639 | 7.1296 | 8.68 | 7.25 | 7.19 | 7.12 | 7.45 | 6.97 | 7.2372 | 7.9922 | 7.5099 |  |  | 16
| 1 | 4 | 7.4423 | 7.2236 | 7.45 | 6.99 | 7.62 | 7.37 | 6.47 | 7.29 | 7.16 | 7.27 | 7.55 | 7.27 | 7.33 | 7.5518 | 7.1818 | 7.356 | 7.45 | 7.23 | 7.11 | 7.46 | 6.97 | 7.33 | 7.9231 | 7.5024 |  |  | 7.3591
| 1 | 5 | 7.6055 | 7.2638 | 7.53 | 7.67 | 7.6701 | 7.438 | 6.58 | 7.6 | 7.26 | 7.2 | 7.21 | 7.32 | 7.29 | 7.5801 | 7.0039 | 7.409 | 7.4 | 7.25 | 7.22 | 7.48 | 6.96 | 7.3786 | 7.8329 | 7.4972 |  |  | 16
| 1 | 6 | 7.6978 | 7.2248 | 7.42 | 7.08 | 7.6001 | 7.398 | 5.8 | 7.52 | 7.26 | 7.31 | 7.82 | 7.28 | 7.33 | 6.6519 | 7.0547 | 7.515 | 7.4 | 7.15 | 7.31 | 7.49 | 7.16 | 7.34 | 7.996 | 7.5107
| Mean |  | 7.4994 | 7.2366 | 7.5067 | 7.0283 | 7.6262 | 7.4417 | 6.0633 | 7.465 | 7.235 | 7.3633 | 7.4133 | 7.3467 | 7.3633 | 7.3678 | 7.1807 | 7.621 | 7.37 | 7.2183 | 7.1567 | 7.495 | 7.0017 | 7.3114 | 7.93 | 7.5038
| Median |  | 7.4646 | 7.2322 | 7.51 | 7.045 | 7.6186 | 7.418 | 6.06 | 7.495 | 7.255 | 7.35 | 7.39 | 7.3 | 7.33 | 7.5256 | 7.1557 | 7.4355 | 7.395 | 7.235 | 7.14 | 7.485 | 6.97 | 7.335 | 7.9577 | 7.5033
| Std Dev. |  | 0.12784 | 0.016156 | 0.06532 | 0.60078 | 0.036822 | 0.08817 | 0.44635 | 0.11005 | 0.042308 | 0.13049 | 0.25508 | 0.12291 | 0.091579 | 0.35979 | 0.15251 | 0.52415 | 0.070143 | 0.040208 | 0.099733 | 0.041352 | 0.077825 | 0.091979 | 0.092028 | 0.0057109
| Rel.Std.Dev. |  | 1.705% | 0.2233% | 0.8702% | 8.548% | 0.4828% | 1.185% | 7.361% | 1.474% | 0.5848% | 1.772% | 3.441% | 1.673% | 1.244% | 4.883% | 2.124% | 6.878% | 0.9517% | 0.557% | 1.394% | 0.5517% | 1.112% | 1.258% | 1.16% | 0.07611%
| PDM3 |  | 1.905% | -1.665% | 2.005% | -4.495% | 3.629% | 1.121% | -17.61% | 1.438% | -1.687% | 0.05692% | 0.7363% | -0.1696% | 0.05692% | 0.1178% | -2.424% | 3.558% | 0.1475% | -1.913% | -2.751% | 1.846% | -4.858% | -0.6488% | 7.757% | 1.965%
| Z-Score (Absolute) |  | 0.68 | 0.66 | 0.72 | 1.72 | 1.33 | 0.39 | 6.64 | 0.51 | 0.67 | 0.01 | 0.24 | 0.1 | 0.01 | 0.01 | 0.94 | 1.3 | 0.02 | 0.75 | 1.07 | 0.66 | 1.86 | 0.28 | 2.88 | 0.7
| Analytical results for As in OREAS 239b (Certified Value 816 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 0
0
| 1 | 1 | 795.7 | 730 | 845 | 840.4945 | 803 | 819 | 750 | 238.3 | 848 | 829 | 828 | 826 | 843.9 | 765.7853 | 861 | 792.3 | 794.7 | 864.2 | 312.2 | 788.9 | 729.4345 | [5 empty] | 1
| 1 | 2 | 814.6 | 744 | 860 | 841.724 | 806 | 814 | 748 | 243.7 | 849 | 802 | 786 | 867 | 847 | 752.7994 | 919 | 806.3 | 816.1 | 867.5 | 289.6 | 802.5 | 688.1264 | [5 empty] | 26
| 1 | 3 | 825.2 | 760 | 867 | 840.93 | 800 | 795 | 784 | 234.5 | 831 | 835 | 789 | 834 | 849.6 | 761.0656 | 833 | 797.6 | 801.6 | 861.7 | 288.8 | 794.2 | 662.5026 | [5 empty] | 16
| 1 | 4 | 798.5 | 737 | 885 | 838.03 | 809 | 823 | 749 | 239.1 | 829 | 862 | 787 | 844 | 848.6 | 749.9724 | 903 | 807.7 | 853.5 | 852.6 | 276.4 | 807.1 | 664.5799 | [5 empty] | 816.0782
| 1 | 5 | 822.7 | 731 | 856 | 838.7512 | 803 | 800 | 758 | 241.8 | 836 | 825 | 757 | 827 | 838.7 | 749.3079 | 872 | 795.7 | 830.4 | 846.9 | 295.1 | 812.3 | 687.2985 | [5 empty] | 17
| 1 | 6 | 846.1 | 730 | 866 | 841.8999 | 810 | 810 | 789 | 242.8 | 804 | 892 | 761 | 839 | 845.6 | 748.4865 | 904 | 800.8 | 813.2 | 857.2 | 316.1 | 791.6 | 648.1288
| Mean |  | 817.1333 | 738.6667 | 863.1667 | 840.3049 | 805.1667 | 810.1667 | 763 | 240.0333 | 832.8333 | 840.8333 | 784.6667 | 839.5 | 845.5667 | 754.5695 | 882 | 800.0667 | 818.25 | 858.35 | 296.3667 | 799.4333 | 680.0118
| Median |  | 818.65 | 734 | 863 | 840.7123 | 804.5 | 812 | 754 | 240.45 | 833.5 | 832 | 786.5 | 836.5 | 846.3 | 751.3859 | 887.5 | 799.2 | 814.65 | 859.45 | 292.35 | 798.35 | 675.9392
| Std Dev. |  | 18.6983 | 11.7927 | 13.3479 | 1.5858 | 3.8687 | 10.8704 | 18.6118 | 3.4268 | 16.4367 | 31.6191 | 25.4139 | 15.136 | 3.9358 | 7.1687 | 32.3605 | 6.052 | 21.2233 | 7.6767 | 15.1223 | 9.3106 | 28.7151
| Rel.Std.Dev. |  | 2.288% | 1.596% | 1.546% | 0.1887% | 0.4805% | 1.342% | 2.439% | 1.428% | 1.974% | 3.76% | 3.239% | 1.803% | 0.4655% | 0.95% | 3.669% | 0.7564% | 2.594% | 0.8943% | 5.103% | 1.165% | 4.223%
| PDM3 |  | 0.1293% | -9.486% | 5.77% | 2.969% | -1.337% | -0.7244% | -6.504% | -70.59% | 2.053% | 3.033% | -3.849% | 2.87% | 3.613% | -7.537% | 8.078% | -1.962% | 0.2661% | 5.18% | -63.68% | -2.04% | -16.67%
| Z-Score (Absolute) |  | 0.15 | 1.57 | 1.17 | 0.66 | 0.11 | 0 | 1.04 | 12.54 | 0.5 | 0.67 | 0.56 | 0.64 | 0.78 | 1.22 | 1.58 | 0.22 | 0.18 | 1.06 | 11.3 | 0.24 | 2.86
| Analytical results for Au in OREAS 239b (Indicative Value 3.48 ppm) | [27 empty] | Indicative
Au | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 2
2
| 1 | 1 | 3.2643 | [25 empty] | 1
| 1 | 2 | 3.456 | [25 empty] | 3
| 1 | 3 | 3.7782 | [25 empty] | 16
| 1 | 4 | 3.581 | [25 empty] | 3.4836
| 1 | 5 | 3.5306 | [25 empty] | 9
| 1 | 6 | 3.2916
| Mean |  | 3.4836
| Median |  | 3.4933
| Std Dev. |  | 0.19192
| Rel.Std.Dev. |  | 5.509%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for B in OREAS 239b (Indicative Value 29.8 ppm) | [27 empty] | Indicative
B | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.07 | [25 empty] | ppm
|  |  | 4A*OES | [25 empty] | 1
1
| 1 | 1 | 28.87 | [25 empty] | 1
| 1 | 2 | 29.63 | [25 empty] | 4
| 1 | 3 | 27.57 | [25 empty] | 16
| 1 | 4 | 32.69 | [25 empty] | 29.835
| 1 | 5 | 33.55 | [25 empty] | 10
| 1 | 6 | 26.7
| Mean |  | 29.835
| Median |  | 29.25
| Std Dev. |  | 2.7526
| Rel.Std.Dev. |  | 9.226%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 239b (Certified Value 678 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 667.4 | 685.4772 | 681 | 627 | 678.601 | 705 | 683 | 691 | 641 | 700 | 680 | 690 | 680 | 683 | 668.3557 | 678 | 635.6 | 674 | 652 | 683 | 685 | 647 | 655.3 | 710 | 701.5404 | 590.2756 | 1
| 1 | 2 | 668.6 | 687.5723 | 683 | 627 | 680.5575 | 708 | 668 | 682 | 639 | 710 | 660 | 670 | 700 | 688 | 655.7805 | 692 | 643.2 | 687 | 655.6 | 683 | 685 | 650 | 669.5 | 731 | 692.1142 | 601.0764 | 27
| 1 | 3 | 675 | 686.0658 | 681 | 609 | 681.82 | 708 | 689 | 707 | 647 | 690 | 700 | 680 | 680 | 696 | 675.9211 | 771 | 664.2 | 656 | 661 | 695 | 674 | 656 | 653.8 | 725 | 676.8947 | 634.993 | 16
| 1 | 4 | 674.6 | 692.6599 | 698 | 601 | 681.21 | 704 | 698 | 684 | 638 | 690 | 710 | 670 | 680 | 687 | 663.6726 | 681 | 653.4 | 699 | 656.2 | 700 | 675 | 663 | 665.8 | 723 | 685.6827 | 629.2184 | 678.2737
| 1 | 5 | 683 | 688.5109 | 689 | 604 | 684.6 | 705 | 688 | 685 | 644 | 690 | 680 | 670 | 670 | 673 | 666.7799 | 681 | 664.5 | 671 | 659.4 | 705 | 677 | 647 | 664.7 | 717 | 704.6825 | 639.8554 | 18
| 1 | 6 | 681.3 | 696.344 | 684 | 621 | 676.9675 | 704 | 684 | 701 | 652 | 690 | 730 | 670 | 680 | 669 | 676.5047 | 693 | 658 | 670 | 657.1 | 687 | 679 | 651 | 663.9 | 730 | 684.3572 | 621.5169
| Mean |  | 674.9833 | 689.4384 | 686 | 614.8333 | 680.626 | 705.6667 | 685 | 691.6667 | 643.5 | 695 | 693.3333 | 675 | 681.6667 | 682.6667 | 667.8357 | 699.3333 | 653.15 | 676.1667 | 656.8833 | 692.1667 | 679.1667 | 652.3333 | 662.1667 | 722.6667 | 690.8786 | 619.4893
| Median |  | 674.8 | 688.0416 | 683.5 | 615 | 680.8837 | 705 | 686 | 688 | 642.5 | 690 | 690 | 670 | 680 | 685 | 667.5678 | 686.5 | 655.7 | 672.5 | 656.65 | 691 | 678 | 650.5 | 664.3 | 724 | 688.8984 | 625.3676
| Std Dev. |  | 6.365 | 4.2302 | 6.5727 | 11.6347 | 2.6476 | 1.8619 | 9.8793 | 10.1915 | 5.3198 | 8.3666 | 25.0333 | 8.3666 | 9.8319 | 10.0532 | 7.8047 | 35.6576 | 11.6749 | 14.9321 | 3.1397 | 9.261 | 4.8339 | 6.186 | 6.2218 | 8.0166 | 10.6848 | 19.7301
| Rel.Std.Dev. |  | 0.943% | 0.6136% | 0.9581% | 1.892% | 0.389% | 0.2638% | 1.442% | 1.473% | 0.8267% | 1.204% | 3.611% | 1.239% | 1.442% | 1.473% | 1.169% | 5.099% | 1.787% | 2.208% | 0.478% | 1.338% | 0.7117% | 0.9483% | 0.9396% | 1.109% | 1.547% | 3.185%
| PDM3 |  | -0.4851% | 1.646% | 1.139% | -9.353% | 0.3468% | 4.039% | 0.9917% | 1.975% | -5.127% | 2.466% | 2.22% | -0.4827% | 0.5002% | 0.6477% | -1.539% | 3.105% | -3.704% | -0.3106% | -3.154% | 2.048% | 0.1317% | -3.824% | -2.375% | 6.545% | 1.858% | -8.667%
| Z-Score (Absolute) |  | 0.36 | 0.48 | 0.28 | 3.85 | 0.03 | 1.42 | 0.22 | 0.61 | 2.19 | 0.81 | 0.71 | 0.36 | 0.03 | 0.09 | 0.77 | 1.06 | 1.63 | 0.29 | 1.41 | 0.64 | 0.12 | 1.67 | 1.1 | 2.41 | 0.57 | 3.58
| Analytical results for Be in OREAS 239b (Certified Value 2.42 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.24 | 2.5 | 2.4 | < 5 | 2.4 | 2.47 | 2.9 | 2 | 2.43 | 2.44 | 2.54 | 2.36 | 3 | 2.2238 | 2.9 | 2.063 | 2.8 | 3.07 | 2.45 | 2.2 | 2 | 2.3419 | 2.8867 | 2.4306 |  |  | 1
| 1 | 2 | 2.43 | 2.5 | 2.5 | < 5 | 2.4 | 2.48 | 2.3 | 2 | 2.57 | 2.57 | 2.27 | 2.41 | 2 | 2.244 | 2.9 | 2.082 | 2.9 | 3.11 | 2.51 | 2.07 | 2 | 2.387 | 2.8867 | 2.3712 |  |  | 28
| 1 | 3 | 2.38 | 2.5 | 2.4 | < 5 | 2.4 | 2.5 | 2.7 | 2 | 2.47 | 2.48 | 2.47 | 2.31 | 3 | 2.3493 | 2.9 | 2.094 | 2.6 | 3.16 | 2.5 | 2.04 | 2 | 2.3309 | 2.2192 | 2.3961 |  |  | 16
| 1 | 4 | 2.46 | 2.5 | 2.4 | < 5 | 2.3 | 2.55 | 1.9 | 2 | 2.51 | 2.57 | 2.2 | 2.26 | 2 | 2.3415 | 3 | 2.074 | 2.6 | 3.1 | 2.43 | 2.23 | 2 | 2.376 | 2.9424 | 2.3282 |  |  | 2.4153
| 1 | 5 | 2.5 | 2.5 | 2.4 | < 5 | 2.4 | 2.55 | 2.9 | 2 | 2.54 | 2.6 | 2.33 | 2.36 | 3 | 2.3205 | 2.9 | 2.171 | 2.8 | 3.08 | 2.61 | 2.25 | 2 | 2.3672 | 2.7199 | 2.4144 |  |  | 19
| 1 | 6 | 2.44 | 2.5 | 2.5 | < 5 | 2.3 | 2.59 | 2.4 | 2 | 2.44 | 2.52 | 2.27 | 2.34 | 3 | 2.31 | 3 | 2.125 | 2.8 | 3.09 | 2.78 | 2.19 | 2 | 2.3672 | 2.8867 | 2.4214
| Mean |  | 2.4083 | 2.5 | 2.4333 |  | 2.3667 | 2.5233 | 2.5167 | 2 | 2.4933 | 2.53 | 2.3467 | 2.34 | 2.6667 | 2.2982 | 2.9333 | 2.1015 | 2.75 | 3.1017 | 2.5467 | 2.1633 | 2 | 2.3617 | 2.7569 | 2.3937
| Median |  | 2.435 | 2.5 | 2.4 |  | 2.4 | 2.525 | 2.55 | 2 | 2.49 | 2.545 | 2.3 | 2.35 | 3 | 2.3152 | 2.9 | 2.088 | 2.8 | 3.095 | 2.505 | 2.195 | 2 | 2.3672 | 2.8867 | 2.4053
| Std Dev. |  | 0.091305 | 0 | 0.05164 |  | 0.05164 | 0.047188 | 0.392 | 0 | 0.056095 | 0.061319 | 0.13125 | 0.05099 | 0.5164 | 0.052142 | 0.05164 | 0.040143 | 0.12247 | 0.031885 | 0.13033 | 0.087101 | 0 | 0.021193 | 0.27402 | 0.03836
| Rel.Std.Dev. |  | 3.791% | 0% | 2.122% |  | 2.182% | 1.87% | 15.58% | 0% | 2.25% | 2.424% | 5.593% | 2.179% | 19.36% | 2.269% | 1.76% | 1.91% | 4.454% | 1.028% | 5.118% | 4.026% | 0% | 0.8974% | 9.939% | 1.603%
| PDM3 |  | -0.2891% | 3.506% | 0.746% |  | -2.014% | 4.472% | 4.196% | -17.2% | 3.23% | 4.748% | -2.842% | -3.118% | 10.41% | -4.85% | 21.45% | -12.99% | 13.86% | 28.42% | 5.438% | -10.43% | -17.2% | -2.22% | 14.14% | -0.8966%
| Z-Score (Absolute) |  | 0.45 | 0.04 | 0.32 | 0.04 | 0.67 | 0.16 | 0.12 | NA | 0 | 0.2 | 0.78 | 0.82 | NA | 1.04 | 2.34 | 2.09 | 1.37 | 3.24 | 0.28 | 1.76 | NA | 0.7 | 1.4 | 0.53
Results from laboratories 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 239b (Certified Value 0.39 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 0.38 | 0.41234 | 0.4 | 0.37 | < 5 | 0.38 | 0.44 | 0.42 | 0.4 | 0.4 | 0.39 | 0.41 | 0.35 | 0.38 | 0.39195 | 0.5 | 0.356 | 0.37 | < 5 | 0.46 | 0.37 | 0.38 | 0.4674 | < 5 | 1.6383 | 0.41291 | 1
| 1 | 2 | 0.42 | 0.37625 | 0.4 | 0.4 | < 5 | 0.4 | 0.39 | 0.34 | 0.4 | 0.38 | 0.42 | 0.39 | 0.4 | 0.37 | 0.3745 | 0.5 | 0.364 | 0.38 | < 5 | 0.38 | 0.33 | 0.41 | 0.429 | < 5 | 1.855 | 0.41553 | 29
| 1 | 3 | 0.38 | 0.39596 | 0.5 | 0.39 | < 5 | 0.4 | 0.43 | 0.36 | 0.4 | 0.4 | 0.4 | 0.38 | 0.36 | 0.36 | 0.40024 | 0.5 | 0.38 | 0.36 | < 5 | 0.4 | 0.33 | 0.38 | 0.4321 | < 5 | 1.855 | 0.38547 | 16
| 1 | 4 | 0.39 | 0.40793 | 0.7 | 0.4 | < 5 | 0.39 | 0.4 | 0.4 | 0.6 | 0.37 | 0.4 | 0.38 | 0.39 | 0.38 | 0.37415 | 0.5 | 0.384 | 0.42 | < 5 | 0.41 | 0.36 | 0.4 | 0.4589 | < 5 | 1.6679 | 0.39822 | 0.39168
| 1 | 5 | 0.4 | 0.38653 | 0.4 | 0.42 | < 5 | 0.39 | 0.42 | 0.34 | 0.4 | 0.4 | 0.38 | 0.39 | 0.39 | 0.39 | 0.40185 | 0.5 | 0.393 | 0.38 | < 5 | 0.4 | 0.31 | 0.39 | 0.4467 | < 5 | 1.5241 | 0.42103 | 20
| 1 | 6 | 0.45 | 0.39402 | 0.4 | 0.38 | < 5 | 0.4 | 0.4 | 0.41 | 0.4 | 0.4 | 0.4 | 0.38 | 0.36 | 0.38 | 0.38104 | 0.5 | 0.377 | 0.48 | < 5 | 0.44 | 0.36 | 0.39 | 0.4193 | < 5 | 1.2837 | 0.39365
| Mean |  | 0.40333 | 0.39551 | 0.46667 | 0.39333 |  | 0.39333 | 0.41333 | 0.37833 | 0.43333 | 0.39167 | 0.39833 | 0.38833 | 0.375 | 0.37667 | 0.38729 | 0.5 | 0.37567 | 0.39833 |  | 0.415 | 0.34333 | 0.39167 | 0.44223 |  | 1.6373 | 0.40447
| Median |  | 0.395 | 0.39499 | 0.4 | 0.395 |  | 0.395 | 0.41 | 0.38 | 0.4 | 0.4 | 0.4 | 0.385 | 0.375 | 0.38 | 0.3865 | 0.5 | 0.3785 | 0.38 |  | 0.405 | 0.345 | 0.39 | 0.4394 |  | 1.6531 | 0.40557
| Std Dev. |  | 0.027325 | 0.013354 | 0.12111 | 0.017512 |  | 0.008165 | 0.019664 | 0.036009 | 0.08165 | 0.013292 | 0.013292 | 0.01169 | 0.020736 | 0.010328 | 0.012463 | 0 | 0.013515 | 0.044907 |  | 0.029496 | 0.023381 | 0.01169 | 0.018627 |  | 0.21613 | 0.014035
| Rel.Std.Dev. |  | 6.775% | 3.376% | 25.95% | 4.452% |  | 2.076% | 4.757% | 9.518% | 18.84% | 3.394% | 3.337% | 3.01% | 5.53% | 2.742% | 3.218% | 0% | 3.598% | 11.27% |  | 7.107% | 6.81% | 2.985% | 4.212% |  | 13.2% | 3.47%
| PDM3 |  | 2.975% | 0.9762% | 19.14% | 0.4216% |  | 0.4216% | 5.528% | -3.408% | 10.63% | -0.00393% | 1.698% | -0.855% | -4.259% | -3.834% | -1.122% | 27.65% | -4.089% | 1.698% |  | 5.953% | -12.34% | -0.00393% | 12.91% |  | 318% | 3.264%
| Z-Score (Absolute) |  | 0.25 | 0.06 | NA | 0.14 | 82.22 | 0.14 | 0.64 | 0.73 | NA | 0.21 | 0.06 | 0.34 | 0.86 | 0.79 | 0.38 | 4.03 | 0.83 | 0.06 | 82.22 | 0.71 | 2.1 | 0.21 | 1.77 | 82.22 | 48.5 | 0.3
Results from laboratories 1.05 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 239b (Certified Value 0.867 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.8617 | 0.85694 | 0.87 | 0.86 | 0.85667 | 0.893 | 0.75 | 0.851 | 0.9 | 0.88 | 0.88 | 0.9 | 0.89 | 0.8721 | 0.85357 | 0.779 | 0.86 | 0.88 | 0.86 | 0.87 | 0.83 | 0.8413 | 0.9259 |  |  |  | 1
| 1 | 2 | 0.862 | 0.84766 | 0.87 | 0.86 | 0.84 | 0.886 | 0.74 | 0.857 | 0.89 | 0.89 | 0.86 | 0.87 | 0.91 | 0.8497 | 0.83726 | 0.793 | 0.87 | 0.87 | 0.86 | 0.87 | 0.81 | 0.86405 | 0.9576 |  |  |  | #N/A
| 1 | 3 | 0.8725 | 0.84327 | 0.88 | 0.78 | 0.89 | 0.889 | 0.8 | 0.845 | 0.89 | 0.88 | 0.93 | 0.89 | 0.89 | 0.867 | 0.87671 | 0.929 | 0.86 | 0.86 | 0.85 | 0.86 | 0.83 | 0.8397 | 0.9489 |  |  |  | 16
| 1 | 4 | 0.8565 | 0.84135 | 0.88 | 0.84 | 0.87 | 0.897 | 0.83 | 0.831 | 0.89 | 0.86 | 0.9 | 0.87 | 0.89 | 0.8741 | 0.83687 | 0.793 | 0.86 | 0.86 | 0.85 | 0.86 | 0.82 | 0.85831 | 0.9398 |  |  |  | 0.86738
| 1 | 5 | 0.8862 | 0.84327 | 0.89 | 0.9 | 0.86 | 0.909 | 0.84 | 0.864 | 0.89 | 0.86 | 0.87 | 0.88 | 0.88 | 0.868 | 0.86013 | 0.8 | 0.86 | 0.87 | 0.86 | 0.86 | 0.82 | 0.85878 | 0.9385 |  |  |  | 21
| 1 | 6 | 0.8748 | 0.84838 | 0.86 | 0.89 | 0.87 | 0.908 | 0.79 | 0.853 | 0.89 | 0.88 | 0.94 | 0.88 | 0.88 | 0.7905 | 0.84425 | 0.808 | 0.86 | 0.87 | 0.88 | 0.86 | 0.84 | 0.85568 | 0.9609
| Mean |  | 0.86895 | 0.84681 | 0.875 | 0.855 | 0.86444 | 0.897 | 0.79167 | 0.85017 | 0.89167 | 0.875 | 0.89667 | 0.88167 | 0.89 | 0.85357 | 0.85146 | 0.817 | 0.86167 | 0.86833 | 0.86 | 0.86333 | 0.825 | 0.85297 | 0.94527
| Median |  | 0.86725 | 0.84547 | 0.875 | 0.86 | 0.865 | 0.895 | 0.795 | 0.852 | 0.89 | 0.88 | 0.89 | 0.88 | 0.89 | 0.8675 | 0.84891 | 0.7965 | 0.86 | 0.87 | 0.86 | 0.86 | 0.825 | 0.857 | 0.94435
| Std Dev. |  | 0.010957 | 0.0056688 | 0.010488 | 0.042778 | 0.016689 | 0.009654 | 0.040702 | 0.011321 | 0.0040825 | 0.012247 | 0.03266 | 0.01169 | 0.010954 | 0.032082 | 0.015388 | 0.055696 | 0.0040825 | 0.0075277 | 0.010954 | 0.005164 | 0.010488 | 0.010047 | 0.013118
| Rel.Std.Dev. |  | 1.261% | 0.6694% | 1.199% | 5.003% | 1.931% | 1.076% | 5.141% | 1.332% | 0.4578% | 1.4% | 3.642% | 1.326% | 1.231% | 3.759% | 1.807% | 6.817% | 0.4738% | 0.8669% | 1.274% | 0.5981% | 1.271% | 1.178% | 1.388%
| PDM3 |  | 0.1807% | -2.372% | 0.8782% | -1.428% | -0.3388% | 3.415% | -8.729% | -1.985% | 2.8% | 0.8782% | 3.376% | 1.647% | 2.608% | -1.593% | -1.835% | -5.809% | -0.659% | 0.1096% | -0.8512% | -0.4669% | -4.886% | -1.661% | 8.979%
| Z-Score (Absolute) |  | 0.32 | 0.94 | 0.66 | 0.47 | 0.06 | 1.91 | 4.07 | 0.75 | 1.61 | 0.66 | 1.89 | 1.04 | 1.51 | 0.55 | 0.67 | 2.62 | 0.09 | 0.28 | 0.19 | 0 | 2.18 | 0.59 | 4.65
| Analytical results for Cd in OREAS 239b (Certified Value 0.093 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.07 | 1.17 | < 0.5 | 0.7 | 0.819 | 0.08 | 0.13 | 0.29 | 0.3 | 0.06 | 0.06 | 0.07 | 0.07 | 0.12 | 0.14182 | 0.1 | 0.105 | 0.09 | 10.66 | 0.06 | 0.08 | 0.08 | 0.1513 | < 5 | 0.2474 | 0.06654 | 1
| 1 | 2 | 0.09 | 1.28 | < 0.5 | 0.7 | 0.84414 | 0.06 | 0.13 | 0.15 | 0.3 | 0.07 | 0.06 | 0.07 | 0.07 | 0.11 | 0.1469 | 0.1 | 0.149 | 0.08 | 10.66 | 0.14 | 0.07 | 0.06 | 0.1299 | < 5 | 0.2385 | 0.08384 | 30
| 1 | 3 | 0.13 | 1.15 | < 0.5 | 1 | 0.84433 | 0.08 | 0.13 | 0.22 | 0.4 | 0.06 | 0.08 | 0.06 | 0.06 | 0.11 | 0.14797 | 0.1 | 0.119 | 0.09 | 10.97 | 0.08 | 0.08 | 0.07 | 0.1392 | < 5 | 0.2314 | 0.07393 | 16
| 1 | 4 | 0.08 | 1.27 | < 0.5 | 0.6 | 0.839 | 0.07 | 0.14 | 0.44 | 0.4 | 0.08 | 0.07 | 0.07 | 0.06 | 0.12 | 0.11848 | 0.1 | 0.104 | 0.08 | 10.08 | 0.05 | 0.08 | 0.04 | 0.158 | < 5 | 0.2403 | 0.08601 | 0.092562
| 1 | 5 | 0.1 | 1.17 | < 0.5 | 0.8 | 0.84033 | 0.07 | 0.14 | 0.37 | 0.4 | 0.09 | 0.08 | 0.06 | 0.06 | 0.16 | 0.13571 | 0.1 | 0.123 | 0.07 | 10.63 | 0.09 | 0.07 | 0.05 | 0.162 | < 5 | 0.2403 | 0.07112 | 22
| 1 | 6 | 0.13 | 1.3 | < 0.5 | 0.7 | 0.85167 | 0.06 | 0.13 | 0.27 | 0.4 | 0.07 | 0.08 | 0.06 | 0.06 | 0.13 | 0.13407 | 0.1 | 0.133 | 0.09 | 10.02 | 0.1 | 0.07 | 0.06 | 0.1419 | < 5 | 0.2367 | 0.13284
| Mean |  | 0.1 | 1.2233 |  | 0.75 | 0.83975 | 0.07 | 0.13333 | 0.29 | 0.36667 | 0.071667 | 0.071667 | 0.065 | 0.063333 | 0.125 | 0.13749 | 0.1 | 0.12217 | 0.083333 | 10.5033 | 0.086667 | 0.075 | 0.06 | 0.14705 |  | 0.2391 | 0.085713
| Median |  | 0.095 | 1.22 |  | 0.7 | 0.84224 | 0.07 | 0.13 | 0.28 | 0.4 | 0.07 | 0.075 | 0.065 | 0.06 | 0.12 | 0.13877 | 0.1 | 0.121 | 0.085 | 10.645 | 0.085 | 0.075 | 0.06 | 0.1466 |  | 0.2394 | 0.078885
| Std Dev. |  | 0.025298 | 0.066833 |  | 0.13784 | 0.011079 | 0.0089443 | 0.005164 | 0.10373 | 0.05164 | 0.01169 | 0.0098319 | 0.0054772 | 0.005164 | 0.018708 | 0.010893 | 1.5202e-17 | 0.017163 | 0.008165 | 0.37302 | 0.032042 | 0.0054772 | 0.014142 | 0.012196 |  | 0.0052379 | 0.024265
| Rel.Std.Dev. |  | 25.3% | 5.463% |  | 18.38% | 1.319% | 12.78% | 3.873% | 35.77% | 14.08% | 16.31% | 13.72% | 8.427% | 8.154% | 14.97% | 7.923% | 1.52e-14% | 14.05% | 9.798% | 3.551% | 36.97% | 7.303% | 23.57% | 8.294% |  | 2.191% | 28.31%
| PDM3 |  | 8.035% | 1222% |  | 710.3% | 807.2% | -24.38% | 44.05% | 213.3% | 296.1% | -22.57% | -22.57% | -29.78% | -31.58% | 35.04% | 48.54% | 8.035% | 31.98% | -9.971% | 11247% | -6.369% | -18.97% | -35.18% | 58.87% |  | 158.3% | -7.399%
| Z-Score (Absolute) |  | 0.3 | 14.7 | 1.71 | NA | 9.58 | 0.7 | 0.15 | 2.24 | NA | 0.67 | 0.67 | 0.76 | 0.79 | 0.04 | 0.2 | NA | 0 | 0.52 | 138.62 | 0.47 | 0.63 | 0.83 | 0.33 | 31.75 | 1.56 | 0.49
Results from laboratories 1.06, 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 239b (Certified Value 78 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 0
1
| 1 | 1 | 79.94 | 87.0341 | 79.9 | 75.1 | 77.68 | 76 | 77.9 | 83.4 | 83.3 | 79.6 | 80.74 | 76.2587 | 66 | 76.16 | 74.25 | 83.14 | 84.17 | 76.89 | 69.5 | 73.0608 | 73.1389 | [5 empty] | 1
| 1 | 2 | 82.12 | 85.7874 | 80 | 70.4 | 76.06 | 78 | 81.6 | 82.8 | 77.8 | 79.6 | 83.55 | 75.0594 | 66 | 76.24 | 77.06 | 82.44 | 80.11 | 75.01 | 70.4 | 72.0728 | 72.1371 | [5 empty] | 31
| 1 | 3 | 81.97 | 87.8016 | 79.2 | 54.6 | 81.07 | 76 | 80.1 | 81.9 | 82.5 | 76.5 | 79.63 | 78.5481 | 72 | 74.97 | 76.05 | 84.49 | 81 | 77.16 | 70 | 71.983 | 73.7749 | [5 empty] | 16
| 1 | 4 | 82.84 | 87.4167 | 80.9 | 75.6 | 82.53 | 76 | 78 | 86.8 | 76.6 | 76.1 | 81.1 | 79.5453 | 64 | 75.64 | 79.25 | 85.27 | 81.71 | 78.93 | 70.8 | 77.8662 | 74.9123 | [5 empty] | 78.3607
| 1 | 5 | 81.32 | 87.9584 | 80.7 | 81.5 | 83.9 | 75 | 79.5 | 84.1 | 79.3 | 77.3 | 79.64 | 78.1087 | 66 | 76.33 | 77.74 | 85.81 | 78.51 | 77.04 | 71.3 | 77.2823 | 74.3824 | [5 empty] | 23
| 1 | 6 | 82.64 | 87.9397 | 80.4 | 79.2 | 81.43 | 77 | 81.1 | 83.8 | 81.3 | 78.5 | 75.14 | 78.1386 | 67 | 76.6 | 80.09 | 84.19 | 83.07 | 76.78 | 71.7 | 72.2524 | 74.8414
| Mean |  | 81.805 | 87.323 | 80.1833 | 72.7333 | 80.445 | 76.3333 | 79.7 | 83.8 | 80.1333 | 77.9333 | 79.9667 | 77.6098 | 66.8333 | 75.99 | 77.4067 | 84.2233 | 81.4283 | 76.9683 | 70.6167 | 74.0863 | 73.8645
| Median |  | 82.045 | 87.6092 | 80.2 | 75.35 | 81.25 | 76 | 79.8 | 83.6 | 80.3 | 77.9 | 80.19 | 78.1237 | 66 | 76.2 | 77.4 | 84.34 | 81.355 | 76.965 | 70.6 | 72.6566 | 74.0787
| Std Dev. |  | 1.0591 | 0.83228 | 0.61779 | 9.6618 | 2.9839 | 1.0328 | 1.5427 | 1.6649 | 2.6688 | 1.5293 | 2.7655 | 1.6423 | 2.7142 | 0.59025 | 2.1277 | 1.268 | 2.0361 | 1.2463 | 0.81833 | 2.7348 | 1.0806
| Rel.Std.Dev. |  | 1.295% | 0.9531% | 0.7705% | 13.28% | 3.709% | 1.353% | 1.936% | 1.987% | 3.33% | 1.962% | 3.458% | 2.116% | 4.061% | 0.7768% | 2.749% | 1.506% | 2.501% | 1.619% | 1.159% | 3.691% | 1.463%
| PDM3 |  | 4.395% | 11.44% | 2.326% | -7.181% | 2.66% | -2.587% | 1.709% | 6.941% | 2.262% | -0.5454% | 2.049% | -0.9583% | -14.71% | -3.025% | -1.218% | 7.482% | 3.915% | -1.777% | -9.883% | -5.455% | -5.738%
| Z-Score (Absolute) |  | 1.04 | 2.52 | 0.6 | 1.4 | 0.67 | 0.43 | 0.47 | 1.58 | 0.59 | 0 | 0.55 | 0.09 | 2.98 | 0.52 | 0.14 | 1.69 | 0.94 | 0.26 | 1.96 | 1.03 | 1.09
| Analytical results for Co in OREAS 239b (Certified Value 16.9 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 16.5 | 15.0805 | 16 | 17 | 21.29 | 16.6 | 16 | 16 | 16.5 | 16.9 | 16.7 | 17.4 | 17.2 | 16.8 | 16.6494 | 16.5 | 14.18 | 17.2 | 22.63 | 16.8 | 16.9 | 17.3 | 19.095 | 20 | 22.3497 | 16.4551 | 1
| 1 | 2 | 16.7 | 14.7517 | 16 | 17 | 20.8167 | 16.9 | 15 | 16 | 16.4 | 17.2 | 16.9 | 16.4 | 17.3 | 16.9 | 17.0477 | 16.7 | 14.25 | 18.2 | 22.21 | 16.8 | 16 | 17.6 | 19.38 | 21 | 20.6126 | 16.4319 | 32
| 1 | 3 | 17 | 14.4122 | 16 | 17 | 22.03 | 16.5 | 16 | 16 | 16.7 | 16.8 | 16.9 | 16.8 | 17.3 | 16.8 | 15.498 | 16.6 | 15.21 | 16.3 | 22.41 | 16.8 | 16.1 | 17.8 | 19.095 | 21 | 21.0135 | 16.3594 | 16
| 1 | 4 | 16.4 | 14.8686 | 16 | 17 | 22.0367 | 16.8 | 16 | 15 | 16.2 | 16.8 | 18.1 | 16.2 | 17.5 | 17.7 | 16.5569 | 16.8 | 14.09 | 17.1 | 22.59 | 16.9 | 16.4 | 17.4 | 18.905 | 20 | 20.2117 | 16.6425 | 16.8745
| 1 | 5 | 17.3 | 15.2275 | 16 | 17 | 21.6433 | 16.7 | 16 | 15 | 16.6 | 17.4 | 17.5 | 16.2 | 17 | 18.7 | 16.5189 | 16.6 | 14.55 | 17.1 | 22.51 | 17 | 15.6 | 17.7 | 18.715 | 20 | 19.8108 | 16.4343 | 24
| 1 | 6 | 17.1 | 15.205 | 16 | 17 | 21.1467 | 16.6 | 16 | 17 | 16.6 | 17.5 | 16.8 | 16.1 | 17.1 | 18.4 | 16.8394 | 16.7 | 14.41 | 16.4 | 22.55 | 17.2 | 16.4 | 18 | 19.57 | 21 | 20.6126 | 16.6655
| Mean |  | 16.8333 | 14.9242 | 16 | 17 | 21.4939 | 16.6833 | 15.8333 | 15.8333 | 16.5 | 17.1 | 17.15 | 16.5167 | 17.2333 | 17.55 | 16.5184 | 16.65 | 14.4483 | 17.05 | 22.4833 | 16.9167 | 16.2333 | 17.6333 | 19.1267 | 20.5 | 20.7685 | 16.4981
| Median |  | 16.85 | 14.9745 | 16 | 17 | 21.4667 | 16.65 | 16 | 16 | 16.55 | 17.05 | 16.9 | 16.3 | 17.25 | 17.3 | 16.6032 | 16.65 | 14.33 | 17.1 | 22.53 | 16.85 | 16.25 | 17.65 | 19.095 | 20.5 | 20.6126 | 16.4447
| Std Dev. |  | 0.3559 | 0.31311 | 0 | 0 | 0.495 | 0.1472 | 0.40825 | 0.75277 | 0.17889 | 0.30984 | 0.54314 | 0.49967 | 0.17512 | 0.85029 | 0.53724 | 0.10488 | 0.40784 | 0.68337 | 0.15371 | 0.16021 | 0.44121 | 0.2582 | 0.31027 | 0.54772 | 0.8759 | 0.12524
| Rel.Std.Dev. |  | 2.114% | 2.098% | 0% | 0% | 2.303% | 0.8823% | 2.578% | 4.754% | 1.084% | 1.812% | 3.167% | 3.025% | 1.016% | 4.845% | 3.252% | 0.6299% | 2.823% | 4.008% | 0.6837% | 0.947% | 2.718% | 1.464% | 1.622% | 2.672% | 4.217% | 0.7591%
| PDM3 |  | -0.2438% | -11.56% | -5.182% | 0.7438% | 27.38% | -1.133% | -6.17% | -6.17% | -2.219% | 1.336% | 1.633% | -2.12% | 2.127% | 4.003% | -2.11% | -1.33% | -14.38% | 1.04% | 33.24% | 0.25% | -3.8% | 4.497% | 13.35% | 21.49% | 23.08% | -2.23%
| Z-Score (Absolute) |  | 0.21 | 2.87 | NA | NA | 6.28 | 0.42 | NA | NA | 0.67 | 0.16 | 0.23 | 0.65 | 0.35 | 0.79 | 0.65 | 0.46 | 3.53 | 0.09 | 7.66 | 0.09 | 1.04 | 0.91 | 2.98 | 4.9 | 5.27 | 0.68
Results from laboratories 1.05, 1.06, 1.09 and 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 239b (Certified Value 123 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 141 | 140 | 158 | 124.55 | 124 | 176 | 113 | 118 | 129 | 127 | 117 | 122 | 108 | 127.3877 | 134 | 110.5 | 132.7 | 122 | 114 | 115 | 131.5 | 119 | 96.842 | 145.702 |  |  | 1
| 1 | 2 | 145 | 130 | 173 | 128.7133 | 111 | 172 | 107 | 109 | 128 | 125 | 119 | 124 | 104 | 129.2622 | 136 | 107.6 | 133 | 121 | 116 | 112 | 134.3 | 121 | 95.892 | 145.2536 |  |  | 33
| 1 | 3 | 112 | 130 | 153 | 123.48 | 111 | 161 | 112 | 113 | 128 | 130 | 122 | 119 | 107 | 127.9193 | 152 | 129.5 | 131.8 | 117 | 114 | 119 | 130.7 | 121 | 79.3858 | 145.1385 |  |  | 16
| 1 | 4 | 153 | 130 | 167 | 125.775 | 120 | 159 | 106 | 112 | 126 | 132 | 122 | 119 | 111 | 129.1487 | 133 | 110.6 | 135.6 | 118 | 114 | 117 | 132.7 | 124 | 90.2395 | 142.9804 |  |  | 123.2527
| 1 | 5 | 151 | 140 | 166 | 128.6833 | 123 | 145 | 94 | 121 | 125 | 130 | 119 | 119 | 107 | 127.6709 | 134 | 116.9 | 136.9 | 119 | 117 | 118 | 132.4 | 117 | 95.5358 | 146.3946 |  |  | 25
| 1 | 6 | 152 | 140 | 157 | 129.57 | 112 | 160 | 108 | 122 | 122 | 139 | 118 | 122 | 113 | 129.1161 | 137 | 111.1 | 134.7 | 124 | 124 | 118 | 132.9 | 125 | 93.9683 | 147.9972
| Mean |  | 142.3333 | 135 | 162.3333 | 126.7953 | 116.8333 | 162.1667 | 106.6667 | 115.8333 | 126.3333 | 130.5 | 119.5 | 120.8333 | 108.3333 | 128.4175 | 137.6667 | 114.3667 | 134.1167 | 120.1667 | 116.5 | 116.5 | 132.4167 | 121.1667 | 91.9772 | 145.5777
| Median |  | 148 | 135 | 162 | 127.2292 | 116 | 160.5 | 107.5 | 115.5 | 127 | 130 | 119 | 120.5 | 107.5 | 128.5177 | 135 | 110.85 | 133.85 | 120 | 115 | 117.5 | 132.55 | 121 | 94.752 | 145.4778
| Std Dev. |  | 15.5649 | 5.4772 | 7.5277 | 2.5305 | 6.1779 | 10.9438 | 6.802 | 5.2694 | 2.582 | 4.8477 | 2.0736 | 2.137 | 3.2042 | 0.84879 | 7.174 | 8.0129 | 1.9447 | 2.6394 | 3.8859 | 2.5884 | 1.2368 | 2.9944 | 6.5906 | 1.6481
| Rel.Std.Dev. |  | 10.94% | 4.057% | 4.637% | 1.996% | 5.288% | 6.748% | 6.377% | 4.549% | 2.044% | 3.715% | 1.735% | 1.769% | 2.958% | 0.661% | 5.211% | 7.006% | 1.45% | 2.196% | 3.336% | 2.222% | 0.934% | 2.471% | 7.166% | 1.132%
| PDM3 |  | 15.48% | 9.531% | 31.71% | 2.874% | -5.208% | 31.57% | -13.46% | -6.02% | 2.499% | 5.88% | -3.045% | -1.963% | -12.1% | 4.19% | 11.69% | -7.21% | 8.814% | -2.504% | -5.479% | -5.479% | 7.435% | -1.692% | -25.38% | 18.11%
| Z-Score (Absolute) |  | 1.51 | 0.91 | 3.14 | 0.25 | 0.56 | 3.12 | 1.39 | 0.64 | 0.21 | 0.55 | 0.35 | 0.24 | 1.25 | 0.38 | 1.13 | 0.76 | 0.84 | 0.29 | 0.59 | 0.59 | 0.7 | 0.21 | 2.58 | 1.78
| Analytical results for Cs in OREAS 239b (Certified Value 9.37 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 9.39 | 8.1584 | 9.1 | 8.93 | 9.8 | 8.01 | 9.6 | 9 | 9.36 | 9.37 | 9.33 | 9.88 | 9.1 | 10.4 | 8.238 | 9 | 9.55 | 9.27 | 9.2 | 9.2754 | 9.1768 | [5 empty] | 1
| 1 | 2 | 9.41 | 8.0297 | 9.1 | 8.87 | 9.7 | 7.94 | 9.6 | 9.3 | 9.47 | 9.07 | 8.94 | 9.87 | 8.3 | 10.4 | 8.208 | 9 | 9.81 | 8.79 | 9.3 | 9.1755 | 9.277 | [5 empty] | 34
| 1 | 3 | 9.49 | 8.1883 | 9.1 | 7.95 | 9.7 | 8.46 | 9.2 | 9.7 | 9.51 | 9.6 | 9.29 | 9.68 | 9.2 | 10.3 | 8.158 | 9 | 9.57 | 8.8 | 9.3 | 9.3544 | 9.2703 | [5 empty] | 16
| 1 | 4 | 9.5 | 8.2243 | 9.3 | 8.92 | 9.7 | 8.5 | 10.2 | 9.8 | 9.39 | 9.46 | 8.81 | 9.75 | 10.4 | 10.4 | 8.101 | 9 | 9.85 | 8.92 | 9.3 | 9.777 | 9.3397 | [5 empty] | 9.3728
| 1 | 5 | 9.74 | 8.1173 | 9.2 | 9.15 | 9.5 | 8.7 | 9.3 | 9.2 | 9.49 | 9.35 | 8.99 | 9.88 | 9.9 | 10.3 | 8.232 | 9 | 9.89 | 8.6 | 9.3 | 9.4988 | 9.4247 | [5 empty] | 26
| 1 | 6 | 9.51 | 7.9357 | 9.3 | 9.2 | 9.5 | 8.68 | 9.7 | 9.6 | 9.17 | 9.32 | 9.16 | 9.57 | 11.9 | 10.4 | 8.204 | 9 | 9.78 | 9.02 | 9.4 | 9.2727 | 9.425
| Mean |  | 9.5067 | 8.1089 | 9.1833 | 8.8367 | 9.65 | 8.3817 | 9.6 | 9.4333 | 9.3983 | 9.3617 | 9.0867 | 9.7717 | 9.8 | 10.3667 | 8.1902 | 9 | 9.7417 | 8.9 | 9.3 | 9.3923 | 9.3189
| Median |  | 9.495 | 8.1379 | 9.15 | 8.925 | 9.7 | 8.48 | 9.6 | 9.45 | 9.43 | 9.36 | 9.075 | 9.81 | 9.55 | 10.4 | 8.206 | 9 | 9.795 | 8.86 | 9.3 | 9.3149 | 9.3084
| Std Dev. |  | 0.12469 | 0.1081 | 0.098319 | 0.45456 | 0.12247 | 0.32975 | 0.35214 | 0.31411 | 0.12624 | 0.1752 | 0.20656 | 0.12859 | 1.2554 | 0.05164 | 0.052025 | 0 | 0.14566 | 0.2297 | 0.063246 | 0.21714 | 0.097151
| Rel.Std.Dev. |  | 1.312% | 1.333% | 1.071% | 5.144% | 1.269% | 3.934% | 3.668% | 3.33% | 1.343% | 1.872% | 2.273% | 1.316% | 12.81% | 0.4981% | 0.6352% | 0% | 1.495% | 2.581% | 0.6801% | 2.312% | 1.043%
| PDM3 |  | 1.429% | -13.48% | -2.021% | -5.72% | 2.958% | -10.57% | 2.424% | 0.6461% | 0.2727% | -0.1185% | -3.053% | 4.256% | 4.558% | 10.6% | -12.62% | -3.977% | 3.936% | -5.044% | -0.7764% | 0.2083% | -0.5746%
| Z-Score (Absolute) |  | 0.31 | 3.04 | 0.46 | 1.29 | 0.65 | 2.38 | 0.53 | 0.13 | 0.05 | 0.04 | 0.7 | 0.94 | 1.01 | 2.37 | 2.84 | NA | 0.87 | 1.14 | 0.18 | 0.04 | 0.14
Results from laboratory 1.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 239b (Certified Value 30.9 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 32.8 | 28.0916 | 32 | 29 | 30.6167 | 31 | 29 | 30 | 30.6 | 29.7 | 31.1 | 32 | 31.1 | 31.7 | 29.6491 | 28 | 28.7 | 32.4 | 33.19 | 31.3 | 31.5 | 30.6 | 32.78 | 34 | 9 | 30.7354 | 1
| 1 | 2 | 32.4 | 27.328 | 32 | 29 | 30.0405 | 32 | 33 | 33 | 32.9 | 30.6 | 30.1 | 29.9 | 31.8 | 30.1 | 29.5699 | 28.9 | 29.08 | 32.5 | 31.46 | 31.3 | 29.3 | 30.8 | 33.31 | 35 | 8 | 30.5319 | 35
| 1 | 3 | 33 | 28.4872 | 32 | 29 | 31.4233 | 31 | 36 | 28 | 30.3 | 29.8 | 30.6 | 32.2 | 30.9 | 30.3 | 29.9454 | 28.5 | 28.95 | 31.9 | 32.72 | 31.6 | 30 | 31.4 | 32.42 | 34 | 9 | 33.7347 | 16
| 1 | 4 | 32.3 | 28.1836 | 32 | 30 | 31.26 | 31 | 33 | 26 | 31.6 | 30.1 | 32 | 30.9 | 30.8 | 30.1 | 31.9018 | 28.2 | 27.96 | 31.7 | 31.34 | 31.9 | 30.1 | 32 | 33.08 | 33 | 7 | 30.4675 | 30.9206
| 1 | 5 | 33.4 | 27.788 | 32 | 29 | 30.8433 | 32 | 30 | 29 | 31 | 30 | 31.1 | 31.4 | 30.5 | 30.2 | 32.0675 | 28.1 | 29.66 | 32.1 | 31.51 | 32 | 29 | 31.5 | 33.27 | 33 | 8 | 31.0074 | 27
| 1 | 6 | 33.8 | 28.4964 | 30 | 36 | 30.9933 | 31 | 29 | 30 | 33.3 | 31 | 30.9 | 31 | 30 | 30.1 | 31.0258 | 28.3 | 29.03 | 33.2 | 32.53 | 32.7 | 31 | 31.8 | 33.05 | 34 | 6 | 30.7633
| Mean |  | 32.95 | 28.0625 | 31.6667 | 30.3333 | 30.8629 | 31.3333 | 31.6667 | 29.3333 | 31.6167 | 30.2 | 30.9667 | 31.2333 | 30.85 | 30.4167 | 30.6933 | 28.3333 | 28.8967 | 32.3 | 32.125 | 31.8 | 30.15 | 31.35 | 32.985 | 33.8333 | 7.8333 | 31.2067
| Median |  | 32.9 | 28.1376 | 32 | 29 | 30.9183 | 31 | 31.5 | 29.5 | 31.3 | 30.05 | 31 | 31.2 | 30.85 | 30.15 | 30.4856 | 28.25 | 28.99 | 32.25 | 32.02 | 31.75 | 30.05 | 31.45 | 33.065 | 34 | 8 | 30.7493
| Std Dev. |  | 0.57879 | 0.44699 | 0.8165 | 2.8048 | 0.49521 | 0.5164 | 2.8048 | 2.3381 | 1.2352 | 0.502 | 0.6314 | 0.83586 | 0.60249 | 0.63377 | 1.1288 | 0.3266 | 0.5573 | 0.53292 | 0.786 | 0.52915 | 0.96073 | 0.55045 | 0.33507 | 0.75277 | 1.169 | 1.2531
| Rel.Std.Dev. |  | 1.757% | 1.593% | 2.578% | 9.246% | 1.605% | 1.648% | 8.857% | 7.971% | 3.907% | 1.662% | 2.039% | 2.676% | 1.953% | 2.084% | 3.678% | 1.153% | 1.929% | 1.65% | 2.447% | 1.664% | 3.186% | 1.756% | 1.016% | 2.225% | 14.92% | 4.015%
| PDM3 |  | 6.563% | -9.243% | 2.413% | -1.899% | -0.1867% | 1.335% | 2.413% | -5.133% | 2.251% | -2.33% | 0.149% | 1.011% | -0.2283% | -1.63% | -0.7352% | -8.367% | -6.546% | 4.461% | 3.895% | 2.844% | -2.492% | 1.389% | 6.676% | 9.42% | -74.67% | 0.9252%
| Z-Score (Absolute) |  | 1.71 | 2.78 | 0.53 | 0.69 | 0.21 | 0.23 | 0.53 | 1.61 | 0.49 | 0.82 | 0.11 | 0.13 | 0.22 | 0.62 | 0.36 | 2.53 | 2.01 | 1.12 | 0.95 | 0.66 | 0.86 | 0.24 | 1.75 | 2.53 | 21.38 | 0.11
| Analytical results for Dy in OREAS 239b (Certified Value 3.47 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 2.9419 | 3.75 | 3.6 | 2.7 | 4.14 | 2.889 | 3.91 | 2.9 | 2.65 | 4.561 | [16 empty] | 1
| 1 | 2 | 2.9612 | 3.8 | 3.7 | 2.7 | 4.24 | 2.927 | 4.18 | 2.9 | 2.7 | 4.6127 | [16 empty] | 9
| 1 | 3 | 2.9413 | 3.85 | 3.2 | 2.8 | 4.26 | 2.79 | 3.98 | 3 | 2.66 | 4.5853 | [16 empty] | 16
| 1 | 4 | 3.0052 | 3.75 | 3.6 | 3.2 | 4.2 | 2.983 | 3.88 | 3 | 2.67 | 4.7572 | [16 empty] | 3.4742
| 1 | 5 | 2.868 | 3.8 | 3.9 | 3.3 | 4.1 | 2.842 | 4.96 | 3 | 2.71 | 4.6049 | [16 empty] | 28
| 1 | 6 | 3.0037 | 3.75 | 3.8 | 3.8 | 4.14 | 3.007 | 3.92 | 2.9 | 2.75 | 4.5782
| Mean |  | 2.9536 | 3.7833 | 3.6333 | 3.0833 | 4.18 | 2.9063 | 4.1383 | 2.95 | 2.69 | 4.6165
| Median |  | 2.9516 | 3.775 | 3.65 | 3 | 4.17 | 2.908 | 3.95 | 2.95 | 2.685 | 4.5951
| Std Dev. |  | 0.050677 | 0.040825 | 0.24221 | 0.43551 | 0.063246 | 0.082959 | 0.41677 | 0.054772 | 0.037417 | 0.071376
| Rel.Std.Dev. |  | 1.716% | 1.079% | 6.666% | 14.12% | 1.513% | 2.854% | 10.07% | 1.857% | 1.391% | 1.546%
| PDM3 |  | -14.99% | 8.897% | 4.58% | -11.25% | 20.31% | -16.35% | 19.12% | -15.09% | -22.57% | 32.88%
| Z-Score (Absolute) |  | 0.62 | 0.65 | 0.42 | 0.42 | 1.26 | 0.7 | 1.2 | 0.63 | 1.03 | 1.93
| Analytical results for Er in OREAS 239b (Certified Value 1.76 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 1.5806 | 1.95 | 1.9 | 1.6 | 2.28 | 1.604 | 2.07 | 1.5 | 1.3348 | 2.6484 | [16 empty] | 1
| 1 | 2 | 1.6162 | 1.95 | 1.9 | 1.5 | 2.3 | 1.52 | 2.08 | 1.5 | 1.3413 | 2.6703 | [16 empty] | 10
| 1 | 3 | 1.5254 | 1.95 | 1.7 | 1.5 | 2.28 | 1.452 | 2.14 | 1.5 | 1.4241 | 2.6465 | [16 empty] | 16
| 1 | 4 | 1.5386 | 2 | 2 | 1.9 | 2.26 | 1.499 | 2.05 | 1.6 | 1.423 | 2.6763 | [16 empty] | 1.7632
| 1 | 5 | 1.5252 | 2 | 1.9 | 1.7 | 2.23 | 1.474 | 2.3 | 1.5 | 1.43 | 2.6315 | [16 empty] | 29
| 1 | 6 | 1.5205 | 2 | 2 | 2 | 2.22 | 1.509 | 2.08 | 1.5 | 1.3478 | 2.6456
| Mean |  | 1.5511 | 1.975 | 1.9 | 1.7 | 2.2617 | 1.5097 | 2.12 | 1.5167 | 1.3835 | 2.6531
| Median |  | 1.532 | 1.975 | 1.9 | 1.65 | 2.27 | 1.504 | 2.08 | 1.5 | 1.3854 | 2.6475
| Std Dev. |  | 0.038825 | 0.027386 | 0.10954 | 0.20976 | 0.031252 | 0.052378 | 0.093167 | 0.040825 | 0.046471 | 0.016858
| Rel.Std.Dev. |  | 2.503% | 1.387% | 5.766% | 12.34% | 1.382% | 3.47% | 4.395% | 2.692% | 3.359% | 0.6354%
| PDM3 |  | -12.03% | 12.01% | 7.76% | -3.583% | 28.27% | -14.38% | 20.24% | -13.98% | -21.53% | 50.47%
| Z-Score (Absolute) |  | 0.59 | 0.41 | 0.24 | 0.24 | 1.09 | 0.68 | 0.75 | 0.67 | 0.98 | 2.01
| Analytical results for Eu in OREAS 239b (Certified Value 1.19 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 1.5645 | 1.2 | 1.13 | 1.1 | 1.31 | 1.27 | 1.27 | 1.1 | 1.0833 | 1.1116 | [16 empty] | 1
| 1 | 2 | 1.5365 | 1.2 | 1.13 | 1 | 1.29 | 1.264 | 1.26 | 1.1 | 1.06 | 1.1522 | [16 empty] | 11
| 1 | 3 | 1.5224 | 1.2 | 0.95 | 1.1 | 1.31 | 1.242 | 1.28 | 1.1 | 1.0794 | 1.1522 | [16 empty] | 16
| 1 | 4 | 1.4552 | 1.2 | 1.11 | 1.2 | 1.28 | 1.244 | 1.3 | 1.2 | 1.1309 | 1.1473 | [16 empty] | 1.1937
| 1 | 5 | 1.5085 | 1.25 | 1.2 | 1.2 | 1.3 | 1.25 | 1.37 | 1.1 | 1.0907 | 1.1554 | [16 empty] | 30
| 1 | 6 | 1.5853 | 1.2 | 1.21 | 1.5 | 1.3 | 1.269 | 1.29 | 1.1 | 1.0837 | 1.1366
| Mean |  | 1.5287 | 1.2083 | 1.1217 | 1.1833 | 1.2983 | 1.2565 | 1.295 | 1.1167 | 1.088 | 1.1425
| Median |  | 1.5294 | 1.2 | 1.13 | 1.15 | 1.3 | 1.257 | 1.285 | 1.1 | 1.0835 | 1.1498
| Std Dev. |  | 0.045587 | 0.020412 | 0.09347 | 0.17224 | 0.01169 | 0.012677 | 0.03937 | 0.040825 | 0.023436 | 0.01655
| Rel.Std.Dev. |  | 2.982% | 1.689% | 8.333% | 14.56% | 0.9004% | 1.009% | 3.04% | 3.656% | 2.154% | 1.449%
| PDM3 |  | 28.07% | 1.225% | -6.036% | -0.8696% | 8.764% | 5.26% | 8.485% | -6.454% | -8.856% | -4.287%
| Z-Score (Absolute) |  | 2.49 | 0 | 0.67 | 0.19 | 0.7 | 0.37 | 0.67 | NA | 0.94 | 0.51
| Analytical results for Fe in OREAS 239b (Certified Value 4.07 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 4.03 | 3.9838 | 4.07 | 4.28 | 4.325 | 4.29 | 3.99 | 4.12 | 3.88 | 4.1 | 4.05 | 4.05 | 3.85 | 4.0912 | 3.9846 | 3.903 | 3.97 | 4.05 | 4.07 | 4.16 | 3.98 | 4.0316 | 4.0799 | 4.1479 |  |  | 1
| 1 | 2 | 4.13 | 3.9423 | 4.1 | 4.25 | 4.24 | 4.34 | 3.96 | 4.14 | 3.84 | 4.09 | 3.94 | 3.89 | 3.96 | 4.0545 | 4.0654 | 3.98 | 4.01 | 4.07 | 4.11 | 4.17 | 4 | 4.1539 | 4.2738 | 4.0959 |  |  | #N/A
| 1 | 3 | 4.12 | 3.9499 | 4.06 | 4.18 | 4.295 | 4.33 | 4.19 | 4.02 | 3.81 | 4.04 | 4.19 | 3.99 | 3.84 | 4.1188 | 4.1074 | 4.644 | 3.97 | 4.15 | 4.09 | 4.11 | 4.01 | 3.9486 | 4.2236 | 4.0586 |  |  | 16
| 1 | 4 | 4.03 | 3.9019 | 4.07 | 4.03 | 4.3 | 4.29 | 4.21 | 3.95 | 3.82 | 4 | 4.2 | 3.91 | 3.86 | 4.0861 | 4.0864 | 3.959 | 4.02 | 4.11 | 4.1 | 4.12 | 3.99 | 4.0546 | 4.1612 | 4.1042 |  |  | 4.0747
| 1 | 5 | 4.19 | 3.9373 | 4.11 | 4.36 | 4.31 | 4.34 | 4.26 | 4.17 | 3.85 | 4 | 4.05 | 3.94 | 3.82 | 4.0616 | 3.9902 | 3.994 | 4.06 | 4.13 | 4.1 | 4.12 | 3.98 | 4.0794 | 4.1822 | 4.1171 |  |  | 31
| 1 | 6 | 4.13 | 3.9931 | 4.07 | 4.41 | 4.295 | 4.36 | 4.22 | 4.11 | 3.88 | 3.99 | 4.35 | 3.94 | 3.84 | 3.9688 | 4.12 | 4.043 | 4.03 | 4.08 | 4.12 | 4.15 | 4.05 | 4.0817 | 4.2801 | 4.0184
| Mean |  | 4.105 | 3.9514 | 4.08 | 4.2517 | 4.2942 | 4.325 | 4.1383 | 4.085 | 3.8467 | 4.0367 | 4.13 | 3.9533 | 3.8617 | 4.0635 | 4.059 | 4.0872 | 4.01 | 4.0983 | 4.0983 | 4.1383 | 4.0017 | 4.0583 | 4.2001 | 4.0903
| Median |  | 4.125 | 3.9461 | 4.07 | 4.265 | 4.2975 | 4.335 | 4.2 | 4.115 | 3.845 | 4.02 | 4.12 | 3.94 | 3.845 | 4.0739 | 4.0759 | 3.987 | 4.015 | 4.095 | 4.1 | 4.135 | 3.995 | 4.067 | 4.2029 | 4.1
| Std Dev. |  | 0.063166 | 0.033221 | 0.02 | 0.13556 | 0.028882 | 0.02881 | 0.12891 | 0.083126 | 0.029439 | 0.048442 | 0.14546 | 0.058195 | 0.049967 | 0.051747 | 0.058526 | 0.27659 | 0.035214 | 0.038166 | 0.017224 | 0.024833 | 0.026394 | 0.067665 | 0.075717 | 0.045684
| Rel.Std.Dev. |  | 1.539% | 0.8407% | 0.4902% | 3.188% | 0.6726% | 0.6661% | 3.115% | 2.035% | 0.7653% | 1.2% | 3.522% | 1.472% | 1.294% | 1.273% | 1.442% | 6.767% | 0.8781% | 0.9313% | 0.4203% | 0.6001% | 0.6596% | 1.667% | 1.803% | 1.117%
| PDM3 |  | 0.7445% | -3.026% | 0.1309% | 4.344% | 5.387% | 6.144% | 1.563% | 0.2536% | -5.596% | -0.9326% | 1.358% | -2.978% | -5.227% | -0.274% | -0.3844% | 0.3068% | -1.587% | 0.5808% | 0.5808% | 1.563% | -1.792% | -0.4016% | 3.079% | 0.3849%
| Z-Score (Absolute) |  | 0.25 | 1.79 | 0.08 | 2.2 | 2.76 | 3.17 | 0.69 | 0.01 | 3.18 | 0.66 | 0.58 | 1.76 | 2.98 | 0.3 | 0.36 | 0.01 | 1.01 | 0.16 | 0.16 | 0.69 | 1.12 | 0.37 | 1.51 | 0.06
| Analytical results for Ga in OREAS 239b (Certified Value 19.3 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 19.47 | 22.6765 | 18.4 | 21.2 | 21.6867 | 19.6 | 19 | 20.4 | 21 | 18.05 | 20.2 | 19.85 | 19.2 | 18.08 | 18.5632 | 16.8 | 19.7 | 19.1 | 18.84 | 17.84 | 19.5938 | 23 | 19.2 |  |  |  | 1
| 1 | 2 | 19.58 | 23.674 | 19 | 21.5 | 21.7567 | 19.4 | 18.98 | 20.7 | 22 | 18.5 | 20.1 | 18.4 | 19.5 | 17.58 | 18.5658 | 16.8 | 20.9 | 18.83 | 17.66 | 17.84 | 19.8418 | 25 | 19.1139 |  |  |  | 39
| 1 | 3 | 19.37 | 23.7025 | 19.2 | 20.8 | 22.6933 | 19.6 | 19.84 | 22.1 | 21 | 18.35 | 18.8 | 19.95 | 19 | 19.16 | 18.3922 | 16.8 | 18.6 | 19.04 | 17.9 | 18.14 | 19.6225 | 23 | 19.0783 |  |  |  | 16
| 1 | 4 | 19.53 | 23.4745 | 18.8 | 21.6 | 22.7433 | 20.1 | 19.87 | 20.3 | 20 | 18.2 | 21.2 | 18.6 | 19.65 | 18.96 | 18.4545 | 17.2 | 19.1 | 19.58 | 19.37 | 18.01 | 20.2821 | 24 | 19.0109 |  |  |  | 19.2786
| 1 | 5 | 19.65 | 24.149 | 18.8 | 22.4 | 22.0667 | 19.2 | 19.98 | 20.8 | 20 | 18.6 | 20.5 | 19.15 | 18.85 | 19.35 | 18.5429 | 16.7 | 19.5 | 19.81 | 17.1 | 18 | 19.904 | 27 | 19.016 |  |  |  | 32
| 1 | 6 | 20.11 | 22.9045 | 18.6 | 22.6 | 22.1767 | 20 | 20.28 | 20.3 | 20 | 18.6 | 20.4 | 19.25 | 18.65 | 20.11 | 18.3026 | 16.8 | 19 | 18.78 | 19.42 | 17.95 | 19.4929 | 27 | 19.0983
| Mean |  | 19.6183 | 23.4302 | 18.8 | 21.6833 | 22.1872 | 19.65 | 19.6583 | 20.7667 | 20.6667 | 18.3833 | 20.2 | 19.2 | 19.1417 | 18.8733 | 18.4702 | 16.85 | 19.4667 | 19.19 | 18.3817 | 17.9633 | 19.7895 | 24.8333 | 19.0862
| Median |  | 19.555 | 23.5742 | 18.8 | 21.55 | 22.1217 | 19.6 | 19.855 | 20.55 | 20.5 | 18.425 | 20.3 | 19.2 | 19.1 | 19.06 | 18.4987 | 16.8 | 19.3 | 19.07 | 18.37 | 17.975 | 19.7321 | 24.5 | 19.0883
| Std Dev. |  | 0.25911 | 0.54711 | 0.28284 | 0.69402 | 0.45077 | 0.34496 | 0.54061 | 0.68605 | 0.8165 | 0.22509 | 0.7874 | 0.63087 | 0.38395 | 0.91086 | 0.10724 | 0.17607 | 0.80166 | 0.41588 | 0.96545 | 0.11431 | 0.28731 | 1.8348 | 0.070028
| Rel.Std.Dev. |  | 1.321% | 2.335% | 1.504% | 3.201% | 2.032% | 1.756% | 2.75% | 3.304% | 3.951% | 1.224% | 3.898% | 3.286% | 2.006% | 4.826% | 0.5806% | 1.045% | 4.118% | 2.167% | 5.252% | 0.6363% | 1.452% | 7.389% | 0.3669%
| PDM3 |  | 1.762% | 21.53% | -2.483% | 12.47% | 15.09% | 1.926% | 1.97% | 7.719% | 7.2% | -4.644% | 4.779% | -0.4079% | -0.7104% | -2.102% | -4.193% | -12.6% | 0.9754% | -0.4597% | -4.653% | -6.823% | 2.65% | 28.81% | -0.998%
| Z-Score (Absolute) |  | 0.14 | 3.64 | 0.61 | 2.04 | 2.5 | 0.17 | 0.18 | 1.2 | 1.1 | 1 | 0.67 | 0.25 | 0.3 | 0.55 | 0.92 | 2.41 | 0 | 0.25 | 1 | 1.38 | 0.3 | 4.93 | 0.35
| Analytical results for Gd in OREAS 239b (Certified Value 4.96 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.74 | 5.2 | 4.8 | 4.5 | 5 | 5.105 | 5.16 | 4.4 | 4.6813 | 0.76113 | [16 empty] | 1
| 1 | 2 | 6.7351 | 5.4 | 4.7 | 4.1 | 5.1 | 5.124 | 5.23 | 4.4 | 4.6423 | 0.74853 | [16 empty] | 12
| 1 | 3 | 6.7161 | 5.2 | 4 | 4.4 | 5 | 4.945 | 5.32 | 4.4 | 4.807 | 0.77521 | [16 empty] | 16
| 1 | 4 | 6.7603 | 5.4 | 5 | 5.4 | 5 | 5.013 | 5.59 | 4.6 | 5.0069 | 0.782 | [16 empty] | 4.9552
| 1 | 5 | 6.814 | 5.2 | 5.2 | 5 | 5 | 5.109 | 5.3 | 4.4 | 4.7654 | 0.77054 | [16 empty] | 33
| 1 | 6 | 6.7214 | 5.2 | 4.9 | 5.9 | 5 | 5.072 | 5.32 | 4.4 | 4.8624 | 0.76543
| Mean |  | 6.7478 | 5.2667 | 4.7667 | 4.8833 | 5.0167 | 5.0613 | 5.32 | 4.4333 | 4.7942 | 0.76714
| Median |  | 6.7376 | 5.2 | 4.85 | 4.75 | 5 | 5.0885 | 5.31 | 4.4 | 4.7862 | 0.76799
| Std Dev. |  | 0.035965 | 0.10328 | 0.41312 | 0.67946 | 0.040825 | 0.069411 | 0.14629 | 0.08165 | 0.13169 | 0.011687
| Rel.Std.Dev. |  | 0.533% | 1.961% | 8.667% | 13.91% | 0.8138% | 1.371% | 2.75% | 1.842% | 2.747% | 1.523%
| PDM3 |  | 36.18% | 6.286% | -3.805% | -1.45% | 1.241% | 2.142% | 7.362% | -10.53% | -3.249% | -84.52%
| Z-Score (Absolute) |  | 4.85 | 0.85 | 0.49 | 0.18 | 0.18 | 0.3 | 1 | 1.39 | 0.42 | 11.28
| Analytical results for Ge in OREAS 239b (Indicative Value 0.2 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 1.6 | 0.3 | 1.2 | 0.18 | 1.4 | 4 | 0.14 | 0.25 | 0.08 | 0.19 | 0.38 | < 0.05 | 0.13 | 0.1885 | [12 empty] | 1
| 1 | 2 | 1.7 | 0.3 | 1.3 | 0.18 | 1.5 | 4 | 0.17 | 0.24 | 0.09 | 0.18 | 0.34 | < 0.05 | 0.11 | 0.27 | [12 empty] | 5
| 1 | 3 | 1.5 | 0.4 | 1.2 | 0.19 | 1.3 | 3.9 | 0.14 | 0.12 | 0.09 | 0.17 | 0.33 | < 0.05 | 0.11 | 0.1213 | [12 empty] | 16
| 1 | 4 | 1.7 | 0.3 | 1.3 | 0.2 | 1.9 | 4.1 | 0.15 | 0.27 | 0.09 | 0.18 | 0.41 | < 0.05 | 0.12 | 0.0959 | [12 empty] | 0.19958
| 1 | 5 | 1.8 | 0.2 | 1.2 | 0.21 | 1.7 | 3.8 | 0.15 | 0.28 | 0.08 | 0.19 | 0.43 | < 0.05 | 0.11 | 0.1486 | [12 empty] | 11
| 1 | 6 | 1.8 | 0.2 | 1.2 | 0.21 | 1.6 | 3.9 | 0.14 | 0.28 | 0.09 | 0.2 | 0.58 | < 0.05 | 0.12 | 0.1416
| Mean |  | 1.6833 | 0.28333 | 1.2333 | 0.195 | 1.5667 | 3.95 | 0.14833 | 0.24 | 0.086667 | 0.185 | 0.41167 |  | 0.11667 | 0.16098
| Median |  | 1.7 | 0.3 | 1.2 | 0.195 | 1.55 | 3.95 | 0.145 | 0.26 | 0.09 | 0.185 | 0.395 |  | 0.115 | 0.1451
| Std Dev. |  | 0.1169 | 0.075277 | 0.05164 | 0.013784 | 0.21602 | 0.10488 | 0.01169 | 0.060992 | 0.005164 | 0.010488 | 0.091086 |  | 0.008165 | 0.061614
| Rel.Std.Dev. |  | 6.945% | 26.57% | 4.187% | 7.069% | 13.79% | 2.655% | 7.881% | 25.41% | 5.958% | 5.669% | 22.13% |  | 6.999% | 38.27%
| PDM3 |  | 743.5% | 41.97% | 518% | -2.293% | 685% | 1879% | -25.68% | 20.26% | -56.57% | -7.303% | 106.3% |  | -41.54% | -19.34%
| Z-Score (Absolute) |  | 8.53 | 0.38 | 5.91 | 0.13 | 7.85 | 21.73 | 0.4 | 0.13 | 0.76 | 0.19 | 1.13 | 1.12 | 0.59 | 0.33
| Analytical results for Hf in OREAS 239b (Certified Value 3.81 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 3.84 | 3.77 | 6.6 | 4.4 | 4 | 3.8 | 4.1 | 3.8 | 3.8 | 3.9 | 3.9 | 3.8 | 3.89 | 4.3 | 3.56 | 4.06 | 3.9 | 3.28 | 3.0849 | 2.3222 | 6.1122 | [5 empty] | 1
| 1 | 2 | 3.86 | 3.7551 | 7.2 | 4.5 | 3.8 | 3.4 | 3.6 | 3.7 | 3.9 | 3.7 | 3.8 | 3.7 | 3.52 | 4.3 | 3.71 | 4.23 | 3.7 | 3.29 | 3.0312 | 2.3222 | 6.0823 | [5 empty] | 25
| 1 | 3 | 4.08 | 3.6809 | 6.6 | 4.5 | 3.8 | 3.4 | 3.6 | 3.8 | 3.8 | 3.8 | 3.9 | 3.6 | 3.86 | 4.3 | 3.62 | 4.11 | 3.7 | 3.29 | 3.0789 | 2.106 | 6.0594 | [5 empty] | 16
| 1 | 4 | 4.02 | 3.8195 | 6.6 | 4.6 | 4 | 3.7 | 3.6 | 3.9 | 3.8 | 4 | 3.7 | 3.6 | 4 | 4.2 | 3.61 | 4.13 | 3.8 | 3.32 | 3.2105 | 2.289 | 5.9947 | [5 empty] | 3.8083
| 1 | 5 | 4.11 | 3.6868 | 7 | 4.6 | 3.8 | 3.6 | 3.7 | 3.8 | 3.9 | 3.7 | 3.7 | 3.6 | 4.01 | 4.2 | 3.5 | 4.29 | 3.7 | 3.24 | 3.0954 | 2.3638 | 6.1717 | [5 empty] | 34
| 1 | 6 | 4.02 | 3.8822 | 6.8 | 4.6 | 3.9 | 3.6 | 4.2 | 3.8 | 3.5 | 3.9 | 3.9 | 3.6 | 4.13 | 4.2 | 3.58 | 4.16 | 3.9 | 3.24 | 3.0866 | 2.2391 | 6.114
| Mean |  | 3.9883 | 3.7657 | 6.8 | 4.5333 | 3.8833 | 3.5833 | 3.8 | 3.8 | 3.7833 | 3.8333 | 3.8167 | 3.65 | 3.9017 | 4.25 | 3.5967 | 4.1633 | 3.7833 | 3.2767 | 3.0979 | 2.2737 | 6.0891
| Median |  | 4.02 | 3.7626 | 6.7 | 4.55 | 3.85 | 3.6 | 3.65 | 3.8 | 3.8 | 3.85 | 3.85 | 3.6 | 3.945 | 4.25 | 3.595 | 4.145 | 3.75 | 3.285 | 3.0857 | 2.3056 | 6.0973
| Std Dev. |  | 0.11286 | 0.077447 | 0.25298 | 0.08165 | 0.098319 | 0.16021 | 0.27568 | 0.063246 | 0.1472 | 0.12111 | 0.098319 | 0.083666 | 0.21028 | 0.054772 | 0.070048 | 0.083825 | 0.098319 | 0.031411 | 0.059652 | 0.092062 | 0.059695
| Rel.Std.Dev. |  | 2.83% | 2.057% | 3.72% | 1.801% | 2.532% | 4.471% | 7.255% | 1.664% | 3.891% | 3.159% | 2.576% | 2.292% | 5.389% | 1.289% | 1.948% | 2.013% | 2.599% | 0.9586% | 1.926% | 4.049% | 0.9804%
| PDM3 |  | 4.728% | -1.117% | 78.56% | 19.04% | 1.971% | -5.907% | -0.2173% | -0.2173% | -0.655% | 0.658% | 0.2203% | -4.156% | 2.452% | 11.6% | -5.557% | 9.323% | -0.655% | -13.96% | -18.65% | -40.3% | 59.89%
| Z-Score (Absolute) |  | 0.67 | 0.12 | 10.74 | 2.63 | 0.3 | 0.78 | 0 | 0 | 0.06 | 0.12 | 0.06 | 0.54 | 0.36 | 1.61 | 0.73 | 1.3 | 0.06 | 1.87 | 2.51 | 5.46 | 8.2
| Analytical results for Hg in OREAS 239b (Indicative Value 0.032 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 1.09 | 1.30 | [24 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [24 empty] | 3
3
| 1 | 1 | < 2 | 0.0265 | [24 empty] | 1
| 1 | 2 | < 2 | 0.0325 | [24 empty] | 6
| 1 | 3 | < 2 | 0.0359 | [24 empty] | 16
| 1 | 4 | < 2 | 0.0296 | [24 empty] | 0.032417
| 1 | 5 | < 2 | 0.0321 | [24 empty] | 12
| 1 | 6 | < 2 | 0.0379
| Mean |  |  | 0.032417
| Median |  |  | 0.0323
| Std Dev. |  |  | 0.0041291
| Rel.Std.Dev. |  |  | 12.74%
| PDM3 |  |  | -8.882e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 239b (Certified Value 0.68 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.55936 | 0.68 | 0.7 | 0.5 | 0.74 | 0.588 | 0.74 | 0.5 | 0.477 | 0.91467 | [16 empty] | 1
| 1 | 2 | 0.55019 | 0.7 | 0.7 | 0.5 | 0.76 | 0.54 | 1.01 | 0.5 | 0.4842 | 0.90995 | [16 empty] | 14
| 1 | 3 | 0.53829 | 0.68 | 0.6 | 0.5 | 0.74 | 0.51 | 0.75 | 0.5 | 0.4798 | 0.92518 | [16 empty] | 16
| 1 | 4 | 0.54749 | 0.68 | 0.7 | 0.6 | 0.76 | 0.553 | 0.76 | 0.5 | 0.5039 | 0.91979 | [16 empty] | 0.67848
| 1 | 5 | 0.54872 | 0.68 | 0.7 | 0.6 | 0.72 | 0.524 | 0.77 | 0.5 | 0.4902 | 0.9181 | [16 empty] | 35
| 1 | 6 | 0.54569 | 0.7 | 0.7 | 0.8 | 0.72 | 0.528 | 0.96 | 0.5 | 0.4886 | 0.90208
| Mean |  | 0.54829 | 0.68667 | 0.68333 | 0.58333 | 0.74 | 0.5405 | 0.83167 | 0.5 | 0.48728 | 0.91496
| Median |  | 0.54811 | 0.68 | 0.7 | 0.55 | 0.74 | 0.534 | 0.765 | 0.5 | 0.4864 | 0.91639
| Std Dev. |  | 0.0068358 | 0.010328 | 0.040825 | 0.1169 | 0.017889 | 0.027465 | 0.12024 | 0 | 0.0095667 | 0.0081082
| Rel.Std.Dev. |  | 1.247% | 1.504% | 5.974% | 20.04% | 2.417% | 5.081% | 14.46% | 0% | 1.963% | 0.8862%
| PDM3 |  | -19.19% | 1.206% | 0.7151% | -14.02% | 9.067% | -20.34% | 22.58% | -26.31% | -28.18% | 34.85%
| Z-Score (Absolute) |  | 0.64 | 0 | NA | NA | 0.25 | 0.68 | 0.67 | NA | 0.93 | 1.06
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 239b (Certified Value 0.068 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.046 | 0.1 | 0.06 | 0.07 | 0.059 | 0.068 | 0.068 | 0.07 | 0.07 | 0.072 | 0.062 | 0.06 | 0.0751 | 0.06685 | [7 empty] | 1
| 1 | 2 | 0.06 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.045 | 0.06 | 0.06 | 0.068 | 0.061 | 0.066 | 0.068 | 0.07 | 0.07 | 0.08 | 0.058 | 0.06 | 0.078 | 0.06528 | [7 empty] | 27
| 1 | 3 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.045 | 0.06 | 0.08 | 0.068 | 0.076 | 0.067 | 0.067 | 0.07 | 0.07 | 0.083 | 0.059 | 0.06 | 0.0757 | 0.0664 | [7 empty] | 16
| 1 | 4 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.07 | 0.046 | 0.08 | 0.07 | 0.069 | 0.073 | 0.062 | 0.066 | 0.07 | 0.07 | 0.068 | 0.064 | 0.06 | 0.0808 | 0.06809 | [7 empty] | 0.068141
| 1 | 5 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.06 | 0.047 | 0.07 | 0.07 | 0.067 | 0.066 | 0.07 | 0.067 | 0.08 | 0.07 | 0.08 | 0.056 | 0.06 | 0.0829 | 0.06725 | [7 empty] | 36
| 1 | 6 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.07 | 0.044 | 0.08 | 0.08 | 0.064 | 0.067 | 0.061 | 0.067 | 0.07 | 0.07 | 0.074 | 0.064 | 0.06 | 0.0744 | 0.0669
| Mean |  | 0.068333 |  |  |  | 0.063333 | 0.0455 | 0.075 | 0.07 | 0.067667 | 0.067 | 0.065667 | 0.067167 | 0.071667 | 0.07 | 0.076167 | 0.0605 | 0.06 | 0.077817 | 0.066795
| Median |  | 0.07 |  |  |  | 0.06 | 0.0455 | 0.075 | 0.07 | 0.068 | 0.0665 | 0.0665 | 0.067 | 0.07 | 0.07 | 0.077 | 0.0605 | 0.06 | 0.07685 | 0.066875
| Std Dev. |  | 0.0040825 |  |  |  | 0.005164 | 0.0010488 | 0.015166 | 0.0089443 | 0.0020656 | 0.006603 | 0.0035024 | 0.00075277 | 0.0040825 | 0 | 0.0057417 | 0.0033317 | 0 | 0.0034149 | 0.00093267
| Rel.Std.Dev. |  | 5.974% |  |  |  | 8.154% | 2.305% | 20.22% | 12.78% | 3.053% | 9.855% | 5.334% | 1.121% | 5.696% | 0% | 7.538% | 5.507% | 0% | 4.388% | 1.396%
| PDM3 |  | 0.2826% |  |  |  | -7.055% | -33.23% | 10.07% | 2.729% | -0.6958% | -1.674% | -3.631% | -1.43% | 5.174% | 2.729% | 11.78% | -11.21% | -11.95% | 14.2% | -1.975%
| Z-Score (Absolute) |  | 0.17 | 6.32 | 2.57 | 139.78 | 0.57 | 3.25 | 1.17 | 0.42 | 0.07 | 0.02 | 0.22 | 0 | 0.67 | 0.42 | 1.35 | 1 | 1.07 | 1.6 | 0.06
| Analytical results for Ir in OREAS 239b (Indicative Value 0 ppm) | [27 empty] | Indicative
Ir | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | < 0.0004 | [25 empty] | 1
| 1 | 2 | 0.0004 | [25 empty] | 3
| 1 | 3 | 0.0004 | [25 empty] | 16
| 1 | 4 | 0.0011 | [25 empty] | 0.0003
| 1 | 5 | < 0.0002 | [25 empty] | 9
| 1 | 6 | < 0.0008
| Mean |  | 0.00063333
| Median |  | 0.0004
| Std Dev. |  | 0.00040415
| Rel.Std.Dev. |  | 63.81%
| PDM3 |  | 111.1%
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 239b (Certified Value 2.67 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 2.807 | 2.7263 | 2.69 | 1.65 | 2.565 | 2.707 | 2.74 | 2.52 | 2.75 | 2.71 | 2.65 | 2.69 | 2.6 | 2.4653 | 2.6608 | 2.6648 | 2.65 | 2.58 | 2.23 | 2.74 | 2.69 | 2.6431 | 2.9759 | 3.1526 |  |  | 1
| 1 | 2 | 2.8128 | 2.6854 | 2.74 | 1.73 | 2.5296 | 2.759 | 2.78 | 2.5 | 2.74 | 2.73 | 2.6 | 2.58 | 2.65 | 2.7846 | 2.6678 | 2.6814 | 2.7 | 2.57 | 2.65 | 2.73 | 2.66 | 2.7039 | 3.084 | 3.1541 |  |  | #N/A
| 1 | 3 | 2.7797 | 2.6705 | 2.67 | 1.53 | 2.545 | 2.77 | 2.86 | 2.7 | 2.76 | 2.68 | 2.86 | 2.64 | 2.55 | 2.7203 | 2.6467 | 3.1463 | 2.63 | 2.6 | 2.54 | 2.7 | 2.71 | 2.6319 | 3.0662 | 3.1542 |  |  | 16
| 1 | 4 | 2.7718 | 2.7084 | 2.71 | 1.51 | 2.5269 | 2.702 | 2.88 | 2.65 | 2.74 | 2.67 | 2.75 | 2.59 | 2.61 | 2.6214 | 2.6033 | 2.6814 | 2.65 | 2.64 | 2.37 | 2.7 | 2.68 | 2.6903 | 3.0407 | 3.1547 |  |  | 2.6694
| 1 | 5 | 2.866 | 2.7091 | 2.7 | 1.47 | 2.5627 | 2.722 | 2.73 | 2.78 | 2.78 | 2.65 | 2.65 | 2.6 | 2.57 | 2.3266 | 2.6425 | 2.7063 | 2.68 | 2.61 | 2.14 | 2.71 | 2.65 | 2.6875 | 3.0433 | 3.1523 |  |  | 37
| 1 | 6 | 2.8303 | 2.7169 | 2.69 | 1.47 | 2.5705 | 2.744 | 2.75 | 2.48 | 2.78 | 2.63 | 2.89 | 2.57 | 2.58 | 2.4072 | 2.6034 | 2.7229 | 2.7 | 2.57 | 2.72 | 2.73 | 2.7 | 2.6798 | 3.0965 | 3.1553
| Mean |  | 2.8113 | 2.7028 | 2.7 | 1.56 | 2.5499 | 2.734 | 2.79 | 2.605 | 2.7583 | 2.6783 | 2.7333 | 2.6117 | 2.5933 | 2.5542 | 2.6374 | 2.7672 | 2.6683 | 2.595 | 2.4417 | 2.7183 | 2.6817 | 2.6728 | 3.0511 | 3.1539
| Median |  | 2.8099 | 2.7088 | 2.695 | 1.52 | 2.5538 | 2.733 | 2.765 | 2.585 | 2.755 | 2.675 | 2.7 | 2.595 | 2.59 | 2.5434 | 2.6446 | 2.6939 | 2.665 | 2.59 | 2.455 | 2.72 | 2.685 | 2.6837 | 3.0548 | 3.1542
| Std Dev. |  | 0.034446 | 0.020835 | 0.023664 | 0.1064 | 0.018887 | 0.027993 | 0.064498 | 0.12292 | 0.018348 | 0.037103 | 0.12044 | 0.04535 | 0.035024 | 0.18248 | 0.027947 | 0.18687 | 0.029269 | 0.027386 | 0.23302 | 0.017224 | 0.023166 | 0.028613 | 0.042885 | 0.001181
| Rel.Std.Dev. |  | 1.225% | 0.7709% | 0.8765% | 6.82% | 0.7407% | 1.024% | 2.312% | 4.719% | 0.6652% | 1.385% | 4.406% | 1.736% | 1.351% | 7.144% | 1.06% | 6.753% | 1.097% | 1.055% | 9.543% | 0.6336% | 0.8639% | 1.071% | 1.406% | 0.03744%
| PDM3 |  | 5.315% | 1.251% | 1.147% | -41.56% | -4.475% | 2.421% | 4.518% | -2.412% | 3.332% | 0.3353% | 2.396% | -2.162% | -2.849% | -4.314% | -1.198% | 3.664% | -0.03936% | -2.787% | -8.531% | 1.834% | 0.4601% | 0.1261% | 14.3% | 18.15%
| Z-Score (Absolute) |  | 1.15 | 0.2 | 0.18 | 9.85 | 1.14 | 0.47 | 0.97 | 0.66 | 0.69 | 0.01 | 0.47 | 0.6 | 0.76 | 1.11 | 0.37 | 0.77 | 0.1 | 0.75 | 2.1 | 0.34 | 0.01 | 0.06 | 3.26 | 4.17
| Analytical results for La in OREAS 239b (Certified Value 37.9 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 39.21 | 43.1022 | 40.7 | 35.3 | 35.77 | 34 | 42 | 38.4 | 40.5 | 41.1 | 39.9 | 37.7 | 35.0279 | 39 | 35.45 | 36.8 | 34.71 | 39.1 | 39.6 | 37.8 | 33.72 | 35.4866 |  |  |  |  | 1
| 1 | 2 | 39.53 | 43.4809 | 42 | 32.4 | 32.5119 | 33.4 | 39 | 40.7 | 39.8 | 37.7 | 40.3 | 41.2 | 35.7557 | 40 | 35.02 | 36 | 35.06 | 38.8 | 37.4 | 37.1 | 34.44 | 35.2816 |  |  |  |  | 28
| 1 | 3 | 39.93 | 42.7673 | 41.6 | 23.1 | 35.03 | 35.9 | 39 | 39.3 | 40.6 | 39.6 | 38.8 | 39.6 | 36.6555 | 47 | 34.41 | 34.6 | 35.98 | 39.4 | 37.7 | 38.2 | 33.83 | 36.5185 |  |  |  |  | 16
| 1 | 4 | 39.69 | 44.2142 | 40.2 | 35.3 | 35.82 | 36.4 | 39 | 38.4 | 41.7 | 37.4 | 38.2 | 41.5 | 35.4005 | 39 | 34.29 | 36.8 | 34.97 | 39.4 | 38.1 | 38.5 | 34.56 | 36.7526 |  |  |  |  | 37.918
| 1 | 5 | 39.72 | 44.0976 | 40.5 | 39.8 | 36 | 37.3 | 39 | 39.7 | 39.7 | 38.4 | 38.7 | 38.4 | 35.9877 | 40 | 34.58 | 35.6 | 34.98 | 39.8 | 36.7 | 37.5 | 34.48 | 36.692 |  |  |  |  | 38
| 1 | 6 | 40.06 | 43.5828 | 41.6 | 38.4 | 34.6587 | 35.8 | 39 | 37.6 | 38.9 | 39.6 | 39.1 | 36.6 | 36.6048 | 40 | 34.97 | 36.2 | 35.1 | 39.6 | 38.9 | 37.5 | 34.79 | 36.9073
| Mean |  | 39.69 | 43.5409 | 41.1 | 34.05 | 34.9651 | 35.4667 | 39.5 | 39.0167 | 40.2 | 38.9667 | 39.1667 | 39.1667 | 35.9053 | 40.8333 | 34.7867 | 36 | 35.1333 | 39.35 | 38.0667 | 37.7667 | 34.3033 | 36.2731
| Median |  | 39.705 | 43.5319 | 41.15 | 35.3 | 35.4 | 35.85 | 39 | 38.85 | 40.15 | 39 | 38.95 | 39 | 35.8717 | 40 | 34.775 | 36.1 | 35.02 | 39.4 | 37.9 | 37.65 | 34.46 | 36.6052
| Std Dev. |  | 0.30047 | 0.55848 | 0.72664 | 5.9621 | 1.3089 | 1.4801 | 1.2247 | 1.1089 | 0.95917 | 1.3952 | 0.78909 | 1.9541 | 0.64899 | 3.0605 | 0.43803 | 0.82946 | 0.43652 | 0.35637 | 1.0482 | 0.51251 | 0.42824 | 0.70277
| Rel.Std.Dev. |  | 0.757% | 1.283% | 1.768% | 17.51% | 3.743% | 4.173% | 3.101% | 2.842% | 2.386% | 3.581% | 2.015% | 4.989% | 1.807% | 7.495% | 1.259% | 2.304% | 1.242% | 0.9056% | 2.754% | 1.357% | 1.248% | 1.937%
| PDM3 |  | 4.673% | 14.83% | 8.392% | -10.2% | -7.788% | -6.465% | 4.172% | 2.898% | 6.018% | 2.766% | 3.293% | 3.293% | -5.308% | 7.689% | -8.258% | -5.058% | -7.344% | 3.777% | 0.3922% | -0.399% | -9.533% | -4.338%
| Z-Score (Absolute) |  | 0.35 | 1.49 | 0.76 | 1.32 | 1.05 | 0.9 | 0.29 | 0.15 | 0.5 | 0.13 | 0.19 | 0.19 | 0.77 | 0.69 | 1.1 | 0.74 | 1 | 0.25 | 0.13 | 0.22 | 1.25 | 0.66
| Analytical results for Li in OREAS 239b (Certified Value 48.3 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 48.5 | 50.5491 | 50 | 53 | 46.73 | 46.1 | 52 | 45.5 | 73.7 | 52.4 | 43.4 | 50.3 | 45.4 | 51.1 | 47.5459 | 49 | 40.02 | 47.2 | 50.11 | 46.8 | 45.5 | 49.5 | 52.6 | 39.9551 |  |  | 1
| 1 | 2 | 48.5 | 51.4947 | 50 | 52 | 48.26 | 46 | 52 | 45.8 | 71.1 | 52.2 | 43.6 | 46.3 | 46.3 | 49.7 | 49.7309 | 49 | 40.69 | 48.9 | 50.31 | 46.7 | 42.7 | 49.4 | 54.2 | 41.2894 |  |  | #N/A
| 1 | 3 | 49 | 50.9322 | 50 | 52 | 49.2267 | 45.9 | 52 | 47.2 | 75.1 | 52.8 | 46.7 | 49.7 | 44.3 | 51.4 | 49.4026 | 58 | 41.71 | 44.8 | 49.67 | 45.1 | 43.6 | 50.1 | 52.5 | 43.8081 |  |  | 16
| 1 | 4 | 49.1 | 50.7908 | 50 | 53 | 48.3967 | 48.1 | 53 | 45.8 | 74.2 | 52.1 | 46.1 | 47.7 | 43 | 50.8 | 49.8219 | 49 | 40.83 | 48.1 | 49.93 | 48.1 | 44.3 | 51.1 | 53.7 | 44.4748 |  |  | 48.3104
| 1 | 5 | 51.2 | 50.6291 | 50 | 53 | 47.5333 | 47.5 | 52 | 44.7 | 75.2 | 54 | 44.8 | 48.4 | 43.6 | 50.5 | 49.5806 | 49 | 42.65 | 48.7 | 50.69 | 47 | 43.9 | 50 | 53.8 | 45.01 |  |  | 39
| 1 | 6 | 50 | 51.0377 | 50 | 53 | 48.15 | 47.2 | 51 | 48.3 | 70.2 | 47.3 | 45.3 | 43.8 | 45.3 | 50.6 | 49.2944 | 50 | 41.93 | 48 | 50.39 | 48.6 | 44.4 | 49.5 | 53.7 | 44.5931
| Mean |  | 49.3833 | 50.9056 | 50 | 52.6667 | 48.0494 | 46.8 | 52 | 46.2167 | 73.25 | 51.8 | 44.9833 | 47.7 | 44.65 | 50.6833 | 49.2294 | 50.6667 | 41.305 | 47.6167 | 50.1833 | 47.05 | 44.0667 | 49.9333 | 53.4167 | 43.1884
| Median |  | 49.05 | 50.8615 | 50 | 53 | 48.205 | 46.65 | 52 | 45.8 | 73.95 | 52.3 | 45.05 | 48.05 | 44.8 | 50.7 | 49.4916 | 49 | 41.27 | 48.05 | 50.21 | 46.9 | 44.1 | 49.75 | 53.7 | 44.1414
| Std Dev. |  | 1.0458 | 0.34116 | 0 | 0.5164 | 0.84477 | 0.9252 | 0.63246 | 1.3014 | 2.1097 | 2.3108 | 1.3227 | 2.3841 | 1.2373 | 0.58452 | 0.84784 | 3.6148 | 0.96112 | 1.5039 | 0.36015 | 1.2243 | 0.93095 | 0.64083 | 0.6969 | 2.0684
| Rel.Std.Dev. |  | 2.118% | 0.6702% | 0% | 0.9805% | 1.758% | 1.977% | 1.216% | 2.816% | 2.88% | 4.461% | 2.941% | 4.998% | 2.771% | 1.153% | 1.722% | 7.134% | 2.327% | 3.158% | 0.7177% | 2.602% | 2.113% | 1.283% | 1.305% | 4.789%
| PDM3 |  | 2.221% | 5.372% | 3.497% | 9.017% | -0.5401% | -3.126% | 7.637% | -4.334% | 51.62% | 7.223% | -6.887% | -1.263% | -7.577% | 4.912% | 1.902% | 4.877% | -14.5% | -1.436% | 3.877% | -2.609% | -8.784% | 3.359% | 10.57% | -10.6%
| Z-Score (Absolute) |  | 0.04 | 0.47 | NA | 0.95 | 0.33 | 0.67 | 0.77 | 0.84 | 6.67 | 0.71 | 1.18 | 0.42 | 1.27 | 0.4 | 0 | 0.4 | 2.2 | 0.45 | 0.26 | 0.6 | 1.43 | 0.2 | 1.16 | 1.68
Results from laboratory 1.05 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 239b (Certified Value 0.28 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.24853 | 0.3 | 0.3 | 0.2 | 0.226 | 0.27 | 0.32 | 0.2 | 0.1105 | 0.3715 | [16 empty] | 1
| 1 | 2 | 0.25516 | 0.3 | 0.3 | 0.2 | 0.224 | 0.27 | 0.34 | 0.2 | 0.1117 | 0.36621 | [16 empty] | 29
| 1 | 3 | 0.25325 | 0.3 | 0.3 | 0.2 | 0.217 | 0.26 | 0.34 | 0.2 | 0.1213 | 0.3721 | [16 empty] | 16
| 1 | 4 | 0.24024 | 0.3 | 0.3 | 0.3 | 0.222 | 0.27 | 0.33 | 0.2 | 0.1061 | 0.36541 | [16 empty] | 0.27896
| 1 | 5 | 0.23122 | 0.3 | 0.3 | 0.3 | 0.221 | 0.28 | 0.36 | 0.2 | 0.1314 | 0.37366 | [16 empty] | 40
| 1 | 6 | 0.25341 | 0.3 | 0.3 | 0.4 | 0.248 | 0.3 | 0.33 | 0.2 | 0.089 | 0.36128
| Mean |  | 0.24697 | 0.3 | 0.3 | 0.26667 | 0.22633 | 0.275 | 0.33667 | 0.2 | 0.11167 | 0.36836
| Median |  | 0.25089 | 0.3 | 0.3 | 0.25 | 0.223 | 0.27 | 0.335 | 0.2 | 0.1111 | 0.36885
| Std Dev. |  | 0.0094223 | 0 | 0 | 0.08165 | 0.011039 | 0.013784 | 0.013663 | 3.0405e-17 | 0.014341 | 0.0048039
| Rel.Std.Dev. |  | 3.815% | 0% | 0% | 30.62% | 4.877% | 5.012% | 4.058% | 1.52e-14% | 12.84% | 1.304%
| PDM3 |  | -11.47% | 7.541% | 7.541% | -4.408% | -18.87% | -1.42% | 20.69% | -28.31% | -59.97% | 32.05%
| Z-Score (Absolute) |  | 0.44 | 0.53 | 0.53 | 0.08 | 0.81 | 0.08 | 1.21 | 1.3 | 2.91 | 1.79
| Analytical results for Mg in OREAS 239b (Certified Value 1.72 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 1.6555 | 1.7652 | 1.73 | 1.77 | 1.78 | 1.746 | 1.7 | 1.75 | 1.71 | 1.7 | 1.67 | 1.69 | 1.64 | 1.7366 | 1.7002 | 1.7609 | 1.69 | 1.73 | 1.78 | 1.72 | 1.7 | 1.7283 | 2.003 | 1.6458 |  |  | 1
| 1 | 2 | 1.6919 | 1.7337 | 1.76 | 1.75 | 1.7201 | 1.763 | 1.68 | 1.76 | 1.71 | 1.71 | 1.62 | 1.62 | 1.66 | 1.7189 | 1.7316 | 1.779 | 1.69 | 1.73 | 1.79 | 1.72 | 1.69 | 1.7637 | 2.0808 | 1.635 |  |  | #N/A
| 1 | 3 | 1.6781 | 1.7483 | 1.7 | 1.59 | 1.75 | 1.769 | 1.72 | 1.73 | 1.71 | 1.68 | 1.74 | 1.66 | 1.62 | 1.7562 | 1.7242 | 2.0987 | 1.69 | 1.74 | 1.78 | 1.69 | 1.7 | 1.7286 | 2.0693 | 1.6598 |  |  | 16
| 1 | 4 | 1.6577 | 1.7511 | 1.72 | 1.71 | 1.76 | 1.742 | 1.73 | 1.71 | 1.79 | 1.66 | 1.73 | 1.63 | 1.62 | 1.7277 | 1.735 | 1.773 | 1.7 | 1.74 | 1.78 | 1.69 | 1.7 | 1.7571 | 2.0496 | 1.6409 |  |  | 1.7167
| 1 | 5 | 1.7351 | 1.7416 | 1.75 | 1.79 | 1.7765 | 1.76 | 1.72 | 1.78 | 1.71 | 1.67 | 1.68 | 1.64 | 1.61 | 1.7277 | 1.7084 | 1.7911 | 1.71 | 1.73 | 1.79 | 1.7 | 1.69 | 1.7484 | 2.0522 | 1.6366 |  |  | 41
| 1 | 6 | 1.7062 | 1.7812 | 1.74 | 1.85 | 1.722 | 1.766 | 1.69 | 1.76 | 1.73 | 1.68 | 1.8 | 1.63 | 1.61 | 1.7101 | 1.6896 | 1.8213 | 1.67 | 1.73 | 1.8 | 1.71 | 1.72 | 1.7469 | 2.0906 | 1.599
| Mean |  | 1.6874 | 1.7535 | 1.7333 | 1.7433 | 1.7514 | 1.7577 | 1.7067 | 1.7483 | 1.7267 | 1.6833 | 1.7067 | 1.645 | 1.6267 | 1.7295 | 1.7148 | 1.8373 | 1.6917 | 1.7333 | 1.7867 | 1.705 | 1.7 | 1.7455 | 2.0576 | 1.6362
| Median |  | 1.685 | 1.7497 | 1.735 | 1.76 | 1.755 | 1.7615 | 1.71 | 1.755 | 1.71 | 1.68 | 1.705 | 1.635 | 1.62 | 1.7277 | 1.7163 | 1.7851 | 1.69 | 1.73 | 1.785 | 1.705 | 1.7 | 1.7476 | 2.0608 | 1.6388
| Std Dev. |  | 0.030447 | 0.01716 | 0.021602 | 0.088242 | 0.025954 | 0.011075 | 0.019664 | 0.024833 | 0.032042 | 0.018619 | 0.06314 | 0.025884 | 0.019664 | 0.015865 | 0.018247 | 0.12968 | 0.013292 | 0.005164 | 0.008165 | 0.013784 | 0.010954 | 0.014544 | 0.031108 | 0.020281
| Rel.Std.Dev. |  | 1.804% | 0.9786% | 1.246% | 5.062% | 1.482% | 0.6301% | 1.152% | 1.42% | 1.856% | 1.106% | 3.7% | 1.574% | 1.209% | 0.9173% | 1.064% | 7.058% | 0.7857% | 0.2979% | 0.457% | 0.8084% | 0.6444% | 0.8332% | 1.512% | 1.24%
| PDM3 |  | -1.708% | 2.14% | 0.9663% | 1.549% | 2.021% | 2.384% | -0.587% | 1.84% | 0.578% | -1.946% | -0.587% | -4.179% | -5.247% | 0.745% | -0.1123% | 7.024% | -1.461% | 0.9663% | 4.073% | -0.6841% | -0.9753% | 1.675% | 19.85% | -4.693%
| Z-Score (Absolute) |  | 1.13 | 0.7 | 0.14 | 0.42 | 0.65 | 0.82 | 0.59 | 0.56 | 0.04 | 1.24 | 0.59 | 2.3 | 2.81 | 0.04 | 0.37 | 3.02 | 1.01 | 0.14 | 1.62 | 0.64 | 0.78 | 0.48 | 9.12 | 2.54
| Analytical results for Mn in OREAS 239b (Certified Value 0.039 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0408 | 0.039168 | 0.0402 | 0.0428 | 0.040269 | 0.0402 | 0.0385 | 0.0368 | 0.0411 | 0.0393 | 0.0385 | 0.039 | 0.0384 | 0.0403 | 0.039413 | 0.0372 | 0.0363 | 0.03686 | 0.0393 | 0.0385 | 0.0393 | 0.040876 | 0.0423 | 0.0391 |  |  | 1
| 1 | 2 | 0.0414 | 0.039153 | 0.0414 | 0.0418 | 0.040315 | 0.0392 | 0.0386 | 0.0369 | 0.0413 | 0.0393 | 0.0373 | 0.0378 | 0.0399 | 0.0396 | 0.039032 | 0.0379 | 0.0365 | 0.03677 | 0.0399 | 0.0385 | 0.0392 | 0.042317 | 0.0433 | 0.0389 |  |  | 30
| 1 | 3 | 0.0411 | 0.039293 | 0.0408 | 0.0426 | 0.041043 | 0.0387 | 0.0382 | 0.0363 | 0.0404 | 0.0384 | 0.0401 | 0.038 | 0.0376 | 0.0404 | 0.039337 | 0.0441 | 0.0358 | 0.03722 | 0.0395 | 0.038 | 0.0391 | 0.040821 | 0.043 | 0.0382 |  |  | 16
| 1 | 4 | 0.0406 | 0.038986 | 0.0404 | 0.0425 | 0.04069 | 0.0404 | 0.039 | 0.0359 | 0.0402 | 0.0383 | 0.0395 | 0.038 | 0.0384 | 0.04 | 0.039713 | 0.0372 | 0.0373 | 0.03651 | 0.0395 | 0.0379 | 0.0392 | 0.041833 | 0.0423 | 0.0388 |  |  | 0.039344
| 1 | 5 | 0.041 | 0.039283 | 0.0414 | 0.0438 | 0.040994 | 0.039 | 0.0381 | 0.0378 | 0.0408 | 0.0386 | 0.0386 | 0.038 | 0.0383 | 0.0398 | 0.039612 | 0.0372 | 0.0368 | 0.03729 | 0.0395 | 0.0382 | 0.039 | 0.041866 | 0.0422 | 0.0391 |  |  | 42
| 1 | 6 | 0.0421 | 0.039049 | 0.0404 | 0.0431 | 0.040763 | 0.0387 | 0.0379 | 0.0369 | 0.041 | 0.0376 | 0.0414 | 0.038 | 0.0384 | 0.0394 | 0.0386 | 0.0379 | 0.0366 | 0.03667 | 0.0399 | 0.0381 | 0.0397 | 0.041503 | 0.0437 | 0.0382
| Mean |  | 0.041167 | 0.039155 | 0.040767 | 0.042767 | 0.040679 | 0.039367 | 0.038383 | 0.036767 | 0.0408 | 0.038583 | 0.039233 | 0.038133 | 0.0385 | 0.039917 | 0.039284 | 0.038583 | 0.03655 | 0.036887 | 0.0396 | 0.0382 | 0.03925 | 0.041536 | 0.0428 | 0.038717
| Median |  | 0.04105 | 0.039161 | 0.0406 | 0.0427 | 0.040726 | 0.0391 | 0.03835 | 0.03685 | 0.0409 | 0.0385 | 0.03905 | 0.038 | 0.0384 | 0.0399 | 0.039375 | 0.03755 | 0.03655 | 0.036815 | 0.0395 | 0.03815 | 0.0392 | 0.041668 | 0.04265 | 0.03885
| Std Dev. |  | 0.00053166 | 0.0001228 | 0.00052789 | 0.00066533 | 0.00032828 | 0.00075011 | 0.00039707 | 0.00064395 | 0.00042426 | 0.00064936 | 0.0014278 | 0.00043205 | 0.00075366 | 0.000392 | 0.00041046 | 0.0027243 | 0.000501 | 0.00030885 | 0.00024495 | 0.00025298 | 0.0002429 | 0.00059245 | 0.0006261 | 0.00041673
| Rel.Std.Dev. |  | 1.291% | 0.3136% | 1.295% | 1.556% | 0.807% | 1.905% | 1.034% | 1.751% | 1.04% | 1.683% | 3.639% | 1.133% | 1.958% | 0.9821% | 1.045% | 7.061% | 1.371% | 0.8373% | 0.6186% | 0.6623% | 0.6189% | 1.426% | 1.463% | 1.076%
| PDM3 |  | 4.634% | -0.4784% | 3.617% | 8.701% | 3.394% | 0.05879% | -2.441% | -6.55% | 3.702% | -1.932% | -0.2801% | -3.076% | -2.144% | 1.457% | -0.1503% | -1.932% | -7.1% | -6.245% | 0.6519% | -2.907% | -0.2377% | 5.573% | 8.785% | -1.593%
| Z-Score (Absolute) |  | 1.37 | 0.06 | 1.08 | 2.5 | 1.02 | 0.09 | 0.61 | 1.76 | 1.11 | 0.47 | 0.01 | 0.79 | 0.53 | 0.48 | 0.03 | 0.46 | 1.91 | 1.67 | 0.25 | 0.74 | 0.01 | 1.63 | 2.53 | 0.37
| Analytical results for Mo in OREAS 239b (Certified Value 0.99 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 1 | 2 | < 1 | < 2 | 0.9 | 0.98 | 0.9 | 1.6 | 0.88 | 1 | 0.98 | 0.95 | 1.08 | 1.0052 | 1.01 | 0.991 | 1.14 | < 1 | 1.03 | 1.1 | 0.98 | 0.795 | < 5 | 2.6524 | 0.97156 |  | 1
| 1 | 2 | 1.1 | 2 | < 1 | < 2 | 1 | 1.09 | 0.8 | 1.7 | 0.9 | 0.88 | 0.95 | 0.94 | 0.89 | 1.0247 | 0.98 | 1.249 | 1.01 | < 1 | 0.99 | 1.04 | 0.91 | 0.7518 | < 5 | 2.569 | 0.99218 |  | 31
| 1 | 3 | 1 | 2 | < 1 | < 2 | 1 | 1.12 | 0.9 | 1.9 | 0.92 | 0.99 | 1.04 | 0.96 | 0.99 | 1.0028 | 1 | 1.54 | 0.98 | < 1 | 1.08 | 1 | 0.97 | 0.7107 | < 5 | 2.569 | 0.9871 |  | 16
| 1 | 4 | 1 | 1.5 | < 1 | < 2 | 0.9 | 1.09 | 0.9 | 1.7 | 0.89 | 0.96 | 0.97 | 0.93 | 0.99 | 0.98977 | 1.06 | 1.581 | 1.02 | < 1 | 0.96 | 1.05 | 0.92 | 0.7777 | < 5 | 2.7358 | 0.95107 |  | 0.98576
| 1 | 5 | 1 | 2 | < 1 | < 2 | 0.9 | 1.1 | 1 | 1.6 | 0.9 | 0.93 | 1 | 0.91 | 1.12 | 1.0207 | 1.02 | 1.161 | 1.04 | < 1 | 1 | 0.96 | 0.88 | 0.7485 | < 5 | 2.8749 | 0.92089 |  | 43
| 1 | 6 | 1 | 2 | < 1 | < 2 | 1 | 1.18 | 0.9 | 1.8 | 0.95 | 0.89 | 0.95 | 0.93 | 1.09 | 0.97427 | 0.97 | 1.208 | 1.07 | < 1 | 1.01 | 1.06 | 1.06 | 0.7831 | < 5 | 2.5133 | 0.97257
| Mean |  | 1.0167 | 1.9167 |  |  | 0.95 | 1.0933 | 0.9 | 1.7167 | 0.90667 | 0.94167 | 0.98167 | 0.93667 | 1.0267 | 1.0029 | 1.0067 | 1.2883 | 1.0433 |  | 1.0117 | 1.035 | 0.95333 | 0.76113 |  | 2.6524 | 0.96589
| Median |  | 1 | 2 |  |  | 0.95 | 1.095 | 0.9 | 1.7 | 0.9 | 0.945 | 0.975 | 0.935 | 1.035 | 1.004 | 1.005 | 1.2285 | 1.03 |  | 1.005 | 1.045 | 0.945 | 0.76475 |  | 2.6107 | 0.97206
| Std Dev. |  | 0.040825 | 0.20412 |  |  | 0.054772 | 0.065013 | 0.063246 | 0.1169 | 0.025033 | 0.050365 | 0.034303 | 0.017512 | 0.085946 | 0.018914 | 0.032042 | 0.22874 | 0.056095 |  | 0.040702 | 0.048888 | 0.064395 | 0.03064 |  | 0.13396 | 0.026315
| Rel.Std.Dev. |  | 4.016% | 10.65% |  |  | 5.766% | 5.946% | 7.027% | 6.81% | 2.761% | 5.349% | 3.494% | 1.87% | 8.371% | 1.886% | 3.183% | 17.75% | 5.377% |  | 4.023% | 4.723% | 6.755% | 4.026% |  | 5.05% | 2.724%
| PDM3 |  | 3.136% | 94.44% |  |  | -3.627% | 10.91% | -8.7% | 74.15% | -8.023% | -4.473% | -0.415% | -4.98% | 4.15% | 1.738% | 2.121% | 30.69% | 5.841% |  | 2.628% | 4.995% | -3.289% | -22.79% |  | 169.1% | -2.015%
| Z-Score (Absolute) |  | 0.15 | 10.06 | 5.54 | 0.03 | 0.58 | 1 | 1.13 | 7.86 | 1.06 | 0.67 | 0.23 | 0.73 | 0.26 | 0 | 0.04 | 3.14 | 0.45 | 5.54 | 0.1 | 0.35 | 0.55 | 2.66 | 16.49 | 18.17 | 0.41
| Analytical results for Na in OREAS 239b (Certified Value 0.682 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.6819 | 0.7064 | 0.7 | 0.76 | 0.749 | 0.691 | 0.69 | 0.704 | 0.68 | 0.68 | 0.67 | 0.69 | 0.69 | 0.68 | 0.64275 | 0.668 | 0.67 | 0.66 | 0.68 | 0.68 | 0.679 | 0.68071 | 0.7173 |  |  |  | 1
| 1 | 2 | 0.6893 | 0.70483 | 0.69 | 0.73 | 0.73 | 0.688 | 0.67 | 0.709 | 0.67 | 0.69 | 0.65 | 0.65 | 0.7 | 0.666 | 0.65721 | 0.675 | 0.68 | 0.67 | 0.68 | 0.68 | 0.68 | 0.69854 | 0.7034 |  |  |  | #N/A
| 1 | 3 | 0.6982 | 0.70863 | 0.69 | 0.71 | 0.7645 | 0.65 | 0.68 | 0.696 | 0.68 | 0.67 | 0.71 | 0.67 | 0.68 | 0.683 | 0.66283 | 0.794 | 0.67 | 0.67 | 0.68 | 0.67 | 0.686 | 0.68492 | 0.7172 |  |  |  | 16
| 1 | 4 | 0.7036 | 0.70372 | 0.69 | 0.75 | 0.7535 | 0.655 | 0.71 | 0.685 | 0.68 | 0.67 | 0.7 | 0.66 | 0.68 | 0.683 | 0.64828 | 0.668 | 0.68 | 0.67 | 0.68 | 0.67 | 0.677 | 0.69521 | 0.7126 |  |  |  | 0.68217
| 1 | 5 | 0.725 | 0.70198 | 0.7 | 0.79 | 0.748 | 0.664 | 0.67 | 0.717 | 0.68 | 0.66 | 0.68 | 0.66 | 0.68 | 0.674 | 0.6394 | 0.675 | 0.68 | 0.67 | 0.68 | 0.67 | 0.672 | 0.6908 | 0.7221 |  |  |  | 44
| 1 | 6 | 0.7274 | 0.70121 | 0.69 | 0.77 | 0.756 | 0.668 | 0.67 | 0.708 | 0.67 | 0.66 | 0.72 | 0.66 | 0.68 | 0.68 | 0.65762 | 0.682 | 0.69 | 0.67 | 0.69 | 0.67 | 0.691 | 0.68767 | 0.7056
| Mean |  | 0.70423 | 0.70446 | 0.69333 | 0.75167 | 0.75017 | 0.66933 | 0.68167 | 0.70317 | 0.67667 | 0.67167 | 0.68833 | 0.665 | 0.685 | 0.67767 | 0.65135 | 0.69367 | 0.67833 | 0.66833 | 0.68167 | 0.67333 | 0.68083 | 0.68964 | 0.71303
| Median |  | 0.7009 | 0.70427 | 0.69 | 0.755 | 0.75125 | 0.666 | 0.675 | 0.706 | 0.68 | 0.67 | 0.69 | 0.66 | 0.68 | 0.68 | 0.65274 | 0.675 | 0.68 | 0.67 | 0.68 | 0.67 | 0.6795 | 0.68923 | 0.7149
| Std Dev. |  | 0.018582 | 0.0027767 | 0.005164 | 0.028577 | 0.011518 | 0.016896 | 0.016021 | 0.011232 | 0.005164 | 0.01169 | 0.026394 | 0.013784 | 0.0083666 | 0.0065929 | 0.0092922 | 0.049431 | 0.0075277 | 0.0040825 | 0.0040825 | 0.005164 | 0.0067355 | 0.0065972 | 0.007294
| Rel.Std.Dev. |  | 2.639% | 0.3942% | 0.7448% | 3.802% | 1.535% | 2.524% | 2.35% | 1.597% | 0.7631% | 1.741% | 3.835% | 2.073% | 1.221% | 0.9729% | 1.427% | 7.126% | 1.11% | 0.6108% | 0.5989% | 0.7669% | 0.9893% | 0.9566% | 1.023%
| PDM3 |  | 3.234% | 3.267% | 1.636% | 10.19% | 9.967% | -1.882% | -0.07429% | 3.077% | -0.8072% | -1.54% | 0.903% | -2.517% | 0.4143% | -0.6607% | -4.519% | 1.685% | -0.5629% | -2.029% | -0.07429% | -1.296% | -0.1964% | 1.095% | 4.524%
| Z-Score (Absolute) |  | 1.3 | 1.32 | 0.67 | 4.05 | 3.96 | 0.71 | 0 | 1.24 | 0.29 | 0.58 | 0.39 | 0.96 | 0.19 | 0.23 | 1.75 | 0.69 | 0.19 | 0.77 | 0 | 0.48 | 0.05 | 0.46 | 1.81
| Analytical results for Nb in OREAS 239b (Certified Value 14.1 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 14.57 | 12.608 | 4.1 | 16.55 | 14 | 11 | 12.4 | 13.4 | 14.4 | 14.7 | 14.2 | 14.7 | 13.81 | 14.4609 | 15 | 13.2 | 15.32 | 14.6 | 11.89 | 10.0026 | 15.233 | [5 empty] | 1
| 1 | 2 | 14.8 | 12.6126 | 5.4 | 15.88 | 13.6 | 10 | 12.1 | 14 | 14.5 | 14.7 | 13.8 | 14.1 | 13.52 | 14.8774 | 14.6 | 13.4 | 15.23 | 14 | 11.92 | 9.8932 | 15.2275 | [5 empty] | 32
| 1 | 3 | 14.31 | 12.4828 | 7.7 | 15.07 | 13.7 | 9 | 13.2 | 14 | 14.3 | 15.2 | 14.4 | 13.7 | 14.53 | 15.445 | 14.9 | 12.8 | 15.24 | 13.8 | 12.16 | 10.09 | 15.1694 | [5 empty] | 16
| 1 | 4 | 15.24 | 12.7718 | 4.4 | 14.64 | 13.7 | 10 | 13 | 14 | 14.4 | 14.8 | 13.3 | 13.7 | 15.3 | 14.1905 | 14.9 | 12.7 | 15.48 | 14.1 | 11.57 | 10.6261 | 15.1883 | [5 empty] | 14.1432
| 1 | 5 | 15.52 | 12.5423 | 1.9 | 16.26 | 14.1 | 10 | 13.6 | 13.9 | 14.5 | 14.7 | 13.4 | 13.9 | 15.81 | 14.3414 | 14.6 | 13 | 15.4 | 13.6 | 11.82 | 10.3136 | 15.2188 | [5 empty] | 45
| 1 | 6 | 15.56 | 12.5296 | 2.3 | 15.34 | 14 | 10 | 13.6 | 13.7 | 14.3 | 14.9 | 13.8 | 13.7 | 15.82 | 14.9995 | 14.9 | 13 | 15.67 | 14.2 | 11.62 | 10.0447 | 15.1201
| Mean |  | 15 | 12.5912 | 4.3 | 15.6233 | 13.85 | 10 | 12.9833 | 13.8333 | 14.4 | 14.8333 | 13.8167 | 13.9667 | 14.7983 | 14.7191 | 14.8167 | 13.0167 | 15.39 | 14.05 | 11.83 | 10.1617 | 15.1928
| Median |  | 15.02 | 12.5752 | 4.25 | 15.61 | 13.85 | 10 | 13.1 | 13.95 | 14.4 | 14.75 | 13.8 | 13.8 | 14.915 | 14.6691 | 14.9 | 13 | 15.36 | 14.05 | 11.855 | 10.0674 | 15.2035
| Std Dev. |  | 0.51819 | 0.10128 | 2.1251 | 0.73258 | 0.20736 | 0.63246 | 0.62102 | 0.24221 | 0.089443 | 0.19664 | 0.43089 | 0.39328 | 1.0002 | 0.47316 | 0.17224 | 0.25626 | 0.16709 | 0.34496 | 0.21559 | 0.26646 | 0.043259
| Rel.Std.Dev. |  | 3.455% | 0.8044% | 49.42% | 4.689% | 1.497% | 6.325% | 4.783% | 1.751% | 0.6211% | 1.326% | 3.119% | 2.816% | 6.759% | 3.215% | 1.162% | 1.969% | 1.086% | 2.455% | 1.822% | 2.622% | 0.2847%
| PDM3 |  | 6.058% | -10.97% | -69.6% | 10.47% | -2.073% | -29.29% | -8.201% | -2.191% | 1.816% | 4.879% | -2.309% | -1.248% | 4.632% | 4.072% | 4.762% | -7.965% | 8.815% | -0.6592% | -16.36% | -28.15% | 7.421%
| Z-Score (Absolute) |  | 0.74 | 1.05 | 7.21 | 1.2 | 0.12 | NA | 0.76 | 0.13 | 0.29 | 0.61 | 0.14 | 0.03 | 0.59 | 0.53 | 0.6 | 0.74 | 1.03 | 0.03 | 1.62 | 2.86 | 0.88
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 239b (Certified Value 33.2 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 1
2
| 1 | 1 | 32.8781 | 34.1 | 31.5 | 28.9 | 33.9 | 32.6 | 33.9 | 33.2 | 29.47 | 32.0231 | [16 empty] | 1
| 1 | 2 | 33.0159 | 34.6 | 30.4 | 28.7 | 34.8 | 32.57 | 33.96 | 32.7 | 30.56 | 31.9784 | [16 empty] | 9
| 1 | 3 | 31.6141 | 34.4 | 24 | 31.3 | 34.4 | 31.84 | 34.12 | 33.8 | 29.71 | 32.9669 | [16 empty] | 16
| 1 | 4 | 31.8281 | 35.2 | 31.9 | 34.8 | 34.1 | 32.2 | 34.54 | 33.9 | 30.5 | 33.1677 | [16 empty] | 33.1793
| 1 | 5 | 32.4605 | 35.3 | 34.4 | 35.1 | 34.3 | 32.8 | 35.41 | 33.2 | 30.54 | 32.7997 | [16 empty] | 46
| 1 | 6 | 33.1604 | 34.8 | 33.6 | 38.3 | 34.3 | 32.68 | 35.51 | 33.2 | 30.95 | 32.741
| Mean |  | 32.4928 | 34.7333 | 30.9667 | 32.85 | 34.3 | 32.4483 | 34.5733 | 33.3333 | 30.2883 | 32.6128
| Median |  | 32.6693 | 34.7 | 31.7 | 33.05 | 34.3 | 32.585 | 34.33 | 33.2 | 30.52 | 32.7704
| Std Dev. |  | 0.64539 | 0.46332 | 3.7076 | 3.8417 | 0.30332 | 0.35969 | 0.72298 | 0.44572 | 0.56989 | 0.49699
| Rel.Std.Dev. |  | 1.986% | 1.334% | 11.97% | 11.69% | 0.8843% | 1.108% | 2.091% | 1.337% | 1.882% | 1.524%
| PDM3 |  | -2.069% | 4.684% | -6.669% | -0.9926% | 3.378% | -2.203% | 4.201% | 0.4641% | -8.713% | -1.707%
| Z-Score (Absolute) |  | 0.15 | 1.24 | 1.1 | 0.07 | 0.97 | 0.18 | 1.14 | 0.37 | 1.52 | 0.07
| Analytical results for Ni in OREAS 239b (Certified Value 65 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 63.2 | 60.42 | 70 | 67 | 60.0433 | 64 | 69 | 63 | 65.1 | 64.6 | 68.5 | 67.9 | 63.9 | 63.1 | 66.1716 | 66 | 60.58 | 68 | 58.65 | 63.4 | 65.9 | 65.7 | 62.92 | 67 | 58.1173 | 62.6041 | 1
| 1 | 2 | 63.3 | 60.08 | 70 | 66 | 60.2817 | 64 | 69 | 64 | 65.5 | 65.2 | 65.4 | 63.8 | 64.3 | 62.1 | 67.174 | 65 | 53.3 | 66 | 57.07 | 62.6 | 62.3 | 68 | 65.56 | 69 | 57.0891 | 63.5193 | 34
| 1 | 3 | 64.1 | 59.75 | 70 | 68 | 61.3867 | 64 | 70 | 62 | 65.5 | 63.6 | 65.1 | 66.1 | 62.7 | 63.4 | 66.5452 | 64.9 | 50.33 | 64 | 57.86 | 62.5 | 62.6 | 67.9 | 63.25 | 69 | 56.6179 | 64.6495 | 16
| 1 | 4 | 63.3 | 59.06 | 70 | 68 | 61.83 | 64 | 64 | 62 | 65.8 | 63.7 | 72.5 | 63.4 | 62.8 | 61.1 | 65.9085 | 66.5 | 48.43 | 65 | 57.74 | 63.7 | 63.5 | 67.1 | 64.9 | 67 | 56.6607 | 62.6963 | 64.8498
| 1 | 5 | 63.7 | 60.14 | 70 | 67 | 60.7067 | 65 | 64 | 64 | 65 | 64.2 | 69.2 | 64 | 63.1 | 62.4 | 66.5301 | 65.7 | 50.24 | 67 | 58.43 | 64.3 | 60.6 | 68.1 | 64.24 | 68 | 55.6754 | 64.3393 | 47
| 1 | 6 | 64.9 | 60.47 | 70 | 67 | 59.2233 | 65 | 65 | 62 | 66.7 | 66 | 67.5 | 64 | 61.1 | 62.9 | 66.7735 | 64.6 | 51.17 | 64 | 58.64 | 64.8 | 64 | 67.6 | 64.46 | 69 | 54.0475 | 64.2361
| Mean |  | 63.75 | 59.9867 | 70 | 67.1667 | 60.5786 | 64.3333 | 66.8333 | 62.8333 | 65.6 | 64.55 | 68.0333 | 64.8667 | 62.9833 | 62.5 | 66.5172 | 65.45 | 52.3417 | 65.6667 | 58.065 | 63.55 | 63.15 | 67.4 | 64.2217 | 68.1667 | 56.368 | 63.6741
| Median |  | 63.5 | 60.11 | 70 | 67 | 60.4942 | 64 | 67 | 62.5 | 65.5 | 64.4 | 68 | 64 | 62.95 | 62.65 | 66.5376 | 65.35 | 50.75 | 65.5 | 58.145 | 63.55 | 63.05 | 67.75 | 64.35 | 68.5 | 56.6393 | 63.8777
| Std Dev. |  | 0.65651 | 0.52313 | 0 | 0.75277 | 0.94306 | 0.5164 | 2.7869 | 0.98319 | 0.61319 | 0.92466 | 2.7333 | 1.7614 | 1.1179 | 0.83187 | 0.44451 | 0.73417 | 4.3342 | 1.633 | 0.6235 | 0.91378 | 1.7852 | 0.90774 | 0.99467 | 0.98319 | 1.3851 | 0.87574
| Rel.Std.Dev. |  | 1.03% | 0.8721% | 0% | 1.121% | 1.557% | 0.8027% | 4.17% | 1.565% | 0.9347% | 1.432% | 4.018% | 2.715% | 1.775% | 1.331% | 0.6683% | 1.122% | 8.281% | 2.487% | 1.074% | 1.438% | 2.827% | 1.347% | 1.549% | 1.442% | 2.457% | 1.375%
| PDM3 |  | -1.696% | -7.499% | 7.942% | 3.573% | -6.586% | -0.7964% | 3.059% | -3.109% | 1.157% | -0.4623% | 4.909% | 0.02604% | -2.878% | -3.623% | 2.571% | 0.9256% | -19.29% | 1.26% | -10.46% | -2.004% | -2.621% | 3.933% | -0.9686% | 5.115% | -13.08% | -1.813%
| Z-Score (Absolute) |  | 0.22 | 1.8 | 2.4 | 1.21 | 1.55 | 0.02 | 1.07 | 0.6 | 0.55 | 0.11 | 1.57 | 0.25 | 0.54 | 0.74 | 0.94 | 0.49 | 5 | 0.58 | 2.6 | 0.3 | 0.47 | 1.31 | 0.02 | 1.63 | 3.31 | 0.25
| Analytical results for P in OREAS 239b (Certified Value 0.065 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0642 | 0.062755 | 0.06 | 0.063 | 0.064559 | 0.068 | 0.063 | 0.065 | 0.067 | 0.068 | 0.067 | 0.067 | 0.065 | 0.0666 | 0.065764 | 0.0681 | 0.07 | 0.07 | 0.0633 | 0.0683 | 0.065 | 0.0649 | 0.0654 | 0.0729 |  |  | 1
| 1 | 2 | 0.0647 | 0.062866 | 0.065 | 0.063 | 0.063935 | 0.07 | 0.0631 | 0.065 | 0.064 | 0.068 | 0.065 | 0.066 | 0.067 | 0.0663 | 0.064014 | 0.0694 | 0.06 | 0.07 | 0.0646 | 0.0684 | 0.066 | 0.0664 | 0.0678 | 0.072 |  |  | #N/A
| 1 | 3 | 0.0661 | 0.062694 | 0.06 | 0.062 | 0.0646 | 0.069 | 0.0648 | 0.065 | 0.062 | 0.066 | 0.068 | 0.065 | 0.065 | 0.0673 | 0.065544 | 0.0698 | 0.06 | 0.07 | 0.0631 | 0.0673 | 0.067 | 0.0653 | 0.0671 | 0.0713 |  |  | 16
| 1 | 4 | 0.065 | 0.062896 | 0.06 | 0.064 | 0.06398 | 0.068 | 0.0662 | 0.063 | 0.061 | 0.066 | 0.069 | 0.065 | 0.065 | 0.0671 | 0.065488 | 0.0689 | 0.06 | 0.07 | 0.0645 | 0.067 | 0.065 | 0.0653 | 0.0664 | 0.072 |  |  | 0.065429
| 1 | 5 | 0.0646 | 0.062034 | 0.06 | 0.063 | 0.064387 | 0.069 | 0.0633 | 0.066 | 0.063 | 0.066 | 0.066 | 0.066 | 0.065 | 0.0663 | 0.066179 | 0.0676 | 0.06 | 0.07 | 0.0639 | 0.068 | 0.065 | 0.0653 | 0.0657 | 0.072 |  |  | 48
| 1 | 6 | 0.0633 | 0.062384 | 0.065 | 0.064 | 0.064598 | 0.069 | 0.0636 | 0.065 | 0.064 | 0.065 | 0.072 | 0.068 | 0.065 | 0.0663 | 0.063009 | 0.0698 | 0.06 | 0.07 | 0.0644 | 0.0677 | 0.064 | 0.0653 | 0.0679 | 0.0709
| Mean |  | 0.06465 | 0.062605 | 0.061667 | 0.063167 | 0.064343 | 0.068833 | 0.064 | 0.064833 | 0.0635 | 0.0665 | 0.067833 | 0.066167 | 0.065333 | 0.06665 | 0.065 | 0.068933 | 0.061667 | 0.07 | 0.063967 | 0.067783 | 0.065333 | 0.065417 | 0.066717 | 0.07185
| Median |  | 0.06465 | 0.062724 | 0.06 | 0.063 | 0.064473 | 0.069 | 0.06345 | 0.065 | 0.0635 | 0.066 | 0.0675 | 0.066 | 0.065 | 0.06645 | 0.065516 | 0.06915 | 0.06 | 0.07 | 0.06415 | 0.06785 | 0.065 | 0.0653 | 0.06675 | 0.072
| Std Dev. |  | 0.0009225 | 0.00033427 | 0.002582 | 0.00075277 | 0.00030909 | 0.00075277 | 0.0012602 | 0.00098319 | 0.0020736 | 0.0012247 | 0.0024833 | 0.001169 | 0.0008165 | 0.00044609 | 0.0012201 | 0.00091579 | 0.0040825 | 0 | 0.00064395 | 0.00055648 | 0.0010328 | 0.00050761 | 0.0010572 | 0.0006892
| Rel.Std.Dev. |  | 1.427% | 0.5339% | 4.187% | 1.192% | 0.4804% | 1.094% | 1.969% | 1.516% | 3.266% | 1.842% | 3.661% | 1.767% | 1.25% | 0.6693% | 1.877% | 1.329% | 6.62% | 0% | 1.007% | 0.821% | 1.581% | 0.776% | 1.585% | 0.9592%
| PDM3 |  | -1.191% | -4.317% | -5.75% | -3.458% | -1.66% | 5.203% | -2.184% | -0.9106% | -2.948% | 1.637% | 3.675% | 1.127% | -0.1464% | 1.866% | -0.6562% | 5.356% | -5.75% | 6.986% | -2.235% | 3.598% | -0.1464% | -0.01903% | 1.968% | 9.813%
| Z-Score (Absolute) |  | 0.34 | 1.34 | 1.8 | 1.06 | 0.49 | 1.72 | 0.65 | 0.25 | 0.9 | 0.57 | 1.23 | 0.41 | 0 | 0.65 | 0.16 | 1.77 | 1.8 | 2.29 | 0.67 | 1.2 | 0 | 0.04 | 0.68 | 3.2
| Analytical results for Pb in OREAS 239b (Certified Value 23.2 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 24.2 | 23 | 22 | 22.5453 | 23 | 23 | 24 | 22 | 24.2 | 22.8 | 24.4 | 23.5 | 23.21 | 24.8888 | 22.2 | 23.29 | 23.9 | 21.43 | 24.7 | 22.7 | 23.2 | 21.3 | 20 | 23.1276 | 23.9557 |  | 1
| 1 | 2 | 24.7 | 23 | 21 | 22.198 | 23 | 21.7 | 23 | 22.5 | 23.5 | 22.3 | 23.2 | 24.3 | 22.42 | 24.9943 | 25.1 | 24.23 | 24.2 | 20.4 | 24.4 | 21.1 | 24.04 | 24.7 | 21 | 22.8292 | 24.1549 |  | 35
| 1 | 3 | 24.4 | 23 | 22 | 22.258 | 23 | 20.1 | 24 | 21.4 | 24 | 23.7 | 23 | 22.6 | 23.58 | 24.6809 | 22.9 | 24.9 | 23.2 | 22.09 | 25.2 | 21.6 | 23.64 | 21.8 | 19 | 23.9733 | 23.4567 |  | 16
| 1 | 4 | 24.4 | 23 | 21 | 23.3 | 24 | 23.5 | 23 | 22.3 | 23.8 | 23.7 | 23.5 | 23.9 | 24.36 | 24.5992 | 23.3 | 22.97 | 25.4 | 21.62 | 25.8 | 23.4 | 23.72 | 22 | 21 | 22.7297 | 24.2205 |  | 23.1693
| 1 | 5 | 24.2 | 23 | 21 | 21.89 | 23 | 19.7 | 23 | 21.1 | 24.2 | 22.8 | 23.3 | 24.3 | 23.91 | 24.8142 | 22.6 | 24.4 | 24.5 | 20.56 | 25.9 | 20.7 | 23.12 | 18.3 | 19 | 22.8789 | 23.4288 |  | 49
| 1 | 6 | 24.9 | 24 | 22 | 23.0287 | 23 | 20.6 | 25 | 21.7 | 26.4 | 22.9 | 23.5 | 23.7 | 24.77 | 24.6991 | 22.1 | 22.94 | 24 | 21.05 | 24.4 | 23.5 | 22.89 | 21.6 | 21 | 23.6748 | 23.9718
| Mean |  | 24.4667 | 23.1667 | 21.5 | 22.5367 | 23.1667 | 21.4333 | 23.6667 | 21.8333 | 24.35 | 23.0333 | 23.4833 | 23.7167 | 23.7083 | 24.7794 | 23.0333 | 23.7883 | 24.2 | 21.1917 | 25.0667 | 22.1667 | 23.435 | 21.6167 | 20.1667 | 23.2022 | 23.8647
| Median |  | 24.4 | 23 | 21.5 | 22.4017 | 23 | 21.15 | 23.5 | 21.85 | 24.1 | 22.85 | 23.4 | 23.8 | 23.745 | 24.7566 | 22.75 | 23.76 | 24.1 | 21.24 | 24.95 | 22.15 | 23.42 | 21.7 | 20.5 | 23.0033 | 23.9637
| Std Dev. |  | 0.28048 | 0.40825 | 0.54772 | 0.53579 | 0.40825 | 1.5667 | 0.8165 | 0.53541 | 1.0387 | 0.55737 | 0.48751 | 0.63377 | 0.8387 | 0.14693 | 1.1057 | 0.82978 | 0.72938 | 0.64701 | 0.67429 | 1.1994 | 0.43376 | 2.039 | 0.98319 | 0.50802 | 0.34267
| Rel.Std.Dev. |  | 1.146% | 1.762% | 2.548% | 2.377% | 1.762% | 7.31% | 3.45% | 2.452% | 4.266% | 2.42% | 2.076% | 2.672% | 3.538% | 0.593% | 4.801% | 3.488% | 3.014% | 3.053% | 2.69% | 5.411% | 1.851% | 9.433% | 4.875% | 2.19% | 1.436%
| PDM3 |  | 5.6% | -0.01134% | -7.205% | -2.73% | -0.01134% | -7.493% | 2.147% | -5.766% | 5.096% | -0.5868% | 1.355% | 2.362% | 2.327% | 6.949% | -0.5868% | 2.672% | 4.449% | -8.536% | 8.189% | -4.327% | 1.147% | -6.701% | -12.96% | 0.1422% | 3.002%
| Z-Score (Absolute) |  | 1.28 | 0.04 | 1.72 | 0.67 | 0.04 | 1.79 | 0.47 | 1.39 | 1.16 | 0.17 | 0.28 | 0.52 | 0.51 | 1.6 | 0.17 | 0.59 | 1.01 | 2.04 | 1.89 | 1.05 | 0.24 | 1.61 | 3.08 | 0 | 0.67
| Analytical results for Pd in OREAS 239b (Indicative Value 11.2 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppb
|  |  | 4A*MS | [25 empty] | 1
1
| 1 | 1 | 1.1 | [25 empty] | 1
| 1 | 2 | < 26.8 | [25 empty] | 1
| 1 | 3 | < 20.3 | [25 empty] | 16
| 1 | 4 | < 15 | [25 empty] | 11.15
| 1 | 5 | 23.6 | [25 empty] | 10
| Mean |  | 12.35
| Median |  | 12.35
| Std Dev. |  | 15.9099
| Rel.Std.Dev. |  | 128.8%
| PDM3 |  | 10.76%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 239b (Certified Value 9.04 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 8.7771 | 9.05 | 8.3 | 8.3 | 9.6 | 8.916 | 9.54 | 9.3 | 7.36 | 8.6038 | [16 empty] | 1
| 1 | 2 | 8.4855 | 9.3 | 8.1 | 7.9 | 9.7 | 8.851 | 9.29 | 8.9 | 7.83 | 8.513 | [16 empty] | 11
| 1 | 3 | 8.6218 | 9.15 | 6.5 | 8.4 | 9.8 | 8.713 | 9.62 | 9.3 | 7.79 | 8.6565 | [16 empty] | 16
| 1 | 4 | 8.5311 | 9.4 | 8.5 | 9.9 | 9.6 | 8.827 | 9.77 | 9.1 | 7.65 | 8.8982 | [16 empty] | 9.0412
| 1 | 5 | 8.4912 | 9.5 | 9.2 | 10 | 9.6 | 8.835 | 9.9 | 9.1 | 7.51 | 8.8213 | [16 empty] | 50
| 1 | 6 | 8.6811 | 9.3 | 8.8 | 11.3 | 9.7 | 8.937 | 9.71 | 9.2 | 7.59 | 8.7576
| Mean |  | 8.598 | 9.2833 | 8.2333 | 9.3 | 9.6667 | 8.8465 | 9.6383 | 9.15 | 7.6217 | 8.7084
| Median |  | 8.5764 | 9.3 | 8.4 | 9.15 | 9.65 | 8.843 | 9.665 | 9.15 | 7.62 | 8.7071
| Std Dev. |  | 0.11669 | 0.1633 | 0.9331 | 1.313 | 0.08165 | 0.079216 | 0.21085 | 0.15166 | 0.17577 | 0.14354
| Rel.Std.Dev. |  | 1.357% | 1.759% | 11.33% | 14.12% | 0.8447% | 0.8954% | 2.188% | 1.657% | 2.306% | 1.648%
| PDM3 |  | -4.903% | 2.678% | -8.936% | 2.862% | 6.917% | -2.154% | 6.604% | 1.203% | -15.7% | -3.681%
| Z-Score (Absolute) |  | 0.77 | 0.55 | 1.47 | 0.58 | 1.28 | 0.29 | 1.23 | 0.29 | 2.64 | 0.56
| Analytical results for Pt in OREAS 239b (Indicative Value 1.47 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppb
|  |  | 4A*MS | [25 empty] | 2
2
| 1 | 1 | 1.6 | [25 empty] | 1
| 1 | 2 | 2.3 | [25 empty] | 1
| 1 | 3 | < 3.1 | [25 empty] | 16
| 1 | 4 | 0.2 | [25 empty] | 1.4667
| 1 | 5 | < 1.7 | [25 empty] | 11
| 1 | 6 | 2.3
| Mean |  | 1.6
| Median |  | 1.95
| Std Dev. |  | 0.98995
| Rel.Std.Dev. |  | 61.87%
| PDM3 |  | 9.091%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 239b (Certified Value 154 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 161.95 | 140.5032 | 155 | 102 | 161 | 160.9 | 136 | 155.7 | 151.5 | 158.5 | 158 | 164.5 | 158.6 | 146.9777 | 153 | 141.6 | 160 | 134.29 | 151 | 155.2 | 161.346 | 124.8288 |  |  |  |  | 1
| 1 | 2 | 163.82 | 139.2741 | 156 | 96.3 | 162.1 | 157.8 | 131 | 156.6 | 161.5 | 153.5 | 146.5 | 163 | 176.2 | 144.3444 | 151 | 142.8 | 176 | 156.91 | 143.2 | 150.7 | 158.9032 | 135.4844 |  |  |  |  | 37
| 1 | 3 | 162.74 | 138.6763 | 155 | 62.1 | 161.6 | 165.9 | 150 | 156.7 | 159 | 168 | 154.5 | 159.5 | 189.4 | 149.4153 | 152 | 145.4 | 156 | 150.03 | 143 | 155.4 | 161.0906 | 146.6725 |  |  |  |  | 16
| 1 | 4 | 160.35 | 135.5612 | 155 | 97.8 | 162.8 | 165.5 | 150 | 153.6 | 153 | 163.5 | 147 | 158 | 203.1 | 148.4453 | 154 | 143.5 | 168 | 143.93 | 145.5 | 155.5 | 169.0457 | 150.3869 |  |  |  |  | 153.5534
| 1 | 5 | 166.97 | 138.8581 | 157 | 112 | 162.3 | 170.8 | 147 | 156.9 | 157 | 154 | 151 | 165.5 | 192.6 | 146.8426 | 153 | 147 | 166 | 128.19 | 138.6 | 155.1 | 162.4902 | 153.1413 |  |  |  |  | 51
| 1 | 6 | 165.96 | 139.1807 | 158 | 112 | 161.5 | 169.7 | 138 | 159.8 | 160 | 152.5 | 150 | 160.5 | 225.9 | 141.3277 | 153 | 144.5 | 172 | 163.44 | 146.2 | 154.2 | 159.9822 | 151.7361
| Mean |  | 163.6317 | 138.6756 | 156 | 97.0333 | 161.8833 | 165.1 | 142 | 156.55 | 157 | 158.3333 | 151.1667 | 161.8333 | 190.9667 | 146.2255 | 152.6667 | 144.1333 | 166.3333 | 146.1317 | 144.5833 | 154.35 | 162.143 | 143.7083
| Median |  | 163.28 | 139.0194 | 155.5 | 99.9 | 161.85 | 165.7 | 142.5 | 156.65 | 158 | 156.25 | 150.5 | 161.75 | 191 | 146.9101 | 153 | 144 | 167 | 146.98 | 144.35 | 155.15 | 161.2183 | 148.5297
| Std Dev. |  | 2.4897 | 1.6545 | 1.2649 | 18.4072 | 0.64317 | 5.0076 | 8.0747 | 2.0067 | 3.9875 | 6.2663 | 4.4347 | 2.9609 | 22.971 | 2.9521 | 1.0328 | 1.9263 | 7.4207 | 13.4009 | 4.1194 | 1.8469 | 3.5965 | 11.2329
| Rel.Std.Dev. |  | 1.522% | 1.193% | 0.8108% | 18.97% | 0.3973% | 3.033% | 5.686% | 1.282% | 2.54% | 3.958% | 2.934% | 1.83% | 12.03% | 2.019% | 0.6765% | 1.336% | 4.461% | 9.17% | 2.849% | 1.197% | 2.218% | 7.816%
| PDM3 |  | 6.563% | -9.689% | 1.593% | -36.81% | 5.425% | 7.52% | -7.524% | 1.951% | 2.245% | 3.113% | -1.554% | 5.392% | 24.36% | -4.772% | -0.5775% | -6.135% | 8.323% | -4.833% | -5.842% | 0.5188% | 5.594% | -6.412%
| Z-Score (Absolute) |  | 0.66 | 1.28 | 0.06 | 4.5 | 0.52 | 0.77 | 1.02 | 0.11 | 0.14 | 0.24 | 0.31 | 0.52 | 2.77 | 0.69 | 0.19 | 0.86 | 0.86 | 0.7 | 0.82 | 0.06 | 0.54 | 0.89
| Analytical results for Re in OREAS 239b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.005 | 0.001 | < 0.002 | < 0.002 | 0.0009 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.005 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.005 | < 0.001 | < 0.002 | < 0.002 | 0.0009 | [10 empty] | 38
| 1 | 3 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.005 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.003 | 0.001 | < 0.002 | < 0.002 | 0.0008 | [10 empty] | 16
| 1 | 4 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.006 | 0.002 | < 0.002 | < 0.002 | 0.0007 | [10 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.004 | 0.001 | < 0.002 | < 0.002 | 0.0011 | [10 empty] | 52
| 1 | 6 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.005 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.008 | < 0.001 | < 0.002 | < 0.002 | 0.0006
| Mean | [5 empty] | 0.0045 |  |  |  |  | 0.002 |  | 0.0051667 | 0.00125 |  |  | 0.00083333
| Median | [5 empty] | 0.0045 |  |  |  |  | 0.002 |  | 0.005 | 0.001 |  |  | 0.00085
| Std Dev. | [5 empty] | 0.00054772 |  |  |  |  | 0 |  | 0.0017224 | 0.0005 |  |  | 0.00017512
| Rel.Std.Dev. | [5 empty] | 12.17% |  |  |  |  | 0% |  | 33.34% | 40% |  |  | 21.01%
PDM3
| Z-Score (Absolute) |  | 0 | 99.12 | 1.01 | 48.55 | 7.08 | 48.55 | 3.03 | 0 | 0 | 1.01 | 0 | 8.43 | 0 | 0 | 0 | 0.34
| Analytical results for Rh in OREAS 239b (Indicative Value 0.001 ppm) | [27 empty] | Indicative
Rh | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0001 | [25 empty] | 1
| 1 | 2 | 0.0014 | [25 empty] | 6
| 1 | 3 | 0.0015 | [25 empty] | 16
| 1 | 4 | 0.0009 | [25 empty] | 0.00098333
| 1 | 5 | 0.0015 | [25 empty] | 12
| 1 | 6 | 0.0005
| Mean |  | 0.00098333
| Median |  | 0.00115
| Std Dev. |  | 0.00058793
| Rel.Std.Dev. |  | 59.79%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 239b (Indicative Value 0.003 ppm) | [27 empty] | Indicative
Ru | Round | Replicate | Lab | [25 empty] | 1
| No | No | 1.30 | [25 empty] | ppm
|  |  | 4A*MS | [25 empty] | 3
3
| 1 | 1 | 0.0061 | [25 empty] | 1
| 1 | 2 | < 0.0012 | [25 empty] | 14
| 1 | 3 | < 0.002 | [25 empty] | 16
| 1 | 4 | 0.0029 | [25 empty] | 0.0032333
| 1 | 5 | < 0.0004 | [25 empty] | 9
| 1 | 6 | 0.0086
| Mean |  | 0.0058667
| Median |  | 0.0061
| Std Dev. |  | 0.0028572
| Rel.Std.Dev. |  | 48.7%
| PDM3 |  | 81.44%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 239b (Certified Value 0.208 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.19 | 0.21 | 0.19 | 0.215 | 0.211 | 0.17 | 0.209 | 0.2 | 0.21 | 0.22 | 0.2 | 0.21 | 0.2132 | 0.208 | 0.21 | 0.19 | 0.2 | 0.21 | 0.2 | 0.19598 | 0.2215 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.215 | 0.2 | 0.2 | 0.21 | 0.17 | 0.211 | 0.21 | 0.21 | 0.21 | 0.2 | 0.21 | 0.2266 | 0.214 | 0.21 | 0.19 | 0.21 | 0.21 | 0.2 | 0.19826 | 0.2046 | [5 empty] | 23
| 1 | 3 | 0.21 | 0.215 | 0.19 | 0.205 | 0.216 | 0.17 | 0.211 | 0.2 | 0.21 | 0.22 | 0.2 | 0.21 | 0.2132 | 0.222 | 0.21 | 0.19 | 0.2 | 0.21 | 0.2 | 0.19646 | 0.2171 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.215 | 0.2 | 0.21 | 0.211 | 0.17 | 0.204 | 0.21 | 0.21 | 0.22 | 0.2 | 0.21 | 0.2225 | 0.209 | 0.21 | 0.19 | 0.21 | 0.21 | 0.2 | 0.20035 | 0.2019 | [5 empty] | 0.20816
| 1 | 5 | 0.22 | 0.215 | 0.2 | 0.215 | 0.212 | 0.17 | 0.212 | 0.2 | 0.21 | 0.22 | 0.2 | 0.21 | 0.2184 | 0.211 | 0.22 | 0.19 | 0.21 | 0.21 | 0.2 | 0.20083 | 0.2014 | [5 empty] | 53
| 1 | 6 | 0.2 | 0.215 | 0.2 | 0.21 | 0.216 | 0.16 | 0.211 | 0.21 | 0.2 | 0.23 | 0.2 | 0.21 | 0.2163 | 0.213 | 0.21 | 0.19 | 0.21 | 0.21 | 0.2 | 0.19874 | 0.1954
| Mean |  | 0.20333 | 0.21417 | 0.19667 | 0.20917 | 0.21267 | 0.16833 | 0.20967 | 0.205 | 0.20833 | 0.22 | 0.2 | 0.21 | 0.21837 | 0.21283 | 0.21167 | 0.19 | 0.20667 | 0.21 | 0.2 | 0.19844 | 0.20698
| Median |  | 0.2 | 0.215 | 0.2 | 0.21 | 0.2115 | 0.17 | 0.211 | 0.205 | 0.21 | 0.22 | 0.2 | 0.21 | 0.21735 | 0.212 | 0.21 | 0.19 | 0.21 | 0.21 | 0.2 | 0.1985 | 0.20325
| Std Dev. |  | 0.010328 | 0.0020412 | 0.005164 | 0.0058452 | 0.0026583 | 0.0040825 | 0.0029439 | 0.0054772 | 0.0040825 | 0.0063246 | 3.0405e-17 | 0 | 0.0053392 | 0.0050365 | 0.0040825 | 3.0405e-17 | 0.005164 | 0 | 3.0405e-17 | 0.0019722 | 0.010099
| Rel.Std.Dev. |  | 5.079% | 0.9531% | 2.626% | 2.795% | 1.25% | 2.425% | 1.404% | 2.672% | 1.96% | 2.875% | 1.52e-14% | 0% | 2.445% | 2.366% | 1.929% | 1.6e-14% | 2.499% | 0% | 1.52e-14% | 0.9938% | 4.879%
| PDM3 |  | -2.32% | 2.884% | -5.523% | 0.4824% | 2.164% | -19.13% | 0.7226% | -1.519% | 0.08203% | 5.687% | -3.921% | 0.8827% | 4.902% | 2.244% | 1.683% | -8.725% | -0.7186% | 0.8827% | -3.921% | -4.671% | -0.5665%
| Z-Score (Absolute) |  | 0.75 | 0.87 | 1.75 | 0.12 | 0.65 | 5.99 | 0.2 | 0.5 | 0 | 1.75 | 1.25 | 0.25 | 1.5 | 0.67 | 0.5 | 2.75 | 0.25 | 0.25 | 1.25 | 1.48 | 0.2
| Analytical results for Sb in OREAS 239b (Certified Value 727 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 773.38 | 53.5 | 707.5278 | 763.1 | 739.04 | 666 | 686.5 | 754 | 818 | 754 | 726 | 745.15 | 739.1012 | 591.6 | 719 | 685.7 | 706.85 | 713.7 | 716.7 | 749.214 | 795 | 371.8896 | 737.1682 |  |  |  | 1
| 1 | 2 | 793.89 | 79.1 | 705.5333 | 764.2 | 725.34 | 656 | 680.8 | 766 | 803 | 723 | 756 | 745.28 | 748.7135 | 593.1 | 728 | 677.9 | 727.87 | 710.6 | 718.25 | 768.581 | 826 | 372.7595 | 747.8283 |  |  |  | 39
| 1 | 3 | 805.61 | 159 | 715.74 | 767.8 | 720.98 | 675 | 698.2 | 748 | 719 | 732 | 730 | 749.38 | 725.0665 | 600.5 | 677 | 690.3 | 712.86 | 703.8 | 732.35 | 731.345 | 860 | 366.3633 | 739.8392 |  |  |  | 16
| 1 | 4 | 785.85 | 61.9 | 716.87 | 776 | 739.21 | 650 | 686.2 | 747 | 845 | 740 | 722 | 749.96 | 709.8273 | 587 | 698 | 684.8 | 762.4 | 701.7 | 723.96 | 746.218 | 818 | 370.7895 | 743.8774 |  |  |  | 727.3939
| 1 | 5 | 811 | 26.7 | 709.782 | 771.1 | 734.91 | 661 | 689.1 | 752 | 828 | 725 | 719 | 743.6 | 738.2687 | 598.6 | 714 | 677.9 | 747.89 | 697 | 715.06 | 763.552 | 814 | 376.0088 | 736.6654 |  |  |  | 54
| 1 | 6 | 813.65 | 24.8 | 708.5841 | 776 | 740.91 | 678 | 687.6 | 728 | 872 | 726 | 721 | 748.57 | 728.2434 | 591.3 | 736 | 686.3 | 744.94 | 692.4 | 716.22 | 755.313 | 826 | 372.5037 | 731.3182
| Mean |  | 797.23 | 67.5 | 710.6729 | 769.7 | 733.3983 | 664.3333 | 688.0667 | 749.1667 | 814.1667 | 733.3333 | 729 | 746.99 | 731.5368 | 593.6833 | 712 | 683.8167 | 733.8017 | 703.2 | 720.4233 | 752.3705 | 823.1667 | 371.7191 | 739.4494
| Median |  | 799.75 | 57.7 | 709.183 | 769.45 | 736.975 | 663.5 | 687.05 | 750 | 823 | 729 | 724 | 746.925 | 733.256 | 592.35 | 716.5 | 685.25 | 736.405 | 702.75 | 717.475 | 752.2635 | 822 | 372.1966 | 738.5037
| Std Dev. |  | 15.7413 | 49.4411 | 4.5944 | 5.6349 | 8.2881 | 10.8566 | 5.7046 | 12.4003 | 52.3122 | 11.8603 | 13.7986 | 2.6394 | 13.5764 | 5.0133 | 21.4756 | 4.9544 | 21.6358 | 8.0337 | 6.6298 | 13.3143 | 21.3393 | 3.1521 | 5.8144
| Rel.Std.Dev. |  | 1.974% | 73.25% | 0.6465% | 0.7321% | 1.13% | 1.634% | 0.8291% | 1.655% | 6.425% | 1.617% | 1.893% | 0.3533% | 1.856% | 0.8444% | 3.016% | 0.7245% | 2.948% | 1.142% | 0.9203% | 1.77% | 2.592% | 0.848% | 0.7863%
| PDM3 |  | 9.601% | -90.72% | -2.299% | 5.816% | 0.8255% | -8.669% | -5.407% | 2.993% | 11.93% | 0.8165% | 0.2208% | 2.694% | 0.5696% | -18.38% | -2.116% | -5.991% | 0.8809% | -3.326% | -0.9583% | 3.434% | 13.17% | -48.9% | 1.657%
| Z-Score (Absolute) |  | 2.12 | 21.46 | 0.67 | 1.23 | 0.06 | 2.17 | 1.4 | 0.57 | 2.67 | 0.06 | 0.08 | 0.5 | 0 | 4.46 | 0.63 | 1.54 | 0.07 | 0.92 | 0.36 | 0.67 | 2.96 | 11.63 | 0.26
| Analytical results for Sc in OREAS 239b (Certified Value 14 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.2 | 13.1795 | 15 | 13 | 12.4247 | 14 | 10.4 | 14 | 14.3 | 14.3 | 15.2 | 14.9 | 14.9 | 13.3 | 14.1581 | 13.5 | 11.64 | 14.9 | 14.4 | 12.9 | 14.4 | 14.3325 | 14.8273 | 10.1171 |  |  | 1
| 1 | 2 | 14.7 | 12.8254 | 16 | 12 | 10.446 | 15 | 10.4 | 14 | 13.3 | 14.9 | 15.1 | 14 | 14.9 | 13.3 | 14.1599 | 13.5 | 11.84 | 14.5 | 14.1 | 12.1 | 13.9 | 14.6475 | 14.8273 | 10.4524 |  |  | 40
| 1 | 3 | 14.3 | 12.7566 | 15 | 10 | 12.172 | 14 | 11.6 | 14 | 13.3 | 14.8 | 15 | 14.7 | 14.7 | 13.7 | 14.13 | 13.4 | 12.19 | 13.7 | 14.6 | 12.2 | 14.6 | 14.301 | 14.4845 | 10.5687 |  |  | 16
| 1 | 4 | 14 | 12.7856 | 16 | 13 | 10.8513 | 15 | 11.7 | 14 | 13.1 | 14.6 | 15.8 | 14 | 14.6 | 13.5 | 14.3862 | 13.7 | 11.98 | 14.4 | 14.5 | 13.1 | 14.6 | 14.5215 | 14.7702 | 10.8458 |  |  | 14.0227
| 1 | 5 | 14.6 | 13.0594 | 16 | 14 | 11.742 | 14 | 11.8 | 14 | 13.3 | 14.7 | 15.4 | 14.2 | 14.9 | 13.4 | 13.9686 | 13.6 | 12.38 | 14.3 | 14.1 | 11.6 | 14.4 | 14.49 | 14.8845 | 10.9351 |  |  | 55
| 1 | 6 | 14.5 | 12.8247 | 16 | 13 | 9.516 | 14 | 11 | 14 | 13.4 | 15 | 15.1 | 13.8 | 14.7 | 12.3 | 13.893 | 13.5 | 12.11 | 13.8 | 15 | 13.1 | 14.9 | 14.4795 | 14.5416 | 10.6997
| Mean |  | 14.3833 | 12.9052 | 15.6667 | 12.5 | 11.192 | 14.3333 | 11.15 | 14 | 13.45 | 14.7167 | 15.2667 | 14.2667 | 14.7833 | 13.25 | 14.116 | 13.5333 | 12.0233 | 14.2667 | 14.45 | 12.5 | 14.4667 | 14.462 | 14.7226 | 10.6031
| Median |  | 14.4 | 12.8251 | 16 | 13 | 11.2967 | 14 | 11.3 | 14 | 13.3 | 14.75 | 15.15 | 14.1 | 14.8 | 13.35 | 14.1441 | 13.5 | 12.045 | 14.35 | 14.45 | 12.55 | 14.5 | 14.4848 | 14.7988 | 10.6342
| Std Dev. |  | 0.26394 | 0.17218 | 0.5164 | 1.3784 | 1.1193 | 0.5164 | 0.6442 | 0 | 0.42778 | 0.24833 | 0.29439 | 0.43665 | 0.13292 | 0.48888 | 0.17229 | 0.10328 | 0.26265 | 0.45019 | 0.33912 | 0.6229 | 0.33267 | 0.12785 | 0.16724 | 0.29618
| Rel.Std.Dev. |  | 1.835% | 1.334% | 3.296% | 11.03% | 10% | 3.603% | 5.778% | 0% | 3.181% | 1.687% | 1.928% | 3.061% | 0.8991% | 3.69% | 1.221% | 0.7631% | 2.185% | 3.156% | 2.347% | 4.983% | 2.3% | 0.8841% | 1.136% | 2.793%
| PDM3 |  | 2.572% | -7.969% | 11.72% | -10.86% | -20.19% | 2.216% | -20.49% | -0.1616% | -4.084% | 4.949% | 8.871% | 1.74% | 5.425% | -5.51% | 0.6655% | -3.49% | -14.26% | 1.74% | 3.048% | -10.86% | 3.166% | 3.133% | 4.991% | -24.39%
| Z-Score (Absolute) |  | 0.21 | 1.39 | 1.59 | 1.82 | 3.24 | 0.15 | 3.28 | 0.21 | 0.8 | 0.57 | 1.16 | 0.08 | 0.64 | 1.02 | 0.08 | 0.71 | 2.34 | 0.08 | 0.28 | 1.82 | 0.3 | 0.29 | 0.57 | 3.87
| Analytical results for Se in OREAS 239b (Indicative Value 0.69 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | < 0.5 | < 5 | 0.9 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | 1 | < 1 | 1 | < 0.1 | < 2 | < 0.5 | < 1 | 0.5 | 0.2438 | 3.3068 | [7 empty] | 1
| 1 | 2 | < 0.5 | < 5 | 1.2 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.8 | < 0.1 | < 2 | < 0.5 | < 1 | 0.4 | 0.2545 | 3.4154 | [7 empty] | 41
| 1 | 3 | < 0.5 | < 5 | 0.7 | < 5 | < 2 | < 1 | 2 | < 1 | < 1 | < 1 | < 1 | 0.9 | < 0.1 | < 2 | < 0.5 | < 1 | < 0.3 | 0.2438 | 3.3535 | [7 empty] | 16
| 1 | 4 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | 1 | 1 | 1 | < 1 | < 1 | 1 | < 0.1 | < 2 | < 0.5 | < 1 | < 0.3 | 0.2652 | 3.3068 | [7 empty] | 0.69457
| 1 | 5 | < 0.5 | < 5 | 1.4 | < 5 | < 2 | < 1 | < 1 | 1 | 1 | < 1 | < 1 | 1.1 | < 0.1 | < 2 | 1.3 | < 1 | < 0.3 | 0.2723 | 3.3973 | [7 empty] | 10
| 1 | 6 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | 2 | < 1 | 1 | < 1 | < 1 | 1.4 | < 0.1 | < 2 | 0.9 | < 1 | 0.3 | 0.1974 | 3.2343
| Mean |  |  |  | 1.0667 |  |  |  | 1.6667 | 1 | 1 | 1 |  | 1.0333 |  |  | 1.1 |  | 0.4 | 0.24617 | 3.3357
| Median |  |  |  | 1.1 |  |  |  | 2 | 1 | 1 | 1 |  | 1 |  |  | 1.1 |  | 0.4 | 0.24915 | 3.3301
| Std Dev. |  |  |  | 0.24221 |  |  |  | 0.57735 | 0 | 0 |  |  | 0.20656 |  |  | 0.28284 |  | 0.1 | 0.026473 | 0.066957
| Rel.Std.Dev. |  |  |  | 22.71% |  |  |  | 34.64% | 0% | 0% |  |  | 19.99% |  |  | 25.71% |  | 25% | 10.75% | 2.007%
| PDM3 |  |  |  | 53.57% |  |  |  | 140% | 43.97% | 43.97% | 43.97% |  | 48.77% |  |  | 58.37% |  | -42.41% | -64.56% | 380.2%
| Z-Score (Absolute) |  | 0.77 | 3.37 | 0.74 | 3.37 | 0.61 | 0.31 | 0.77 | 0 | 0.46 | 0.15 | 0.31 | 0.67 | 1.13 | 0.61 | 0.25 | 0.31 | 0.72 | 0.77 | 4.91
| Analytical results for Sm in OREAS 239b (Certified Value 6.31 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.1291 | 6.7 | 6.1 | 5.5 | 6.7 | 5.99 | 6.59 | 6.2 | 7.3508 | 6.0811 | [16 empty] | 1
| 1 | 2 | 6.211 | 6.65 | 6.1 | 5.2 | 6.8 | 5.957 | 6.59 | 6.1 | 7.3226 | 6.1432 | [16 empty] | 10
| 1 | 3 | 6.0968 | 6.55 | 5 | 5.6 | 6.8 | 5.819 | 6.71 | 6.2 | 7.2662 | 6.1989 | [16 empty] | 16
| 1 | 4 | 6.2531 | 6.65 | 5.6 | 6.1 | 6.7 | 5.901 | 6.77 | 6.3 | 7.2192 | 6.2394 | [16 empty] | 6.3094
| 1 | 5 | 6.2627 | 6.65 | 6.5 | 6.2 | 6.8 | 5.985 | 6.64 | 6.2 | 7.2004 | 6.3652 | [16 empty] | 56
| 1 | 6 | 6.4754 | 6.7 | 6.6 | 7.1 | 6.6 | 6.016 | 6.74 | 6.2 | 7.4824 | 6.2622
| Mean |  | 6.238 | 6.65 | 5.9833 | 5.95 | 6.7333 | 5.9447 | 6.6733 | 6.2 | 7.3069 | 6.215
| Median |  | 6.232 | 6.65 | 6.1 | 5.85 | 6.75 | 5.971 | 6.675 | 6.2 | 7.2944 | 6.2191
| Std Dev. |  | 0.13397 | 0.054772 | 0.59805 | 0.6775 | 0.08165 | 0.072968 | 0.077632 | 0.063246 | 0.1036 | 0.098718
| Rel.Std.Dev. |  | 2.148% | 0.8236% | 9.995% | 11.39% | 1.213% | 1.227% | 1.163% | 1.02% | 1.418% | 1.588%
| PDM3 |  | -1.131% | 5.399% | -5.168% | -5.696% | 6.72% | -5.78% | 5.769% | -1.733% | 15.81% | -1.496%
| Z-Score (Absolute) |  | 0.03 | 1.02 | 0.59 | 0.67 | 1.22 | 0.68 | 1.08 | 0.06 | 2.61 | 0.03
| Analytical results for Sn in OREAS 239b (Certified Value 3.64 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.8 | 3.5 | 2 | < 10 | 3.8 | 4 | 3.8 | 3.9 | 3.5 | 3.9 | 3.8 | 3.6 | 3.7 | 3.7179 | 3.72 | 3.406 | 3.8 | < 5 | 3.7 | 3.7 | 3.4 | 2.8309 | < 5 | 3.0141 | 3.5559 |  | 1
| 1 | 2 | 3.8 | 3 | 2 | < 10 | 3.7 | 3.7 | 3.1 | 4 | 3.6 | 3.8 | 3.6 | 3.6 | 3.4 | 3.8193 | 3.5 | 3.472 | 3.9 | < 5 | 3.8 | 3.5 | 3.3 | 2.7776 | < 5 | 3.2905 | 3.6168 |  | 26
| 1 | 3 | 3.8 | 2.9 | 2 | < 10 | 3.7 | 3.7 | 3.7 | 4.1 | 3.6 | 3.6 | 3.8 | 3.4 | 3.7 | 3.7263 | 3.49 | 3.761 | 3.7 | < 5 | 3.6 | 3.5 | 3.5 | 2.8755 | < 5 | 3.6002 | 3.3824 |  | 16
| 1 | 4 | 3.8 | 2.9 | 2 | < 10 | 3.8 | 4.1 | 2.8 | 4 | 3.6 | 4 | 3.6 | 3.5 | 3.9 | 3.7183 | 3.46 | 3.437 | 3.6 | < 5 | 3.8 | 3.6 | 3.4 | 3.0857 | < 5 | 3.346 | 3.6704 |  | 3.6443
| 1 | 5 | 4 | 3.4 | 1 | < 10 | 3.7 | 3.9 | 3.7 | 4 | 3.6 | 4 | 3.7 | 3.5 | 4 | 3.8048 | 3.48 | 3.531 | 3.7 | < 5 | 3.8 | 3.4 | 3.4 | 2.8618 | < 5 | 3.265 | 3.6282 |  | 57
| 1 | 6 | 3.9 | 3.6 | 1 | < 10 | 3.8 | 3.9 | 4 | 3.9 | 3.6 | 3.8 | 3.6 | 3.5 | 3.9 | 3.6392 | 3.65 | 3.514 | 3.8 | < 5 | 3.9 | 3.6 | 3.4 | 2.9162 | < 5 | 3.3394 | 3.7343
| Mean |  | 3.85 | 3.2167 | 1.6667 |  | 3.75 | 3.8833 | 3.5167 | 3.9833 | 3.5833 | 3.85 | 3.6833 | 3.5167 | 3.7667 | 3.7376 | 3.55 | 3.5202 | 3.75 |  | 3.7667 | 3.55 | 3.4 | 2.8913 |  | 3.3092 | 3.598
| Median |  | 3.8 | 3.2 | 2 |  | 3.75 | 3.9 | 3.7 | 4 | 3.6 | 3.85 | 3.65 | 3.5 | 3.8 | 3.7223 | 3.495 | 3.493 | 3.75 |  | 3.8 | 3.55 | 3.4 | 2.8687 |  | 3.315 | 3.6225
| Std Dev. |  | 0.083666 | 0.31885 | 0.5164 |  | 0.054772 | 0.16021 | 0.46224 | 0.075277 | 0.040825 | 0.15166 | 0.098319 | 0.075277 | 0.21602 | 0.065979 | 0.1077 | 0.12683 | 0.10488 |  | 0.10328 | 0.10488 | 0.063246 | 0.10593 |  | 0.18779 | 0.12112
| Rel.Std.Dev. |  | 2.173% | 9.913% | 30.98% |  | 1.461% | 4.126% | 13.14% | 1.89% | 1.139% | 3.939% | 2.669% | 2.141% | 5.735% | 1.765% | 3.034% | 3.603% | 2.797% |  | 2.742% | 2.954% | 1.86% | 3.664% |  | 5.675% | 3.366%
| PDM3 |  | 5.645% | -11.73% | -54.27% |  | 2.9% | 6.559% | -3.502% | 9.303% | -1.673% | 5.645% | 1.071% | -3.502% | 3.358% | 2.561% | -2.588% | -3.406% | 2.9% |  | 3.358% | -2.588% | -6.704% | -20.66% |  | -9.195% | -1.271%
| Z-Score (Absolute) |  | 0.99 | 1.43 | NA | 5.4 | 0.61 | 1.12 | 0.28 | 1.5 | 0.03 | 0.99 | 0.36 | 0.28 | 0.67 | 0.56 | 0.16 | 0.27 | 0.61 | 4.18 | 0.67 | 0.16 | 0.73 | 2.68 | 4.18 | 1.08 | 0.03
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 239b (Certified Value 121 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 122.43 | 125.342 | 119 | 118 | 124.76 | 122.7 | 118.8 | 117 | 119 | 122 | 123 | 125 | 120.5 | 126 | 116.3076 | 119 | 110.7 | 116 | 117.8 | 121.1 | 123.5 | 117 | 139.2 | 132 | 120.6866 | 113.0015 | 1
| 1 | 2 | 121.91 | 123.48 | 124 | 117 | 122.2308 | 123.2 | 116.7 | 117 | 119 | 124.5 | 119 | 118 | 122 | 125 | 117.901 | 122 | 112.1 | 118 | 118.5 | 120.7 | 124 | 116 | 143.1 | 135 | 119.4249 | 114.7202 | 27
| 1 | 3 | 122.32 | 123.4408 | 118 | 113 | 121.8497 | 121.5 | 128.3 | 121 | 119 | 122 | 125.5 | 121 | 119.5 | 126 | 118.6163 | 143 | 114.5 | 117 | 116.8 | 120.3 | 122.9 | 119 | 138 | 133 | 118.1631 | 115.3226 | 16
| 1 | 4 | 122.57 | 123.137 | 124 | 119 | 121.9263 | 122.7 | 126.5 | 120 | 120 | 122 | 125.5 | 118 | 119.5 | 126 | 116.0439 | 121 | 112.5 | 118 | 116.8 | 120.8 | 121.4 | 119 | 141.3 | 133 | 119.1725 | 114.7251 | 120.6895
| 1 | 5 | 122.3 | 124.4306 | 120 | 126 | 121.3844 | 122.4 | 128.8 | 113 | 117 | 121.5 | 121.5 | 119.5 | 120 | 125 | 116.5562 | 123 | 114.7 | 119 | 119 | 124.1 | 123.1 | 120 | 141.2 | 133 | 121.1409 | 115.6434 | 58
| 1 | 6 | 122.3 | 125.734 | 120 | 124 | 120.7482 | 121.5 | 123.6 | 121 | 121 | 121.5 | 126 | 120.5 | 119.5 | 114 | 115.3292 | 124 | 113.5 | 121 | 117 | 124.4 | 123.6 | 118 | 141 | 135 | 118.6678 | 113.54
| Mean |  | 122.305 | 124.2607 | 120.8333 | 119.5 | 122.1499 | 122.3333 | 123.7833 | 118.1667 | 119.1667 | 122.25 | 123.4167 | 120.3333 | 120.1667 | 123.6667 | 116.7924 | 125.3333 | 113 | 118.1667 | 117.65 | 121.9 | 123.0833 | 118.1667 | 140.6333 | 133.5 | 119.5426 | 114.4921
| Median |  | 122.31 | 123.9553 | 120 | 118.5 | 121.888 | 122.55 | 125.05 | 118.5 | 119 | 122 | 124.25 | 120 | 119.75 | 125.5 | 116.4319 | 122.5 | 113 | 118 | 117.4 | 120.95 | 123.3 | 118.5 | 141.1 | 133 | 119.2987 | 114.7227
| Std Dev. |  | 0.22007 | 1.0875 | 2.5626 | 4.7645 | 1.3791 | 0.69474 | 5.0594 | 3.1252 | 1.3292 | 1.1292 | 2.7824 | 2.6013 | 0.98319 | 4.761 | 1.2287 | 8.8242 | 1.5323 | 1.7224 | 0.94181 | 1.8407 | 0.91086 | 1.472 | 1.7874 | 1.2247 | 1.1561 | 1.0247
| Rel.Std.Dev. |  | 0.1799% | 0.8752% | 2.121% | 3.987% | 1.129% | 0.5679% | 4.087% | 2.645% | 1.115% | 0.9236% | 2.254% | 2.162% | 0.8182% | 3.85% | 1.052% | 7.041% | 1.356% | 1.458% | 0.8005% | 1.51% | 0.74% | 1.246% | 1.271% | 0.9174% | 0.9671% | 0.895%
| PDM3 |  | 1.339% | 2.959% | 0.1191% | -0.9856% | 1.21% | 1.362% | 2.563% | -2.09% | -1.262% | 1.293% | 2.26% | -0.2951% | -0.4332% | 2.467% | -3.229% | 3.848% | -6.371% | -2.09% | -2.518% | 1.003% | 1.983% | -2.09% | 16.52% | 10.61% | -0.9503% | -5.135%
| Z-Score (Absolute) |  | 0.28 | 0.87 | 0.16 | 0.56 | 0.23 | 0.29 | 0.72 | 0.96 | 0.66 | 0.26 | 0.61 | 0.31 | 0.36 | 0.69 | 1.37 | 1.19 | 2.51 | 0.96 | 1.11 | 0.16 | 0.51 | 0.96 | 5.77 | 3.64 | 0.55 | 2.06
| Analytical results for Ta in OREAS 239b (Certified Value 1.05 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 1.06 | 0.90375 | 0.2 | < 5 | 1.05 | 1.02 | 0.99 | 0.6 | 1.05 | 1.09 | 1.06 | 1.06 | 1 | 1.2304 | 0.96 | 1.03 | 1.18 | 1.1 | 0.8 | 0.5913 | 1.1743 | [5 empty] | 1
| 1 | 2 | 1.08 | 0.90356 | 0.3 | < 5 | 1.03 | 0.87 | 0.85 | 0.7 | 1.07 | 1.02 | 1.02 | 1.05 | 1 | 1.1089 | 0.99 | 0.97 | 1.19 | 1.06 | 0.8 | 0.5928 | 1.1628 | [5 empty] | 28
| 1 | 3 | 1.04 | 0.87124 | 0.5 | < 5 | 1.03 | 0.94 | 0.93 | 0.6 | 1.07 | 1.11 | 1.04 | 1 | 1 | 1.2313 | 0.99 | 0.99 | 1.18 | 1.05 | 0.8 | 0.5785 | 1.1757 | [5 empty] | 16
| 1 | 4 | 1.12 | 0.88966 | 0.2 | 5.2567 | 1.05 | 1.07 | 0.99 | 0.7 | 1.05 | 1.13 | 0.98 | 1.04 | 1.1 | 1.0805 | 1 | 0.99 | 1.16 | 1.08 | 0.8 | 0.6009 | 1.1752 | [5 empty] | 1.0509
| 1 | 5 | 1.11 | 0.87456 | < 0.1 | < 5 | 1.06 | 0.94 | 0.97 | 0.6 | 1.09 | 1.08 | 1.02 | 1.06 | 1.1 | 1.0854 | 0.98 | 1.03 | 1.2 | 1.02 | 0.8 | 0.646 | 1.1735 | [5 empty] | 59
| 1 | 6 | 1.13 | 0.90849 | < 0.1 | < 5 | 1.05 | 0.97 | 1.02 | 0.7 | 0.97 | 1.09 | 1.04 | 1.04 | 1.2 | 1.1849 | 1 | 0.96 | 1.15 | 1.1 | 0.8 | 0.6591 | 1.168
| Mean |  | 1.09 | 0.89188 | 0.3 | 5.2567 | 1.045 | 0.96833 | 0.95833 | 0.65 | 1.05 | 1.0867 | 1.0267 | 1.0417 | 1.0667 | 1.1536 | 0.98667 | 0.995 | 1.1767 | 1.0683 | 0.8 | 0.61143 | 1.1716
| Median |  | 1.095 | 0.89661 | 0.25 | 5.2567 | 1.05 | 0.955 | 0.98 | 0.65 | 1.06 | 1.09 | 1.03 | 1.045 | 1.05 | 1.1469 | 0.99 | 0.99 | 1.18 | 1.07 | 0.8 | 0.59685 | 1.1739
| Std Dev. |  | 0.035777 | 0.016026 | 0.14142 |  | 0.012247 | 0.06969 | 0.0608 | 0.054772 | 0.041952 | 0.037238 | 0.027325 | 0.022286 | 0.08165 | 0.070606 | 0.015055 | 0.029496 | 0.018619 | 0.031252 | 1.2162e-16 | 0.032909 | 0.0051386
| Rel.Std.Dev. |  | 3.282% | 1.797% | 47.14% |  | 1.172% | 7.197% | 6.344% | 8.427% | 3.995% | 3.427% | 2.662% | 2.139% | 7.655% | 6.121% | 1.526% | 2.964% | 1.582% | 2.925% | 1.52e-14% | 5.382% | 0.4386%
| PDM3 |  | 3.719% | -15.13% | -71.45% | 400.2% | -0.5631% | -7.858% | -8.81% | -38.15% | -0.0873% | 3.402% | -2.308% | -0.8803% | 1.499% | 9.766% | -6.114% | -5.321% | 11.97% | 1.657% | -23.88% | -41.82% | 11.48%
| Z-Score (Absolute) |  | 0.44 | 1.38 | 7.59 | 17.63 | 0.03 | 0.67 | 0.77 | 3.6 | 0.08 | 0.41 | 0.14 | 0 | 0.23 | 1.03 | 0.51 | 0.43 | 1.24 | 0.25 | 2.22 | 3.96 | 1.19
| Analytical results for Tb in OREAS 239b (Certified Value 0.67 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.34 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.6751 | 0.68 | 0.7 | 0.5 | 0.69 | 0.64 | 0.63 | 0.71 | 0.6 | 2.27 | 5.0584 | [15 empty] | 1
| 1 | 2 | 0.68128 | 0.66 | 0.7 | 0.5 | 0.71 | 0.655 | 0.63 | 0.71 | 0.6 | 2.04 | 5.1532 | [15 empty] | 11
| 1 | 3 | 0.66567 | 0.68 | 0.6 | 0.6 | 0.7 | 0.615 | 0.6 | 0.78 | 0.6 | 2.28 | 5.279 | [15 empty] | 16
| 1 | 4 | 0.67732 | 0.68 | 0.7 | 0.6 | 0.69 | 0.658 | 0.64 | 0.74 | 0.6 | 2.21 | 5.1833 | [15 empty] | 0.67045
| 1 | 5 | 0.67123 | 0.7 | 0.7 | 0.6 | 0.67 | 0.641 | 0.6 | 0.71 | 0.6 | 2.38 | 5.1466 | [15 empty] | 60
| 1 | 6 | 0.67371 | 0.68 | 0.7 | 0.9 | 0.68 | 0.637 | 0.63 | 0.71 | 0.6 | 2.38 | 5.1711
| Mean |  | 0.67405 | 0.68 | 0.68333 | 0.61667 | 0.69 | 0.641 | 0.62167 | 0.72667 | 0.6 | 2.26 | 5.1653
| Median |  | 0.6744 | 0.68 | 0.7 | 0.6 | 0.69 | 0.6405 | 0.63 | 0.71 | 0.6 | 2.275 | 5.1621
| Std Dev. |  | 0.0053366 | 0.012649 | 0.040825 | 0.1472 | 0.014142 | 0.015349 | 0.017224 | 0.028752 | 0 | 0.12665 | 0.071028
| Rel.Std.Dev. |  | 0.7917% | 1.86% | 5.974% | 23.87% | 2.05% | 2.395% | 2.771% | 3.957% | 0% | 5.604% | 1.375%
| PDM3 |  | 0.5368% | 1.424% | 1.921% | -8.022% | 2.915% | -4.393% | -7.277% | 8.384% | -10.51% | 237.1% | 670.4%
| Z-Score (Absolute) |  | 0.17 | 0.08 | NA | NA | 0.08 | 0.69 | 1 | 0.66 | NA | 24.79 | 70.53
| Analytical results for Te in OREAS 239b (Indicative Value 0.049 ppm) | [27 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | < 0.1 | 9.9487 | < 0.1 | 0.11 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.06 | 0.06 | < 0.05 | 0.03003 | [8 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.1 | 9.0573 | < 0.1 | 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | 0.03 | 0.06 | 0.06 | 0.04499 | [8 empty] | 5
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 9.3547 | < 0.1 | 0.11 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | 0.05 | 0.05 | < 0.05 | < 0.05 | 0.03 | 0.05 | < 0.05 | 0.02987 | [8 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 9.5273 | < 0.1 | 0.13 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.09 | 0.08 | < 0.05 | 0.0235 | [8 empty] | 0.049326
| 1 | 5 | < 0.2 | < 0.2 | < 0.1 | 9.6033 | < 0.1 | 0.11 | 0.2 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.04 | < 0.05 | < 0.05 | 0.0243 | [8 empty] | 11
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 9.5147 | < 0.1 | 0.09 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.07 | 0.05 | 0.01496
| Mean |  |  |  |  | 9.501 |  | 0.10833 | 0.2 |  |  |  | 0.05 | 0.051667 |  |  | 0.05 | 0.064 | 0.055 | 0.027942
| Median |  |  |  |  | 9.521 |  | 0.11 | 0.2 |  |  |  | 0.05 | 0.05 |  |  | 0.04 | 0.06 | 0.055 | 0.027085
| Std Dev. |  |  |  |  | 0.29326 |  | 0.013292 | [5 empty] | 0.0040825 |  |  | 0.025495 | 0.011402 | 0.0070711 | 0.010002
| Rel.Std.Dev. |  |  |  |  | 3.087% |  | 12.27% | [5 empty] | 7.902% |  |  | 50.99% | 17.82% | 12.86% | 35.8%
| PDM3 |  |  |  |  | 19162% |  | 119.6% | 305.5% |  |  |  | 1.366% | 4.745% |  |  | 1.366% | 29.75% | 11.5% | -43.35%
| Z-Score (Absolute) |  | 1.43 | 1.43 | 0 | 270.85 | 0 | 1.67 | 0.72 | 0 | 0.72 | 0.72 | 0.6 | 0.05 | 0.72 | 0.72 | 0.17 | 0.21 | 0.43 | 0.63
| Analytical results for Th in OREAS 239b (Certified Value 14.2 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.42 | 14.926 | 16 | 12.9 | 14.72 | 11.68 | 15.2 | 13.9 | 14.05 | 13.8 | 14.55 | 14.95 | 13.4 | 14.2862 | 13.4 | 13.69 | 13.9 | 13.7 | 16 | 13.8 | 13.0583 | 14.6137 | 13.7898 |  |  |  | 1
| 1 | 2 | 14.56 | 14.7389 | 16.2 | 11.9 | 14.69 | 11.23 | 14.5 | 13.1 | 14.25 | 15.35 | 13.75 | 15 | 13.9 | 14.8418 | 13.6 | 13.9 | 14.3 | 14 | 14.9 | 13.7 | 12.8356 | 14.6728 | 13.8188 |  |  |  | 31
| 1 | 3 | 14.49 | 14.1882 | 16.7 | 8 | 14.47 | 12.38 | 15.3 | 14 | 13.9 | 14.85 | 13.85 | 14.35 | 14.3 | 14.4884 | 13.5 | 13.89 | 14.2 | 14.1 | 15.1 | 13.8 | 13.1118 | 14.3378 | 14.2932 |  |  |  | 16
| 1 | 4 | 14.56 | 14.7087 | 16.2 | 12.9 | 14.72 | 12.56 | 14.9 | 14 | 13.55 | 16 | 13.25 | 14.75 | 17.2 | 14.2117 | 13.5 | 13.64 | 14.2 | 13.9 | 15.3 | 14.1 | 13.6582 | 14.791 | 14.2785 |  |  |  | 14.2122
| 1 | 5 | 14.78 | 14.5746 | 16.8 | 14.8 | 14.64 | 12.62 | 14.3 | 14 | 13.9 | 13.85 | 13.7 | 15.3 | 17.1 | 14.1885 | 13.3 | 13.78 | 14.1 | 14.1 | 14.6 | 13.8 | 13.1132 | 14.7713 | 14.016 |  |  |  | 61
| 1 | 6 | 14.82 | 14.6533 | 16.2 | 14 | 14.49 | 11.86 | 15.1 | 14.2 | 15.7 | 14.95 | 13.9 | 14.7 | 21.3 | 13.92 | 13.3 | 13.87 | 14 | 14.1 | 15.6 | 13.5 | 13.1211 | 14.0225 | 14.2355
| Mean |  | 14.605 | 14.6316 | 16.35 | 12.4167 | 14.6217 | 12.055 | 14.8833 | 13.8667 | 14.225 | 14.8 | 13.8333 | 14.8417 | 16.2 | 14.3228 | 13.4333 | 13.795 | 14.1167 | 13.9833 | 15.25 | 13.7833 | 13.1497 | 14.5349 | 14.072
| Median |  | 14.56 | 14.681 | 16.2 | 12.9 | 14.665 | 12.12 | 15 | 14 | 13.975 | 14.9 | 13.8 | 14.85 | 15.7 | 14.249 | 13.45 | 13.825 | 14.15 | 14.05 | 15.2 | 13.8 | 13.1125 | 14.6433 | 14.1258
| Std Dev. |  | 0.16022 | 0.24676 | 0.32094 | 2.3845 | 0.11374 | 0.55482 | 0.40208 | 0.38816 | 0.75812 | 0.85674 | 0.42032 | 0.32158 | 2.9853 | 0.3133 | 0.12111 | 0.11041 | 0.1472 | 0.16021 | 0.501 | 0.19408 | 0.27171 | 0.29937 | 0.2303
| Rel.Std.Dev. |  | 1.097% | 1.686% | 1.963% | 19.2% | 0.7779% | 4.602% | 2.702% | 2.799% | 5.33% | 5.789% | 3.038% | 2.167% | 18.43% | 2.187% | 0.9015% | 0.8004% | 1.043% | 1.146% | 3.285% | 1.408% | 2.066% | 2.06% | 1.637%
| PDM3 |  | 2.764% | 2.951% | 15.04% | -12.63% | 2.881% | -15.18% | 4.722% | -2.431% | 0.08987% | 4.136% | -2.666% | 4.429% | 13.99% | 0.7777% | -5.48% | -2.936% | -0.6724% | -1.611% | 7.302% | -3.018% | -7.476% | 2.27% | -0.9869%
| Z-Score (Absolute) |  | 0.6 | 0.64 | 3.33 | 2.84 | 0.62 | 3.4 | 1.03 | 0.56 | 0 | 0.9 | 0.61 | 0.97 | 3.1 | 0.15 | 1.24 | 0.67 | 0.17 | 0.38 | 1.61 | 0.69 | 1.69 | 0.49 | 0.24
| Analytical results for Ti in OREAS 239b (Certified Value 0.46 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.4487 | 0.44297 | 0.46 | 0.437 | 0.34 | 0.491 | 0.465 | 0.439 | 0.466 | 0.463 | 0.456 | 0.451 | 0.462 | 0.4651 | 0.45841 | 0.5108 | 0.45 | 0.45 | 0.45 | 0.5 | 0.441 | 0.36746 | 0.4926 | 0.388 | 0.46899 |  | 1
| 1 | 2 | 0.4521 | 0.44132 | 0.465 | 0.435 | 0.32667 | 0.497 | 0.466 | 0.438 | 0.456 | 0.465 | 0.444 | 0.438 | 0.459 | 0.4529 | 0.45726 | 0.5216 | 0.45 | 0.44 | 0.45 | 0.5 | 0.441 | 0.37778 | 0.5062 | 0.3885 | 0.45451 |  | 32
| 1 | 3 | 0.4424 | 0.43932 | 0.46 | 0.454 | 0.35 | 0.511 | 0.482 | 0.435 | 0.451 | 0.458 | 0.49 | 0.451 | 0.451 | 0.4651 | 0.46356 | 0.6055 | 0.44 | 0.45 | 0.45 | 0.49 | 0.441 | 0.37008 | 0.5042 | 0.3708 | 0.46493 |  | 16
| 1 | 4 | 0.4601 | 0.43823 | 0.455 | 0.445 | 0.34333 | 0.495 | 0.47 | 0.433 | 0.454 | 0.458 | 0.472 | 0.438 | 0.448 | 0.4641 | 0.45433 | 0.5156 | 0.44 | 0.44 | 0.44 | 0.49 | 0.437 | 0.37668 | 0.4949 | 0.396 | 0.45485 |  | 0.45955
| 1 | 5 | 0.4715 | 0.44256 | 0.455 | 0.433 | 0.34667 | 0.504 | 0.472 | 0.463 | 0.45 | 0.45 | 0.457 | 0.44 | 0.449 | 0.4621 | 0.4642 | 0.521 | 0.45 | 0.44 | 0.45 | 0.5 | 0.439 | 0.37633 | 0.4957 | 0.4005 | 0.46317 |  | 62
| 1 | 6 | 0.4683 | 0.4436 | 0.45 | 0.448 | 0.33667 | 0.509 | 0.488 | 0.448 | 0.453 | 0.44 | 0.486 | 0.448 | 0.452 | 0.4661 | 0.45096 | 0.5275 | 0.45 | 0.45 | 0.46 | 0.5 | 0.443 | 0.37237 | 0.5071 | 0.382 | 0.4744
| Mean |  | 0.45718 | 0.44133 | 0.4575 | 0.442 | 0.34056 | 0.50117 | 0.47383 | 0.44267 | 0.455 | 0.45567 | 0.4675 | 0.44433 | 0.4535 | 0.46257 | 0.45812 | 0.53367 | 0.44667 | 0.445 | 0.45 | 0.49667 | 0.44033 | 0.37345 | 0.50012 | 0.38763 | 0.46347
| Median |  | 0.4561 | 0.44194 | 0.4575 | 0.441 | 0.34167 | 0.5005 | 0.471 | 0.4385 | 0.4535 | 0.458 | 0.4645 | 0.444 | 0.4515 | 0.4646 | 0.45784 | 0.5213 | 0.45 | 0.445 | 0.45 | 0.5 | 0.441 | 0.37435 | 0.49995 | 0.38825 | 0.46405
| Std Dev. |  | 0.011432 | 0.002145 | 0.005244 | 0.0082946 | 0.0082776 | 0.0080602 | 0.0092177 | 0.011219 | 0.0057966 | 0.0092664 | 0.01824 | 0.0063456 | 0.0056833 | 0.0049261 | 0.0051603 | 0.035647 | 0.005164 | 0.0054772 | 0.0063246 | 0.005164 | 0.0020656 | 0.0041468 | 0.0064136 | 0.010503 | 0.0078332
| Rel.Std.Dev. |  | 2.5% | 0.486% | 1.146% | 1.877% | 2.431% | 1.608% | 1.945% | 2.534% | 1.274% | 2.034% | 3.902% | 1.428% | 1.253% | 1.065% | 1.126% | 6.68% | 1.156% | 1.231% | 1.405% | 1.04% | 0.4691% | 1.11% | 1.282% | 2.709% | 1.69%
| PDM3 |  | -0.5151% | -3.964% | -0.4462% | -3.819% | -25.89% | 9.056% | 3.108% | -3.674% | -0.9902% | -0.8451% | 1.73% | -3.311% | -1.317% | 0.6563% | -0.3113% | 16.13% | -2.804% | -3.166% | -2.078% | 8.077% | -4.182% | -18.74% | 8.827% | -15.65% | 0.8539%
| Z-Score (Absolute) |  | 0.12 | 0.75 | 0.14 | 0.71 | 6.26 | 2.52 | 1.03 | 0.67 | 0 | 0.04 | 0.68 | 0.58 | 0.08 | 0.41 | 0.17 | 4.3 | 0.46 | 0.55 | 0.27 | 2.28 | 0.8 | 4.46 | 2.47 | 3.68 | 0.46
| Analytical results for Tl in OREAS 239b (Certified Value 0.83 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.85 | 0.80749 | 0.8 | 0.84 | 0.85 | 1.03 | 0.8 | 0.9 | 0.89 | 0.8 | 0.82 | 0.82 | 0.8 | 0.878 | 0.772 | 0.8 | 0.83 | 0.8 | 0.8 | 0.94259 | 0.9059 | 0.85016 |  |  |  |  | 1
| 1 | 2 | 0.88 | 0.75582 | 0.8 | 0.86 | 0.83 | 1.02 | 0.7 | 0.8 | 0.9 | 0.8 | 0.75 | 0.82 | 0.76 | 0.83984 | 0.78 | 0.83 | 0.84 | 0.76 | 0.8 | 0.92972 | 0.8699 | 0.84759 |  |  |  |  | 33
| 1 | 3 | 0.88 | 0.77182 | 0.8 | 0.83 | 0.86 | 1.04 | 0.8 | 0.9 | 0.89 | 0.82 | 0.77 | 0.79 | 0.79 | 0.88272 | 0.795 | 0.8 | 0.86 | 0.75 | 0.81 | 0.92895 | 0.8785 | 0.878 |  |  |  |  | 16
| 1 | 4 | 0.86 | 0.80921 | 0.8 | 0.87 | 0.85 | 1.03 | 0.76 | 0.8 | 0.86 | 0.86 | 0.79 | 0.78 | 0.83 | 0.87997 | 0.798 | 0.82 | 0.83 | 0.82 | 0.82 | 0.97031 | 0.8815 | 0.87266 |  |  |  |  | 0.83153
| 1 | 5 | 0.87 | 0.80984 | 0.8 | 0.87 | 0.85 | 0.96 | 0.77 | 0.8 | 0.89 | 0.76 | 0.8 | 0.82 | 0.84 | 0.82403 | 0.805 | 0.81 | 0.88 | 0.73 | 0.79 | 0.93731 | 0.8888 | 0.85741 |  |  |  |  | 63
| 1 | 6 | 0.88 | 0.82315 | 0.8 | 0.88 | 0.83 | 0.93 | 0.82 | 0.8 | 0.81 | 0.87 | 0.81 | 0.81 | 0.87 | 0.83799 | 0.788 | 0.78 | 0.89 | 0.81 | 0.81 | 0.93412 | 0.8626 | 0.8589
| Mean |  | 0.87 | 0.79622 | 0.8 | 0.85833 | 0.845 | 1.0017 | 0.775 | 0.83333 | 0.87333 | 0.81833 | 0.79 | 0.80667 | 0.815 | 0.85709 | 0.78967 | 0.80667 | 0.855 | 0.77833 | 0.805 | 0.9405 | 0.8812 | 0.86079
| Median |  | 0.875 | 0.80835 | 0.8 | 0.865 | 0.85 | 1.025 | 0.785 | 0.8 | 0.89 | 0.81 | 0.795 | 0.815 | 0.815 | 0.85892 | 0.7915 | 0.805 | 0.85 | 0.78 | 0.805 | 0.93572 | 0.88 | 0.85816
| Std Dev. |  | 0.012649 | 0.026209 | 1.2162e-16 | 0.019408 | 0.012247 | 0.04535 | 0.042778 | 0.05164 | 0.033862 | 0.041191 | 0.026077 | 0.017512 | 0.03937 | 0.025973 | 0.012176 | 0.017512 | 0.025884 | 0.03656 | 0.010488 | 0.01545 | 0.015158 | 0.01216
| Rel.Std.Dev. |  | 1.454% | 3.292% | 1.52e-14% | 2.261% | 1.449% | 4.528% | 5.52% | 6.197% | 3.877% | 5.033% | 3.301% | 2.171% | 4.831% | 3.03% | 1.542% | 2.171% | 3.027% | 4.697% | 1.303% | 1.643% | 1.72% | 1.413%
| PDM3 |  | 4.626% | -4.246% | -3.792% | 3.223% | 1.62% | 20.46% | -6.799% | 0.2166% | 5.027% | -1.587% | -4.995% | -2.99% | -1.988% | 3.074% | -5.035% | -2.99% | 2.822% | -6.398% | -3.191% | 13.1% | 5.973% | 3.518%
| Z-Score (Absolute) |  | 0.93 | 0.63 | 0.55 | 0.69 | 0.41 | 3.72 | 1.08 | 0.16 | 1 | 0.16 | 0.76 | 0.41 | 0.23 | 0.66 | 0.77 | 0.41 | 0.62 | 1 | 0.44 | 2.43 | 1.17 | 0.74
| Analytical results for Tm in OREAS 239b (Certified Value 0.26 ppm) | [27 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.20 | 1.26 | 1.28 | 1.30 | 1.34 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [17 empty] | 2
2
| 1 | 1 | 0.22681 | 0.3 | 0.3 | 0.2 | 0.208 | 0.3 | 0.2 | 0.1999 | 0.38073 | [17 empty] | 1
| 1 | 2 | 0.2299 | 0.3 | 0.3 | 0.2 | 0.211 | 0.28 | 0.2 | 0.1972 | 0.39059 | [17 empty] | 13
| 1 | 3 | 0.22072 | 0.25 | 0.3 | 0.3 | 0.199 | 0.28 | 0.2 | 0.2018 | 0.39579 | [17 empty] | 16
| 1 | 4 | 0.22563 | 0.3 | 0.3 | 0.2 | 0.205 | 0.27 | 0.2 | 0.2094 | 0.38963 | [17 empty] | 0.26147
| 1 | 5 | 0.22394 | 0.3 | 0.3 | 0.2 | 0.202 | 0.29 | 0.2 | 0.2022 | 0.39255 | [17 empty] | 64
| 1 | 6 | 0.2305 | 0.3 | 0.3 | 0.4 | 0.207 | 0.28 | 0.2 | 0.1994 | 0.38007
| Mean |  | 0.22625 | 0.29167 | 0.3 | 0.25 | 0.20533 | 0.28333 | 0.2 | 0.20165 | 0.38823
| Median |  | 0.22622 | 0.3 | 0.3 | 0.2 | 0.206 | 0.28 | 0.2 | 0.20085 | 0.39011
| Std Dev. |  | 0.0036879 | 0.020412 | 0 | 0.083666 | 0.0043205 | 0.010328 | 3.0405e-17 | 0.0042027 | 0.0064215
| Rel.Std.Dev. |  | 1.63% | 6.999% | 0% | 33.47% | 2.104% | 3.645% | 1.52e-14% | 2.084% | 1.654%
| PDM3 |  | -13.47% | 11.55% | 14.74% | -4.387% | -21.47% | 8.361% | -23.51% | -22.88% | 48.48%
| Z-Score (Absolute) |  | 0.36 | 0.63 | 0.75 | 0 | 0.67 | 0.5 | 0.75 | 0.73 | 2.09
| Analytical results for U in OREAS 239b (Certified Value 2.66 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.65 | 2.6879 | 2.7 | 2.7 | 2.77 | 2.6 | 2.92 | 2.7 | 2.6 | 2.7 | 2.9 | 2.6 | 2.6 | 2.6706 | 2.62 | 2.497 | 2.61 | 2.71 | 2.7 | 2.5 | 2.5621 | 4.0376 | 3.0062 |  |  |  | 1
| 1 | 2 | 2.71 | 2.6732 | 2.7 | 2.7 | 2.73 | 2.5 | 2.8 | 2.6 | 2.6 | 2.7 | 2.6 | 2.6 | 2.5 | 2.7506 | 2.73 | 2.638 | 2.7 | 2.83 | 2.6 | 2.5 | 2.5468 | 4.0108 | 3.008 |  |  |  | 35
| 1 | 3 | 2.73 | 2.6182 | 2.6 | 2.5 | 2.8 | 2.6 | 3.07 | 2.7 | 2.6 | 2.7 | 2.7 | 2.5 | 2.6 | 2.6011 | 2.62 | 2.535 | 2.6 | 2.8 | 2.6 | 2.6 | 2.5867 | 4.2234 | 3.0551 |  |  |  | 16
| 1 | 4 | 2.71 | 2.6327 | 2.7 | 2.8 | 2.8 | 2.7 | 2.88 | 2.8 | 2.6 | 2.8 | 2.6 | 2.6 | 2.9 | 2.6463 | 2.58 | 2.54 | 2.8 | 2.77 | 2.8 | 2.6 | 2.7144 | 4.4093 | 3.0582 |  |  |  | 2.6641
| 1 | 5 | 2.74 | 2.6887 | 2.7 | 2.9 | 2.76 | 2.7 | 2.8 | 2.7 | 2.5 | 2.6 | 2.7 | 2.6 | 2.8 | 2.6332 | 2.56 | 2.6 | 2.63 | 2.86 | 2.5 | 2.5 | 2.5959 | 4.1447 | 3.038 |  |  |  | 65
| 1 | 6 | 2.78 | 2.6851 | 2.7 | 2.8 | 2.74 | 2.8 | 2.94 | 2.8 | 2.6 | 2.7 | 2.7 | 2.5 | 2.9 | 2.7527 | 2.58 | 2.574 | 2.77 | 2.76 | 2.8 | 2.5 | 2.5921 | 4.0738 | 3.0559
| Mean |  | 2.72 | 2.6643 | 2.6833 | 2.7333 | 2.7667 | 2.65 | 2.9017 | 2.7167 | 2.5833 | 2.7 | 2.7 | 2.5667 | 2.7167 | 2.6758 | 2.615 | 2.564 | 2.685 | 2.7883 | 2.6667 | 2.5333 | 2.5997 | 4.1499 | 3.0369
| Median |  | 2.72 | 2.6792 | 2.7 | 2.75 | 2.765 | 2.65 | 2.9 | 2.7 | 2.6 | 2.7 | 2.7 | 2.6 | 2.7 | 2.6584 | 2.6 | 2.557 | 2.665 | 2.785 | 2.65 | 2.5 | 2.5894 | 4.1093 | 3.0466
| Std Dev. |  | 0.042895 | 0.030949 | 0.040825 | 0.13663 | 0.029439 | 0.10488 | 0.10128 | 0.075277 | 0.040825 | 0.063246 | 0.10954 | 0.05164 | 0.17224 | 0.062924 | 0.061237 | 0.050553 | 0.085499 | 0.053448 | 0.12111 | 0.05164 | 0.059346 | 0.14864 | 0.024204
| Rel.Std.Dev. |  | 1.577% | 1.162% | 1.521% | 4.999% | 1.064% | 3.958% | 3.49% | 2.771% | 1.58% | 2.342% | 4.057% | 2.012% | 6.34% | 2.352% | 2.342% | 1.972% | 3.184% | 1.917% | 4.541% | 2.038% | 2.283% | 3.582% | 0.797%
| PDM3 |  | 2.097% | 0.006566% | 0.7204% | 2.597% | 3.848% | -0.5307% | 8.916% | 1.972% | -3.033% | 1.346% | 1.346% | -3.659% | 1.972% | 0.4362% | -1.844% | -3.759% | 0.783% | 4.662% | 0.09485% | -4.91% | -2.42% | 55.77% | 13.99%
| Z-Score (Absolute) |  | 0.49 | 0.29 | 0.02 | 0.67 | 1.14 | 0.49 | 3.02 | 0.44 | 1.42 | 0.21 | 0.21 | 1.65 | 0.44 | 0.13 | 0.98 | 1.69 | 0 | 1.44 | 0.26 | 2.12 | 1.19 | 20.44 | 4.91
| Analytical results for V in OREAS 239b (Certified Value 105 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 108 | 88.14 | 105 | 118 | 117.3269 | 105 | 103 | 101 | 104 | 107 | 108 | 103 | 100 | 108 | 104.9266 | 104 | 108 | 100.8 | 106 | 110 | 105 | 98.45 | 108 | 106.5971 | 106.8236 |  | 1
| 1 | 2 | 106 | 87.7 | 105 | 120 | 116.8891 | 110 | 102 | 102 | 102 | 107 | 105 | 100 | 104 | 106 | 106.8755 | 107 | 109 | 100.3 | 108 | 110 | 106 | 100.18 | 112 | 78.3644 | 105.4147 |  | 36
| 1 | 3 | 106 | 88.47 | 105 | 116 | 118.2249 | 110 | 105 | 98 | 104 | 105 | 110 | 104 | 102 | 108 | 108.2605 | 118 | 106 | 100.5 | 105 | 108 | 106 | 98.23 | 112 | 76.434 | 105.3247 |  | 16
| 1 | 4 | 110 | 88.86 | 105 | 118 | 117.3421 | 106 | 105 | 90 | 104 | 105 | 113 | 100 | 102 | 106 | 104.3715 | 104 | 108 | 99.78 | 106 | 109 | 105 | 99.66 | 110 | 106.6776 | 104.4798 |  | 105.2798
| 1 | 5 | 110 | 88.82 | 105 | 119 | 116.8693 | 105 | 104 | 97 | 104 | 105 | 109 | 102 | 102 | 107 | 105.0369 | 105 | 107 | 100.1 | 106 | 109 | 106 | 99.66 | 109 | 109.171 | 104.825 |  | 66
| 1 | 6 | 109 | 89.18 | 105 | 122 | 117.154 | 106 | 105 | 104 | 106 | 104 | 116 | 102 | 101 | 108 | 106.0707 | 107 | 107 | 100.5 | 107 | 109 | 106 | 99.77 | 112 | 79.0079 | 107.115
| Mean |  | 108.1667 | 88.5283 | 105 | 118.8333 | 117.301 | 107 | 104 | 98.6667 | 104 | 105.5 | 110.1667 | 101.8333 | 101.8333 | 107.1667 | 105.9236 | 107.5 | 107.5 | 100.33 | 106.3333 | 109.1667 | 105.6667 | 99.325 | 110.5 | 92.7087 | 105.6638
| Median |  | 108.5 | 88.645 | 105 | 118.5 | 117.2404 | 106 | 104.5 | 99.5 | 104 | 105 | 109.5 | 102 | 102 | 107.5 | 105.5538 | 106 | 107.5 | 100.4 | 106 | 109 | 106 | 99.66 | 111 | 92.8025 | 105.3697
| Std Dev. |  | 1.8348 | 0.54039 | 0 | 2.0412 | 0.4967 | 2.3664 | 1.2649 | 4.9666 | 1.2649 | 1.2247 | 3.8687 | 1.6021 | 1.3292 | 0.98319 | 1.4545 | 5.3198 | 1.0488 | 0.35637 | 1.0328 | 0.75277 | 0.5164 | 0.7897 | 1.7607 | 16.2318 | 1.0708
| Rel.Std.Dev. |  | 1.696% | 0.6104% | 0% | 1.718% | 0.4234% | 2.212% | 1.216% | 5.034% | 1.216% | 1.161% | 3.512% | 1.573% | 1.305% | 0.9174% | 1.373% | 4.949% | 0.9756% | 0.3552% | 0.9713% | 0.6896% | 0.4887% | 0.7951% | 1.593% | 17.51% | 1.013%
| PDM3 |  | 2.742% | -15.91% | -0.2658% | 12.87% | 11.42% | 1.634% | -1.216% | -6.281% | -1.216% | 0.2092% | 4.642% | -3.274% | -3.274% | 1.792% | 0.6115% | 2.109% | 2.109% | -4.702% | 1.001% | 3.692% | 0.3675% | -5.656% | 4.958% | -11.94% | 0.3647%
| Z-Score (Absolute) |  | 0.67 | 4.62 | 0.18 | 3.55 | 3.14 | 0.36 | 0.45 | 1.89 | 0.45 | 0.04 | 1.21 | 1.03 | 1.03 | 0.4 | 0.07 | 0.49 | 0.49 | 1.44 | 0.18 | 0.94 | 0 | 1.71 | 1.3 | 3.5 | 0
| Analytical results for W in OREAS 239b (Certified Value 2.94 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 3 | 3.0037 | 2.5 | 0.5 | 19.91 | 2.9 | 2.9 | 1.7 | 2.6 | 3.1 | 2.9 | 3 | 2.9 | 3.1 | 2.912 | 4.1 | 2.8 | < 5 | 3.5 | 3.6 | 2.4 | 1.8658 | < 5 | 3.3251 |  |  | 1
| 1 | 2 | 3 | 2.9865 | 2.5 | 0.1 | 22.4633 | 2.9 | 2.6 | 2 | 2.8 | 3.2 | 2.9 | 2.9 | 3 | 2.9 | 2.9696 | 4.1 | 2.9 | < 5 | 3.45 | 3.4 | 2.4 | 1.8529 | < 5 | 3.3691 |  |  | 37
| 1 | 3 | 2.8 | 3.1176 | 2.5 | 0.5 | 20.094 | 2.9 | 2.3 | 3.2 | 3.5 | 3.1 | 3.1 | 2.9 | 2.7 | 3 | 2.8904 | 4.1 | 2.6 | < 5 | 3.48 | 3.4 | 2.2 | 1.8776 | < 5 | 3.4053 |  |  | 16
| 1 | 4 | 3.3 | 3.0157 | 3 | 0.1 | 20.5515 | 3 | 2.8 | 1.9 | 2.5 | 3.2 | 3 | 2.8 | 2.8 | 3.2 | 2.9029 | 4.1 | 2.7 | < 5 | 3.67 | 3.5 | 2.3 | 1.9349 | < 5 | 3.3611 |  |  | 2.9352
| 1 | 5 | 3.2 | 3.0365 | 3 | 0.3 | 19.5466 | 3 | 2.4 | 2.6 | 2.7 | 3.1 | 2.9 | 2.9 | 2.7 | 3.1 | 2.9194 | 4 | 2.7 | < 5 | 3.6 | 3.3 | 2.2 | 1.8531 | < 5 | 3.3756 |  |  | 67
| 1 | 6 | 3.3 | 3.0878 | 3 | 0.1 | 19.53 | 3.1 | 2.7 | 2.7 | 2.9 | 2.9 | 2.9 | 3.3 | 2.7 | 3.4 | 2.8747 | 4.1 | 2.8 | < 5 | 3.41 | 3.6 | 2.3 | 1.8859 | < 5 | 3.3628
| Mean |  | 3.1 | 3.0413 | 2.75 | 0.26667 | 20.3492 | 2.9667 | 2.6167 | 2.35 | 2.8333 | 3.1 | 2.95 | 2.9667 | 2.8 | 3.1167 | 2.9115 | 4.0833 | 2.75 |  | 3.5183 | 3.4667 | 2.3 | 1.8784 |  | 3.3665
| Median |  | 3.1 | 3.0261 | 2.75 | 0.2 | 20.002 | 2.95 | 2.65 | 2.3 | 2.75 | 3.1 | 2.9 | 2.9 | 2.75 | 3.1 | 2.9074 | 4.1 | 2.75 |  | 3.49 | 3.45 | 2.3 | 1.8717 |  | 3.366
| Std Dev. |  | 0.2 | 0.051139 | 0.27386 | 0.19664 | 1.1031 | 0.08165 | 0.23166 | 0.57533 | 0.3559 | 0.10954 | 0.083666 | 0.17512 | 0.12649 | 0.17224 | 0.03257 | 0.040825 | 0.10488 |  | 0.09786 | 0.12111 | 0.089443 | 0.030654 |  | 0.025906
| Rel.Std.Dev. |  | 6.452% | 1.681% | 9.959% | 73.74% | 5.421% | 2.752% | 8.853% | 24.48% | 12.56% | 3.534% | 2.836% | 5.903% | 4.518% | 5.526% | 1.119% | 0.9998% | 3.814% |  | 2.781% | 3.493% | 3.889% | 1.632% |  | 0.7695%
| PDM3 |  | 5.613% | 3.613% | -6.311% | -90.91% | 593.3% | 1.071% | -10.85% | -19.94% | -3.472% | 5.613% | 0.5029% | 1.071% | -4.607% | 6.181% | -0.809% | 39.11% | -6.311% |  | 19.87% | 18.11% | -21.64% | -36.01% |  | 14.69%
| Z-Score (Absolute) |  | 0.46 | 0.3 | 0.49 | 7.19 | 46.98 | 0.1 | 0.85 | 1.57 | 0.26 | 0.46 | 0.05 | 0.1 | 0.35 | 0.5 | 0.05 | 3.11 | 0.49 | 1.16 | 1.58 | 1.45 | 1.7 | 2.84 | 1.16 | 1.18
| Analytical results for Y in OREAS 239b (Certified Value 15.6 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 15.73 | 13.2747 | 18.2 | 18.7 | 11.93 | 16.8 | 12.5 | 16.3 | 14.5 | 17 | 16.1 | 16.1 | 17 | 14.9 | 15.4016 | 20.7 | 12.54 | 14.6 | 18.77 | 16.8 | 14.1 | 12.35 | 13.9597 | 23.4962 |  |  | 1
| 1 | 2 | 15.8 | 13.1724 | 18.1 | 17.2 | 10.67 | 18.2 | 12.2 | 15.9 | 14.5 | 17.7 | 16 | 15.8 | 16.5 | 15.3 | 15.7092 | 21.2 | 12.41 | 15.6 | 18.87 | 15.7 | 14.6 | 12.74 | 14.165 | 23.8033 |  |  | 38
| 1 | 3 | 16.25 | 12.679 | 18 | 14.7 | 11.715 | 16.9 | 13.3 | 17 | 14.4 | 17 | 16.8 | 16 | 16.4 | 15.1 | 16.1385 | 20.8 | 11.8 | 14.2 | 18.72 | 16 | 15 | 12.32 | 14.0623 | 23.6907 |  |  | 16
| 1 | 4 | 15.75 | 12.7504 | 18.5 | 18.3 | 11.065 | 17.1 | 13.7 | 18 | 14.2 | 17.1 | 17 | 18 | 21.9 | 15.6 | 15.8301 | 21.2 | 12 | 14.5 | 19.08 | 16 | 14.8 | 12.54 | 14.7808 | 24.2321 |  |  | 15.6328
| 1 | 5 | 16.51 | 12.9009 | 18.5 | 19.9 | 11.31 | 17.2 | 13.8 | 17.5 | 14.3 | 17.3 | 16.4 | 15.8 | 16.7 | 15.2 | 16.0917 | 20.4 | 12.02 | 15.2 | 19.18 | 15.5 | 14.6 | 12.54 | 14.4729 | 23.6602 |  |  | 68
| 1 | 6 | 16.33 | 12.5949 | 18 | 19.7 | 10.9217 | 18.4 | 13 | 17.5 | 14.3 | 16.9 | 16.6 | 16 | 16.2 | 13.4 | 15.3948 | 20.1 | 12.22 | 15.2 | 25.59 | 16.1 | 14.2 | 12.69 | 14.2676 | 23.5997
| Mean |  | 16.0617 | 12.8954 | 18.2167 | 18.0833 | 11.2686 | 17.4333 | 13.0833 | 17.0333 | 14.3667 | 17.1667 | 16.4833 | 16.2833 | 17.45 | 14.9167 | 15.761 | 20.7333 | 12.165 | 14.8833 | 20.035 | 16.0167 | 14.55 | 12.53 | 14.2847 | 23.747
| Median |  | 16.025 | 12.8256 | 18.15 | 18.5 | 11.1875 | 17.15 | 13.15 | 17.25 | 14.35 | 17.05 | 16.5 | 16 | 16.6 | 15.15 | 15.7697 | 20.75 | 12.12 | 14.9 | 18.975 | 16 | 14.6 | 12.54 | 14.2163 | 23.6754
| Std Dev. |  | 0.34178 | 0.27526 | 0.23166 | 1.9271 | 0.48137 | 0.68896 | 0.64317 | 0.79917 | 0.12111 | 0.29439 | 0.392 | 0.84951 | 2.197 | 0.77825 | 0.32331 | 0.43665 | 0.27754 | 0.53072 | 2.7272 | 0.4446 | 0.34496 | 0.17111 | 0.30041 | 0.25838
| Rel.Std.Dev. |  | 2.128% | 2.135% | 1.272% | 10.66% | 4.272% | 3.952% | 4.916% | 4.692% | 0.843% | 1.715% | 2.378% | 5.217% | 12.59% | 5.217% | 2.051% | 2.106% | 2.281% | 3.566% | 13.61% | 2.776% | 2.371% | 1.366% | 2.103% | 1.088%
| PDM3 |  | 2.744% | -17.51% | 16.53% | 15.68% | -27.92% | 11.52% | -16.31% | 8.959% | -8.099% | 9.812% | 5.441% | 4.162% | 11.62% | -4.581% | 0.8203% | 32.63% | -22.18% | -4.794% | 28.16% | 2.456% | -6.926% | -19.85% | -8.623% | 51.91%
| Z-Score (Absolute) |  | 0.01 | 1.34 | 0.93 | 0.87 | 2.03 | 0.59 | 1.26 | 0.42 | 0.71 | 0.48 | 0.19 | 0.1 | 0.6 | 0.48 | 0.12 | 2 | 1.65 | 0.49 | 1.7 | 0.01 | 0.64 | 1.5 | 0.75 | 3.29
| Analytical results for Yb in OREAS 239b (Certified Value 1.7 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.03 | 1.05 | 1.06 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.28 | 1.30 | 1.32 | 1.34 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | [14 empty] | 2
2
| 1 | 1 | 1.541 | 1.9 | 2.1 | 1.5 | 2.09 | 1.5 | 1.5 | 1.98 | 1.5 | 1.34 | 1.5574 | 2.522 | [14 empty] | 1
| 1 | 2 | 1.5577 | 1.85 | 2.1 | 1.4 | 2.07 | 1.491 | 1.5 | 2.09 | 1.5 | 1.37 | 1.5149 | 2.5319 | [14 empty] | 39
| 1 | 3 | 1.4805 | 1.9 | 1.9 | 1.5 | 2.09 | 1.419 | 1.5 | 2.34 | 1.5 | 1.37 | 1.4724 | 2.55 | [14 empty] | 16
| 1 | 4 | 1.4489 | 1.95 | 2.1 | 1.7 | 2.09 | 1.468 | 1.5 | 2.09 | 1.5 | 1.36 | 1.6423 | 2.5372 | [14 empty] | 1.7028
| 1 | 5 | 1.5309 | 1.95 | 2.1 | 1.7 | 2.04 | 1.452 | 1.6 | 2.16 | 1.5 | 1.38 | 1.5998 | 2.5463 | [14 empty] | 69
| 1 | 6 | 1.5086 | 1.95 | 2.1 | 2.2 | 2.01 | 1.49 | 1.6 | 2.1 | 1.5 | 1.38 | 1.5149 | 2.5348
| Mean |  | 1.5113 | 1.9167 | 2.0667 | 1.6667 | 2.065 | 1.47 | 1.5333 | 2.1267 | 1.5 | 1.3667 | 1.5503 | 2.5371
| Median |  | 1.5198 | 1.925 | 2.1 | 1.6 | 2.08 | 1.479 | 1.5 | 2.095 | 1.5 | 1.37 | 1.5361 | 2.536
| Std Dev. |  | 0.040675 | 0.040825 | 0.08165 | 0.28752 | 0.033317 | 0.030561 | 0.05164 | 0.11961 | 0 | 0.015055 | 0.062513 | 0.010109
| Rel.Std.Dev. |  | 2.691% | 2.13% | 3.951% | 17.25% | 1.613% | 2.079% | 3.368% | 5.624% | 0% | 1.102% | 4.032% | 0.3985%
| PDM3 |  | -11.25% | 12.56% | 21.37% | -2.121% | 21.27% | -13.67% | -9.951% | 24.89% | -11.91% | -19.74% | -8.956% | 48.99%
| Z-Score (Absolute) |  | 0.34 | 1.09 | 1.62 | 0.21 | 1.62 | 0.49 | 0.27 | 1.84 | 0.38 | 0.86 | 0.21 | 3.29
| Analytical results for Zn in OREAS 239b (Certified Value 98 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 100 | 93.3 | 92 | 104 | 96.2495 | 98 | 104 | 110 | 127 | 100 | 99 | 94 | 96 | 99.5 | 94.5331 | 98 | 96.24 | 98 | 107.2 | 93 | 98 | 97.1 | 100.221 | 102 | 89.3903 | 94.7768 | 1
| 1 | 2 | 99 | 93.9992 | 94 | 104 | 95.11 | 97 | 98 | 95 | 128 | 100 | 96 | 91 | 98 | 97.1 | 95.9143 | 100 | 97.78 | 94 | 100.3 | 95 | 98 | 92 | 101.607 | 103 | 83.6397 | 94.1621 | 40
| 1 | 3 | 101 | 91.8936 | 92 | 103 | 98.3106 | 96 | 101 | 96 | 124 | 97 | 105 | 94 | 96 | 99.5 | 96.3443 | 99 | 101.4 | 94 | 102 | 95 | 97 | 98.4 | 100.925 | 103 | 91.2009 | 92.428 | 16
| 1 | 4 | 100 | 92.684 | 94 | 104 | 98.943 | 98 | 105 | 94 | 119 | 97 | 103 | 92 | 95 | 98.8 | 93.3802 | 99 | 97.44 | 97 | 103.6 | 93 | 97 | 96.5 | 101.739 | 101 | 83.7266 | 94.0835 | 97.8706
| 1 | 5 | 103 | 93.3152 | 92 | 103 | 98.8095 | 99 | 105 | 97 | 120 | 98 | 99 | 93 | 96 | 98.4 | 95.4771 | 98 | 104.4 | 94 | 104.1 | 94 | 97 | 97.8 | 101.409 | 101 | 83.828 | 94.6185 | 70
| 1 | 6 | 102 | 93.5232 | 92 | 105 | 96.2981 | 98 | 102 | 97 | 125 | 99 | 105 | 92 | 96 | 99 | 93.3287 | 99 | 99.52 | 95 | 102.9 | 94 | 97 | 95.6 | 101.002 | 104 | 82.9734 | 95.5153
| Mean |  | 100.8333 | 93.1192 | 92.6667 | 103.8333 | 97.2868 | 97.6667 | 102.5 | 98.1667 | 123.8333 | 98.5 | 101.1667 | 92.6667 | 96.1667 | 98.7167 | 94.8296 | 98.8333 | 99.4633 | 95.3333 | 103.35 | 94 | 97.3333 | 96.2333 | 101.1505 | 102.3333 | 85.7932 | 94.264
| Median |  | 100.5 | 93.3076 | 92 | 104 | 97.3043 | 98 | 103 | 96.5 | 124.5 | 98.5 | 101 | 92.5 | 96 | 98.9 | 95.0051 | 99 | 98.65 | 94.5 | 103.25 | 94 | 97 | 96.8 | 101.2055 | 102.5 | 83.7773 | 94.3903
| Std Dev. |  | 1.472 | 0.73488 | 1.0328 | 0.75277 | 1.6064 | 1.0328 | 2.7386 | 5.9133 | 3.656 | 1.3784 | 3.7103 | 1.2111 | 0.98319 | 0.89759 | 1.2908 | 0.75277 | 3.0147 | 1.7512 | 2.3141 | 0.89443 | 0.5164 | 2.2932 | 0.55829 | 1.2111 | 3.5469 | 1.0362
| Rel.Std.Dev. |  | 1.46% | 0.7892% | 1.115% | 0.725% | 1.651% | 1.057% | 2.672% | 6.024% | 2.952% | 1.399% | 3.668% | 1.307% | 1.022% | 0.9093% | 1.361% | 0.7617% | 3.031% | 1.837% | 2.239% | 0.9515% | 0.5305% | 2.383% | 0.5519% | 1.183% | 4.134% | 1.099%
| PDM3 |  | 3.027% | -4.855% | -5.317% | 6.093% | -0.5965% | -0.2083% | 4.73% | 0.3025% | 26.53% | 0.6431% | 3.368% | -5.317% | -1.741% | 0.8645% | -3.107% | 0.9837% | 1.627% | -2.592% | 5.599% | -3.955% | -0.5489% | -1.673% | 3.351% | 4.56% | -12.34% | -3.685%
| Z-Score (Absolute) |  | 0.62 | 1.02 | 1.12 | 1.26 | 0.13 | 0.05 | 0.98 | 0.05 | 5.53 | 0.12 | 0.69 | 1.12 | 0.37 | 0.17 | 0.66 | 0.2 | 0.33 | 0.55 | 1.16 | 0.84 | 0.12 | 0.36 | 0.69 | 0.94 | 2.59 | 0.78
| Analytical results for Zr in OREAS 239b (Certified Value 131 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 135.5 | 128.068 | 127 | 156 | 115.91 | 128.4 | 123 | 135 | 132 | 132.5 | 134 | 131 | 142 | 128.2 | 127.2117 | 144 | 129 | 135.5 | 136 | 129.5 | 118 | 114.697 | 148 | 120.3568 | 220.819 |  | 1
| 1 | 2 | 139.2 | 124.9237 | 130 | 157 | 112.8339 | 128.8 | 119 | 126 | 139.5 | 135.5 | 133 | 127.5 | 138.5 | 127.3 | 128.9882 | 144 | 129 | 134.4 | 134.5 | 127.3 | 117.9 | 117.469 | 154 | 119.2993 | 222.6751 |  | 41
| 1 | 3 | 136.2 | 127.5793 | 125 | 160 | 113.95 | 128.4 | 120 | 133 | 137.9 | 135.5 | 144.5 | 134 | 134 | 129.1 | 128.9871 | 170 | 124 | 132.8 | 135.5 | 126.7 | 120.3 | 114.939 | 152 | 113.0553 | 221.2907 |  | 16
| 1 | 4 | 141.2 | 123.0426 | 131 | 159 | 113.9762 | 131.1 | 128 | 135 | 140.4 | 135.5 | 141 | 128 | 132.5 | 128.4 | 125.5857 | 143 | 128 | 130.4 | 138.2 | 127.3 | 118.6 | 117.975 | 153 | 120.3568 | 217.7948 |  | 130.5435
| 1 | 5 | 144.8 | 126.1777 | 132 | 162 | 116.62 | 128.2 | 124 | 135 | 136.7 | 135 | 135.5 | 132 | 134 | 127.3 | 125.4496 | 142 | 126 | 133.9 | 138.1 | 128.9 | 116.9 | 117.161 | 148 | 123.1766 | 219.2721 |  | 71
| 1 | 6 | 143.7 | 127.9112 | 126 | 163 | 114.2906 | 131 | 122 | 138 | 137.5 | 133 | 140 | 136.5 | 135 | 125.6 | 128.3152 | 145 | 126 | 132.4 | 139.9 | 128.6 | 117.8 | 117.018 | 151 | 116.0766 | 223.088
| Mean |  | 140.1 | 126.2837 | 128.5 | 159.5 | 114.5968 | 129.3167 | 122.6667 | 133.6667 | 137.3333 | 134.5 | 138 | 131.5 | 136 | 127.65 | 127.4229 | 148 | 127 | 133.2333 | 137.0333 | 128.05 | 118.25 | 116.5432 | 151 | 118.7202 | 220.8233
| Median |  | 140.2 | 126.8785 | 128.5 | 159.5 | 114.1334 | 128.6 | 122.5 | 135 | 137.7 | 135.25 | 137.75 | 131.5 | 134.5 | 127.75 | 127.7635 | 144 | 127 | 133.35 | 137.05 | 127.95 | 117.95 | 117.0895 | 151.5 | 119.8281 | 221.0548
| Std Dev. |  | 3.8325 | 1.9938 | 2.881 | 2.7386 | 1.4016 | 1.3571 | 3.2042 | 4.0825 | 2.9426 | 1.3784 | 4.5277 | 3.4641 | 3.5637 | 1.2178 | 1.6129 | 10.8259 | 2 | 1.7806 | 2.0275 | 1.1023 | 1.1432 | 1.3781 | 2.5298 | 3.5953 | 2.0176
| Rel.Std.Dev. |  | 2.736% | 1.579% | 2.242% | 1.717% | 1.223% | 1.049% | 2.612% | 3.054% | 2.143% | 1.025% | 3.281% | 2.634% | 2.62% | 0.954% | 1.266% | 7.315% | 1.575% | 1.336% | 1.48% | 0.8608% | 0.9668% | 1.183% | 1.675% | 3.028% | 0.9137%
| PDM3 |  | 7.321% | -3.263% | -1.565% | 22.18% | -12.22% | -0.9398% | -6.034% | 2.392% | 5.201% | 3.031% | 5.712% | 0.7327% | 4.18% | -2.216% | -2.39% | 13.37% | -2.714% | 2.06% | 4.971% | -1.91% | -9.417% | -10.72% | 15.67% | -9.057% | 69.16%
| Z-Score (Absolute) |  | 1.05 | 0.64 | 0.37 | 3.41 | 2.06 | 0.27 | 1.08 | 0.26 | 0.71 | 0.37 | 0.79 | 0 | 0.55 | 0.47 | 0.5 | 2.01 | 0.55 | 0.21 | 0.67 | 0.42 | 1.61 | 1.82 | 2.38 | 1.56 | 10.89
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C870: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C906: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C942: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1015: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1033: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1051: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1069: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1087: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1105: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1123: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1159: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1177: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1213: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1231: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 239b (Certified Value 0.266 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.26 | 0.284 | < 0.5 | 0.25 | < 0.5 | 0.29 | 0.2 | 0.27 | 0.29 | 0.25 | 0.26 | 0.239 | < 1 | 0.3 | 0.26 | 0.25 | < 0.5 | 0.27 | 0.31 | < 1 | 1.339 |  |  |  | 1
| 1 | 2 | 0.25 | 0.278 | < 0.5 | 0.27 | < 0.5 | 0.27 | 0.2 | 0.27 | 0.26 | 0.27 | 0.26 | 0.243 | < 1 | 0.3 | 0.27 | 0.19 | < 0.5 | 0.27 | 0.31 | < 1 | 1.1535 |  |  |  | 25
| 1 | 3 | 0.26 | 0.292 | < 0.5 | 0.26 | < 0.5 | 0.28 | 0.2 | 0.27 | 0.28 | 0.25 | 0.27 | 0.242 | < 1 | 0.3 | 0.28 | 0.18 | < 0.5 | 0.27 | 0.28 | < 1 | 1.0183 |  |  |  | 16
| 1 | 4 | 0.27 | 0.293 | < 0.5 | 0.26 | < 0.5 | 0.28 | 0.3 | 0.27 | 0.27 | 0.26 | 0.27 | 0.244 | < 1 | 0.3 | 0.27 | 0.21 | < 0.5 | 0.27 | 0.31 | < 1 | 1.1026 |  |  |  | 0.26607
| 1 | 5 | 0.26 | 0.278 | < 0.5 | 0.25 | < 0.5 | 0.28 | 0.2 | 0.26 | 0.33 | 0.27 | 0.27 | 0.263 | < 1 | 0.3 | 0.27 | 0.17 | < 0.5 | 0.27 | 0.31 | < 1 | 1.5157 |  |  |  | 73
| 1 | 6 | 0.26 | 0.265 | < 0.5 | 0.27 | < 0.5 | 0.29 | 0.2 | 0.25 | 0.27 | 0.26 | 0.25 | 0.254 | < 1 | 0.3 | 0.25 | 0.17 | < 0.5 | 0.27 | 0.29 | < 1 | 1.2709
| Mean |  | 0.26 | 0.28167 |  | 0.26 |  | 0.28167 | 0.21667 | 0.265 | 0.28333 | 0.26 | 0.26333 | 0.2475 |  | 0.3 | 0.26667 | 0.195 |  | 0.27 | 0.30167 |  | 1.2333
| Median |  | 0.26 | 0.281 |  | 0.26 |  | 0.28 | 0.2 | 0.27 | 0.275 | 0.26 | 0.265 | 0.2435 |  | 0.3 | 0.27 | 0.185 |  | 0.27 | 0.31 |  | 1.2122
| Std Dev. |  | 0.0063246 | 0.010443 |  | 0.0089443 |  | 0.0075277 | 0.040825 | 0.0083666 | 0.025033 | 0.0089443 | 0.008165 | 0.0091378 |  | 0 | 0.010328 | 0.030822 |  | 0 | 0.013292 |  | 0.18001
| Rel.Std.Dev. |  | 2.433% | 3.708% |  | 3.44% |  | 2.673% | 18.84% | 3.157% | 8.835% | 3.44% | 3.101% | 3.692% |  | 0% | 3.873% | 15.81% |  | 0% | 4.406% |  | 14.6%
| PDM3 |  | -2.28% | 5.863% |  | -2.28% |  | 5.863% | -18.57% | -0.4009% | 6.49% | -2.28% | -1.027% | -6.978% |  | 12.75% | 0.2255% | -26.71% |  | 1.478% | 13.38% |  | 363.5%
| Z-Score (Absolute) |  | 0.22 | 0.75 | 0.67 | 0.22 | 0.67 | 0.75 | NA | 0 | 0.82 | 0.22 | 0.07 | 0.79 | 10.56 | NA | 0.07 | 3.15 | 0.67 | 0.22 | 1.65 | 10.56 | 43.53
Results from laboratories 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 239b (Certified Value 2.58 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 2.388 | 3.0466 | 2.71 | 2.92 | 2.742 | 2.73 | 3.2 | 2.55 | 2.51 | 2.78 | 2.6 | 2.64 | 2.54 | 2.6317 | 2.387 | 2.47 | 2.65 | 2.38 | 2.3614 | 4.9848 |  |  |  |  | 1
| 1 | 2 | 2.459 | 3.0478 | 2.73 | 2.92 | 2.729 | 2.67 | 3.15 | 2.56 | 2.52 | 2.65 | 2.56 | 2.66 | 2.59 | 2.5734 | 2.419 | 2.47 | 2.7 | 2.39 | 2.3375 | 4.9866 |  |  |  |  | #N/A
| 1 | 3 | 2.453 | 3.0389 | 2.72 | 2.89 | 2.747 | 2.7 | 3.22 | 2.55 | 2.48 | 2.75 | 2.57 | 2.62 | 2.6 | 2.5548 | 2.376 | 2.47 | 2.67 | 2.39 | 2.3045 | 4.9878 |  |  |  |  | 16
| 1 | 4 | 2.418 | 3.0642 | 2.92 | 2.89 | 2.712 | 2.68 | 3.22 | 2.58 | 2.53 | 2.64 | 2.58 | 2.67 | 2.56 | 2.5282 | 2.371 | 2.56 | 2.69 | 2.42 | 2.3557 | 4.9855 |  |  |  |  | 2.5807
| 1 | 5 | 2.386 | 3.0269 | 2.73 | 2.9 | 2.71 | 2.69 | 3.19 | 2.56 | 2.49 | 2.6 | 2.63 | 2.66 | 2.48 | 2.6183 | 2.382 | 2.36 | 2.64 | 2.45 | 2.3722 | 4.9839 |  |  |  |  | 74
| 1 | 6 | 2.403 | 3.0169 | 2.74 | 2.9 | 2.7 | 2.73 | 3.27 | 2.56 | 2.54 | 2.62 | 2.59 | 2.67 | 2.49 | 2.6469 | 2.387 | 2.37 | 2.67 | 2.52 | 2.349 | 4.9871
| Mean |  | 2.4178 | 3.0402 | 2.7583 | 2.9033 | 2.7233 | 2.7 | 3.2083 | 2.56 | 2.5117 | 2.6733 | 2.5883 | 2.6533 | 2.5433 | 2.5922 | 2.387 | 2.45 | 2.67 | 2.425 | 2.3467 | 4.9859
| Median |  | 2.4105 | 3.0427 | 2.73 | 2.9 | 2.7205 | 2.695 | 3.21 | 2.56 | 2.515 | 2.645 | 2.585 | 2.66 | 2.55 | 2.5959 | 2.3845 | 2.47 | 2.67 | 2.405 | 2.3524 | 4.9861
| Std Dev. |  | 0.031796 | 0.01671 | 0.079854 | 0.013663 | 0.018928 | 0.025298 | 0.039707 | 0.010954 | 0.023166 | 0.073666 | 0.024833 | 0.019664 | 0.050067 | 0.047094 | 0.016888 | 0.074565 | 0.022804 | 0.053198 | 0.023741 | 0.0014761
| Rel.Std.Dev. |  | 1.315% | 0.5496% | 2.895% | 0.4706% | 0.695% | 0.937% | 1.238% | 0.4279% | 0.9223% | 2.756% | 0.9594% | 0.7411% | 1.969% | 1.817% | 0.7075% | 3.043% | 0.8541% | 2.194% | 1.012% | 0.02961%
| PDM3 |  | -6.31% | 17.81% | 6.885% | 12.5% | 5.528% | 4.624% | 24.32% | -0.8008% | -2.674% | 3.591% | 0.2971% | 2.816% | -1.447% | 0.4479% | -7.504% | -5.063% | 3.462% | -6.032% | -9.065% | 93.2%
| Z-Score (Absolute) |  | 1.12 | 2.28 | 0.74 | 1.53 | 0.55 | 0.42 | 3.2 | 0.34 | 0.61 | 0.28 | 0.19 | 0.17 | 0.43 | 0.17 | 1.29 | 0.94 | 0.26 | 1.08 | 1.51 | 12.92
| Analytical results for As in OREAS 239b (Certified Value 811 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 798.08 | 846.04 | 776 | 798.06 | 801 | 837.7 | 795 | 864.1 | 813 | 854 | 834 | 845 | 822.4 | 815.4538 | 786 | 753 | 954.96 | 836.9 | 745.8 | 789 | 792 | 623.2135 | 789.9317 |  | 1
| 1 | 2 | 806.1 | 845.46 | 770 | 808.417 | 810 | 844.1 | 824 | 858.3 | 816 | 829 | 827 | 841 | 825.4 | 824.0638 | 770 | 750 | 965.66 | 832.1 | 769.1 | 773.7 | 802 | 685.1828 | 802.1963 |  | 26
| 1 | 3 | 806.57 | 851.15 | 783 | 798.9218 | 818 | 835.4 | 826 | 856 | 811 | 854 | 827 | 832 | 860.8 | 798.0682 | 775 | 742 | 937.72 | 837.6 | 760.2 | 773.2 | 797 | 592.091 | 786.3734 |  | 16
| 1 | 4 | 815.21 | 851.67 | 785 | 804.1154 | 827 | 840.8 | 822 | 881.4 | 811 | 817 | 838 | 846 | 815.2 | 799.805 | 803 | 725 | 947.62 | 834.4 | 761.9 | 795 | 815 | 524.8538 | 786.5095 |  | 810.5157
| 1 | 5 | 796.35 | 846.75 | 734 | 809.82 | 799 | 860.2 | 791 | 854.2 | 801 | 814 | 839 | 845 | 795.7 | 803.0862 | 790 | 732 | 917.26 | 831.9 | 771 | 806.7 | 808 | 573.6376 | 783.5244 |  | 75
| 1 | 6 | 812.95 | 846.81 | 778 | 807.0233 | 790 | 833.9 | 830 | 869.2 | 813 | 824 | 834 | 851 | 825.1 | 818.689 | 787 | 727 | 900.58 | 837.8 | 779.4 | 775.9 | 813 | 656.2306 | 789.8841
| Mean |  | 805.8767 | 847.98 | 771 | 804.3929 | 807.5 | 842.0167 | 814.6667 | 863.8667 | 810.8333 | 832 | 833.1667 | 843.3333 | 824.1 | 809.861 | 785.1667 | 738.1667 | 937.3 | 835.1167 | 764.5667 | 785.5833 | 804.5 | 609.2015 | 789.7366
| Median |  | 806.335 | 846.78 | 777 | 805.5693 | 805.5 | 839.25 | 823 | 861.2 | 812 | 826.5 | 834 | 845 | 823.75 | 809.27 | 786.5 | 737 | 942.67 | 835.65 | 765.5 | 782.45 | 805 | 607.6522 | 788.1968
| Std Dev. |  | 7.6055 | 2.7078 | 18.8892 | 4.9534 | 13.5462 | 9.641 | 17.0372 | 10.2199 | 5.1543 | 17.8326 | 5.1929 | 6.4083 | 21.1577 | 10.9272 | 11.6519 | 11.9234 | 24.3836 | 2.7022 | 11.492 | 13.6718 | 9.094 | 58.0574 | 6.5669
| Rel.Std.Dev. |  | 0.9438% | 0.3193% | 2.45% | 0.6158% | 1.678% | 1.145% | 2.091% | 1.183% | 0.6357% | 2.143% | 0.6233% | 0.7599% | 2.567% | 1.349% | 1.484% | 1.615% | 2.601% | 0.3236% | 1.503% | 1.74% | 1.13% | 9.53% | 0.8315%
| PDM3 |  | -0.5724% | 4.622% | -4.875% | -0.7554% | -0.3721% | 3.887% | 0.5121% | 6.582% | 0.03918% | 2.651% | 2.795% | 4.049% | 1.676% | -0.08078% | -3.128% | -8.926% | 15.64% | 3.035% | -5.669% | -3.076% | -0.7422% | -24.84% | -2.564%
| Z-Score (Absolute) |  | 0.11 | 1.06 | 1.08 | 0.15 | 0.07 | 0.89 | 0.13 | 1.5 | 0.03 | 0.61 | 0.65 | 0.93 | 0.4 | 0 | 0.69 | 1.99 | 3.54 | 0.7 | 1.26 | 0.67 | 0.15 | 5.57 | 0.56
| Analytical results for B in OREAS 239b (Indicative Value 15.7 ppm) | [25 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.27 | 1.30 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [13 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [13 empty] | 1
| 1 | 1 | < 5 | 5 | 27.25 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 21 | [13 empty] | 1
| 1 | 2 | < 5 | 5 | 26.965 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 21 | [13 empty] | 4
| 1 | 3 | < 5 | 5 | 27.235 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 20 | [13 empty] | 16
| 1 | 4 | < 5 | 6 | 26.51 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 21 | [13 empty] | 15.7323
| 1 | 5 | < 5 | 5 | 26.975 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 21 | [13 empty] | 14
| 1 | 6 | < 5 | 5 | 25.64 | < 10 | < 10 | 10 | < 10 | < 10 | < 20 | < 10 | 22
| Mean |  |  | 5.1667 | 26.7625 |  |  | 10 |  |  |  |  | 21
| Median |  |  | 5 | 26.97 |  |  | 10 |  |  |  |  | 21
| Std Dev. |  |  | 0.40825 | 0.61169 |  |  | 0 |  |  |  |  | 0.63246
| Rel.Std.Dev. |  |  | 7.902% | 2.286% |  |  | 0% |  |  |  |  | 3.012%
| PDM3 |  |  | -67.16% | 70.11% |  |  | -36.44% |  |  |  |  | 33.48%
| Z-Score (Absolute) |  | 10.11 | 0.67 | 88.05 | 0 | 0 | 20.23 | 0 | 0 | 20.23 | 0 | 64.73
| Analytical results for Ba in OREAS 239b (Certified Value 107 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 98.68 | 116.1123 | 127 | 117.475 | 112 | 109 | 138 | 109 | 100 | 100 | 110 | 110 | 104.3 | 107.3592 | 100 | 116.8 | 84 | 120.06 | 118.6 | 106 | 102.67 | 597.8267 |  |  | 1
| 1 | 2 | 101.62 | 115.5083 | 60.4 | 117.31 | 108 | 103 | 136 | 109 | 100 | 100 | 110 | 110 | 104 | 104.1112 | 101 | 114.3 | 83 | 114.14 | 119.1 | 107 | 102.46 | 593.8544 |  |  | 27
| 1 | 3 | 101.88 | 115.0355 | 97.4 | 117.365 | 107 | 105 | 144 | 105 | 100 | 100 | 110 | 100 | 104.7 | 103.495 | 99 | 109.3 | 84 | 117.06 | 116.5 | 107 | 101 | 587.3494 |  |  | 16
| 1 | 4 | 101.33 | 116.2458 | 56.6 | 117.455 | 112 | 107 | 138 | 113 | 100 | 100 | 110 | 110 | 103.7 | 104.4909 | 100 | 109.6 | 86 | 118.12 | 117.8 | 108 | 104.29 | 591.0363 |  |  | 106.9023
| 1 | 5 | 99.66 | 116.5065 | 52.7 | 117.005 | 108 | 110 | 144 | 108 | 100 | 90 | 110 | 110 | 100.9 | 111.317 | 99 | 117.8 | 81 | 118.23 | 118.2 | 109 | 104.37 | 599.0558 |  |  | 76
| 1 | 6 | 100.59 | 113.4574 | 120 | 118.635 | 105 | 107 | 141 | 109 | 100 | 90 | 110 | 110 | 101.6 | 104.5107 | 101 | 113.3 | 81 | 116 | 117.2 | 113 | 103.35 | 590.3864
| Mean |  | 100.6267 | 115.4776 | 85.6833 | 117.5408 | 108.6667 | 106.8333 | 140.1667 | 108.8333 | 100 | 96.6667 | 110 | 108.3333 | 103.2 | 105.8807 | 100 | 113.5167 | 83.1667 | 117.2683 | 117.9 | 108.3333 | 103.0233 | 593.2515
| Median |  | 100.96 | 115.8103 | 78.9 | 117.41 | 108 | 107 | 139.5 | 109 | 100 | 100 | 110 | 110 | 103.85 | 104.5008 | 100 | 113.8 | 83.5 | 117.59 | 118 | 107.5 | 103.01 | 592.4453
| Std Dev. |  | 1.2472 | 1.1251 | 33.4514 | 0.56218 | 2.8048 | 2.5626 | 3.3714 | 2.5626 | 0 | 5.164 | 0 | 4.0825 | 1.562 | 2.979 | 0.89443 | 3.5471 | 1.9408 | 2.0436 | 0.94657 | 2.5033 | 1.2697 | 4.5373
| Rel.Std.Dev. |  | 1.239% | 0.9743% | 39.04% | 0.4783% | 2.581% | 2.399% | 2.405% | 2.355% | 0% | 5.342% | 0% | 3.768% | 1.514% | 2.814% | 0.8944% | 3.125% | 2.334% | 1.743% | 0.8029% | 2.311% | 1.232% | 0.7648%
| PDM3 |  | -5.87% | 8.022% | -19.85% | 9.952% | 1.65% | -0.06455% | 31.12% | 1.806% | -6.457% | -9.575% | 2.898% | 1.339% | -3.463% | -0.9557% | -6.457% | 6.187% | -22.2% | 9.697% | 10.29% | 1.339% | -3.629% | 454.9%
| Z-Score (Absolute) |  | 0.7 | 0.65 | 2.06 | 0.84 | 0.03 | 0.14 | 2.89 | 0.05 | 0.76 | 1.06 | 0.15 | 0 | 0.47 | 0.22 | 0.76 | 0.47 | 2.29 | 0.81 | 0.87 | 0 | 0.48 | 44.04
| Analytical results for Be in OREAS 239b (Certified Value 1.19 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.14 | 1.4322 | 1.1 | < 5 | 1.1 | 1.18 | 1.43 | < 1 | 1.3 | 1.2 | 1.23 | 1.18 | 1.2195 | 1.2 | 1.173 | 1.3 | 1.25 | < 2 | 0.99 | 0.955 | 1.1246 | 1.1555 |  |  | 1
| 1 | 2 | 1.16 | 1.2545 | 1.2 | < 5 | 1.1 | 1.18 | 1.37 | 1 | 1.27 | 1.1 | 1.32 | 1.18 | 1.3251 | 1.2 | 1.147 | 1.3 | 1.28 | < 2 | 1.03 | 0.947 | 1.1247 | 1.1392 |  |  | 28
| 1 | 3 | 1.17 | 1.2754 | 1.1 | < 5 | 1.1 | 1.19 | 1.52 | 1 | 1.21 | 1.2 | 1.22 | 1.18 | 1.2131 | 1.2 | 1.201 | 1.2 | 1.22 | < 2 | 1.04 | 0.931 | 0.8645 | 1.1344 |  |  | 16
| 1 | 4 | 1.18 | 1.5054 | 1.2 | < 5 | 1.1 | 1.2 | 1.44 | 1 | 1.23 | 1.16 | 1.28 | 1.21 | 1.2023 | 1.2 | 1.119 | 1.2 | 1.19 | < 2 | 0.98 | 0.959 | 1.146 | 1.2039 |  |  | 1.1903
| 1 | 5 | 1.13 | 1.2918 | 1.2 | < 5 | 1.1 | 1.25 | 1.38 | < 1 | 1.18 | 1.12 | 1.23 | 1.23 | 1.321 | 1.3 | 1.222 | 1.3 | 1.2 | < 2 | 1.03 | 0.958 | 1.0592 | 1.1795 |  |  | 77
| 1 | 6 | 1.15 | 1.4217 | 1.3 | < 5 | 1.1 | 1.18 | 1.37 | 1 | 1.3 | 1.12 | 1.2 | 1.15 | 1.2295 | 1.2 | 1.196 | 1.3 | 1.19 | < 2 | 1.02 | 0.953 | 1.124 | 1.2566
| Mean |  | 1.155 | 1.3635 | 1.1833 |  | 1.1 | 1.1967 | 1.4183 | 1 | 1.2483 | 1.15 | 1.2467 | 1.1883 | 1.2518 | 1.2167 | 1.1763 | 1.2667 | 1.2217 |  | 1.015 | 0.9505 | 1.0738 | 1.1782
| Median |  | 1.155 | 1.3567 | 1.2 |  | 1.1 | 1.185 | 1.405 | 1 | 1.25 | 1.14 | 1.23 | 1.18 | 1.2245 | 1.2 | 1.1845 | 1.3 | 1.21 |  | 1.025 | 0.954 | 1.1243 | 1.1675
| Std Dev. |  | 0.018708 | 0.10297 | 0.075277 |  | 2.4324e-16 | 0.027325 | 0.058452 | 0 | 0.049565 | 0.043359 | 0.044572 | 0.027869 | 0.055935 | 0.040825 | 0.038051 | 0.05164 | 0.03656 |  | 0.02429 | 0.010464 | 0.1067 | 0.046416
| Rel.Std.Dev. |  | 1.62% | 7.552% | 6.361% |  | 2.211e-14% | 2.283% | 4.121% | 0% | 3.97% | 3.77% | 3.575% | 2.345% | 4.469% | 3.355% | 3.235% | 4.077% | 2.993% |  | 2.393% | 1.101% | 9.936% | 3.94%
| PDM3 |  | -2.964% | 14.55% | -0.5835% |  | -7.585% | 0.5367% | 19.16% | -15.99% | 4.877% | -3.384% | 4.737% | -0.1634% | 5.165% | 2.217% | -1.172% | 6.418% | 2.637% |  | -14.73% | -20.14% | -9.783% | -1.017%
| Z-Score (Absolute) |  | 0.37 | 1.97 | 0.06 | 14.74 | 0.99 | 0.09 | 2.58 | NA | 0.67 | 0.43 | 0.66 | 0 | 0.71 | 0.32 | 0.13 | 0.88 | 0.37 | 2.12 | 1.95 | 2.67 | 1.29 | 0.11
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 239b (Certified Value 0.37 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.37 | 0.35229 | 0.42 | < 2 | 0.48 | 0.36 | 0.38 | 0.4 | 0.36 | 0.37 | 0.35 | 0.4 | 0.34 | 0.39864 | 3 | 0.479 | 0.38 | 41.18 | < 5 | 0.35 | 0.5481 | 1.6692 |  |  | 1
| 1 | 2 | 0.37 | 0.36896 | 0.39 | < 2 | 0.47 | 0.38 | 0.37 | 0.4 | 0.36 | 0.35 | 0.37 | 0.39 | 0.37 | 0.4257 | 5 | 0.494 | 0.37 | 43.72 | < 5 | 0.36 | 0.5687 | 1.8902 |  |  | 29
| 1 | 3 | 0.36 | 0.36523 | 0.41 | < 2 | 0.47 | 0.37 | 0.39 | 0.4 | 0.36 | 0.39 | 0.37 | 0.41 | 0.35 | 0.38139 | 3 | 0.439 | 0.37 | 21.69 | < 5 | 0.35 | 0.5267 | 2.0405 |  |  | 16
| 1 | 4 | 0.37 | 0.32906 | 0.4 | < 2 | 0.46 | 0.43 | 0.36 | 0.4 | 0.48 | 0.37 | 0.35 | 0.42 | 0.35 | 0.36731 | 3 | 0.437 | 0.38 | 22.69 | < 5 | 0.33 | 0.5119 | 1.699 |  |  | 0.37393
| 1 | 5 | 0.36 | 0.37361 | 0.38 | < 2 | 0.48 | 0.39 | 0.38 | 0.4 | 0.35 | 0.36 | 0.38 | 0.41 | 0.36 | 0.40752 | 3 | 0.461 | 0.35 | 39.04 | < 5 | 0.34 | 0.3944 | 1.5523 |  |  | 78
| 1 | 6 | 0.37 | 0.36082 | 0.41 | < 2 | 0.46 | 0.39 | 0.4 | 0.4 | 0.35 | 0.37 | 0.37 | 0.4 | 0.34 | 0.39465 | 3 | 0.48 | 0.4 | 45.77 | < 5 | 0.34 | 0.5107 | 1.3073
| Mean |  | 0.36667 | 0.35833 | 0.40167 |  | 0.47 | 0.38667 | 0.38 | 0.4 | 0.37667 | 0.36833 | 0.365 | 0.405 | 0.35167 | 0.39587 | 3.3333 | 0.465 | 0.375 | 35.6817 |  | 0.345 | 0.51008 | 1.6931
| Median |  | 0.37 | 0.36302 | 0.405 |  | 0.47 | 0.385 | 0.38 | 0.4 | 0.36 | 0.37 | 0.37 | 0.405 | 0.35 | 0.39664 | 3 | 0.47 | 0.375 | 40.11 |  | 0.345 | 0.5193 | 1.6841
| Std Dev. |  | 0.005164 | 0.016085 | 0.01472 |  | 0.0089443 | 0.024221 | 0.014142 | 6.0809e-17 | 0.050859 | 0.013292 | 0.012247 | 0.010488 | 0.01169 | 0.020298 | 0.8165 | 0.023401 | 0.016432 | 10.7 |  | 0.010488 | 0.060899 | 0.25644
| Rel.Std.Dev. |  | 1.408% | 4.489% | 3.665% |  | 1.903% | 6.264% | 3.722% | 1.52e-14% | 13.5% | 3.609% | 3.355% | 2.59% | 3.324% | 5.128% | 24.49% | 5.032% | 4.382% | 29.99% |  | 3.04% | 11.94% | 15.15%
| PDM3 |  | -1.942% | -4.171% | 7.418% |  | 25.69% | 3.407% | 1.624% | 6.973% | 0.7327% | -1.496% | -2.387% | 8.31% | -5.953% | 5.867% | 791.4% | 24.36% | 0.287% | 9442% |  | -7.736% | 36.41% | 352.8%
| Z-Score (Absolute) |  | 0.56 | 0.75 | 0.24 | 13.87 | 1.79 | 0.1 | 0.26 | NA | 0.33 | 0.52 | 0.6 | 0.31 | 0.9 | 0.1 | NA | 1.68 | 0.37 | 803.87 | 48.03 | 1.05 | 2.71 | 29.65
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 239b (Certified Value 0.333 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.315 | 0.34685 | 0.39 | 0.38 | 0.34 | 0.34 | 0.35 | 0.35 | 0.31 | 0.35 | 0.35 | 0.33 | 0.33 | 0.35134 | 0.314 | 0.29 | 0.35128 | 0.31 | 0.32 | 0.2897 |  |  |  |  | 1
| 1 | 2 | 0.32 | 0.34081 | 0.38 | 0.39 | 0.34 | 0.34 | 0.34 | 0.36 | 0.32 | 0.33 | 0.34 | 0.33 | 0.33 | 0.35129 | 0.322 | 0.29 | 0.35263 | 0.31 | 0.32 | 0.2868 |  |  |  |  | #N/A
| 1 | 3 | 0.327 | 0.35738 | 0.39 | 0.381 | 0.35 | 0.34 | 0.35 | 0.36 | 0.31 | 0.34 | 0.34 | 0.33 | 0.34 | 0.35343 | 0.314 | 0.3 | 0.34894 | 0.31 | 0.32 | 0.2877 |  |  |  |  | 16
| 1 | 4 | 0.324 | 0.34375 | 0.4 | 0.39 | 0.35 | 0.33 | 0.35 | 0.36 | 0.32 | 0.33 | 0.34 | 0.34 | 0.33 | 0.35156 | 0.322 | 0.3 | 0.35326 | 0.31 | 0.32 | 0.2931 |  |  |  |  | 0.33321
| 1 | 5 | 0.313 | 0.3481 | 0.39 | 0.38633 | 0.35 | 0.33 | 0.35 | 0.35 | 0.31 | 0.33 | 0.34 | 0.33 | 0.32 | 0.34936 | 0.322 | 0.28 | 0.3523 | 0.31 | 0.32 | 0.2946 |  |  |  |  | 79
| 1 | 6 | 0.321 | 0.34991 | 0.41 | 0.39 | 0.34 | 0.34 | 0.36 | 0.36 | 0.32 | 0.33 | 0.34 | 0.34 | 0.32 | 0.35481 | 0.329 | 0.28 | 0.34555 | 0.31 | 0.33 | 0.2913
| Mean |  | 0.32 | 0.3478 | 0.39333 | 0.38622 | 0.345 | 0.33667 | 0.35 | 0.35667 | 0.315 | 0.335 | 0.34167 | 0.33333 | 0.32833 | 0.35196 | 0.3205 | 0.29 | 0.35066 | 0.31 | 0.32167 | 0.29053
| Median |  | 0.3205 | 0.34747 | 0.39 | 0.38817 | 0.345 | 0.34 | 0.35 | 0.36 | 0.315 | 0.33 | 0.34 | 0.33 | 0.33 | 0.35145 | 0.322 | 0.29 | 0.35179 | 0.31 | 0.32 | 0.2905
| Std Dev. |  | 0.0052915 | 0.0057021 | 0.010328 | 0.0046651 | 0.0054772 | 0.005164 | 0.0063246 | 0.005164 | 0.0054772 | 0.0083666 | 0.0040825 | 0.005164 | 0.0075277 | 0.001901 | 0.0057184 | 0.0089443 | 0.0029241 | 0 | 0.0040825 | 0.0030454
| Rel.Std.Dev. |  | 1.654% | 1.639% | 2.626% | 1.208% | 1.588% | 1.534% | 1.807% | 1.448% | 1.739% | 2.497% | 1.195% | 1.549% | 2.293% | 0.5401% | 1.784% | 3.084% | 0.8339% | 0% | 1.269% | 1.048%
| PDM3 |  | -3.964% | 4.38% | 18.04% | 15.91% | 3.539% | 1.038% | 5.04% | 7.041% | -5.464% | 0.5381% | 2.539% | 0.0379% | -1.463% | 5.629% | -3.814% | -12.97% | 5.238% | -6.965% | -3.463% | -12.81%
| Z-Score (Absolute) |  | 0.71 | 0.53 | 2.56 | 2.25 | 0.41 | 0.04 | 0.63 | 0.93 | 0.93 | 0.04 | 0.26 | 0.11 | 0.33 | 0.72 | 0.69 | 2.04 | 0.66 | 1.15 | 0.63 | 2.02
| Analytical results for Cd in OREAS 239b (Certified Value 0.072 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.054 | 1.24 | 0.06 | 0.67077 | 0.09 | 0.08 | 0.11 | 0.2 | 0.06 | 0.07 | 0.05 | 0.06 | 0.07 | 0.08565 | < 1 | < 0.1 | 0.06 | 28.5 | 15.4 | 0.06 | 0.97 | 0.3534 |  |  | 1
| 1 | 2 | 0.062 | 1.29 | 0.02 | 0.68157 | 0.09 | 0.1 | 0.1 | < 0.1 | 0.07 | 0.06 | 0.06 | 0.06 | 0.07 | 0.080529 | < 1 | < 0.1 | 0.06 | 28.74 | 15.1 | 0.06 | 0.9 | 0.3407 |  |  | 30
| 1 | 3 | 0.055 | 1.22 | 0.11 | 0.66753 | 0.09 | 0.08 | 0.1 | < 0.1 | 0.07 | 0.06 | 0.06 | 0.06 | 0.07 | 0.079558 | < 1 | < 0.1 | 0.06 | 28.44 | 15.8 | 0.07 | 0.78 | 0.3305 |  |  | 16
| 1 | 4 | 0.06 | 1.28 | 0.08 | 0.67893 | 0.09 | 0.08 | 0.1 | 0.2 | 0.06 | 0.06 | 0.06 | 0.06 | 0.08 | 0.083462 | < 1 | < 0.1 | 0.06 | 28 | 15.5 | 0.07 | 0.98 | 0.3431 |  |  | 0.072145
| 1 | 5 | 0.067 | 1.2 | 0.05 | 0.67467 | 0.09 | 0.09 | 0.1 | < 0.1 | 0.06 | 0.06 | 0.06 | 0.07 | 0.07 | 0.081147 | < 1 | < 0.1 | 0.06 | 27.58 | 15.7 | 0.06 | 0.85 | 0.3431 |  |  | 80
| 1 | 6 | 0.065 | 1.24 | 0.04 | 0.6732 | 0.09 | 0.09 | 0.11 | < 0.1 | 0.07 | 0.06 | 0.06 | 0.08 | 0.08 | 0.085953 | < 1 | < 0.1 | 0.06 | 27.29 | 14.9 | 0.06 | 0.89 | 0.3379
| Mean |  | 0.0605 | 1.245 | 0.06 | 0.67444 | 0.09 | 0.086667 | 0.10333 | 0.2 | 0.065 | 0.061667 | 0.058333 | 0.065 | 0.073333 | 0.082716 |  |  | 0.06 | 28.0917 | 15.4 | 0.063333 | 0.895 | 0.34145
| Median |  | 0.061 | 1.24 | 0.055 | 0.67393 | 0.09 | 0.085 | 0.1 | 0.2 | 0.065 | 0.06 | 0.06 | 0.06 | 0.07 | 0.082304 |  |  | 0.06 | 28.22 | 15.45 | 0.06 | 0.895 | 0.3419
| Std Dev. |  | 0.005244 | 0.034496 | 0.031623 | 0.0051723 | 1.5202e-17 | 0.008165 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0083666 | 0.005164 | 0.0027149 |  |  | 0 | 0.56944 | 0.34641 | 0.005164 | 0.075033 | 0.0075009
| Rel.Std.Dev. |  | 8.668% | 2.771% | 52.7% | 0.7669% | 1.689e-14% | 9.421% | 4.997% | 0% | 8.427% | 6.62% | 6.999% | 12.87% | 7.042% | 3.282% |  |  | 0% | 2.027% | 2.249% | 8.154% | 8.384% | 2.197%
| PDM3 |  | -16.14% | 1626% | -16.83% | 834.8% | 24.75% | 20.13% | 43.23% | 177.2% | -9.903% | -14.52% | -19.14% | -9.903% | 1.647% | 14.65% |  |  | -16.83% | 38838% | 21246% | -12.21% | 1141% | 373.3%
| Z-Score (Absolute) |  | 0.66 | 34.5 | 0.67 | 17.56 | 0.22 | 0.12 | 0.61 | NA | 0.53 | 0.62 | 0.72 | 0.53 | 0.28 | 0 | 12.39 | 0.97 | 0.67 | 831.41 | 454.67 | 0.58 | 24.11 | 7.68
| Analytical results for Ce in OREAS 239b (Certified Value 54 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 55.767 | 74.91 | 56.7 | 66.79 | 50.77 | 69.5 | 56 | 53.8 | 55.4 | 56.7 | 56.7 | 53.1065 | 56.62 | 55.69 | 47.92 | 40.7869 | 0.2189 | [7 empty] | 1
| 1 | 2 | 55.625 | 74.9643 | 51.3 | 65.91 | 47.77 | 73.7 | 56 | 55.1 | 47.3 | 57.5 | 58.2 | 53.4663 | 59.94 | 54.22 | 49.49 | 40.7014 | 0.216 | [7 empty] | 31
| 1 | 3 | 55.843 | 75.1588 | 54.5 | 65.82 | 48.2 | 72.3 | 55 | 53.8 | 53 | 56.2 | 55.7 | 53.8407 | 59.8 | 52.58 | 49.78 | 37.9806 | 0.2157 | [7 empty] | 16
| 1 | 4 | 57.918 | 75.1229 | 55.2 | 67.01 | 48.09 | 74.6 | 55 | 55.3 | 52 | 55.9 | 58.3 | 53.5533 | 64.08 | 56.22 | 47.42 | 39.9218 | 0.2333 | [7 empty] | 54.4618
| 1 | 5 | 57.113 | 75.4137 | 50.8 | 66.53 | 50.17 | 68.9 | 55 | 53 | 53.6 | 55.9 | 58.5 | 53.9143 | 60.01 | 55.2 | 48.35 | 32.6344 | 0.2315 | [7 empty] | 81
| 1 | 6 | 55.151 | 75.0231 | 54.3 | 66.26 | 53.33 | 74.6 | 56 | 56.4 | 54.1 | 55.9 | 55.6 | 52.8627 | 57.55 | 55.07 | 48.8 | 40.7923 | 0.2164
| Mean |  | 56.2362 | 75.0988 | 53.8 | 66.3867 | 49.7217 | 72.2667 | 55.5 | 54.5667 | 52.5667 | 56.35 | 57.1667 | 53.4573 | 59.6667 | 54.83 | 48.6267 | 38.8029 | 0.22197
| Median |  | 55.805 | 75.073 | 54.4 | 66.395 | 49.185 | 73 | 55.5 | 54.45 | 53.3 | 56.05 | 57.45 | 53.5098 | 59.87 | 55.135 | 48.575 | 40.3116 | 0.21765
| Std Dev. |  | 1.051 | 0.18044 | 2.2961 | 0.47668 | 2.1497 | 2.5272 | 0.54772 | 1.2501 | 2.8176 | 0.6442 | 1.3382 | 0.41032 | 2.5865 | 1.2885 | 0.90967 | 3.2086 | 0.0081804
| Rel.Std.Dev. |  | 1.869% | 0.2403% | 4.268% | 0.718% | 4.324% | 3.497% | 0.9869% | 2.291% | 5.36% | 1.143% | 2.341% | 0.7676% | 4.335% | 2.35% | 1.871% | 8.269% | 3.685%
| PDM3 |  | 3.258% | 37.89% | -1.215% | 21.9% | -8.704% | 32.69% | 1.906% | 0.1925% | -3.48% | 3.467% | 4.967% | -1.844% | 9.557% | 0.676% | -10.71% | -28.75% | -99.59%
| Z-Score (Absolute) |  | 0.41 | 5.85 | 0.3 | 3.33 | 1.47 | 5.03 | 0.19 | 0.08 | 0.65 | 0.44 | 0.67 | 0.4 | 1.4 | 0 | 1.79 | 4.63 | 15.76
| Analytical results for Co in OREAS 239b (Certified Value 15.2 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 14.67 | 16.902 | 16.1 | 16.9041 | 15.4 | 15.3 | 15 | 13.9 | 15.9 | 15.8 | 15 | 14.8 | 15.8 | 15.7857 | 14 | 12.8 | 15.4 | 24.47 | 19 | 14.5 | 16.245 | 15.4043 | 14.965 |  | 1
| 1 | 2 | 14.74 | 15.768 | 16.6 | 17.1087 | 15 | 15.1 | 15.8 | 13.8 | 15.9 | 14.7 | 15.6 | 15 | 15.6 | 14.8175 | 13 | 29.99 | 15.8 | 19.44 | 19.3 | 14.6 | 15.865 | 14.2085 | 14.7647 |  | 32
| 1 | 3 | 14.84 | 17.199 | 16.5 | 16.6799 | 14.9 | 15.2 | 15.6 | 13.6 | 15.8 | 15.6 | 15.1 | 14.7 | 16.8 | 15.0951 | 13 | 12.74 | 15.6 | 20.84 | 18.7 | 14.6 | 15.96 | 14.4826 | 15.0975 |  | 16
| 1 | 4 | 14.67 | 16.11 | 17 | 16.9305 | 15.4 | 15.5 | 15.6 | 13.7 | 16 | 15 | 15 | 14.6 | 17.2 | 16.0702 | 13 | 19.64 | 15.5 | 24.52 | 18.3 | 14.6 | 16.34 | 13.9279 | 15.0341 |  | 15.1646
| 1 | 5 | 14.72 | 15.84 | 15.5 | 16.9767 | 15 | 15.9 | 15.2 | 13.5 | 15.4 | 14.7 | 15.2 | 15 | 15.5 | 15.2652 | 13 | 15.29 | 15.6 | 20.15 | 19 | 14.8 | 16.435 | 13.6496 | 14.9582 |  | 82
| 1 | 6 | 14.7 | 16.668 | 16.7 | 16.9093 | 14.9 | 15.6 | 15.8 | 13.8 | 15.6 | 14.6 | 15 | 14.4 | 15.2 | 15.7555 | 14 | 12.68 | 14.8 | 19.43 | 18.6 | 14.3 | 16.34 | 14.1999 | 15.0412
| Mean |  | 14.7233 | 16.4145 | 16.4 | 16.9182 | 15.1 | 15.4333 | 15.5 | 13.7167 | 15.7667 | 15.0667 | 15.15 | 14.75 | 16.0167 | 15.4649 | 13.3333 | 17.19 | 15.45 | 21.475 | 18.8167 | 14.5667 | 16.1975 | 14.3121 | 14.9768
| Median |  | 14.71 | 16.389 | 16.55 | 16.9199 | 15 | 15.4 | 15.6 | 13.75 | 15.85 | 14.85 | 15.05 | 14.75 | 15.7 | 15.5104 | 13 | 14.045 | 15.55 | 20.495 | 18.85 | 14.6 | 16.2925 | 14.2042 | 14.9995
| Std Dev. |  | 0.063456 | 0.59297 | 0.52915 | 0.13923 | 0.23664 | 0.29439 | 0.32863 | 0.1472 | 0.22509 | 0.51251 | 0.23452 | 0.23452 | 0.79603 | 0.47949 | 0.5164 | 6.8245 | 0.34496 | 2.3968 | 0.35449 | 0.1633 | 0.23075 | 0.60528 | 0.11617
| Rel.Std.Dev. |  | 0.431% | 3.613% | 3.227% | 0.823% | 1.567% | 1.908% | 2.12% | 1.073% | 1.428% | 3.402% | 1.548% | 1.59% | 4.97% | 3.101% | 3.873% | 39.7% | 2.233% | 11.16% | 1.884% | 1.121% | 1.425% | 4.229% | 0.7757%
| PDM3 |  | -2.91% | 8.242% | 8.147% | 11.56% | -0.4259% | 1.772% | 2.212% | -9.548% | 3.97% | -0.6457% | -0.09621% | -2.734% | 5.619% | 1.98% | -12.08% | 13.36% | 1.882% | 41.61% | 24.08% | -3.943% | 6.811% | -5.621% | -1.238%
| Z-Score (Absolute) |  | 0.67 | 0.9 | 0.88 | 1.36 | 0.32 | 0.02 | 0.05 | 1.61 | 0.29 | 0.36 | 0.28 | 0.65 | 0.53 | 0.01 | 1.96 | 1.61 | 0 | 5.59 | 3.12 | 0.82 | 0.69 | 1.06 | 0.44
| Analytical results for Cr in OREAS 239b (Certified Value 107 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 107.1 | 108.3665 | 96 | 114.11 | 100 | 111 | 112 | 110 | 104 | 113 | 107 | 103 | 103.7 | 97.9505 | 103 | 113.7 | 97 | 110.68 | 105 | 100 | 92.8661 | 108.8338 |  |  | 1
| 1 | 2 | 109.2 | 105.1725 | 97 | 115.1433 | 99 | 110 | 111 | 107 | 107 | 109 | 107 | 104 | 105.1 | 107.3426 | 105 | 113.4 | 98 | 104.35 | 106 | 99.3 | 91.9643 | 106.0429 |  |  | 33
| 1 | 3 | 110.6 | 109.0695 | 96 | 114.4867 | 98 | 111 | 112 | 104 | 104 | 112 | 106 | 103 | 105.4 | 104.3904 | 103 | 114.5 | 97 | 106.07 | 105 | 99.1 | 76.1227 | 107.1198 |  |  | 16
| 1 | 4 | 109.7 | 102.581 | 100 | 114.7733 | 100 | 111 | 112 | 110 | 105 | 108 | 108 | 104 | 104.3 | 104.0267 | 105 | 114.3 | 100 | 107.08 | 107 | 101.8 | 86.5173 | 107.4684 |  |  | 106.5348
| 1 | 5 | 108.4 | 105.5075 | 94 | 114.5167 | 98 | 109 | 111 | 106 | 103 | 107 | 108 | 104 | 101.7 | 102.3174 | 106 | 117.4 | 92 | 107.33 | 108 | 102.1 | 91.5813 | 108.4731 |  |  | 83
| 1 | 6 | 108.5 | 107.1695 | 99 | 115.8967 | 97 | 111 | 115 | 108 | 104 | 108 | 107 | 105 | 104.1 | 102.9092 | 108 | 118.8 | 91 | 108.46 | 110 | 100.4 | 90.0652 | 107.9436
| Mean |  | 108.9167 | 106.3111 | 97 | 114.8211 | 98.6667 | 110.5 | 112.1667 | 107.5 | 104.5 | 109.5 | 107.1667 | 103.8333 | 104.05 | 103.1561 | 105 | 115.35 | 95.8333 | 107.3283 | 106.8333 | 100.45 | 88.1861 | 107.6469
| Median |  | 108.85 | 106.3385 | 96.5 | 114.645 | 98.5 | 111 | 112 | 107.5 | 104 | 108.5 | 107 | 104 | 104.2 | 103.4679 | 105 | 114.4 | 97 | 107.205 | 106.5 | 100.2 | 90.8233 | 107.706
| Std Dev. |  | 1.2057 | 2.3844 | 2.1909 | 0.6282 | 1.2111 | 0.83666 | 1.472 | 2.3452 | 1.3784 | 2.429 | 0.75277 | 0.75277 | 1.3142 | 3.0878 | 1.8974 | 2.2116 | 3.5449 | 2.1462 | 1.9408 | 1.2566 | 6.3174 | 1.0062
| Rel.Std.Dev. |  | 1.107% | 2.243% | 2.259% | 0.5471% | 1.227% | 0.7572% | 1.312% | 2.182% | 1.319% | 2.218% | 0.7024% | 0.725% | 1.263% | 2.993% | 1.807% | 1.917% | 3.699% | 2% | 1.817% | 1.251% | 7.164% | 0.9347%
| PDM3 |  | 2.236% | -0.21% | -8.95% | 7.778% | -7.386% | 3.722% | 5.286% | 0.906% | -1.91% | 2.783% | 0.5931% | -2.536% | -2.332% | -3.171% | -1.441% | 8.274% | -10.05% | 0.7448% | 0.2802% | -5.712% | -17.22% | 1.044%
| Z-Score (Absolute) |  | 0.56 | 0.06 | 2.28 | 1.96 | 1.88 | 0.93 | 1.33 | 0.22 | 0.49 | 0.7 | 0.14 | 0.65 | 0.6 | 0.81 | 0.37 | 2.09 | 2.56 | 0.18 | 0.06 | 1.46 | 4.38 | 0.26
| Analytical results for Cs in OREAS 239b (Certified Value 6.71 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.20 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 6.56 | 5.2367 | 6.69 | 6.94 | 6.49 | 7.58 | 6.6 | 6.9 | 7.07 | 6.65 | 6.78 | 6.989 | 6.86 | 6.5 | 7.3361 | [9 empty] | 1
| 1 | 2 | 6.66 | 5.224 | 6.42 | 6.75 | 6.39 | 8.07 | 6.7 | 6.78 | 6.49 | 7.02 | 6.91 | 6.536 | 7.11 | 6.66 | 7.361 | [9 empty] | 34
| 1 | 3 | 6.64 | 5.3091 | 6.75 | 6.68 | 6.47 | 7.96 | 6.5 | 6.8 | 7.06 | 6.62 | 6.85 | 6.561 | 6.77 | 6.58 | 7.2143 | [9 empty] | 16
| 1 | 4 | 6.63 | 5.4303 | 6.77 | 6.86 | 6.63 | 7.89 | 6.6 | 6.78 | 6.78 | 6.8 | 6.97 | 6.473 | 6.9 | 6.35 | 7.3803 | [9 empty] | 6.7143
| 1 | 5 | 6.59 | 5.2289 | 6.25 | 6.78 | 6.6 | 7.58 | 6.6 | 6.74 | 6.7 | 6.67 | 7.07 | 6.634 | 6.97 | 6.43 | 6.0545 | [9 empty] | 84
| 1 | 6 | 6.63 | 5.2886 | 6.72 | 6.6 | 6.58 | 8.12 | 6.7 | 6.69 | 6.83 | 6.79 | 6.7 | 7.024 | 7.02 | 6.41 | 7.392
| Mean |  | 6.6183 | 5.2863 | 6.6 | 6.7683 | 6.5267 | 7.8667 | 6.6167 | 6.7817 | 6.8217 | 6.7583 | 6.88 | 6.7028 | 6.9383 | 6.4883 | 7.123
| Median |  | 6.63 | 5.2626 | 6.705 | 6.765 | 6.535 | 7.925 | 6.6 | 6.78 | 6.805 | 6.73 | 6.88 | 6.5975 | 6.935 | 6.465 | 7.3485
| Std Dev. |  | 0.03656 | 0.078614 | 0.21392 | 0.12205 | 0.091797 | 0.23628 | 0.075277 | 0.069976 | 0.2214 | 0.14825 | 0.13297 | 0.24106 | 0.1209 | 0.11548 | 0.52738
| Rel.Std.Dev. |  | 0.5524% | 1.487% | 3.241% | 1.803% | 1.406% | 3.004% | 1.138% | 1.032% | 3.245% | 2.194% | 1.933% | 3.596% | 1.742% | 1.78% | 7.404%
| PDM3 |  | -1.429% | -21.27% | -1.702% | 0.8053% | -2.794% | 17.16% | -1.454% | 1.004% | 1.6% | 0.6564% | 2.468% | -0.1702% | 3.337% | -3.365% | 6.088%
| Z-Score (Absolute) |  | 0.67 | 7.01 | 0.75 | 0.05 | 1.1 | 5.28 | 0.67 | 0.11 | 0.3 | 0 | 0.58 | 0.26 | 0.86 | 1.29 | 1.74
| Analytical results for Cu in OREAS 239b (Certified Value 29.6 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 29.61 | 27.508 | 31.2 | 28.63 | 30 | 29 | 28 | 29.5 | 30.4 | 30.7 | 29.4 | 30.9 | 29.11 | 32.0035 | 31 | 27.44 | 27.7 | 31.01 | 31.9 | 30.7 | 31.312 | 30 | 1 | 29.7129 | 1
| 1 | 2 | 28.62 | 27.023 | 30.4 | 29.2133 | 30 | 28 | 27 | 28.4 | 31 | 28.2 | 31.1 | 31 | 29.84 | 30.683 | 32 | 26.81 | 27.6 | 28.55 | 32.1 | 30.6 | 30.552 | 30 | 1 | 29.3158 | 35
| 1 | 3 | 29.3 | 28.337 | 31.5 | 28.82 | 31 | 27 | 28 | 27.3 | 30.2 | 30.6 | 29.6 | 30.8 | 29.83 | 31.5538 | 30 | 26.76 | 27.2 | 29.18 | 31.9 | 30.1 | 31.027 | 30 | 2 | 29.5645 | 16
| 1 | 4 | 29.41 | 26.744 | 31 | 28.6667 | 31 | 29 | 29 | 28.2 | 31 | 29.2 | 29.6 | 31.2 | 29.24 | 31.3384 | 31 | 27.22 | 28.2 | 29.38 | 32 | 30.5 | 31.236 | 30 | 1 | 29.6154 | 29.557
| 1 | 5 | 28.7 | 27.249 | 29.7 | 28.5967 | 31 | 26 | 30 | 27.5 | 29.7 | 29.1 | 29.4 | 31.2 | 28.44 | 30.862 | 30 | 25.93 | 26 | 29.42 | 31.7 | 30.7 | 31.274 | 31 | 1 | 29.2567 | 85
| 1 | 6 | 29.85 | 28.04 | 31.2 | 28.9567 | 31 | 29 | 30 | 28.3 | 30.2 | 28.7 | 29.8 | 29.8 | 29.64 | 31.114 | 31 | 39.32 | 25.7 | 29.37 | 31.4 | 29.7 | 31.046 | 30 | 1 | 29.5705
| Mean |  | 29.2483 | 27.4835 | 30.8333 | 28.8139 | 30.6667 | 28 | 28.6667 | 28.2 | 30.4167 | 29.4167 | 29.8167 | 30.8167 | 29.35 | 31.2591 | 30.8333 | 28.9133 | 27.0667 | 29.485 | 31.8333 | 30.3833 | 31.0745 | 30.1667 | 1.1667 | 29.506
| Median |  | 29.355 | 27.3785 | 31.1 | 28.7433 | 31 | 28.5 | 28.5 | 28.25 | 30.3 | 29.15 | 29.6 | 30.95 | 29.44 | 31.2262 | 31 | 27.015 | 27.4 | 29.375 | 31.9 | 30.55 | 31.141 | 30 | 1 | 29.5675
| Std Dev. |  | 0.49337 | 0.60871 | 0.66533 | 0.23776 | 0.5164 | 1.2649 | 1.2111 | 0.77974 | 0.50761 | 1.0187 | 0.64627 | 0.52313 | 0.53896 | 0.48124 | 0.75277 | 5.1244 | 0.99933 | 0.81503 | 0.25033 | 0.40208 | 0.28226 | 0.40825 | 0.40825 | 0.17927
| Rel.Std.Dev. |  | 1.687% | 2.215% | 2.158% | 0.8252% | 1.684% | 4.518% | 4.225% | 2.765% | 1.669% | 3.463% | 2.167% | 1.698% | 1.836% | 1.54% | 2.441% | 17.72% | 3.692% | 2.764% | 0.7864% | 1.323% | 0.9083% | 1.353% | 34.99% | 0.6076%
| PDM3 |  | -1.044% | -7.015% | 4.318% | -2.514% | 3.754% | -5.268% | -3.012% | -4.591% | 2.909% | -0.4747% | 0.8786% | 4.262% | -0.7003% | 5.759% | 4.318% | -2.178% | -8.425% | -0.2435% | 7.702% | 2.796% | 5.134% | 2.063% | -96.05% | -0.1725%
| Z-Score (Absolute) |  | 0.16 | 1.3 | 0.86 | 0.44 | 0.75 | 0.96 | 0.53 | 0.84 | 0.59 | 0.05 | 0.21 | 0.85 | 0.09 | 1.14 | 0.86 | 0.38 | 1.57 | 0.01 | 1.51 | 0.57 | 1.02 | 0.43 | 18.26 | 0.01
| Analytical results for Dy in OREAS 239b (Indicative Value 2.02 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 2.1367 | 2.1 | 2.021 | 1.9538 | [20 empty] | 1
| 1 | 2 | 2.1732 | 2 | 2.134 | 1.9707 | [20 empty] | 9
| 1 | 3 | 2.109 | 2 | 2.128 | 1.9029 | [20 empty] | 16
| 1 | 4 | 2.1567 | 2 | 2.251 | 1.939 | [20 empty] | 2.0209
| 1 | 5 | 2.1325 | 1.7 | 2.123 | 1.5613 | [20 empty] | 15
| 1 | 6 | 2.1258 | 1.9 | 2.048 | 1.9345
| Mean |  | 2.139 | 1.95 | 2.1175 | 1.877
| Median |  | 2.1346 | 2 | 2.1255 | 1.9368
| Std Dev. |  | 0.022826 | 0.13784 | 0.08047 | 0.15631
| Rel.Std.Dev. |  | 1.067% | 7.069% | 3.8% | 8.328%
| PDM3 |  | 5.845% | -3.508% | 4.781% | -7.118%
| Z-Score (Absolute) |  | 0.75 | 0.6 | 0.6 | 1.12
| Analytical results for Er in OREAS 239b (Indicative Value 0.91 ppm) | [25 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.94428 | 0.9 | 0.913 | 0.8702 | [20 empty] | 1
| 1 | 2 | 0.86803 | 1 | 0.976 | 0.8698 | [20 empty] | 10
| 1 | 3 | 0.92658 | 0.9 | 0.953 | 0.846 | [20 empty] | 16
| 1 | 4 | 0.86928 | 0.9 | 1.006 | 0.8698 | [20 empty] | 0.91439
| 1 | 5 | 0.88646 | 1 | 0.966 | 0.7022 | [20 empty] | 16
| 1 | 6 | 0.91746 | 0.9 | 0.926 | 0.872
| Mean |  | 0.90202 | 0.93333 | 0.95667 | 0.83833
| Median |  | 0.90196 | 0.9 | 0.9595 | 0.8698
| Std Dev. |  | 0.031926 | 0.05164 | 0.033927 | 0.06741
| Rel.Std.Dev. |  | 3.539% | 5.533% | 3.546% | 8.041%
| PDM3 |  | -1.354% | 2.071% | 4.623% | -8.318%
| Z-Score (Absolute) |  | 0.39 | 0.39 | 0.96 | 1.96
| Analytical results for Eu in OREAS 239b (Indicative Value 0.64 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.71105 | 0.7 | 0.642 | 0.5801 | [20 empty] | 1
| 1 | 2 | 0.70381 | 0.6 | 0.679 | 0.5688 | [20 empty] | 11
| 1 | 3 | 0.71062 | 0.7 | 0.674 | 0.5384 | [20 empty] | 16
| 1 | 4 | 0.69567 | 0.6 | 0.715 | 0.5692 | [20 empty] | 0.64426
| 1 | 5 | 0.69566 | 0.6 | 0.683 | 0.4682 | [20 empty] | 17
| 1 | 6 | 0.69902 | 0.6 | 0.648 | 0.5814
| Mean |  | 0.70264 | 0.63333 | 0.6735 | 0.55102
| Median |  | 0.70142 | 0.6 | 0.6765 | 0.569
| Std Dev. |  | 0.0070165 | 0.05164 | 0.026403 | 0.043437
| Rel.Std.Dev. |  | 0.9986% | 8.154% | 3.92% | 7.883%
| PDM3 |  | 9.061% | -1.696% | 4.538% | -14.47%
| Z-Score (Absolute) |  | 0.96 | 0.39 | 0.39 | 1.99
| Analytical results for Fe in OREAS 239b (Certified Value 3.51 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.341 | 3.5873 | 3.53 | 3.685 | 3.47 | 3.54 | 3.89 | 3.54 | 3.39 | 3.62 | 3.45 | 3.37 | 3.53 | 3.5694 | 3.455 | 3.21 | 3.6054 | 3.67 | 3.5 | 3.29 | 4.1433 |  |  |  | 1
| 1 | 2 | 3.379 | 3.4991 | 3.54 | 3.7167 | 3.48 | 3.51 | 3.85 | 3.56 | 3.45 | 3.49 | 3.43 | 3.36 | 3.56 | 3.6479 | 3.497 | 3.24 | 3.5983 | 3.69 | 3.53 | 3.2821 | 4.0918 |  |  |  | #N/A
| 1 | 3 | 3.41 | 3.485 | 3.45 | 3.72 | 3.51 | 3.5 | 3.86 | 3.55 | 3.4 | 3.6 | 3.43 | 3.34 | 3.57 | 3.6048 | 3.469 | 3.2 | 3.5771 | 3.66 | 3.52 | 3.2729 | 4.0539 |  |  |  | 16
| 1 | 4 | 3.397 | 3.515 | 3.63 | 3.6967 | 3.51 | 3.57 | 3.89 | 3.57 | 3.46 | 3.45 | 3.47 | 3.39 | 3.5 | 3.6448 | 3.525 | 3.32 | 3.6106 | 3.69 | 3.56 | 3.3353 | 4.0988 |  |  |  | 3.5146
| 1 | 5 | 3.362 | 3.4479 | 3.35 | 3.69 | 3.49 | 3.54 | 3.87 | 3.55 | 3.39 | 3.44 | 3.5 | 3.37 | 3.45 | 3.6643 | 3.532 | 3 | 3.6136 | 3.68 | 3.6 | 3.3564 | 4.1111 |  |  |  | 86
| 1 | 6 | 3.397 | 3.6315 | 3.6 | 3.6933 | 3.5 | 3.54 | 4.02 | 3.55 | 3.42 | 3.46 | 3.47 | 3.41 | 3.49 | 3.6641 | 3.546 | 2.98 | 3.5608 | 3.67 | 3.67 | 3.3174 | 4.012
| Mean |  | 3.381 | 3.5276 | 3.5167 | 3.7003 | 3.4933 | 3.5333 | 3.8967 | 3.5533 | 3.4183 | 3.51 | 3.4583 | 3.3733 | 3.5167 | 3.6326 | 3.504 | 3.1583 | 3.5943 | 3.6767 | 3.5633 | 3.309 | 4.0851
| Median |  | 3.388 | 3.507 | 3.535 | 3.695 | 3.495 | 3.54 | 3.88 | 3.55 | 3.41 | 3.475 | 3.46 | 3.37 | 3.515 | 3.6464 | 3.511 | 3.205 | 3.6018 | 3.675 | 3.545 | 3.3037 | 4.0953
| Std Dev. |  | 0.02576 | 0.068557 | 0.1027 | 0.014546 | 0.01633 | 0.025033 | 0.062503 | 0.010328 | 0.030605 | 0.079498 | 0.027142 | 0.024221 | 0.045461 | 0.037801 | 0.036508 | 0.13717 | 0.020972 | 0.012111 | 0.062823 | 0.032854 | 0.046071
| Rel.Std.Dev. |  | 0.7619% | 1.943% | 2.92% | 0.3931% | 0.4675% | 0.7085% | 1.604% | 0.2907% | 0.8953% | 2.265% | 0.7848% | 0.718% | 1.293% | 1.041% | 1.042% | 4.343% | 0.5835% | 0.3294% | 1.763% | 0.9929% | 1.128%
| PDM3 |  | -3.801% | 0.3708% | 0.05956% | 5.284% | -0.6043% | 0.5338% | 10.87% | 1.103% | -2.738% | -0.1301% | -1.6% | -4.019% | 0.05956% | 3.357% | -0.3008% | -10.14% | 2.268% | 4.612% | 1.387% | -5.849% | 16.23%
| Z-Score (Absolute) |  | 1.18 | 0.1 | 0 | 1.6 | 0.2 | 0.14 | 3.3 | 0.32 | 0.85 | 0.06 | 0.51 | 1.25 | 0 | 1.01 | 0.11 | 3.11 | 0.67 | 1.39 | 0.41 | 1.8 | 4.94
| Analytical results for Ga in OREAS 239b (Certified Value 8.09 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.27 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 8.2 | 6.669 | 7.17 | 8.8773 | 8.4 | 9.9 | 7.4 | 9 | 8.08 | 8.33 | 8.06 | 7.68 | 7.9 | 8.3958 | 8.4 | 11.2 | 8.53 | 7.9099 | 7.6211 | [5 empty] | 1
| 1 | 2 | 8.3 | 6.5953 | 7.36 | 9.0693 | 8.2 | 10.03 | 7.8 | 9 | 7.99 | 7.56 | 8.43 | 7.81 | 8.6 | 8.2694 | 8.1 | 11 | 8.74 | 7.8126 | 7.5252 | [5 empty] | 39
| 1 | 3 | 8.3 | 6.5527 | 7.21 | 9.04 | 8.2 | 10.32 | 7.7 | 9 | 7.86 | 8.12 | 8.06 | 7.73 | 8.6 | 8.2964 | 8.2 | 10.1 | 8.8 | 7.6522 | 7.3622 | [5 empty] | 16
| 1 | 4 | 8.3 | 6.8182 | 7.55 | 8.9733 | 8.4 | 10.35 | 7.7 | 9 | 7.89 | 7.9 | 8.14 | 7.85 | 8.7 | 8.5446 | 8.2 | 11.7 | 8.64 | 8.0131 | 7.7478 | [5 empty] | 8.0876
| 1 | 5 | 8.2 | 6.5482 | 7.14 | 9.0907 | 8.2 | 10.61 | 7.4 | 9 | 7.77 | 7.86 | 8.03 | 7.85 | 8.3 | 8.388 | 8.2 | 10.6 | 8.79 | 6.5689 | 7.8973 | [5 empty] | 87
| 1 | 6 | 8.3 | 6.707 | 7.23 | 8.888 | 8.2 | 10.71 | 7.8 | 9 | 7.89 | 7.91 | 8.11 | 7.6 | 8.2 | 8.4179 | 8 | 11.6 | 8.81 | 8.075 | 7.9192
| Mean |  | 8.2667 | 6.6484 | 7.2767 | 8.9898 | 8.2667 | 10.32 | 7.6333 | 9 | 7.9133 | 7.9467 | 8.1383 | 7.7533 | 8.3833 | 8.3853 | 8.1833 | 11.0333 | 8.7183 | 7.6719 | 7.6788
| Median |  | 8.3 | 6.6322 | 7.22 | 9.0067 | 8.2 | 10.335 | 7.7 | 9 | 7.89 | 7.905 | 8.085 | 7.77 | 8.45 | 8.3919 | 8.2 | 11.1 | 8.765 | 7.8613 | 7.6844
| Std Dev. |  | 0.05164 | 0.10459 | 0.15384 | 0.091981 | 0.10328 | 0.31547 | 0.18619 | 0 | 0.10783 | 0.25982 | 0.14825 | 0.10132 | 0.30605 | 0.097869 | 0.13292 | 0.60882 | 0.11161 | 0.56078 | 0.21793
| Rel.Std.Dev. |  | 0.6247% | 1.573% | 2.114% | 1.023% | 1.249% | 3.057% | 2.439% | 0% | 1.363% | 3.27% | 1.822% | 1.307% | 3.651% | 1.167% | 1.624% | 5.518% | 1.28% | 7.309% | 2.838%
| PDM3 |  | 2.215% | -17.79% | -10.03% | 11.16% | 2.215% | 27.6% | -5.616% | 11.28% | -2.154% | -1.742% | 0.6277% | -4.133% | 3.657% | 3.682% | 1.184% | 36.42% | 7.799% | -5.139% | -5.054%
| Z-Score (Absolute) |  | 0.16 | 2.29 | 1.34 | 1.26 | 0.16 | 3.27 | 0.8 | NA | 0.38 | 0.32 | 0.03 | 0.62 | 0.34 | 0.34 | 0.03 | 4.35 | 0.85 | 0.74 | 0.73
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 239b (Indicative Value 4 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 5.2832 | 3.5 | 3.709 | 3.6439 | [20 empty] | 1
| 1 | 2 | 5.1831 | 3.2 | 3.924 | 3.6654 | [20 empty] | 12
| 1 | 3 | 5.3905 | 3.4 | 3.888 | 3.4706 | [20 empty] | 16
| 1 | 4 | 5.1666 | 3.2 | 4.147 | 3.5795 | [20 empty] | 4.0014
| 1 | 5 | 5.2442 | 3.1 | 3.922 | 2.9293 | [20 empty] | 18
| 1 | 6 | 5.2403 | 3.2 | 3.73 | 3.645
| Mean |  | 5.2513 | 3.2667 | 3.8867 | 3.489
| Median |  | 5.2422 | 3.2 | 3.905 | 3.6117
| Std Dev. |  | 0.080494 | 0.15055 | 0.15908 | 0.28328
| Rel.Std.Dev. |  | 1.533% | 4.609% | 4.093% | 8.119%
| PDM3 |  | 31.24% | -18.36% | -2.867% | -12.81%
| Z-Score (Absolute) |  | 3.4 | 0.92 | 0.43 | 0.43
| Analytical results for Ge in OREAS 239b (Certified Value 0.1 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.05 | < 0.1 | 0.12 | 1.7 | 0.2 | 0.11 | 0.07 | 0.11 | 0.2 | 0.11 | 0.0637 | [13 empty] | 1
| 1 | 2 | < 0.05 | < 0.1 | 0.13 | 1.7 | 0.21 | 0.1 | 0.07 | 0.11 | 0.2 | 0.13 | 0.0698 | [13 empty] | 5
| 1 | 3 | < 0.05 | < 0.1 | 0.13 | 1.7 | 0.19 | 0.11 | 0.07 | 0.1 | 0.2 | 0.12 | 0.0776 | [13 empty] | 16
| 1 | 4 | < 0.05 | < 0.1 | 0.13 | 1.7 | 0.19 | 0.11 | 0.07 | 0.11 | 0.2 | 0.13 | 0.0728 | [13 empty] | 0.099892
| 1 | 5 | < 0.05 | < 0.1 | 0.13 | 1.7 | 0.19 | 0.11 | 0.07 | 0.09 | 0.2 | 0.12 | 0.0602 | [13 empty] | 88
| 1 | 6 | < 0.05 | < 0.1 | 0.13 | 1.7 | 0.2 | 0.1 | 0.07 | 0.1 | 0.2 | 0.11 | 0.082
| Mean |  |  |  | 0.12833 | 1.7 | 0.19667 | 0.10667 | 0.07 | 0.10333 | 0.2 | 0.12 | 0.071017
| Median |  |  |  | 0.13 | 1.7 | 0.195 | 0.11 | 0.07 | 0.105 | 0.2 | 0.12 | 0.0713
| Std Dev. |  |  |  | 0.0040825 | 0 | 0.008165 | 0.005164 | 0 | 0.008165 | 3.0405e-17 | 0.0089443 | 0.0082361
| Rel.Std.Dev. |  |  |  | 3.181% | 0% | 4.152% | 4.841% | 0% | 7.902% | 1.52e-14% | 7.454% | 11.6%
| PDM3 |  |  |  | 28.47% | 1602% | 96.88% | 6.782% | -29.92% | 3.445% | 100.2% | 20.13% | -28.91%
| Z-Score (Absolute) |  | 1.63 | 1.11 | 0.52 | NA | 1.95 | 0.07 | 0.7 | 0 | NA | 0.35 | 0.67
Results from laboratories 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 239b (Certified Value 0.54 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.6 | 0.68977 | 0.2 | 0.7 | 0.59 | 0.63 | 0.5 | 0.56 | 0.58 | 0.53 | 0.54 | 0.64 | 0.29 | 0.3639 | 0.3827 | [9 empty] | 1
| 1 | 2 | 0.59 | 0.68184 | 0.1 | 0.65 | 0.49 | 0.67 | 0.5 | 0.58 | 0.5 | 0.54 | 0.58 | 0.68 | 0.29 | 0.3669 | 0.3828 | [9 empty] | 25
| 1 | 3 | 0.58 | 0.71616 | 0.1 | 0.67 | 0.52 | 0.64 | 0.5 | 0.57 | 0.54 | 0.55 | 0.55 | 0.64 | 0.29 | 0.3758 | 0.3471 | [9 empty] | 16
| 1 | 4 | 0.56 | 0.68269 | 0.1 | 0.7 | 0.48 | 0.71 | 0.5 | 0.58 | 0.53 | 0.52 | 0.58 | 0.68 | 0.27 | 0.3993 | 0.3772 | [9 empty] | 0.54012
| 1 | 5 | 0.56 | 0.67435 | < 0.1 | 0.67 | 0.53 | 0.62 | 0.5 | 0.55 | 0.52 | 0.54 | 0.55 | 0.67 | 0.27 | 0.3156 | 0.3895 | [9 empty] | 89
| 1 | 6 | 0.58 | 0.71338 | 0.2 | 0.65 | 0.54 | 0.68 | 0.5 | 0.58 | 0.53 | 0.54 | 0.54 | 0.68 | 0.25 | 0.3692 | 0.3689
| Mean |  | 0.57833 | 0.69303 | 0.14 | 0.67333 | 0.525 | 0.65833 | 0.5 | 0.57 | 0.53333 | 0.53667 | 0.55667 | 0.665 | 0.27667 | 0.36512 | 0.3747
| Median |  | 0.58 | 0.68623 | 0.1 | 0.67 | 0.525 | 0.655 | 0.5 | 0.575 | 0.53 | 0.54 | 0.55 | 0.675 | 0.28 | 0.36805 | 0.37995
| Std Dev. |  | 0.016021 | 0.017556 | 0.054772 | 0.022509 | 0.03937 | 0.034303 | 0 | 0.012649 | 0.026583 | 0.010328 | 0.018619 | 0.019748 | 0.01633 | 0.027408 | 0.015168
| Rel.Std.Dev. |  | 2.77% | 2.533% | 39.12% | 3.343% | 7.499% | 5.211% | 0% | 2.219% | 4.984% | 1.924% | 3.345% | 2.97% | 5.902% | 7.507% | 4.048%
| PDM3 |  | 7.074% | 28.31% | -74.08% | 24.66% | -2.8% | 21.89% | -7.429% | 5.531% | -1.257% | -0.64% | 3.063% | 23.12% | -48.78% | -32.4% | -30.63%
| Z-Score (Absolute) |  | 0.14 | 0.9 | NA | 0.77 | 0.21 | 0.67 | NA | 0.09 | 0.15 | 0.13 | 0 | 0.72 | 1.86 | 1.27 | 1.21
Results from laboratories 1.06 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 239b (Indicative Value 0.016 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.01 | 0.06 | < 0.05 | < 2 | 0.102 | 0.01 | 0.02 | 0.01 | 0.01 | 0.03 | < 1 | < 0.01 | < 5 | 0.015 | 0.0241 | [9 empty] | 1
| 1 | 2 | 0.01 | 0.09 | < 0.05 | < 2 | 0.092 | 0.02 | 0.02 | 0.02 | 0.01 | 0.028 | < 1 | 0.01 | < 5 | 0.01 | 0.0191 | [9 empty] | 6
| 1 | 3 | 0.01 | 0.07 | < 0.05 | < 2 | 0.103 | 0.01 | 0.02 | 0.02 | 0.01 | 0.029 | < 1 | < 0.01 | < 5 | 0.014 | 0.0237 | [9 empty] | 16
| 1 | 4 | 0.01 | 0.07 | < 0.05 | < 2 | 0.101 | 0.01 | 0.02 | 0.02 | 0.01 | 0.031 | < 1 | 0.02 | < 5 | 0.01 | 0.0251 | [9 empty] | 0.016441
| 1 | 5 | 0.01 | 0.03 | < 0.05 | < 2 | 0.094 | 0.01 | 0.02 | 0.02 | 0.01 | 0.024 | < 1 | 0.02 | < 5 | 0.015 | 0.0175 | [9 empty] | 19
| 1 | 6 | 0.01 | 0.03 | < 0.05 | < 2 | 0.093 | 0.01 | 0.02 | 0.01 | 0.01 | 0.028 | < 1 | 0.02 | < 5 | 0.01 | 0.0193
| Mean |  | 0.01 | 0.058333 |  |  | 0.0975 | 0.011667 | 0.02 | 0.016667 | 0.01 | 0.028333 |  | 0.0175 |  | 0.012333 | 0.021467
| Median |  | 0.01 | 0.065 |  |  | 0.0975 | 0.01 | 0.02 | 0.02 | 0.01 | 0.0285 |  | 0.02 |  | 0.012 | 0.0215
| Std Dev. |  | 0 | 0.024014 |  |  | 0.00501 | 0.0040825 | 0 | 0.005164 | 0 | 0.0024221 |  | 0.005 |  | 0.002582 | 0.0031985
| Rel.Std.Dev. |  | 0% | 41.17% |  |  | 5.138% | 34.99% | 0% | 30.98% | 0% | 8.549% |  | 28.57% |  | 20.94% | 14.9%
| PDM3 |  | -39.18% | 254.8% |  |  | 493% | -29.04% | 21.65% | 1.374% | -39.18% | 72.34% |  | 6.443% |  | -24.98% | 30.57%
| Z-Score (Absolute) |  | 0.79 | 2.54 | 0.24 | 67.33 | 5.23 | 0.67 | 0.1 | 0.33 | 0.79 | 0.47 | 32.93 | 0.56 | 170.54 | 0.63 | 0
| Analytical results for Ho in OREAS 239b (Indicative Value 0.33 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.35665 | 0.3 | 0.344 | 0.3361 | [20 empty] | 1
| 1 | 2 | 0.33621 | 0.3 | 0.362 | 0.3346 | [20 empty] | 14
| 1 | 3 | 0.34696 | 0.3 | 0.36 | 0.3258 | [20 empty] | 16
| 1 | 4 | 0.34274 | 0.3 | 0.378 | 0.3211 | [20 empty] | 0.33384
| 1 | 5 | 0.33919 | 0.3 | 0.361 | 0.2728 | [20 empty] | 20
| 1 | 6 | 0.34975 | 0.3 | 0.349 | 0.338
| Mean |  | 0.34525 | 0.3 | 0.359 | 0.3214
| Median |  | 0.34485 | 0.3 | 0.3605 | 0.3302
| Std Dev. |  | 0.0074525 | 0 | 0.011832 | 0.024687
| Rel.Std.Dev. |  | 2.159% | 0% | 3.296% | 7.681%
| PDM3 |  | 3.417% | -10.14% | 7.536% | -3.727%
| Z-Score (Absolute) |  | 0.43 | 1.2 | 0.92 | 0.43
| Analytical results for In in OREAS 239b (Certified Value 0.03 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.03 | 0.03 | < 5 | < 0.05 | 0.026 | < 0.05 | 0.03 | 0.032 | 0.03 | 0.03 | 0.029 | 0.03 | 0.028 | 0.0368 | [10 empty] | 1
| 1 | 2 | 0.03 | 0.03 | < 5 | < 0.05 | 0.03 | < 0.05 | 0.03 | 0.03 | 0.029 | 0.03 | 0.029 | 0.03 | 0.03 | 0.0353 | [10 empty] | 27
| 1 | 3 | 0.03 | 0.03 | < 5 | < 0.05 | 0.028 | < 0.05 | 0.03 | 0.03 | 0.029 | 0.028 | 0.029 | 0.03 | 0.028 | 0.0337 | [10 empty] | 16
| 1 | 4 | 0.03 | 0.03 | < 5 | < 0.05 | 0.024 | < 0.05 | 0.03 | 0.033 | 0.028 | 0.029 | 0.03 | 0.03 | 0.026 | 0.0295 | [10 empty] | 0.029758
| 1 | 5 | 0.03 | 0.03 | < 5 | < 0.05 | 0.021 | < 0.05 | 0.03 | 0.03 | 0.029 | 0.028 | 0.032 | 0.03 | 0.026 | 0.0297 | [10 empty] | 90
| 1 | 6 | 0.03 | 0.03 | < 5 | < 0.05 | 0.029 | < 0.05 | 0.03 | 0.03 | 0.028 | 0.029 | 0.03 | 0.03 | 0.028 | 0.0309
| Mean |  | 0.03 | 0.03 |  |  | 0.026333 |  | 0.03 | 0.030833 | 0.028833 | 0.029 | 0.029833 | 0.03 | 0.027667 | 0.03265
| Median |  | 0.03 | 0.03 |  |  | 0.027 |  | 0.03 | 0.03 | 0.029 | 0.029 | 0.0295 | 0.03 | 0.028 | 0.0323
| Std Dev. |  | 0 | 0 |  |  | 0.0033862 |  | 0 | 0.0013292 | 0.00075277 | 0.00089443 | 0.001169 | 0 | 0.0015055 | 0.0030671
| Rel.Std.Dev. |  | 0% | 0% |  |  | 12.86% |  | 0% | 4.311% | 2.611% | 3.084% | 3.919% | 0% | 5.442% | 9.394%
| PDM3 |  | 0.8121% | 0.8121% |  |  | -11.51% |  | 0.8121% | 3.612% | -3.108% | -2.548% | 0.252% | 0.8121% | -7.029% | 9.717%
| Z-Score (Absolute) |  | 0.06 | 0.06 | 1665.6 | 3.32 | 2.42 | 3.32 | 0.06 | 0.62 | 0.73 | 0.62 | 0.06 | 0.06 | 1.52 | 1.84
| Analytical results for Ir in OREAS 239b (Indicative Value < 0.0019 ppm) | [25 empty] | Indicative
Ir | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | < 0.0013 | [23 empty] | 1
| 1 | 2 | < 0.0015 | [23 empty] | 3
| 1 | 3 | < 0.002 | [23 empty] | 16
| 1 | 4 | < 0.0019 | [23 empty] | < 0.0019
| 1 | 5 | < 0.002 | [23 empty] | 14
| 1 | 6 | < 0.0019
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 239b (Certified Value 0.864 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.815 | 1.0587 | 0.77 | 0.9015 | 0.908 | 0.92 | 0.987 | 0.84 | 0.84 | 0.94 | 0.9 | 0.85 | 0.87 | 0.85766 | 0.838 | 0.8 | 0.8747 | 0.88 | 0.87 | 0.8803 | 1.0077 |  |  |  | 1
| 1 | 2 | 0.849 | 0.99997 | 0.78 | 0.9195 | 0.906 | 0.89 | 0.974 | 0.84 | 0.85 | 0.91 | 0.88 | 0.86 | 0.87 | 0.86741 | 0.855 | 0.8 | 0.86984 | 0.89 | 0.87 | 0.8775 | 1.0083 |  |  |  | #N/A
| 1 | 3 | 0.845 | 1.0791 | 0.79 | 0.9 | 0.904 | 0.9 | 0.991 | 0.84 | 0.84 | 0.93 | 0.89 | 0.84 | 0.87 | 0.87717 | 0.838 | 0.8 | 0.86829 | 0.88 | 0.87 | 0.8641 | 1.0082 |  |  |  | 16
| 1 | 4 | 0.839 | 1.0135 | 0.84 | 0.9055 | 0.909 | 0.89 | 0.988 | 0.85 | 0.85 | 0.9 | 0.89 | 0.86 | 0.86 | 0.85264 | 0.83 | 0.83 | 0.87276 | 0.88 | 0.88 | 0.8871 | 1.0082 |  |  |  | 0.86443
| 1 | 5 | 0.831 | 1.0041 | 0.8 | 0.901 | 0.901 | 0.88 | 0.984 | 0.85 | 0.83 | 0.89 | 0.91 | 0.86 | 0.84 | 0.84438 | 0.83 | 0.76 | 0.87196 | 0.89 | 0.89 | 0.8878 | 1.0073 |  |  |  | 91
| 1 | 6 | 0.836 | 1.004 | 0.82 | 0.9145 | 0.9 | 0.9 | 1.01 | 0.84 | 0.85 | 0.91 | 0.89 | 0.87 | 0.85 | 0.87601 | 0.83 | 0.76 | 0.85842 | 0.88 | 0.91 | 0.883 | 1.0081
| Mean |  | 0.83583 | 1.0265 | 0.8 | 0.907 | 0.90467 | 0.89667 | 0.989 | 0.84333 | 0.84333 | 0.91333 | 0.89333 | 0.85667 | 0.86 | 0.86255 | 0.83683 | 0.79167 | 0.86933 | 0.88333 | 0.88167 | 0.87997 | 1.008
| Median |  | 0.8375 | 1.0088 | 0.795 | 0.9035 | 0.905 | 0.895 | 0.9875 | 0.84 | 0.845 | 0.91 | 0.89 | 0.86 | 0.865 | 0.86254 | 0.834 | 0.8 | 0.8709 | 0.88 | 0.875 | 0.88165 | 1.0082
| Std Dev. |  | 0.01204 | 0.03373 | 0.026077 | 0.008124 | 0.0036697 | 0.013663 | 0.011832 | 0.005164 | 0.008165 | 0.018619 | 0.010328 | 0.010328 | 0.012649 | 0.013197 | 0.0097245 | 0.027142 | 0.0057933 | 0.005164 | 0.016021 | 0.0087108 | 0.00038816
| Rel.Std.Dev. |  | 1.441% | 3.286% | 3.26% | 0.8957% | 0.4056% | 1.524% | 1.196% | 0.6123% | 0.9682% | 2.039% | 1.156% | 1.206% | 1.471% | 1.53% | 1.162% | 3.428% | 0.6664% | 0.5846% | 1.817% | 0.9899% | 0.03851%
| PDM3 |  | -3.308% | 18.75% | -7.453% | 4.925% | 4.655% | 3.729% | 14.41% | -2.441% | -2.441% | 5.657% | 3.344% | -0.8981% | -0.5125% | -0.2181% | -3.192% | -8.418% | 0.5669% | 2.187% | 1.994% | 1.797% | 16.6%
| Z-Score (Absolute) |  | 1.1 | 3.66 | 1.99 | 0.67 | 0.62 | 0.42 | 2.72 | 0.91 | 0.91 | 0.83 | 0.33 | 0.58 | 0.5 | 0.43 | 1.07 | 2.2 | 0.27 | 0.08 | 0.04 | 0 | 3.19
| Analytical results for La in OREAS 239b (Certified Value 27.1 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 27.553 | 33.6652 | 27.3 | 25.8567 | 33.62 | 26 | 33.7 | 28 | 26.8 | 26.1 | 28 | 28 | 26.1 | 26.1655 | 18 | 27.78 | 27.1 | 32.7 | 20.7 | 23.1067 |  |  |  |  | 1
| 1 | 2 | 27.332 | 32.638 | 24.1 | 26.2667 | 33.27 | 25 | 35.9 | 28 | 27.4 | 22.9 | 28.4 | 28.9 | 26.2 | 26.2933 | 18 | 29.33 | 26.3 | 33.5 | 21.3 | 23.1087 |  |  |  |  | 28
| 1 | 3 | 28.011 | 33.4396 | 26 | 25.96 | 33.26 | 26 | 34.9 | 29 | 26.9 | 26.1 | 27.3 | 27.7 | 26.1 | 26.3307 | 17 | 29.19 | 27.9 | 32.4 | 21.4 | 21.649 |  |  |  |  | 16
| 1 | 4 | 27.512 | 32.5635 | 26.3 | 26.37 | 33.67 | 25 | 36.2 | 28 | 27.3 | 24.9 | 28 | 29.1 | 26.3 | 26.687 | 17 | 31.39 | 27.4 | 31.5 | 20.5 | 22.6499 |  |  |  |  | 27.1269
| 1 | 5 | 28.438 | 33.3156 | 23.3 | 25.97 | 33.41 | 25 | 33.5 | 29 | 26.3 | 25.5 | 28 | 29.1 | 25.2 | 26.742 | 17 | 29.46 | 27.5 | 31.8 | 20.6 | 18.5232 |  |  |  |  | 92
| 1 | 6 | 27.482 | 32.7972 | 25.7 | 26.76 | 33.24 | 28 | 36.8 | 29 | 28 | 25.8 | 27.6 | 27.7 | 24.7 | 26.1976 | 18 | 28.26 | 27.5 | 33 | 21.1 | 23.0876
| Mean |  | 27.7213 | 33.0698 | 25.45 | 26.1972 | 33.4117 | 25.8333 | 35.1667 | 28.5 | 27.1167 | 25.2167 | 27.8833 | 28.4167 | 25.7667 | 26.4027 | 17.5 | 29.235 | 27.2833 | 32.4833 | 20.9333 | 22.0208
| Median |  | 27.5325 | 33.0564 | 25.85 | 26.1183 | 33.34 | 25.5 | 35.4 | 28.5 | 27.1 | 25.65 | 28 | 28.45 | 26.1 | 26.312 | 17.5 | 29.26 | 27.45 | 32.55 | 20.9 | 22.8687
| Std Dev. |  | 0.41917 | 0.46236 | 1.4802 | 0.3392 | 0.19115 | 1.169 | 1.3619 | 0.54772 | 0.58452 | 1.2205 | 0.38166 | 0.68823 | 0.65625 | 0.24956 | 0.54772 | 1.2469 | 0.54559 | 0.74677 | 0.38297 | 1.8038
| Rel.Std.Dev. |  | 1.512% | 1.398% | 5.816% | 1.295% | 0.5721% | 4.525% | 3.873% | 1.922% | 2.156% | 4.84% | 1.369% | 2.422% | 2.547% | 0.9452% | 3.13% | 4.265% | 2% | 2.299% | 1.829% | 8.191%
| PDM3 |  | 2.191% | 21.91% | -6.182% | -3.427% | 23.17% | -4.769% | 29.64% | 5.062% | -0.03767% | -7.042% | 2.789% | 4.755% | -5.014% | -2.67% | -35.49% | 7.771% | 0.5767% | 19.75% | -22.83% | -18.82%
| Z-Score (Absolute) |  | 0.19 | 2.57 | 0.82 | 0.48 | 2.72 | 0.64 | 3.5 | 0.54 | 0.07 | 0.92 | 0.27 | 0.5 | 0.67 | 0.39 | NA | 0.87 | 0 | 2.31 | 2.82 | 2.34
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 239b (Certified Value 38.2 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 36.83 | 40.0485 | 40.6 | 41.275 | 35.7 | 41 | 39.9 | 38.4 | 40.9 | 38.7 | 38.4 | 39 | 38.5664 | 36.46 | 36 | 34.7 | 34.4 | 36.6424 | [6 empty] | 1
| 1 | 2 | 37.73 | 39.9851 | 40.6 | 41.59 | 36.2 | 41 | 41.1 | 40.4 | 40.3 | 35 | 39.6 | 38.2 | 38.6662 | 36.67 | 36 | 35.4 | 34 | 36.5743 | [6 empty] | #N/A
| 1 | 3 | 37.18 | 40.2344 | 40.1 | 40.91 | 37.7 | 41 | 41.6 | 37.3 | 39.9 | 37.5 | 38.9 | 38.4 | 37.503 | 36.2 | 36 | 35.2 | 33.5 | 36.1598 | [6 empty] | 16
| 1 | 4 | 37.69 | 39.6606 | 41.5 | 41.475 | 37.9 | 41 | 41.4 | 39.8 | 40.2 | 36.8 | 38.1 | 38.5 | 39.0348 | 35.58 | 37 | 35 | 34.6 | 37.5808 | [6 empty] | 38.1829
| 1 | 5 | 37.6 | 38.5351 | 39.8 | 41.375 | 34.8 | 40 | 40.1 | 38.8 | 39.4 | 36.2 | 37.9 | 39.1 | 38.0387 | 37.1 | 34 | 35.2 | 34.7 | 37.3589 | [6 empty] | 93
| 1 | 6 | 36.6 | 40.3193 | 42.3 | 41.39 | 36.1 | 41 | 41.8 | 38.8 | 40 | 35.7 | 37.8 | 38 | 38.6024 | 37.94 | 34 | 36.4 | 34.2 | 37.9347
| Mean |  | 37.2717 | 39.7972 | 40.8167 | 41.3358 | 36.4 | 40.8333 | 40.9833 | 38.9167 | 40.1167 | 36.65 | 38.45 | 38.5333 | 38.4019 | 36.6583 | 35.5 | 35.3167 | 34.2333 | 37.0418
| Median |  | 37.39 | 40.0168 | 40.6 | 41.3825 | 36.15 | 41 | 41.25 | 38.8 | 40.1 | 36.5 | 38.25 | 38.45 | 38.5844 | 36.565 | 36 | 35.2 | 34.3 | 37.0007
| Std Dev. |  | 0.47906 | 0.65936 | 0.92826 | 0.23387 | 1.1933 | 0.40825 | 0.79854 | 1.0852 | 0.49565 | 1.3248 | 0.6892 | 0.43665 | 0.5436 | 0.80624 | 1.2247 | 0.58109 | 0.44121 | 0.68473
| Rel.Std.Dev. |  | 1.285% | 1.657% | 2.274% | 0.5658% | 3.278% | 0.9998% | 1.948% | 2.789% | 1.236% | 3.615% | 1.792% | 1.133% | 1.416% | 2.199% | 3.45% | 1.645% | 1.289% | 1.849%
| PDM3 |  | -2.387% | 4.228% | 6.898% | 8.257% | -4.669% | 6.941% | 7.334% | 1.922% | 5.064% | -4.015% | 0.6995% | 0.9177% | 0.5735% | -3.993% | -7.026% | -7.507% | -10.34% | -2.988%
| Z-Score (Absolute) |  | 0.44 | 0.52 | 0.91 | 1.11 | 0.77 | 0.92 | 0.97 | 0.19 | 0.64 | 0.68 | 0.01 | 0.04 | 0.01 | 0.67 | 1.11 | 1.18 | 1.6 | 0.53
| Analytical results for Lu in OREAS 239b (Certified Value 0.12 ppm) | [25 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.11235 | 0.1 | 0.101 | 0.11 | 0.168 | [19 empty] | 1
| 1 | 2 | 0.12065 | 0.1 | 0.109 | 0.1 | 0.1723 | [19 empty] | 29
| 1 | 3 | 0.11896 | 0.1 | 0.108 | 0.11 | 0.1413 | [19 empty] | 16
| 1 | 4 | 0.11354 | 0.1 | 0.113 | 0.1 | 0.16 | [19 empty] | 0.11511
| 1 | 5 | 0.10594 | < 0.1 | 0.107 | 0.1 | 0.163 | [19 empty] | 94
| 1 | 6 | 0.11267 | < 0.1 | 0.104 | 0.1 | 0.2026
| Mean |  | 0.11402 | 0.1 | 0.107 | 0.10333 | 0.16787
| Median |  | 0.11311 | 0.1 | 0.1075 | 0.1 | 0.1655
| Std Dev. |  | 0.0052611 | 0 | 0.0041473 | 0.005164 | 0.020087
| Rel.Std.Dev. |  | 4.614% | 0% | 3.876% | 4.997% | 11.97%
| PDM3 |  | -0.9474% | -13.13% | -7.046% | -10.23% | 45.83%
| Z-Score (Absolute) |  | 0.67 | 2.27 | 0 | 0.35 | 5.85
| Analytical results for Mg in OREAS 239b (Certified Value 1.41 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.205 | 1.3021 | 1.45 | 1.4267 | 1.36 | 1.44 | 1.51 | 1.33 | 1.36 | 1.44 | 1.4 | 1.35 | 1.42 | 1.4597 | 1.441 | 1.23 | 1.3795 | 1.42 | 1.38 | 1.5231 | 1.4878 |  |  |  | 1
| 1 | 2 | 1.245 | 1.2964 | 1.49 | 1.44 | 1.36 | 1.45 | 1.51 | 1.33 | 1.36 | 1.39 | 1.39 | 1.35 | 1.43 | 1.474 | 1.465 | 1.24 | 1.38 | 1.43 | 1.39 | 1.5169 | 1.4782 |  |  |  | #N/A
| 1 | 3 | 1.259 | 1.3805 | 1.47 | 1.4167 | 1.37 | 1.45 | 1.47 | 1.33 | 1.35 | 1.43 | 1.39 | 1.34 | 1.43 | 1.5006 | 1.441 | 1.22 | 1.3787 | 1.4 | 1.38 | 1.5095 | 1.5004 |  |  |  | 16
| 1 | 4 | 1.241 | 1.3103 | 1.47 | 1.4367 | 1.36 | 1.48 | 1.5 | 1.33 | 1.37 | 1.38 | 1.41 | 1.36 | 1.41 | 1.484 | 1.447 | 1.27 | 1.3883 | 1.4 | 1.39 | 1.5428 | 1.4831 |  |  |  | 1.4078
| 1 | 5 | 1.225 | 1.2931 | 1.49 | 1.44 | 1.36 | 1.51 | 1.5 | 1.33 | 1.34 | 1.36 | 1.41 | 1.36 | 1.39 | 1.5088 | 1.447 | 1.15 | 1.3805 | 1.43 | 1.41 | 1.5437 | 1.479 |  |  |  | 95
| 1 | 6 | 1.23 | 1.3141 | 1.49 | 1.4267 | 1.36 | 1.49 | 1.5 | 1.33 | 1.36 | 1.38 | 1.41 | 1.37 | 1.4 | 1.51 | 1.435 | 1.15 | 1.3604 | 1.42 | 1.44 | 1.5305 | 1.4448
| Mean |  | 1.2342 | 1.3161 | 1.4767 | 1.4311 | 1.3617 | 1.47 | 1.4983 | 1.33 | 1.3567 | 1.3967 | 1.4017 | 1.355 | 1.4133 | 1.4895 | 1.446 | 1.21 | 1.3779 | 1.4167 | 1.3983 | 1.5277 | 1.4789
| Median |  | 1.2355 | 1.3062 | 1.48 | 1.4317 | 1.36 | 1.465 | 1.5 | 1.33 | 1.36 | 1.385 | 1.405 | 1.355 | 1.415 | 1.4923 | 1.444 | 1.225 | 1.3798 | 1.42 | 1.39 | 1.5268 | 1.4811
| Std Dev. |  | 0.018616 | 0.032544 | 0.01633 | 0.0093492 | 0.0040825 | 0.027568 | 0.01472 | 0 | 0.010328 | 0.031411 | 0.0098319 | 0.010488 | 0.01633 | 0.020359 | 0.010334 | 0.049396 | 0.0092823 | 0.013663 | 0.023166 | 0.013863 | 0.01856
| Rel.Std.Dev. |  | 1.508% | 2.473% | 1.106% | 0.6533% | 0.2998% | 1.875% | 0.9824% | 0% | 0.7613% | 2.249% | 0.7014% | 0.774% | 1.155% | 1.367% | 0.7147% | 4.082% | 0.6736% | 0.9644% | 1.657% | 0.9074% | 1.255%
| PDM3 |  | -12.33% | -6.514% | 4.894% | 1.658% | -3.275% | 4.42% | 6.433% | -5.525% | -3.631% | -0.7892% | -0.434% | -3.749% | 0.3947% | 5.805% | 2.715% | -14.05% | -2.122% | 0.6315% | -0.6708% | 8.522% | 5.051%
| Z-Score (Absolute) |  | 2.42 | 1.24 | 1.08 | 0.43 | 0.58 | 0.99 | 1.4 | 1.04 | 0.65 | 0.07 | 0 | 0.67 | 0.17 | 1.27 | 0.65 | 2.77 | 0.34 | 0.22 | 0.05 | 1.82 | 1.12
| Analytical results for Mn in OREAS 239b (Certified Value 0.028 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02836 | 0.027351 | 0.0287 | 0.030022 | 0.03 | 0.0283 | 0.028 | 0.0312 | 0.0272 | 0.0289 | 0.0279 | 0.0264 | 0.0275 | 0.029355 | 0.0256 | 0.0268 | 0.027895 | 0.03199 | 0.0273 | 0.02601 | 0.0322 |  |  |  | 1
| 1 | 2 | 0.02817 | 0.027576 | 0.0289 | 0.030144 | 0.0293 | 0.028 | 0.0276 | 0.0319 | 0.0275 | 0.028 | 0.0279 | 0.0264 | 0.0277 | 0.029932 | 0.0263 | 0.0269 | 0.027746 | 0.03223 | 0.0275 | 0.02637 | 0.0321 |  |  |  | 30
| 1 | 3 | 0.02888 | 0.027301 | 0.0288 | 0.02992 | 0.0291 | 0.028 | 0.0278 | 0.0314 | 0.0273 | 0.0288 | 0.0273 | 0.0263 | 0.0277 | 0.029435 | 0.0256 | 0.027 | 0.027361 | 0.0315 | 0.0274 | 0.02597 | 0.0315 |  |  |  | 16
| 1 | 4 | 0.02871 | 0.027567 | 0.03 | 0.030001 | 0.0299 | 0.028 | 0.028 | 0.0327 | 0.0276 | 0.0278 | 0.0282 | 0.0266 | 0.0273 | 0.029842 | 0.0256 | 0.0278 | 0.027647 | 0.03188 | 0.0276 | 0.02663 | 0.032 |  |  |  | 0.028042
| 1 | 5 | 0.02805 | 0.027148 | 0.028 | 0.030136 | 0.0291 | 0.0277 | 0.0278 | 0.0312 | 0.0272 | 0.0277 | 0.0283 | 0.0267 | 0.0268 | 0.029602 | 0.0263 | 0.0259 | 0.027826 | 0.03192 | 0.0279 | 0.02626 | 0.0322 |  |  |  | 96
| 1 | 6 | 0.02813 | 0.027281 | 0.0296 | 0.030054 | 0.0287 | 0.0284 | 0.0285 | 0.0316 | 0.0275 | 0.0281 | 0.0279 | 0.0269 | 0.0271 | 0.029492 | 0.0263 | 0.026 | 0.027277 | 0.03162 | 0.0285 | 0.02654 | 0.0314
| Mean |  | 0.028383 | 0.027371 | 0.029 | 0.030046 | 0.02935 | 0.028067 | 0.02795 | 0.031667 | 0.027383 | 0.028217 | 0.027917 | 0.02655 | 0.02735 | 0.02961 | 0.02595 | 0.026733 | 0.027625 | 0.031857 | 0.0277 | 0.026297 | 0.0319
| Median |  | 0.028265 | 0.027326 | 0.02885 | 0.030038 | 0.0292 | 0.028 | 0.0279 | 0.0315 | 0.0274 | 0.02805 | 0.0279 | 0.0265 | 0.0274 | 0.029547 | 0.02595 | 0.02685 | 0.027696 | 0.0319 | 0.02755 | 0.026315 | 0.03205
| Std Dev. |  | 0.00033904 | 0.00016941 | 0.00070711 | 8.5039e-05 | 0.00050498 | 0.00025033 | 0.00030822 | 0.00057155 | 0.00017224 | 0.00051153 | 0.00034881 | 0.00022583 | 0.00035637 | 0.00023114 | 0.00038341 | 0.00070333 | 0.00025268 | 0.00026265 | 0.00044272 | 0.00027053 | 0.00035777
| Rel.Std.Dev. |  | 1.194% | 0.6189% | 2.438% | 0.283% | 1.721% | 0.8919% | 1.103% | 1.805% | 0.629% | 1.813% | 1.249% | 0.8506% | 1.303% | 0.7806% | 1.477% | 2.631% | 0.9147% | 0.8245% | 1.598% | 1.029% | 1.122%
| PDM3 |  | 1.218% | -2.393% | 3.417% | 7.148% | 4.665% | 0.08898% | -0.3271% | 12.93% | -2.348% | 0.6239% | -0.4459% | -5.32% | -2.467% | 5.591% | -7.459% | -4.666% | -1.485% | 13.6% | -1.219% | -6.223% | 13.76%
| Z-Score (Absolute) |  | 0.28 | 0.37 | 0.67 | 1.35 | 0.9 | 0.07 | 0 | 2.39 | 0.36 | 0.17 | 0.02 | 0.9 | 0.39 | 1.07 | 1.29 | 0.78 | 0.21 | 2.51 | 0.16 | 1.06 | 2.54
| Analytical results for Mo in OREAS 239b (Certified Value 0.9 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.89 | 0.83 | < 2 | 0.9 | 0.93 | 1 | 0.9 | 0.91 | 0.94 | 0.8 | 0.87 | 0.86 | 0.93077 | < 1 | 1.05 | 0.87 | < 1 | 0.91 | 0.734 | 2.1228 |  |  |  |  | 1
| 1 | 2 | 0.9 | 0.8 | < 2 | 0.8 | 0.86 | 1 | 0.9 | 0.97 | 0.9 | 0.84 | 0.89 | 0.91 | 0.95086 | < 1 | 1.041 | 0.9 | < 1 | 0.91 | 0.7172 | 2.3532 |  |  |  |  | 31
| 1 | 3 | 0.92 | 0.87 | < 2 | 0.8 | 0.96 | 1 | 0.9 | 0.88 | 0.93 | 0.82 | 0.85 | 0.93 | 0.96097 | < 1 | 1.044 | 0.87 | < 1 | 0.9 | 0.7989 | 2.5837 |  |  |  |  | 16
| 1 | 4 | 0.9 | 0.89 | < 2 | 0.8 | 0.9 | 0.9 | 0.9 | 0.87 | 0.9 | 0.78 | 0.89 | 0.9 | 0.93938 | < 1 | 1.064 | 0.87 | < 1 | 0.93 | 0.7574 | 2.5026 |  |  |  |  | 0.89612
| 1 | 5 | 0.89 | 0.82 | < 2 | 0.8 | 0.85 | 1 | 0.9 | 0.95 | 0.9 | 0.82 | 0.89 | 0.89 | 0.96577 | < 1 | 1.077 | 0.9 | < 1 | 0.94 | 0.6393 | 2.6519 |  |  |  |  | 97
| 1 | 6 | 0.88 | 0.87 | < 2 | 0.8 | 0.83 | 0.9 | 0.9 | 0.9 | 0.89 | 0.82 | 0.83 | 0.9 | 0.89625 | < 1 | 1.042 | 0.88 | < 1 | 0.88 | 0.7625 | 2.8418
| Mean |  | 0.89667 | 0.84667 |  | 0.81667 | 0.88833 | 0.96667 | 0.9 | 0.91333 | 0.91 | 0.81333 | 0.87 | 0.89833 | 0.94067 |  | 1.053 | 0.88167 |  | 0.91167 | 0.73488 | 2.5093
| Median |  | 0.895 | 0.85 |  | 0.8 | 0.88 | 1 | 0.9 | 0.905 | 0.9 | 0.82 | 0.88 | 0.9 | 0.94512 |  | 1.047 | 0.875 |  | 0.91 | 0.7457 | 2.5431
| Std Dev. |  | 0.013663 | 0.035024 |  | 0.040825 | 0.050365 | 0.05164 | 0 | 0.039328 | 0.02 | 0.020656 | 0.025298 | 0.023166 | 0.025369 |  | 0.014505 | 0.01472 |  | 0.02137 | 0.054438 | 0.24909
| Rel.Std.Dev. |  | 1.524% | 4.137% |  | 4.999% | 5.67% | 5.342% | 0% | 4.306% | 2.198% | 2.54% | 2.908% | 2.579% | 2.697% |  | 1.378% | 1.67% |  | 2.344% | 7.408% | 9.926%
| PDM3 |  | 0.06087% | -5.519% |  | -8.866% | -0.8691% | 7.872% | 0.4328% | 1.921% | 1.549% | -9.238% | -2.915% | 0.2469% | 4.971% |  | 17.51% | -1.613% |  | 1.735% | -17.99% | 180%
| Z-Score (Absolute) |  | 0 | 0.77 | 1.58 | 1.23 | 0.13 | NA | 0.05 | 0.26 | 0.2 | 1.28 | 0.41 | 0.03 | 0.67 | 6.08 | 2.4 | 0.23 | 6.08 | 0.23 | 2.48 | 24.71
Results from laboratory 1.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 239b (Certified Value 0.091 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.087 | 0.12602 | 0.106 | 0.083333 | 0.1 | 0.09 | 0.12 | 0.13 | 0.09 | 0.1 | 0.09 | 0.08 | 0.091 | 0.09169 | 0.074 | 0.08 | 0.097672 | 0.08 | 0.09 | 0.079 | 0.0609 |  |  |  | 1
| 1 | 2 | 0.09 | 0.12911 | 0.109 | 0.083333 | 0.1 | 0.1 | 0.11 | 0.14 | 0.09 | 0.09 | 0.09 | 0.08 | 0.091 | 0.088571 | 0.082 | 0.08 | 0.093968 | 0.08 | 0.09 | 0.0783 | 0.0597 |  |  |  | #N/A
| 1 | 3 | 0.089 | 0.12625 | 0.109 | 0.083333 | 0.1 | 0.1 | 0.12 | 0.14 | 0.09 | 0.09 | 0.09 | 0.08 | 0.092 | 0.084443 | 0.074 | 0.08 | 0.095867 | 0.08 | 0.09 | 0.0777 | 0.0609 |  |  |  | 16
| 1 | 4 | 0.089 | 0.12543 | 0.118 | 0.083333 | 0.1 | 0.1 | 0.12 | 0.12 | 0.09 | 0.09 | 0.09 | 0.08 | 0.09 | 0.091339 | 0.074 | 0.08 | 0.094728 | 0.08 | 0.1 | 0.0798 | 0.0605 |  |  |  | 0.090925
| 1 | 5 | 0.088 | 0.12623 | 0.102 | 0.086667 | 0.1 | 0.1 | 0.12 | 0.12 | 0.09 | 0.09 | 0.09 | 0.08 | 0.088 | 0.087464 | 0.074 | 0.08 | 0.096067 | 0.08 | 0.1 | 0.0796 | 0.0613 |  |  |  | 98
| 1 | 6 | 0.088 | 0.12308 | 0.11 | 0.083333 | 0.1 | 0.1 | 0.12 | 0.12 | 0.09 | 0.09 | 0.09 | 0.08 | 0.088 | 0.084518 | 0.074 | 0.08 | 0.094138 | 0.08 | 0.1 | 0.0793 | 0.0599
| Mean |  | 0.0885 | 0.12602 | 0.109 | 0.083889 | 0.1 | 0.098333 | 0.11833 | 0.12833 | 0.09 | 0.091667 | 0.09 | 0.08 | 0.09 | 0.088004 | 0.075333 | 0.08 | 0.095407 | 0.08 | 0.095 | 0.07895 | 0.060533
| Median |  | 0.0885 | 0.12612 | 0.109 | 0.083333 | 0.1 | 0.1 | 0.12 | 0.125 | 0.09 | 0.09 | 0.09 | 0.08 | 0.0905 | 0.088018 | 0.074 | 0.08 | 0.095298 | 0.08 | 0.095 | 0.07915 | 0.0607
| Std Dev. |  | 0.0010488 | 0.0019315 | 0.0052915 | 0.0013608 | 1.5202e-17 | 0.0040825 | 0.0040825 | 0.0098319 | 1.5202e-17 | 0.0040825 | 1.5202e-17 | 0 | 0.0016733 | 0.0031671 | 0.003266 | 0 | 0.0014085 | 0 | 0.0054772 | 0.00080685 | 0.00062503
| Rel.Std.Dev. |  | 1.185% | 1.533% | 4.855% | 1.622% | 1.52e-14% | 4.152% | 3.45% | 7.661% | 1.689e-14% | 4.454% | 1.689e-14% | 0% | 1.859% | 3.599% | 4.335% | 0% | 1.476% | 0% | 5.766% | 1.022% | 1.033%
| PDM3 |  | -2.667% | 38.6% | 19.88% | -7.738% | 9.981% | 8.148% | 30.14% | 41.14% | -1.017% | 0.8156% | -1.017% | -12.02% | -1.017% | -3.212% | -17.15% | -12.02% | 4.929% | -12.02% | 4.482% | -13.17% | -33.43%
| Z-Score (Absolute) |  | 0.1 | 2.43 | 1.28 | 0.41 | 0.67 | 0.56 | 1.91 | 2.58 | 0 | 0.11 | 0 | 0.67 | 0 | 0.13 | 0.98 | 0.67 | 0.36 | 0.67 | 0.34 | 0.75 | 1.99
| Analytical results for Nb in OREAS 239b (Certified Value 0.4 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.53 | 0.6 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.37 | 0.38 | 0.32 | 0.38 | 0.53569 | 0.49 | 0.26 | 0.2866 | [9 empty] | 1
| 1 | 2 | 0.52 | 1 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.42 | 0.36 | 0.31 | 0.41 | 0.56323 | 0.52 | 0.27 | 0.3053 | [9 empty] | 32
| 1 | 3 | 0.5 | 0.7 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.42 | 0.38 | 0.29 | 0.34 | 0.55128 | 0.55 | 0.28 | 0.298 | [9 empty] | 16
| 1 | 4 | 0.51 | 0.9 | < 5 | 0.7 | < 1 | 0.5 | 0.2 | 0.42 | 0.39 | 0.32 | 0.4 | 0.55178 | 0.52 | 0.27 | 0.2762 | [9 empty] | 0.40136
| 1 | 5 | 0.53 | 0.9 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.38 | 0.45 | 0.29 | 0.37 | 0.59476 | 0.49 | 0.26 | 0.2284 | [9 empty] | 99
| 1 | 6 | 0.51 | 0.5 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.41 | 0.42 | 0.31 | 0.38 | 0.57353 | 0.44 | 0.27 | 0.2689
| Mean |  | 0.51667 | 0.76667 |  | 0.7 |  | 0.41667 | 0.2 | 0.40333 | 0.39667 | 0.30667 | 0.38 | 0.56171 | 0.50167 | 0.26833 | 0.27723
| Median |  | 0.515 | 0.8 |  | 0.7 |  | 0.4 | 0.2 | 0.415 | 0.385 | 0.31 | 0.38 | 0.55751 | 0.505 | 0.27 | 0.2814
| Std Dev. |  | 0.012111 | 0.19664 |  | 1.2162e-16 |  | 0.040825 | 3.0405e-17 | 0.022509 | 0.03266 | 0.013663 | 0.024495 | 0.020586 | 0.037639 | 0.0075277 | 0.02743
| Rel.Std.Dev. |  | 2.344% | 25.65% |  | 1.737e-14% |  | 9.798% | 1.52e-14% | 5.581% | 8.234% | 4.455% | 6.446% | 3.665% | 7.503% | 2.805% | 9.894%
| PDM3 |  | 28.73% | 91.02% |  | 74.41% |  | 3.813% | -50.17% | 0.4906% | -1.17% | -23.59% | -5.323% | 39.95% | 24.99% | -33.14% | -30.93%
| Z-Score (Absolute) |  | 0.78 | NA | 14.38 | NA | 0.66 | NA | NA | 0 | 0.05 | 0.66 | 0.16 | 1.09 | 0.67 | 0.93 | 0.86
Results from laboratories 1.06, 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 239b (Indicative Value 23.1 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 1
| 1 | 1 | 25.2141 | 23.3 | 24.22 | 20.0504 | [20 empty] | 1
| 1 | 2 | 25.2549 | 21.1 | 25.49 | 20.2426 | [20 empty] | 9
| 1 | 3 | 25.6427 | 21.9 | 25.52 | 18.8108 | [20 empty] | 16
| 1 | 4 | 24.8125 | 22.6 | 27.25 | 19.6252 | [20 empty] | 23.1274
| 1 | 5 | 25.2472 | 21 | 25.72 | 16.0603 | [20 empty] | 15
| 1 | 6 | 25.5055 | 22.1 | 24.49 | 20.1805
| Mean |  | 25.2795 | 22 | 25.4483 | 19.1616
| Median |  | 25.251 | 22 | 25.505 | 19.8378
| Std Dev. |  | 0.28531 | 0.88091 | 1.0727 | 1.6095
| Rel.Std.Dev. |  | 1.129% | 4.004% | 4.215% | 8.4%
| PDM3 |  | 9.305% | -4.875% | 10.04% | -17.15%
| Z-Score (Absolute) |  | 0.64 | 0.64 | 0.71 | 1.75
| Analytical results for Ni in OREAS 239b (Certified Value 59 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 1
| 1 | 1 | 56.7 | 51.6535 | 59.3 | 54.94 | 61 | 64 | 66 | 63.6 | 60 | 62.4 | 61.2 | 58.2 | 59.6 | 63.4591 | 59 | 42.72 | 57 | 74.08 | 56.73 | 63.5 | 54.245 | 55.7312 | 59.2586 |  | 1
| 1 | 2 | 57.3 | 51.4995 | 62.4 | 55.9933 | 59 | 63 | 65 | 64 | 59.9 | 57.6 | 62.5 | 57.1 | 60.7 | 62.8114 | 60 | 41.7 | 59.9 | 68.98 | 57.17 | 64.2 | 53.77 | 54.7507 | 60.2867 |  | 34
| 1 | 3 | 57 | 51.891 | 60.6 | 55.0867 | 59 | 65 | 64 | 62.2 | 59.2 | 61.3 | 61.6 | 57.4 | 60.8 | 62.0063 | 59 | 42.4 | 58 | 70.76 | 55.78 | 63.5 | 54.34 | 54.2906 | 60.8435 |  | 16
| 1 | 4 | 57.1 | 51.87 | 62.2 | 55.6133 | 61 | 62 | 66 | 63.9 | 58.8 | 59.5 | 62.1 | 57.2 | 60.1 | 61.233 | 60 | 41.78 | 60.4 | 73.9 | 55.4 | 63.4 | 54.625 | 54.3236 | 59.7601 |  | 59.3099
| 1 | 5 | 56.1 | 51.357 | 59.6 | 55.3933 | 59 | 64 | 66 | 63.1 | 58.4 | 58.8 | 60.8 | 58.3 | 58.1 | 63.2914 | 60 | 45.14 | 60.8 | 72.01 | 56.99 | 65 | 55.005 | 53.3709 | 60.3324 |  | 100
| 1 | 6 | 56.4 | 51.0645 | 60.6 | 55.5133 | 59 | 65 | 68 | 63.5 | 58 | 58.4 | 62.8 | 55.6 | 60.2 | 61.7135 | 62 | 45.33 | 58.4 | 75.19 | 56.09 | 62.9 | 54.245 | 51.8025 | 60.1399
| Mean |  | 56.7667 | 51.5559 | 60.7833 | 55.4233 | 59.6667 | 63.8333 | 65.8333 | 63.3833 | 59.05 | 59.6667 | 61.8333 | 57.3 | 59.9167 | 62.4191 | 60 | 43.1783 | 59.0833 | 72.4867 | 56.36 | 63.75 | 54.3717 | 54.0449 | 60.1035
| Median |  | 56.85 | 51.5765 | 60.6 | 55.4533 | 59 | 64 | 66 | 63.55 | 59 | 59.15 | 61.85 | 57.3 | 60.15 | 62.4089 | 60 | 42.56 | 59.15 | 72.955 | 56.41 | 63.5 | 54.2925 | 54.3071 | 60.2133
| Std Dev. |  | 0.45461 | 0.31771 | 1.2875 | 0.37867 | 1.0328 | 1.169 | 1.3292 | 0.66156 | 0.80436 | 1.8327 | 0.77115 | 0.9757 | 0.99079 | 0.90241 | 1.0954 | 1.6392 | 1.5052 | 2.3388 | 0.71004 | 0.73959 | 0.41482 | 1.3388 | 0.5416
| Rel.Std.Dev. |  | 0.8008% | 0.6162% | 2.118% | 0.6832% | 1.731% | 1.831% | 2.019% | 1.044% | 1.362% | 3.072% | 1.247% | 1.703% | 1.654% | 1.446% | 1.826% | 3.796% | 2.548% | 3.227% | 1.26% | 1.16% | 0.7629% | 2.477% | 0.9011%
| PDM3 |  | -4.288% | -13.07% | 2.484% | -6.553% | 0.6015% | 7.627% | 11% | 6.868% | -0.4383% | 0.6015% | 4.255% | -3.389% | 1.023% | 5.242% | 1.164% | -27.2% | -0.3821% | 22.22% | -4.974% | 7.486% | -8.326% | -8.877% | 1.338%
| Z-Score (Absolute) |  | 0.67 | 1.89 | 0.26 | 0.99 | 0 | 0.97 | 1.43 | 0.86 | 0.14 | 0 | 0.5 | 0.55 | 0.06 | 0.64 | 0.08 | 3.83 | 0.14 | 2.98 | 0.77 | 0.95 | 1.23 | 1.31 | 0.1
| Analytical results for P in OREAS 239b (Certified Value 0.058 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.0537 | 0.056 | 0.059222 | 0.057 | 0.0607 | 0.0592 | 0.063 | 0.057 | 0.06 | 0.059 | 0.059 | 0.058 | 0.057628 | 0.0545 | 0.05 | 0.057531 | 0.06 | 0.0566 | 0.0641 | 0.0644 |  |  |  |  | 1
| 1 | 2 | 0.0545 | 0.055 | 0.059817 | 0.058 | 0.0606 | 0.058 | 0.061 | 0.058 | 0.058 | 0.058 | 0.059 | 0.059 | 0.058588 | 0.058 | 0.05 | 0.060215 | 0.06 | 0.0569 | 0.0634 | 0.0636 |  |  |  |  | #N/A
| 1 | 3 | 0.0539 | 0.055 | 0.059166 | 0.058 | 0.06 | 0.0584 | 0.06 | 0.057 | 0.06 | 0.058 | 0.058 | 0.059 | 0.061392 | 0.0554 | 0.05 | 0.058818 | 0.06 | 0.0569 | 0.0642 | 0.0631 |  |  |  |  | 16
| 1 | 4 | 0.0543 | 0.055 | 0.059375 | 0.058 | 0.0598 | 0.0592 | 0.062 | 0.058 | 0.057 | 0.06 | 0.059 | 0.057 | 0.06218 | 0.0559 | 0.05 | 0.05934 | 0.06 | 0.0568 | 0.0649 | 0.0636 |  |  |  |  | 0.058434
| 1 | 5 | 0.0533 | 0.053 | 0.059824 | 0.058 | 0.0586 | 0.059 | 0.061 | 0.056 | 0.057 | 0.059 | 0.059 | 0.056 | 0.059143 | 0.055 | 0.05 | 0.057657 | 0.06 | 0.0581 | 0.0645 | 0.0636 |  |  |  |  | 101
| 1 | 6 | 0.054 | 0.057 | 0.059429 | 0.057 | 0.06 | 0.0603 | 0.062 | 0.057 | 0.058 | 0.059 | 0.059 | 0.057 | 0.058819 | 0.0607 | 0.05 | 0.056386 | 0.06 | 0.0594 | 0.0642 | 0.0626
| Mean |  | 0.05395 | 0.055167 | 0.059472 | 0.057667 | 0.05995 | 0.059017 | 0.0615 | 0.057167 | 0.058333 | 0.058833 | 0.058833 | 0.057667 | 0.059625 | 0.056583 | 0.05 | 0.058325 | 0.06 | 0.05745 | 0.064217 | 0.063483
| Median |  | 0.05395 | 0.055 | 0.059402 | 0.058 | 0.06 | 0.0591 | 0.0615 | 0.057 | 0.058 | 0.059 | 0.059 | 0.0575 | 0.058981 | 0.05565 | 0.05 | 0.058238 | 0.06 | 0.0569 | 0.0642 | 0.0636
| Std Dev. |  | 0.00042778 | 0.0013292 | 0.0002865 | 0.0005164 | 0.00075299 | 0.00079099 | 0.0010488 | 0.00075277 | 0.0013663 | 0.00075277 | 0.00040825 | 0.0012111 | 0.001766 | 0.0023524 | 7.6012e-18 | 0.0013915 | 0 | 0.0010932 | 0.00049565 | 0.00060139
| Rel.Std.Dev. |  | 0.7929% | 2.409% | 0.4817% | 0.8955% | 1.256% | 1.34% | 1.705% | 1.317% | 2.342% | 1.28% | 0.6939% | 2.1% | 2.962% | 4.157% | 1.52e-14% | 2.386% | 0% | 1.903% | 0.7718% | 0.9473%
| PDM3 |  | -7.674% | -5.592% | 1.776% | -1.314% | 2.594% | 0.9966% | 5.246% | -2.169% | -0.1728% | 0.6829% | 0.6829% | -1.314% | 2.037% | -3.168% | -14.43% | -0.1879% | 2.679% | -1.684% | 9.896% | 8.641%
| Z-Score (Absolute) |  | 2.5 | 1.84 | 0.48 | 0.49 | 0.74 | 0.23 | 1.57 | 0.76 | 0.13 | 0.13 | 0.13 | 0.49 | 0.56 | 1.08 | 4.63 | 0.14 | 0.76 | 0.61 | 3.04 | 2.64
| Analytical results for Pb in OREAS 239b (Certified Value 14.4 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  | 1
| 1 | 1 | 14.2 | 16.2 | 14.5 | 15.3 | 12.5 | 12.3 | 15.7 | 14.5 | 14.2 | 14.8 | 15.1 | 13.87 | 15.1328 | 13 | 27.96 | 14.3 | 16.22 | 12.1 | 12.2 | 16.4 | 16 | 9.1254 |  |  | 1
| 1 | 2 | 14.7 | 15.2 | 14.0167 | 14.4 | 11.6 | 13.2 | 15.6 | 14.9 | 13.4 | 15.7 | 14.9 | 13.94 | 14.9317 | 12 | 24.46 | 14.6 | 4.13 | 13.1 | 12.4 | 17.1 | 17 | 9.0085 |  |  | 35
| 1 | 3 | 14.9 | 15.6 | 13.3533 | 14.6 | 11.3 | 12.8 | 15.3 | 14.3 | 14.2 | 15.1 | 15 | 14.78 | 15.2872 | 13 | 20.34 | 14.6 | 12.14 | 12.1 | 12.6 | 16.2 | 17 | 9.4586 |  |  | 16
| 1 | 4 | 14.9 | 15.6 | 14.19 | 14.9 | 12.1 | 13.2 | 14.7 | 14.7 | 13.7 | 14.8 | 15 | 14.32 | 16.1196 | 14 | 23.93 | 14.1 | 7.35 | 12.9 | 11.9 | 17.2 | 18 | 8.9665 |  |  | 14.434
| 1 | 5 | 14.5 | 14.6 | 14.0367 | 14.6 | 11.6 | 12.6 | 15.3 | 14.4 | 14 | 15.3 | 15.3 | 17.15 | 14.9681 | 13 | 25.14 | 14.5 | 11.1 | 13.6 | 11.9 | 16.9 | 17 | 9.0241 |  |  | 102
| 1 | 6 | 14.6 | 16.2 | 13.95 | 15.4 | 12 | 13.2 | 15.2 | 14.6 | 13.7 | 15.4 | 14.4 | 14.86 | 16.501 | 13 | 23.17 | 15.1 | 9.61 | 12.9 | 12.5 | 17.5 | 18 | 9.3366
| Mean |  | 14.6333 | 15.5667 | 14.0078 | 14.8667 | 11.85 | 12.8833 | 15.3 | 14.5667 | 13.8667 | 15.1833 | 14.95 | 14.82 | 15.4901 | 13 | 24.1667 | 14.5333 | 10.0917 | 12.7833 | 12.25 | 16.8833 | 17.1667 | 9.1533
| Median |  | 14.65 | 15.6 | 14.0267 | 14.75 | 11.8 | 13 | 15.3 | 14.55 | 13.85 | 15.2 | 15 | 14.55 | 15.21 | 13 | 24.195 | 14.55 | 10.355 | 12.9 | 12.3 | 17 | 17 | 9.0748
| Std Dev. |  | 0.26583 | 0.6121 | 0.37631 | 0.40825 | 0.43243 | 0.38166 | 0.35214 | 0.21602 | 0.32042 | 0.35449 | 0.30166 | 1.2132 | 0.65898 | 0.63246 | 2.4946 | 0.33862 | 4.1487 | 0.58793 | 0.30166 | 0.49565 | 0.75277 | 0.20007
| Rel.Std.Dev. |  | 1.817% | 3.932% | 2.686% | 2.746% | 3.649% | 2.962% | 2.302% | 1.483% | 2.311% | 2.335% | 2.018% | 8.186% | 4.254% | 4.865% | 10.32% | 2.33% | 41.11% | 4.599% | 2.463% | 2.936% | 4.385% | 2.186%
| PDM3 |  | 1.381% | 7.848% | -2.953% | 2.998% | -17.9% | -10.74% | 6% | 0.9194% | -3.93% | 5.192% | 3.575% | 2.675% | 7.317% | -9.935% | 67.43% | 0.6885% | -30.08% | -11.44% | -15.13% | 16.97% | 18.93% | -36.59%
| Z-Score (Absolute) |  | 0.02 | 0.7 | 0.43 | 0.19 | 2 | 1.25 | 0.51 | 0.02 | 0.53 | 0.42 | 0.25 | 0.16 | 0.65 | 1.16 | 6.95 | 0.05 | 3.27 | 1.32 | 1.71 | 1.66 | 1.86 | 3.96
| Analytical results for Pd in OREAS 239b (Indicative Value 15.4 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.01 | 1.08 | 1.30 | [21 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 1
|  |  | 0.5g | 01g | 0.5g | [21 empty] | 1
| 1 | 1 | < 1 | < 10 | 32.3 | [21 empty] | 1
| 1 | 2 | 1 | < 10 | 34.1 | [21 empty] | 1
| 1 | 3 | 1 | < 10 | 25.1 | [21 empty] | 16
| 1 | 4 | < 1 | < 10 | 34.1 | [21 empty] | 15.3583
| 1 | 5 | < 1 | < 10 | 30.5 | [21 empty] | 16
| 1 | 6 | < 1 | < 10 | 24.2
| Mean |  | 1 |  | 30.05
| Median |  | 1 |  | 31.4
| Std Dev. |  | 0 |  | 4.3999
| Rel.Std.Dev. |  | 0% |  | 14.64%
| PDM3 |  | -93.49% |  | 95.66%
| Z-Score (Absolute) |  | 0.67 | 0 | 3.9
| Analytical results for Pr in OREAS 239b (Indicative Value 6.2 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 6.3904 | 6.6 | 6.575 | 5.3523 | [20 empty] | 1
| 1 | 2 | 6.4286 | 6.1 | 6.68 | 5.3654 | [20 empty] | 11
| 1 | 3 | 6.6177 | 6.4 | 6.911 | 4.9966 | [20 empty] | 16
| 1 | 4 | 6.6805 | 6.5 | 7.458 | 5.229 | [20 empty] | 6.1983
| 1 | 5 | 6.3098 | 5.9 | 6.853 | 4.2868 | [20 empty] | 17
| 1 | 6 | 6.712 | 6.4 | 6.657 | 5.3561
| Mean |  | 6.5232 | 6.3167 | 6.8557 | 5.0977
| Median |  | 6.5231 | 6.4 | 6.7665 | 5.2906
| Std Dev. |  | 0.16821 | 0.26394 | 0.32093 | 0.42149
| Rel.Std.Dev. |  | 2.579% | 4.179% | 4.681% | 8.268%
| PDM3 |  | 5.241% | 1.91% | 10.61% | -17.76%
| Z-Score (Absolute) |  | 0.26 | 0.26 | 1.09 | 3.31
| Analytical results for Pt in OREAS 239b (Indicative Value 2.9 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.01 | 1.08 | 1.30 | [21 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 2
|  |  | 0.5g | 01g | 0.5g | [21 empty] | 2
| 1 | 1 | 2 | < 5 | 3.9 | [21 empty] | 1
| 1 | 2 | 2 | < 5 | 2.7 | [21 empty] | 1
| 1 | 3 | 2 | < 5 | 3.9 | [21 empty] | 16
| 1 | 4 | 2 | < 5 | 4.4 | [21 empty] | 2.9
| 1 | 5 | 2 | < 5 | 3.7 | [21 empty] | 18
| 1 | 6 | 2 | < 5 | 4.2
| Mean |  | 2 |  | 3.8
| Median |  | 2 |  | 3.9
| Std Dev. |  | 0 |  | 0.5933
| Rel.Std.Dev. |  | 0% |  | 15.61%
| PDM3 |  | -31.03% |  | 31.03%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.75
| Analytical results for Rb in OREAS 239b (Certified Value 83 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 83.11 | 66.4577 | 67.2 | 85.47 | 82.1 | 93.7 | 79.5 | 85.3 | 85.4 | 83.7 | 80.3 | 83.363 | 68.88 | 82.4 | 84 | 88.0782 | 82.1398 | [7 empty] | 1
| 1 | 2 | 84.32 | 66.3376 | 67.2 | 83.29 | 77.8 | 99.2 | 81.5 | 84.8 | 79.2 | 87.1 | 80.4 | 82.5942 | 69.22 | 82.5 | 85.2 | 87.3096 | 81.9184 | [7 empty] | 37
| 1 | 3 | 83.34 | 68.1634 | 66.8 | 83.42 | 78.3 | 99.1 | 78.9 | 84.1 | 84.7 | 82.4 | 81.6 | 83.5663 | 69.16 | 85.2 | 83.7 | 86.0424 | 82.8739 | [7 empty] | 16
| 1 | 4 | 84.47 | 66.8156 | 68.5 | 85.23 | 80.7 | 98.2 | 80.5 | 85.3 | 81.4 | 84.3 | 81.3 | 83.9888 | 68.81 | 85.6 | 83.4 | 88.1756 | 83.6143 | [7 empty] | 83.3284
| 1 | 5 | 82.63 | 66.6935 | 64.2 | 83.02 | 81.9 | 91.9 | 80.3 | 83.1 | 82.9 | 83.7 | 84 | 84.1605 | 68.42 | 88.4 | 85.7 | 71.8503 | 83.1416 | [7 empty] | 103
| 1 | 6 | 82.57 | 66.3753 | 66.1 | 83.15 | 80.4 | 98.8 | 80.7 | 84.6 | 79.7 | 84.9 | 78.6 | 83.6027 | 65.99 | 84.1 | 83.9 | 89.118 | 84.6602
| Mean |  | 83.4067 | 66.8072 | 66.6667 | 83.93 | 80.2 | 96.8167 | 80.2333 | 84.5333 | 82.2167 | 84.35 | 81.0333 | 83.5459 | 68.4133 | 84.7 | 84.3167 | 85.0957 | 83.058
| Median |  | 83.225 | 66.5756 | 67 | 83.355 | 80.55 | 98.5 | 80.4 | 84.7 | 82.15 | 84 | 80.85 | 83.5845 | 68.845 | 84.65 | 83.95 | 87.6939 | 83.0078
| Std Dev. |  | 0.81982 | 0.69014 | 1.439 | 1.1107 | 1.7978 | 3.1821 | 0.91797 | 0.83586 | 2.5639 | 1.5821 | 1.7918 | 0.55078 | 1.2211 | 2.2468 | 0.91524 | 6.5696 | 1.0059
| Rel.Std.Dev. |  | 0.9829% | 1.033% | 2.158% | 1.323% | 2.242% | 3.287% | 1.144% | 0.9888% | 3.118% | 1.876% | 2.211% | 0.6593% | 1.785% | 2.653% | 1.085% | 7.72% | 1.211%
| PDM3 |  | 0.09397% | -19.83% | -20% | 0.722% | -3.754% | 16.19% | -3.714% | 1.446% | -1.334% | 1.226% | -2.754% | 0.2611% | -17.9% | 1.646% | 1.186% | 2.121% | -0.3244%
| Z-Score (Absolute) |  | 0 | 8.65 | 8.73 | 0.27 | 1.67 | 6.99 | 1.65 | 0.59 | 0.62 | 0.49 | 1.24 | 0.07 | 7.82 | 0.67 | 0.47 | 0.88 | 0.18
| Analytical results for Re in OREAS 239b (Certified Value < 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0012 | [12 empty] | 1
| 1 | 2 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0005 | [12 empty] | 38
| 1 | 3 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.001 | [12 empty] | 16
| 1 | 4 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0012 | [12 empty] | < 0.001
| 1 | 5 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0006 | [12 empty] | 104
| 1 | 6 | < 0.001 | < 0.001 | < 0.05 | < 0.002 | < 0.05 | < 0.005 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.0012
| Mean | [7 empty] | 0.001 |  |  |  |  | 0.00095
| Median | [7 empty] | 0.001 |  |  |  |  | 0.0011
| Std Dev. | [7 empty] | 0 |  |  |  |  | 0.00032094
| Rel.Std.Dev. | [7 empty] | 0% |  |  |  |  | 33.78%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 131.49 | 2.02 | 131.49 | 10.11 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 1.75
| Analytical results for Rh in OREAS 239b (Indicative Value 0.001 ppm) | [25 empty] | Indicative
Rh | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.0008 | [23 empty] | 1
| 1 | 2 | 0.0014 | [23 empty] | 6
| 1 | 3 | 0.0009 | [23 empty] | 16
| 1 | 4 | 0.0016 | [23 empty] | 0.0012
| 1 | 5 | 0.0008 | [23 empty] | 19
| 1 | 6 | 0.0017
| Mean |  | 0.0012
| Median |  | 0.00115
| Std Dev. |  | 0.00041473
| Rel.Std.Dev. |  | 34.56%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 239b (Indicative Value < 0.0056 ppm) | [25 empty] | Indicative
Ru | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | ppm
|  |  | AR*MS | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | < 0.0056 | [23 empty] | 1
| 1 | 2 | < 0.0047 | [23 empty] | 14
| 1 | 3 | < 0.0037 | [23 empty] | 16
| 1 | 4 | < 0.0028 | [23 empty] | < 0.0056
| 1 | 5 | < 0.0056 | [23 empty] | 20
| 1 | 6 | 0.0009
| Mean |  | 0.0009
| Median |  | 0.0009
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 239b (Certified Value 0.207 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.19 | 0.21935 | 0.197 | 0.176 | 0.21 | 0.21 | 0.21 | 0.22 | 0.22 | 0.22 | 0.21 | 0.2 | 0.21 | 0.168 | 0.19 | 0.005641 | 0.21 | 0.21 | 0.2436 | 0.2202 |  |  |  |  | 1
| 1 | 2 | 0.2 | 0.21515 | 0.196 | 0.17867 | 0.21 | 0.21 | 0.21 | 0.22 | 0.22 | 0.22 | 0.21 | 0.2 | 0.21 | 0.178 | 0.19 | 0.005798 | 0.2 | 0.21 | 0.2432 | 0.2034 |  |  |  |  | 23
| 1 | 3 | 0.2 | 0.21712 | 0.196 | 0.17333 | 0.21 | 0.2 | 0.21 | 0.22 | 0.22 | 0.22 | 0.21 | 0.2 | 0.22 | 0.168 | 0.19 | 0.005332 | 0.21 | 0.21 | 0.2437 | 0.2157 |  |  |  |  | 16
| 1 | 4 | 0.19 | 0.21529 | 0.193 | 0.17333 | 0.21 | 0.2 | 0.21 | 0.22 | 0.22 | 0.21 | 0.22 | 0.2 | 0.2 | 0.169 | 0.19 | 0.005529 | 0.2 | 0.21 | 0.2471 | 0.2006 |  |  |  |  | 0.20722
| 1 | 5 | 0.19 | 0.2128 | 0.191 | 0.17867 | 0.21 | 0.2 | 0.21 | 0.22 | 0.22 | 0.21 | 0.22 | 0.2 | 0.21 | 0.169 | 0.18 | 0.005412 | 0.21 | 0.21 | 0.2491 | 0.2001 |  |  |  |  | 105
| 1 | 6 | 0.19 | 0.21557 | 0.201 | 0.176 | 0.21 | 0.2 | 0.21 | 0.22 | 0.22 | 0.21 | 0.21 | 0.21 | 0.21 | 0.182 | 0.18 | 0.005444 | 0.2 | 0.21 | 0.2462 | 0.1941
| Mean |  | 0.19333 | 0.21588 | 0.19567 | 0.176 | 0.21 | 0.20333 | 0.21 | 0.22 | 0.22 | 0.215 | 0.21333 | 0.20167 | 0.21 | 0.17233 | 0.18667 | 0.005526 | 0.205 | 0.21 | 0.24548 | 0.20568
| Median |  | 0.19 | 0.21543 | 0.196 | 0.176 | 0.21 | 0.2 | 0.21 | 0.22 | 0.22 | 0.215 | 0.21 | 0.2 | 0.21 | 0.169 | 0.19 | 0.0054865 | 0.205 | 0.21 | 0.24495 | 0.202
| Std Dev. |  | 0.005164 | 0.0021939 | 0.0034448 | 0.0023851 | 0 | 0.005164 | 0 | 0 | 0 | 0.0054772 | 0.005164 | 0.0040825 | 0.0063246 | 0.0060882 | 0.005164 | 0.00017004 | 0.0054772 | 0 | 0.0023727 | 0.010074
| Rel.Std.Dev. |  | 2.671% | 1.016% | 1.761% | 1.355% | 0% | 2.54% | 0% | 0% | 0% | 2.548% | 2.421% | 2.024% | 3.012% | 3.533% | 2.766% | 3.077% | 2.672% | 0% | 0.9665% | 4.898%
| PDM3 |  | -6.703% | 4.177% | -5.577% | -15.07% | 1.34% | -1.877% | 1.34% | 6.166% | 6.166% | 3.753% | 2.949% | -2.681% | 1.34% | -16.84% | -9.92% | -97.33% | -1.073% | 1.34% | 18.46% | -0.7428%
| Z-Score (Absolute) |  | 1.29 | 0.71 | 1.08 | 2.83 | 0.19 | 0.4 | 0.19 | 1.08 | 1.08 | 0.64 | 0.49 | 0.55 | 0.19 | 3.15 | 1.88 | 17.96 | 0.25 | 0.19 | 3.34 | 0.19
| Analytical results for Sb in OREAS 239b (Certified Value 554 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 0
| 1 | 1 | 625.4 | > 500 | 602.605 | 488.86 | 320.13 | 526 | 546.61 | 567 | 623 | 607 | 580 | 652.52 | 633.4405 | 453 | 641 | 465 | 660.72 | 663.3 | 492.94 | 419.1 | 161.9285 |  |  |  | 1
| 1 | 2 | 623.25 | > 500 | 606.99 | 472.68 | 311.32 | 556 | 541.56 | 585 | 606 | 600 | 589 | 669.73 | 624.7542 | 489 | 660.1 | 472 | 653.55 | 653.4 | 507.48 | 420.4 | 161.9101 |  |  |  | 39
| 1 | 3 | 619.28 | > 500 | 606.33 | 481.23 | 314.82 | 549 | 574.08 | 584 | 631 | 588 | 579 | 694.7 | 633.7519 | 473 | 668.4 | 498 | 643.16 | 653.5 | 498.75 | 416.7 | 161.8178 |  |  |  | 16
| 1 | 4 | 625.6 | 481 | 599.45 | 493.55 | 314.81 | 553 | 546.6 | 582 | 604 | 608 | 599 | 656.49 | 624.3316 | 458 | 624 | 493 | 668.99 | 650.1 | 481.74 | 428.4 | 161.8916 |  |  |  | 553.6866
| 1 | 5 | 615.87 | 455 | 600.275 | 489.29 | 318.19 | 522 | 550.38 | 567 | 607 | 615 | 591 | 642.55 | 629.0987 | 474 | 636.5 | 457 | 634.4 | 656.2 | 481.94 | 436.7 | 161.8916 |  |  |  | 106
| 1 | 6 | 624.76 | 496 | 604.33 | 476.45 | 317.93 | 540 | 555.82 | 582 | 604 | 615 | 601 | 664.46 | 627.3967 | 485 | 549.7 | 460 | 636.47 | 651.1 | 467.57 | 415.2 | 161.9193
| Mean |  | 622.36 | 477.3333 | 603.33 | 483.6767 | 316.2 | 541 | 552.5083 | 577.8333 | 612.5 | 605.5 | 589.8333 | 663.4083 | 628.7956 | 472 | 629.95 | 474.1667 | 649.5483 | 654.6 | 488.4033 | 422.75 | 161.8932
| Median |  | 624.005 | 481 | 603.4675 | 485.045 | 316.375 | 544.5 | 548.495 | 582 | 606.5 | 607.5 | 590 | 660.475 | 628.2477 | 473.5 | 638.75 | 468.5 | 648.355 | 653.45 | 487.44 | 419.75 | 161.9008
| Std Dev. |  | 3.9471 | 20.7445 | 3.1083 | 8.1853 | 3.1638 | 14.2829 | 11.5793 | 8.4715 | 11.5715 | 10.2518 | 9.2177 | 18.0096 | 4.1097 | 14.2829 | 42.4921 | 17.3599 | 13.8573 | 4.7624 | 14.225 | 8.2338 | 0.039757
| Rel.Std.Dev. |  | 0.6342% | 4.346% | 0.5152% | 1.692% | 1.001% | 2.64% | 2.096% | 1.466% | 1.889% | 1.693% | 1.563% | 2.715% | 0.6536% | 3.026% | 6.745% | 3.661% | 2.133% | 0.7275% | 2.913% | 1.948% | 0.02456%
| PDM3 |  | 12.4% | -13.79% | 8.966% | -12.64% | -42.89% | -2.291% | -0.2128% | 4.361% | 10.62% | 9.358% | 6.528% | 19.82% | 13.57% | -14.75% | 13.77% | -14.36% | 17.31% | 18.23% | -11.79% | -23.65% | -70.76%
| Z-Score (Absolute) |  | 0.42 | 0.94 | 0.24 | 0.89 | 2.46 | 0.35 | 0.24 | 0 | 0.33 | 0.26 | 0.11 | 0.8 | 0.48 | 1 | 0.49 | 0.97 | 0.67 | 0.72 | 0.84 | 1.46 | 3.91
| Analytical results for Sc in OREAS 239b (Certified Value 6.33 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 6.26 | 5.7495 | 6 | 6.0733 | 6 | 7 | 7 | 6.3 | 7.5 | 7.4 | 6.3 | 6.9 | 5.8 | 7.0124 | 6 | 5.623 | 6.7 | 6.4 | 5.8 | 5.96 | 10.6639 | 6.3572 |  |  | 1
| 1 | 2 | 6.39 | 5.5027 | 5.8 | 6.2 | 6 | 7 | 7 | 6.3 | 7.4 | 6.8 | 6.6 | 6.8 | 5.8 | 6.8267 | 6 | 5.945 | 6.5 | 6.4 | 6 | 5.92 | 10.665 | 6.2554 |  |  | 40
| 1 | 3 | 6.43 | 5.6793 | 6 | 6.05 | 6 | 7 | 7 | 6.1 | 7.4 | 7.2 | 6.4 | 6.8 | 5.9 | 6.7903 | 5 | 6.042 | 6.8 | 6.4 | 6 | 5.85 | 10.4168 | 6.328 |  |  | 16
| 1 | 4 | 6.33 | 5.3527 | 6.6 | 6.14 | 6 | 7 | 7 | 6.3 | 7.4 | 7 | 6.3 | 6.8 | 5.8 | 6.868 | 6 | 6.471 | 6.5 | 6.2 | 6 | 5.99 | 10.6207 | 6.5241 |  |  | 6.3277
| 1 | 5 | 6.23 | 5.5634 | 5.9 | 6.1167 | 6 | 6 | 7 | 6.2 | 7.3 | 6.8 | 6.3 | 7 | 5.7 | 6.9162 | 6 | 5.851 | 6.2 | 6.4 | 6.1 | 5.98 | 10.7013 | 6.4407 |  |  | 107
| 1 | 6 | 6.34 | 5.4827 | 6.2 | 6.1767 | 6 | 7 | 7 | 6.3 | 7.4 | 6.9 | 6.3 | 6.6 | 5.7 | 7.0283 | 5 | 5.91 | 6.3 | 6.4 | 6.1 | 5.95 | 10.4532 | 6.5208
| Mean |  | 6.33 | 5.5551 | 6.0833 | 6.1261 | 6 | 6.8333 | 7 | 6.25 | 7.4 | 7.0167 | 6.3667 | 6.8167 | 5.7833 | 6.907 | 5.6667 | 5.9737 | 6.5 | 6.3667 | 6 | 5.9417 | 10.5868 | 6.4044
| Median |  | 6.335 | 5.5331 | 6 | 6.1283 | 6 | 7 | 7 | 6.3 | 7.4 | 6.95 | 6.3 | 6.8 | 5.8 | 6.8921 | 6 | 5.9275 | 6.5 | 6.4 | 6 | 5.955 | 10.6423 | 6.399
| Std Dev. |  | 0.075631 | 0.14301 | 0.28577 | 0.058096 | 0 | 0.40825 | 0 | 0.083666 | 0.063246 | 0.24014 | 0.12111 | 0.13292 | 0.075277 | 0.09745 | 0.5164 | 0.28103 | 0.22804 | 0.08165 | 0.10954 | 0.051153 | 0.12089 | 0.10904
| Rel.Std.Dev. |  | 1.195% | 2.574% | 4.698% | 0.9483% | 0% | 5.974% | 0% | 1.339% | 0.8547% | 3.422% | 1.902% | 1.95% | 1.302% | 1.411% | 9.113% | 4.705% | 3.508% | 1.282% | 1.826% | 0.8609% | 1.142% | 1.703%
| PDM3 |  | 0.03558% | -12.21% | -3.863% | -3.187% | -5.18% | 7.99% | 10.62% | -1.229% | 16.95% | 10.89% | 0.615% | 7.727% | -8.604% | 9.154% | -10.45% | -5.596% | 2.722% | 0.615% | -5.18% | -6.101% | 67.31% | 1.211%
| Z-Score (Absolute) |  | 0.03 | 1.48 | 0.49 | 0.41 | NA | NA | NA | 0.18 | 1.96 | 1.25 | 0.03 | 0.87 | 1.05 | 1.04 | NA | 0.7 | 0.28 | 0.03 | 0.65 | 0.76 | 7.91 | 0.1
| Analytical results for Se in OREAS 239b (Indicative Value 0.33 ppm) | [25 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.21 | 0.6 | < 5 | < 1 | 0.8 | < 1 | 2.1 | 0.3 | 0.4 | 0.2 | 0.3 | 0.5 | < 1 | 0.25 | < 1 | 0.3 | 0.269 | 2.7096 | 0.32863 | [5 empty] | 1
| 1 | 2 | 0.23 | 0.3 | < 5 | < 1 | 0.8 | < 1 | 2.4 | 0.3 | 0.4 | 0.4 | 0.5 | 0.3 | < 1 | 0.26 | < 1 | 0.3 | 0.2428 | 3.1049 | 0.34004 | [5 empty] | 41
| 1 | 3 | 0.22 | 0.4 | < 5 | < 1 | 0.8 | < 1 | 2.3 | 0.6 | 0.4 | 0.2 | 0.3 | 0.4 | < 1 | 0.27 | < 1 | 0.3 | 0.2621 | 2.6985 | 0.33764 | [5 empty] | 16
| 1 | 4 | 0.21 | 0.5 | < 5 | < 1 | 0.9 | < 1 | 2.3 | 0.3 | 0.4 | 0.3 | 0.4 | 0.3 | < 1 | 0.28 | < 1 | 0.4 | 0.2566 | 3.0052 | 0.35738 | [5 empty] | 0.32977
| 1 | 5 | 0.21 | 0.4 | < 5 | < 1 | 1 | < 1 | 2.5 | 0.4 | 0.3 | 0.2 | 0.4 | 0.6 | < 1 | 0.29 | < 1 | 0.3 | 0.298 | 2.588 | 0.36347 | [5 empty] | 14
| 1 | 6 | 0.2 | 0.4 | < 5 | < 1 | 0.8 | < 1 | 2.4 | 0.5 | 0.3 | 0.2 | 0.3 | 0.3 | < 1 | 0.3 | < 1 | 0.3 | 0.258 | 2.9385 | 0.36178
| Mean |  | 0.21333 | 0.43333 |  |  | 0.85 |  | 2.3333 | 0.4 | 0.36667 | 0.25 | 0.36667 | 0.4 |  | 0.275 |  | 0.31667 | 0.26442 | 2.8408 | 0.34816
| Median |  | 0.21 | 0.4 |  |  | 0.8 |  | 2.35 | 0.35 | 0.4 | 0.2 | 0.35 | 0.35 |  | 0.275 |  | 0.3 | 0.26005 | 2.824 | 0.34871
| Std Dev. |  | 0.010328 | 0.10328 |  |  | 0.083666 |  | 0.13663 | 0.12649 | 0.05164 | 0.083666 | 0.08165 | 0.12649 |  | 0.018708 |  | 0.040825 | 0.018568 | 0.20381 | 0.01458
| Rel.Std.Dev. |  | 4.841% | 23.83% |  |  | 9.843% |  | 5.855% | 31.62% | 14.08% | 33.47% | 22.27% | 31.62% |  | 6.803% |  | 12.89% | 7.022% | 7.174% | 4.188%
| PDM3 |  | -35.31% | 31.4% |  |  | 157.8% |  | 607.6% | 21.29% | 11.19% | -24.19% | 11.19% | 21.29% |  | -16.61% |  | -3.975% | -19.82% | 761.4% | 5.574%
| Z-Score (Absolute) |  | 1.26 | 0.22 | 14.16 | 0.67 | 3.03 | 0.67 | 13.04 | 0 | 0.22 | 1.01 | 0.22 | 0 | 0.67 | 0.84 | 0.67 | 0.56 | 0.91 | 16.46 | 0.35
| Analytical results for Si in OREAS 239b (Indicative Value 0.409 wt.%) | [25 empty] | Indicative
Si | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.30 | [23 empty] | wt.%
|  |  | AR*OES | [23 empty] | 3
|  |  | 0.5g | [23 empty] | 3
| 1 | 1 | 0.438 | [23 empty] | 1
| 1 | 2 | 0.4004 | [23 empty] | 9
| 1 | 3 | 0.3963 | [23 empty] | 16
| 1 | 4 | 0.4007 | [23 empty] | 0.40938
| 1 | 5 | 0.4114 | [23 empty] | 15
| 1 | 6 | 0.4095
| Mean |  | 0.40938
| Median |  | 0.4051
| Std Dev. |  | 0.015168
| Rel.Std.Dev. |  | 3.705%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 239b (Indicative Value 4.26 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 4.8369 | 4 | 4.429 | 3.9194 | [20 empty] | 1
| 1 | 2 | 4.5761 | 3.9 | 4.695 | 3.9552 | [20 empty] | 10
| 1 | 3 | 4.8793 | 4.3 | 4.675 | 3.6936 | [20 empty] | 16
| 1 | 4 | 4.6582 | 3.7 | 5.01 | 3.8093 | [20 empty] | 4.2562
| 1 | 5 | 4.569 | 3.3 | 4.739 | 3.0548 | [20 empty] | 16
| 1 | 6 | 4.7311 | 3.6 | 4.5 | 3.8309
| Mean |  | 4.7084 | 3.8 | 4.6747 | 3.7105
| Median |  | 4.6946 | 3.8 | 4.685 | 3.8201
| Std Dev. |  | 0.13092 | 0.34641 | 0.20389 | 0.33407
| Rel.Std.Dev. |  | 2.781% | 9.116% | 4.362% | 9.003%
| PDM3 |  | 10.63% | -10.72% | 9.832% | -12.82%
| Z-Score (Absolute) |  | 0.7 | 0.65 | 0.65 | 0.78
| Analytical results for Sn in OREAS 239b (Certified Value 1.45 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.51 | 2.3171 | 1.48 | < 10 | 1.4 | 1.4 | 1.9 | 1.5 | 1.4 | 1.5 | 1.4 | 1.5 | 1.547 | < 1 | 1.4 | < 5 | 1.3 | 1.3041 | 2.1979 | [5 empty] | 1
| 1 | 2 | 1.5 | 2.2662 | 1.48 | < 10 | 1.4 | 1.7 | 1.7 | 1.5 | 1.5 | 1.3 | 1.5 | 1.5 | 1.5395 | < 1 | 1.5 | < 5 | 1.4 | 1.3072 | 2.3997 | [5 empty] | 26
| 1 | 3 | 1.52 | 2.2949 | 1.49 | < 10 | 1.4 | 1.5 | 1.9 | 1.4 | 1.4 | 1.4 | 1.4 | 1.5 | 1.5583 | < 1 | 1.5 | < 5 | 1.4 | 1.2459 | 2.6251 | [5 empty] | 16
| 1 | 4 | 1.51 | 2.2721 | 1.47 | < 10 | 1.4 | 1.6 | 1.8 | 1.6 | 1.5 | 1.4 | 1.5 | 1.5 | 1.5603 | < 1 | 1.4 | < 5 | 1.4 | 1.3667 | 2.4394 | [5 empty] | 1.4471
| 1 | 5 | 1.52 | 2.2411 | 1.34 | < 10 | 1.4 | 1.6 | 1.9 | 1.5 | 1.4 | 1.4 | 1.4 | 1.5 | 1.5287 | < 1 | 1.3 | < 5 | 1.3 | 1.0298 | 2.38 | [5 empty] | 108
| 1 | 6 | 1.51 | 2.2311 | 1.48 | < 10 | 1.4 | 1.5 | 1.8 | 1.5 | 1.5 | 1.4 | 1.4 | 1.5 | 1.5556 | < 1 | 1.2 | < 5 | 1.3 | 1.3041 | 2.4338
| Mean |  | 1.5117 | 2.2704 | 1.4567 |  | 1.4 | 1.55 | 1.8333 | 1.5 | 1.45 | 1.4 | 1.4333 | 1.5 | 1.5482 |  | 1.3833 |  | 1.35 | 1.2596 | 2.4126
| Median |  | 1.51 | 2.2692 | 1.48 |  | 1.4 | 1.55 | 1.85 | 1.5 | 1.45 | 1.4 | 1.4 | 1.5 | 1.5513 |  | 1.4 |  | 1.35 | 1.3041 | 2.4168
| Std Dev. |  | 0.0075277 | 0.032268 | 0.057504 |  | 2.4324e-16 | 0.10488 | 0.08165 | 0.063246 | 0.054772 | 0.063246 | 0.05164 | 0 | 0.012307 |  | 0.1169 |  | 0.054772 | 0.11891 | 0.13686
| Rel.Std.Dev. |  | 0.498% | 1.421% | 3.948% |  | 1.737e-14% | 6.767% | 4.454% | 4.216% | 3.777% | 4.518% | 3.603% | 0% | 0.7949% |  | 8.451% |  | 4.057% | 9.44% | 5.672%
| PDM3 |  | 4.463% | 56.9% | 0.6618% |  | -3.254% | 7.112% | 26.69% | 3.656% | 0.2011% | -3.254% | -0.9506% | 3.656% | 6.989% |  | -4.406% |  | -6.709% | -12.95% | 66.72%
| Z-Score (Absolute) |  | 0.08 | 5.2 | 0.29 | 23.6 | 0.67 | 0.34 | 2.25 | 0 | 0.34 | 0.67 | 0.45 | 0 | 0.33 | 6.74 | 0.79 | 6.74 | 1.01 | 1.62 | 6.15
| Analytical results for Sr in OREAS 239b (Certified Value 28.9 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 29.39 | 29.2518 | 36.9 | 40.2333 | 30.8 | 31 | 36.6 | 29 | 30.4 | 29.4 | 28.4 | 29.9 | 28.8 | 29.7495 | 28 | 26.76 | 27.1 | 27.6 | 29.8 | 22.4 | 121.5495 | 28.0462 |  |  | 1
| 1 | 2 | 29.88 | 28.8636 | 36.9 | 41.3667 | 29.7 | 30 | 36.5 | 29 | 29.8 | 26.7 | 30 | 29.9 | 28.9 | 30.9367 | 28 | 27.11 | 26.8 | 28.3 | 29.9 | 22.1 | 121.2125 | 27.8269 |  |  | 27
| 1 | 3 | 29.68 | 28.9339 | 36.7 | 40.4433 | 30 | 31 | 36 | 29 | 29.7 | 28.8 | 28.8 | 30.5 | 29.3 | 30.4828 | 28 | 25.53 | 27 | 27.6 | 29.5 | 21.7 | 120.8674 | 26.5425 |  |  | 16
| 1 | 4 | 29.67 | 28.4702 | 39.2 | 40.96 | 30.9 | 30 | 36.4 | 29 | 30.2 | 27.8 | 29.1 | 30.2 | 28.6 | 30.3516 | 28 | 25.1 | 28 | 25.4 | 29.8 | 22.4 | 121.1348 | 30.1775 |  |  | 28.8958
| 1 | 5 | 29.66 | 28.4538 | 34.6 | 40.45 | 29.9 | 30 | 34.4 | 29 | 29.1 | 28.3 | 29 | 31.4 | 27.7 | 29.6492 | 28 | 26.64 | 26.5 | 27.1 | 30.4 | 22.6 | 121.6606 | 29.6343 |  |  | 109
| 1 | 6 | 29.6 | 27.5758 | 34 | 41.2767 | 29.2 | 31 | 36.5 | 29 | 30.2 | 27.9 | 28.7 | 29.3 | 27.9 | 30.6024 | 28 | 26.01 | 26.9 | 28 | 29.7 | 22.3 | 120.9891 | 30.0095
| Mean |  | 29.6467 | 28.5915 | 36.3833 | 40.7883 | 30.0833 | 30.5 | 36.0667 | 29 | 29.9 | 28.15 | 29 | 30.2 | 28.5333 | 30.2954 | 28 | 26.1917 | 27.05 | 27.3333 | 29.85 | 22.25 | 121.2357 | 28.7062
| Median |  | 29.665 | 28.6669 | 36.8 | 40.705 | 29.95 | 30.5 | 36.45 | 29 | 30 | 28.1 | 28.9 | 30.05 | 28.7 | 30.4172 | 28 | 26.325 | 26.95 | 27.6 | 29.8 | 22.35 | 121.1737 | 28.8403
| Std Dev. |  | 0.15769 | 0.58184 | 1.867 | 0.47832 | 0.65549 | 0.54772 | 0.84301 | 0 | 0.47329 | 0.92682 | 0.54772 | 0.70993 | 0.61536 | 0.50187 | 0 | 0.77845 | 0.50892 | 1.0309 | 0.30166 | 0.31464 | 0.31183 | 1.4571
| Rel.Std.Dev. |  | 0.5319% | 2.035% | 5.131% | 1.173% | 2.179% | 1.796% | 2.337% | 0% | 1.583% | 3.292% | 1.889% | 2.351% | 2.157% | 1.657% | 0% | 2.972% | 1.881% | 3.771% | 1.011% | 1.414% | 0.2572% | 5.076%
| PDM3 |  | 2.599% | -1.053% | 25.91% | 41.16% | 4.11% | 5.552% | 24.82% | 0.3607% | 3.475% | -2.581% | 0.3607% | 4.514% | -1.254% | 4.844% | -3.1% | -9.358% | -6.388% | -5.407% | 3.302% | -23% | 319.6% | -0.6562%
| Z-Score (Absolute) |  | 0.2 | 0.46 | 4.44 | 7.21 | 0.48 | 0.74 | 4.24 | 0.2 | 0.36 | 0.74 | 0.2 | 0.55 | 0.5 | 0.61 | 0.83 | 1.97 | 1.43 | 1.25 | 0.33 | 4.45 | 57.78 | 0.39
| Analytical results for Ta in OREAS 239b (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.1201 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0467 | [10 empty] | 28
| 1 | 3 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.3873 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0433 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.01 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.1486 | [10 empty] | 110
| 1 | 6 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0211
| Mean | [5 empty] | 0.018333 |  |  | 0.01 | 0.01 |  |  |  |  | 0.12785
| Median | [5 empty] | 0.02 |  |  | 0.01 | 0.01 |  |  |  |  | 0.0834
| Std Dev. | [5 empty] | 0.0040825 |  |  | 0 | 0 |  |  |  |  | 0.13633
| Rel.Std.Dev. | [5 empty] | 22.27% |  |  | 0% | 0% |  |  |  |  | 106.6%
PDM3
| Z-Score (Absolute) |  | 0.77 | 0.15 | 114.59 | 0.15 | 0.15 | 0.15 | 1.31 | 0.58 | 0.54 | 0.77 | 0.77 | 0.15 | 0.77 | 4.91
| Analytical results for Tb in OREAS 239b (Indicative Value 0.49 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | [19 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.5116 | 0.4 | 0.449 | 0.51 | 1.41 | [19 empty] | 1
| 1 | 2 | 0.47298 | 0.4 | 0.474 | 0.51 | 1.41 | [19 empty] | 11
| 1 | 3 | 0.50754 | 0.4 | 0.471 | 0.51 | 1.41 | [19 empty] | 16
| 1 | 4 | 0.48738 | 0.4 | 0.503 | 0.51 | 1.4 | [19 empty] | 0.49301
| 1 | 5 | 0.48297 | 0.4 | 0.472 | 0.52 | 1.42 | [19 empty] | 17
| 1 | 6 | 0.50763 | 0.4 | 0.455 | 0.52 | 1.4
| Mean |  | 0.49502 | 0.4 | 0.47067 | 0.51333 | 1.4083
| Median |  | 0.49746 | 0.4 | 0.4715 | 0.51 | 1.41
| Std Dev. |  | 0.015999 | 6.0809e-17 | 0.018833 | 0.005164 | 0.0075277
| Rel.Std.Dev. |  | 3.232% | 1.52e-14% | 4.001% | 1.006% | 0.5345%
| PDM3 |  | 0.408% | -18.87% | -4.531% | 4.123% | 185.7%
| Z-Score (Absolute) |  | 0 | 2.63 | 0.67 | 0.51 | 25.29
| Analytical results for Te in OREAS 239b (Indicative Value 0.027 ppm) | [25 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.02 | 0.03 | < 5 | < 0.05 | 0.08 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.03 | < 0.05 | 0.04 | 0.012 | [9 empty] | 1
| 1 | 2 | 0.02 | 0.06 | 5.9352 | < 0.05 | 0.09 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.01 | 0.02 | < 0.05 | 0.04 | 0.0358 | [9 empty] | 5
| 1 | 3 | < 0.02 | 0.04 | < 5 | < 0.05 | 0.09 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.01 | 0.03 | < 0.05 | 0.05 | [10 empty] | 16
| 1 | 4 | < 0.02 | 0.03 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | 0.03 | 0.02 | 0.02 | 0.01 | 0.02 | < 0.05 | 0.04 | 0.0478 | [9 empty] | 0.026504
| 1 | 5 | 0.02 | 0.03 | 5.0638 | < 0.05 | 0.06 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.03 | < 0.05 | 0.05 | [10 empty] | 18
| 1 | 6 | < 0.02 | 0.05 | 5.5235 | < 0.05 | 0.07 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.04 | < 0.0119
| Mean |  | 0.02 | 0.04 | 5.5075 |  | 0.076667 |  |  | 0.021667 | 0.02 | 0.025 | 0.011667 | 0.025 |  | 0.043333 | 0.031867
| Median |  | 0.02 | 0.035 | 5.5235 |  | 0.075 |  |  | 0.02 | 0.02 | 0.025 | 0.01 | 0.025 |  | 0.04 | 0.0358
| Std Dev. |  | 0 | 0.012649 | 0.43589 |  | 0.012111 |  |  | 0.0040825 | 0 | 0.0054772 | 0.0040825 | 0.0054772 |  | 0.005164 | 0.018221
| Rel.Std.Dev. |  | 0% | 31.62% | 7.914% |  | 15.8% |  |  | 18.84% | 0% | 21.91% | 34.99% | 21.91% |  | 11.92% | 57.18%
| PDM3 |  | -24.54% | 50.92% | 20680% |  | 189.3% |  |  | -18.25% | -24.54% | -5.674% | -55.98% | -5.674% |  | 63.5% | 20.23%
| Z-Score (Absolute) |  | 1.57 | 2.02 | 536.58 | 0 | 6.97 | 0 | 3.37 | 0.45 | 0.67 | 0 | 1.8 | 0 | 0 | 2.47 | 0.05
| Analytical results for Th in OREAS 239b (Certified Value 11.8 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 11.443 | 11.4132 | 13 | 13.16 | 11.6 | 14.6 | 12.4 | 11.9 | 11.8 | 11.9 | 11.7 | 10.8 | 12.7429 | 11.47 | 12.1 | 10.98 | 15.3 | 12.3 | 9.4098 | 10.8752 |  |  |  |  | 1
| 1 | 2 | 11.511 | 11.5951 | 12.4 | 12.75 | 10.6 | 14.6 | 12.7 | 12 | 10.4 | 12 | 11.9 | 11.1 | 13.5341 | 12.2 | 11.7 | 9 | 16.4 | 12.3 | 9.3944 | 10.9203 |  |  |  |  | 31
| 1 | 3 | 11.599 | 11.5559 | 12.6 | 12.8 | 11 | 15.1 | 12.5 | 11.4 | 11.5 | 11.5 | 11.9 | 11.2 | 12.5871 | 11.89 | 12.3 | 6.36 | 15.6 | 12.1 | 9.2892 | 10.6694 |  |  |  |  | 16
| 1 | 4 | 11.696 | 11.575 | 12.6 | 13.13 | 11.2 | 14 | 12.9 | 12 | 11.4 | 11.7 | 12 | 11.1 | 12.8742 | 12.57 | 11.9 | 9.3 | 15.1 | 12.5 | 9.2812 | 11.005 |  |  |  |  | 11.7951
| 1 | 5 | 11.629 | 11.1815 | 12 | 12.81 | 11.6 | 14.6 | 12.4 | 11.5 | 11.4 | 12 | 12 | 11.1 | 13.2134 | 11.91 | 12.4 | 9.36 | 15.9 | 12.8 | 7.2644 | 10.9887 |  |  |  |  | 111
| 1 | 6 | 11.49 | 11.4484 | 12.8 | 12.82 | 11.7 | 14.5 | 12.6 | 12 | 11.4 | 11.9 | 11.4 | 11 | 12.8341 | 11.6 | 12.6 | 8.84 | 16.2 | 12.3 | 9.4301 | 10.4301
| Mean |  | 11.5613 | 11.4615 | 12.5667 | 12.9117 | 11.2833 | 14.5667 | 12.5833 | 11.8 | 11.3167 | 11.8333 | 11.8167 | 11.05 | 12.9643 | 11.94 | 12.1667 | 8.9733 | 15.75 | 12.3833 | 9.0115 | 10.8148
| Median |  | 11.555 | 11.5021 | 12.6 | 12.815 | 11.4 | 14.6 | 12.55 | 11.95 | 11.4 | 11.9 | 11.9 | 11.1 | 12.8541 | 11.9 | 12.2 | 9.15 | 15.75 | 12.3 | 9.3418 | 10.8977
| Std Dev. |  | 0.095663 | 0.15528 | 0.34448 | 0.18258 | 0.43089 | 0.35024 | 0.19408 | 0.27568 | 0.47504 | 0.19664 | 0.23166 | 0.13784 | 0.34729 | 0.4011 | 0.33267 | 1.492 | 0.50892 | 0.24014 | 0.85822 | 0.22375
| Rel.Std.Dev. |  | 0.8274% | 1.355% | 2.741% | 1.414% | 3.819% | 2.404% | 1.542% | 2.336% | 4.198% | 1.662% | 1.96% | 1.247% | 2.679% | 3.359% | 2.734% | 16.63% | 3.231% | 1.939% | 9.524% | 2.069%
| PDM3 |  | -1.982% | -2.828% | 6.541% | 9.466% | -4.339% | 23.5% | 6.683% | 0.04143% | -4.056% | 0.324% | 0.1827% | -6.317% | 9.912% | 1.228% | 3.15% | -23.92% | 33.53% | 4.987% | -23.6% | -8.311%
| Z-Score (Absolute) |  | 0.27 | 0.38 | 0.77 | 1.13 | 0.56 | 2.84 | 0.79 | 0.03 | 0.53 | 0.01 | 0.01 | 0.8 | 1.18 | 0.12 | 0.35 | 2.96 | 4.07 | 0.58 | 2.92 | 1.05
| Analytical results for Ti in OREAS 239b (Certified Value 0.142 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1279 | 0.12 | 0.12867 | 0.161 | 0.151 | 0.185 | 0.135 | 0.145 | 0.159 | 0.148 | 0.147 | 0.136 | 0.15755 | 0.138 | 0.13 | 0.14302 | 0.25 | 0.15 | 0.13592 | 0.2675 | 0.14235 |  |  |  | 1
| 1 | 2 | 0.1323 | 0.12 | 0.132 | 0.1546 | 0.144 | 0.181 | 0.136 | 0.148 | 0.153 | 0.144 | 0.148 | 0.139 | 0.15337 | 0.138 | 0.13 | 0.14207 | 0.24 | 0.151 | 0.13441 | 0.2678 | 0.14135 |  |  |  | 32
| 1 | 3 | 0.132 | 0.12 | 0.12867 | 0.1539 | 0.148 | 0.184 | 0.136 | 0.144 | 0.16 | 0.146 | 0.146 | 0.139 | 0.15426 | 0.138 | 0.13 | 0.13857 | 0.24 | 0.152 | 0.13243 | 0.2555 | 0.13569 |  |  |  | 16
| 1 | 4 | 0.1318 | 0.13 | 0.12867 | 0.1614 | 0.146 | 0.185 | 0.137 | 0.147 | 0.15 | 0.147 | 0.15 | 0.138 | 0.15463 | 0.138 | 0.13 | 0.14015 | 0.24 | 0.153 | 0.13631 | 0.2729 | 0.15455 |  |  |  | 0.14198
| 1 | 5 | 0.1303 | 0.12 | 0.12867 | 0.1555 | 0.153 | 0.183 | 0.136 | 0.144 | 0.151 | 0.148 | 0.148 | 0.133 | 0.15503 | 0.138 | 0.14 | 0.14069 | 0.24 | 0.155 | 0.13582 | 0.2759 | 0.14697 |  |  |  | 112
| 1 | 6 | 0.1303 | 0.13 | 0.12867 | 0.1531 | 0.152 | 0.188 | 0.137 | 0.149 | 0.153 | 0.145 | 0.15 | 0.135 | 0.15955 | 0.138 | 0.13 | 0.13834 | 0.25 | 0.158 | 0.13598 | 0.3269 | 0.14796
| Mean |  | 0.13077 | 0.12333 | 0.12922 | 0.15658 | 0.149 | 0.18433 | 0.13617 | 0.14617 | 0.15433 | 0.14633 | 0.14817 | 0.13667 | 0.15573 | 0.138 | 0.13167 | 0.14047 | 0.24333 | 0.15317 | 0.13514 | 0.27775 | 0.14481
| Median |  | 0.13105 | 0.12 | 0.12867 | 0.15505 | 0.1495 | 0.1845 | 0.136 | 0.146 | 0.153 | 0.1465 | 0.148 | 0.137 | 0.15483 | 0.138 | 0.13 | 0.14042 | 0.24 | 0.1525 | 0.13587 | 0.27035 | 0.14466
| Std Dev. |  | 0.0016488 | 0.005164 | 0.0013608 | 0.0036647 | 0.0035777 | 0.0023381 | 0.00075277 | 0.002137 | 0.0041793 | 0.001633 | 0.0016021 | 0.0024221 | 0.0023388 | 0 | 0.0040825 | 0.0018634 | 0.005164 | 0.0029269 | 0.0014847 | 0.025067 | 0.0064885
| Rel.Std.Dev. |  | 1.261% | 4.187% | 1.053% | 2.34% | 2.401% | 1.268% | 0.5528% | 1.462% | 2.708% | 1.116% | 1.081% | 1.772% | 1.502% | 0% | 3.101% | 1.327% | 2.122% | 1.911% | 1.099% | 9.025% | 4.481%
| PDM3 |  | -7.897% | -13.13% | -8.985% | 10.29% | 4.945% | 29.83% | -4.093% | 2.95% | 8.702% | 3.067% | 4.358% | -3.741% | 9.686% | -2.802% | -7.263% | -1.06% | 71.39% | 7.88% | -4.813% | 95.63% | 1.995%
| Z-Score (Absolute) |  | 1.09 | 1.61 | 1.19 | 0.73 | 0.2 | 2.69 | 0.71 | 0 | 0.58 | 0.01 | 0.14 | 0.67 | 0.67 | 0.58 | 1.02 | 0.4 | 6.85 | 0.49 | 0.78 | 9.28 | 0.1
| Analytical results for Tl in OREAS 239b (Certified Value 0.52 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.51 | 0.52 | < 10 | 0.56 | 0.55 | 0.55 | 0.54 | 0.51 | 0.51 | 0.47 | 0.53 | 0.49 | 0.56233 | 0.75 | 0.48 | < 5 | 0.46 | 0.4685 | 0.6425 | [5 empty] | 1
| 1 | 2 | 0.5 | 0.52 | < 10 | 0.56 | 0.49 | 0.56 | 0.54 | 0.51 | 0.45 | 0.49 | 0.52 | 0.49 | 0.58542 | 0.754 | 0.49 | < 5 | 0.48 | 0.4715 | 0.617 | [5 empty] | 33
| 1 | 3 | 0.5 | 0.52 | < 10 | 0.56 | 0.53 | 0.58 | 0.53 | 0.47 | 0.48 | 0.47 | 0.54 | 0.5 | 0.57105 | 0.747 | 0.47 | < 5 | 0.48 | 0.47 | 0.623 | [5 empty] | 16
| 1 | 4 | 0.51 | 0.53 | < 10 | 0.56 | 0.47 | 0.54 | 0.55 | 0.5 | 0.47 | 0.48 | 0.53 | 0.52 | 0.57467 | 0.751 | 0.47 | < 5 | 0.46 | 0.4649 | 0.6251 | [5 empty] | 0.51737
| 1 | 5 | 0.51 | 0.51 | < 10 | 0.56 | 0.48 | 0.56 | 0.55 | 0.48 | 0.49 | 0.49 | 0.54 | 0.51 | 0.58078 | 0.77 | 0.45 | < 5 | 0.47 | 0.3776 | 0.6302 | [5 empty] | 113
| 1 | 6 | 0.5 | 0.52 | < 10 | 0.56 | 0.47 | 0.56 | 0.54 | 0.51 | 0.47 | 0.47 | 0.52 | 0.5 | 0.57304 | 0.755 | 0.45 | < 5 | 0.48 | 0.4762 | 0.6115
| Mean |  | 0.505 | 0.52 |  | 0.56 | 0.49833 | 0.55833 | 0.54167 | 0.49667 | 0.47833 | 0.47833 | 0.53 | 0.50167 | 0.57455 | 0.7545 | 0.46833 |  | 0.47167 | 0.45478 | 0.62488
| Median |  | 0.505 | 0.52 |  | 0.56 | 0.485 | 0.56 | 0.54 | 0.505 | 0.475 | 0.475 | 0.53 | 0.5 | 0.57386 | 0.7525 | 0.47 |  | 0.475 | 0.46925 | 0.62405
| Std Dev. |  | 0.0054772 | 0.0063246 |  | 0 | 0.033714 | 0.013292 | 0.0075277 | 0.017512 | 0.020412 | 0.0098319 | 0.0089443 | 0.01169 | 0.0080108 | 0.0081179 | 0.016021 |  | 0.0098319 | 0.037993 | 0.010802
| Rel.Std.Dev. |  | 1.085% | 1.216% |  | 0% | 6.765% | 2.381% | 1.39% | 3.526% | 4.267% | 2.055% | 1.688% | 2.33% | 1.394% | 1.076% | 3.421% |  | 2.085% | 8.354% | 1.729%
| PDM3 |  | -2.392% | 0.5075% |  | 8.239% | -3.68% | 7.917% | 4.695% | -4.002% | -7.546% | -7.546% | 2.44% | -3.036% | 11.05% | 45.83% | -9.479% |  | -8.835% | -12.1% | 20.78%
| Z-Score (Absolute) |  | 0.24 | 0 | 72.5 | 0.65 | 0.35 | 0.62 | 0.35 | 0.38 | 0.67 | 0.67 | 0.16 | 0.3 | 0.88 | 3.8 | 0.84 | 32.04 | 0.78 | 1.06 | 1.7
| Analytical results for Tm in OREAS 239b (Indicative Value 0.11 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.30 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | [20 empty] | 2
| 1 | 1 | 0.12527 | 0.1 | 0.114 | 0.114 | [20 empty] | 1
| 1 | 2 | 0.11517 | 0.1 | 0.122 | 0.1136 | [20 empty] | 13
| 1 | 3 | 0.12285 | 0.1 | 0.121 | 0.109 | [20 empty] | 16
| 1 | 4 | 0.12266 | 0.1 | 0.125 | 0.1119 | [20 empty] | 0.11356
| 1 | 5 | 0.1246 | 0.1 | 0.12 | 0.0911 | [20 empty] | 19
| 1 | 6 | 0.12055 | 0.1 | 0.116 | 0.1152
| Mean |  | 0.12185 | 0.1 | 0.11967 | 0.10913
| Median |  | 0.12276 | 0.1 | 0.1205 | 0.11275
| Std Dev. |  | 0.0036652 | 1.5202e-17 | 0.0040332 | 0.0090919
| Rel.Std.Dev. |  | 3.008% | 1.52e-14% | 3.37% | 8.331%
| PDM3 |  | 7.297% | -11.94% | 5.373% | -3.902%
| Z-Score (Absolute) |  | 0.79 | 1.53 | 0.56 | 0.56
| Analytical results for U in OREAS 239b (Certified Value 1.45 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.376 | 1.4866 | 1.5 | 1.66 | 1.35 | 1.97 | 1.5 | 1.36 | 1.46 | 1.36 | 1.46 | 1.37 | 1.5074 | 1 | 1.621 | 1.39 | 1.53 | 1.2973 | 3.5895 | [5 empty] | 1
| 1 | 2 | 1.411 | 1.3615 | 1.5 | 1.61 | 1.32 | 1.91 | 1.5 | 1.36 | 1.26 | 1.34 | 1.5 | 1.4 | 1.6166 | 1 | 1.661 | 1.38 | 1.59 | 1.2604 | 3.566 | [5 empty] | 35
| 1 | 3 | 1.409 | 1.4992 | 1.5 | 1.67 | 1.32 | 1.91 | 1.5 | 1.3 | 1.39 | 1.32 | 1.51 | 1.45 | 1.6114 | 1 | 1.591 | 1.42 | 1.61 | 1.2429 | 3.7545 | [5 empty] | 16
| 1 | 4 | 1.417 | 1.4406 | 1.5 | 1.69 | 1.33 | 1.89 | 1.6 | 1.36 | 1.38 | 1.32 | 1.54 | 1.38 | 1.5419 | 1 | 1.521 | 1.46 | 1.57 | 1.2556 | 3.9192 | [5 empty] | 1.4506
| 1 | 5 | 1.384 | 1.3779 | 1.4 | 1.61 | 1.35 | 1.99 | 1.5 | 1.32 | 1.36 | 1.36 | 1.52 | 1.39 | 1.5596 | 1 | 1.669 | 1.36 | 1.58 | 1.0248 | 3.6834 | [5 empty] | 114
| 1 | 6 | 1.4 | 1.4577 | 1.5 | 1.6 | 1.36 | 1.92 | 1.6 | 1.36 | 1.36 | 1.33 | 1.46 | 1.38 | 1.6079 | 1 | 1.59 | 1.4 | 1.64 | 1.2698 | 3.6199
| Mean |  | 1.3995 | 1.4373 | 1.4833 | 1.64 | 1.3383 | 1.9317 | 1.5333 | 1.3433 | 1.3683 | 1.3383 | 1.4983 | 1.395 | 1.5741 | 1 | 1.6088 | 1.4017 | 1.5867 | 1.2251 | 3.6888
| Median |  | 1.4045 | 1.4492 | 1.5 | 1.635 | 1.34 | 1.915 | 1.5 | 1.36 | 1.37 | 1.335 | 1.505 | 1.385 | 1.5838 | 1 | 1.606 | 1.395 | 1.585 | 1.258 | 3.6517
| Std Dev. |  | 0.016257 | 0.056528 | 0.040825 | 0.037947 | 0.017224 | 0.0392 | 0.05164 | 0.026583 | 0.064627 | 0.018348 | 0.032506 | 0.02881 | 0.044803 | 0 | 0.054547 | 0.034881 | 0.037238 | 0.099824 | 0.132
| Rel.Std.Dev. |  | 1.162% | 3.933% | 2.752% | 2.314% | 1.287% | 2.029% | 3.368% | 1.979% | 4.723% | 1.371% | 2.17% | 2.065% | 2.846% | 0% | 3.39% | 2.489% | 2.347% | 8.148% | 3.578%
| PDM3 |  | -3.521% | -0.9182% | 2.258% | 13.06% | -7.738% | 33.17% | 5.705% | -7.393% | -5.67% | -7.738% | 3.292% | -3.832% | 8.518% | -31.06% | 10.91% | -3.372% | 9.382% | -15.54% | 154.3%
| Z-Score (Absolute) |  | 0.26 | 0 | 0.31 | 1.38 | 0.67 | 3.37 | 0.65 | 0.64 | 0.47 | 0.67 | 0.42 | 0.29 | 0.93 | 2.98 | 1.17 | 0.24 | 1.02 | 1.45 | 15.35
| Analytical results for V in OREAS 239b (Certified Value 67 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 68.6 | 71.73 | 58 | 66.4067 | 64 | 69 | 71 | 71 | 66 | 69 | 68 | 65 | 66 | 67.7919 | 73 | 64 | 67.8 | 67.7 | 61 | 61.6 | 94.7671 | 67.8062 |  |  | 1
| 1 | 2 | 69.7 | 71.02 | 59 | 66.8667 | 66 | 66 | 69 | 71 | 66 | 66 | 67 | 65 | 67 | 68.5722 | 74 | 65 | 68.11 | 67.5 | 61 | 61.5 | 95.2152 | 67.2319 |  |  | 36
| 1 | 3 | 69.7 | 71.07 | 57 | 66.0067 | 65 | 67 | 70 | 70 | 65 | 68 | 66 | 64 | 67 | 68.3091 | 73 | 64 | 68.17 | 67.1 | 61 | 61.2 | 95.2152 | 66.6158 |  |  | 16
| 1 | 4 | 69.3 | 71.41 | 60 | 66.45 | 66 | 68 | 71 | 71 | 66 | 66 | 68 | 65 | 67 | 65.912 | 74 | 66 | 67.9 | 66.5 | 62 | 62.5 | 94.8196 | 68.2394 |  |  | 67.0747
| 1 | 5 | 67.4 | 70.34 | 57 | 66.58 | 63 | 69 | 70 | 70 | 65 | 65 | 68 | 65 | 64 | 68.8811 | 70 | 61 | 67.41 | 67.2 | 62 | 62.4 | 97.0213 | 67.026 |  |  | 115
| 1 | 6 | 68.8 | 71.52 | 61 | 66.4433 | 64 | 69 | 72 | 71 | 66 | 66 | 67 | 65 | 66 | 67.184 | 75 | 60 | 67.64 | 67.7 | 63 | 62 | 95.2152 | 68.1211
| Mean |  | 68.9167 | 71.1817 | 58.6667 | 66.4589 | 64.6667 | 68 | 70.5 | 70.6667 | 65.6667 | 66.6667 | 67.3333 | 64.8333 | 66.1667 | 67.775 | 73.1667 | 63.3333 | 67.8383 | 67.2833 | 61.6667 | 61.8667 | 95.3756 | 67.5067
| Median |  | 69.05 | 71.24 | 58.5 | 66.4467 | 64.5 | 68.5 | 70.5 | 71 | 66 | 66 | 67.5 | 65 | 66.5 | 68.0505 | 73.5 | 64 | 67.85 | 67.35 | 61.5 | 61.8 | 95.2152 | 67.519
| Std Dev. |  | 0.87044 | 0.4928 | 1.633 | 0.27875 | 1.2111 | 1.2649 | 1.0488 | 0.5164 | 0.5164 | 1.5055 | 0.8165 | 0.40825 | 1.169 | 1.0922 | 1.7224 | 2.3381 | 0.28701 | 0.45789 | 0.8165 | 0.52026 | 0.83246 | 0.64873
| Rel.Std.Dev. |  | 1.263% | 0.6923% | 2.784% | 0.4194% | 1.873% | 1.86% | 1.488% | 0.7308% | 0.7864% | 2.258% | 1.213% | 0.6297% | 1.767% | 1.612% | 2.354% | 3.692% | 0.4231% | 0.6805% | 1.324% | 0.8409% | 0.8728% | 0.961%
| PDM3 |  | 2.746% | 6.123% | -12.54% | -0.9181% | -3.59% | 1.38% | 5.107% | 5.355% | -2.099% | -0.6083% | 0.3856% | -3.342% | -1.354% | 1.044% | 9.082% | -5.578% | 1.138% | 0.311% | -8.063% | -7.765% | 42.19% | 0.6441%
| Z-Score (Absolute) |  | 0.53 | 1.27 | 2.83 | 0.28 | 0.87 | 0.23 | 1.05 | 1.1 | 0.54 | 0.21 | 0.01 | 0.81 | 0.37 | 0.15 | 1.92 | 1.3 | 0.17 | 0.01 | 1.85 | 1.78 | 9.19 | 0.06
| Analytical results for W in OREAS 239b (Certified Value 0.56 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.56 | 0.5 | 8.14 | 0.7 | 0.47 | 0.5 | 0.6 | 0.58 | 0.65 | 0.53 | 0.61 | 0.55 | < 1 | < 1 | 0.5 | < 5 | 0.46 | 0.5125 | [6 empty] | 1
| 1 | 2 | 0.56 | 0.5 | 8.452 | 0.7 | 0.48 | 0.5 | 0.6 | 0.63 | 0.74 | 0.57 | 0.6 | 0.51 | < 1 | < 1 | 0.5 | < 5 | 0.47 | 0.5117 | [6 empty] | 37
| 1 | 3 | 0.58 | 0.5 | 8.568 | 0.7 | 0.51 | 0.6 | 0.5 | 0.61 | 0.73 | 0.54 | 0.6 | 0.52 | < 1 | < 1 | 0.5 | < 5 | 0.48 | 0.5478 | [6 empty] | 16
| 1 | 4 | 0.58 | 0.5 | 6.296 | 0.7 | 0.49 | 0.5 | 0.6 | 0.59 | 0.68 | 0.57 | 0.62 | 0.52 | < 1 | < 1 | 0.5 | < 5 | 0.43 | 0.5817 | [6 empty] | 0.55586
| 1 | 5 | 0.56 | 0.4 | 6.144 | 0.7 | 0.5 | 0.5 | 0.5 | 0.61 | 0.65 | 0.53 | 0.59 | 0.51 | < 1 | < 1 | 0.5 | < 5 | 0.44 | 0.456 | [6 empty] | 116
| 1 | 6 | 0.58 | 0.6 | 8.188 | 0.7 | 0.59 | 0.6 | 0.6 | 0.61 | 0.66 | 0.53 | 0.58 | 0.51 | < 1 | < 1 | 0.5 | < 5 | 0.41 | 0.5384
| Mean |  | 0.57 | 0.5 | 7.6313 | 0.7 | 0.50667 | 0.53333 | 0.56667 | 0.605 | 0.685 | 0.545 | 0.6 | 0.52 |  |  | 0.5 |  | 0.44833 | 0.52468
| Median |  | 0.57 | 0.5 | 8.164 | 0.7 | 0.495 | 0.5 | 0.6 | 0.61 | 0.67 | 0.535 | 0.6 | 0.515 |  |  | 0.5 |  | 0.45 | 0.52545
| Std Dev. |  | 0.010954 | 0.063246 | 1.1059 | 1.2162e-16 | 0.043205 | 0.05164 | 0.05164 | 0.017607 | 0.040373 | 0.019748 | 0.014142 | 0.015492 |  |  | 0 |  | 0.026394 | 0.042445
| Rel.Std.Dev. |  | 1.922% | 12.65% | 14.49% | 1.737e-14% | 8.527% | 9.682% | 9.113% | 2.91% | 5.894% | 3.624% | 2.357% | 2.979% |  |  | 0% |  | 5.887% | 8.09%
| PDM3 |  | 2.544% | -10.05% | 1273% | 25.93% | -8.85% | -4.052% | 1.944% | 8.841% | 23.23% | -1.953% | 7.941% | -6.451% |  |  | -10.05% |  | -19.34% | -5.608%
| Z-Score (Absolute) |  | 0.53 | 0.67 | 122.1 | 2.77 | 0.56 | 0.1 | 0.47 | 1.13 | 2.51 | 0.1 | 1.05 | 0.33 | 0.67 | 0.67 | 0.67 | 33.76 | 1.56 | 0.25
| Analytical results for Y in OREAS 239b (Certified Value 8.94 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 9.02 | 7.6759 | 8.09 | 8.7833 | 10.16 | 8.01 | 9.97 | 9 | 9.15 | 9.55 | 8.84 | 9.15 | 9.7228 | 9 | 8.063 | 9.07 | 8.61 | 7.6095 | 9.6913 | 8.7964 |  |  |  |  | 1
| 1 | 2 | 8.93 | 7.5242 | 8.11 | 8.8467 | 9.73 | 8.18 | 9.87 | 9.1 | 9.2 | 8.56 | 9.25 | 9.14 | 9.4517 | 9 | 8.509 | 8.87 | 8.9 | 7.562 | 9.8348 | 8.7311 |  |  |  |  | 38
| 1 | 3 | 9.01 | 7.3822 | 8.08 | 8.6633 | 9.87 | 8.16 | 9.85 | 8.9 | 9.01 | 9.49 | 8.9 | 9.07 | 9.5053 | 9 | 8.672 | 9.26 | 8.99 | 7.4385 | 9.7621 | 8.2916 |  |  |  |  | 16
| 1 | 4 | 9.13 | 7.6126 | 8.15 | 8.7833 | 10.13 | 8.3 | 9.7 | 9.1 | 9.19 | 9.16 | 8.89 | 9.11 | 9.6297 | 9 | 8.859 | 9 | 8.76 | 7.6665 | 10.2593 | 9.0054 |  |  |  |  | 8.9394
| 1 | 5 | 9 | 7.6954 | 7.7 | 8.77 | 9.73 | 8.53 | 9.29 | 9 | 8.91 | 9.21 | 8.91 | 9.25 | 9.7621 | 9 | 8.38 | 8.33 | 8.85 | 7.6285 | 10.0441 | 8.9225 |  |  |  |  | 117
| 1 | 6 | 9.06 | 7.3932 | 7.88 | 8.7767 | 9.79 | 8.6 | 9.79 | 9.1 | 9.21 | 9.46 | 8.95 | 8.84 | 9.6855 | 9 | 8.325 | 8.49 | 9.01 | 7.6285 | 9.9001 | 9.2373
| Mean |  | 9.025 | 7.5473 | 8.0017 | 8.7706 | 9.9017 | 8.2967 | 9.745 | 9.0333 | 9.1117 | 9.2383 | 8.9567 | 9.0933 | 9.6262 | 9 | 8.468 | 8.8367 | 8.8533 | 7.5889 | 9.9153 | 8.8307
| Median |  | 9.015 | 7.5684 | 8.085 | 8.78 | 9.83 | 8.24 | 9.82 | 9.05 | 9.17 | 9.335 | 8.905 | 9.125 | 9.6576 | 9 | 8.4445 | 8.935 | 8.875 | 7.619 | 9.8674 | 8.8595
| Std Dev. |  | 0.066558 | 0.13731 | 0.1752 | 0.05942 | 0.19559 | 0.22844 | 0.24015 | 0.08165 | 0.12336 | 0.36788 | 0.14801 | 0.13779 | 0.1236 | 0 | 0.27871 | 0.3573 | 0.15055 | 0.08115 | 0.2075 | 0.31813
| Rel.Std.Dev. |  | 0.7375% | 1.819% | 2.19% | 0.6775% | 1.975% | 2.753% | 2.464% | 0.9039% | 1.354% | 3.982% | 1.653% | 1.515% | 1.284% | 0% | 3.291% | 4.043% | 1.701% | 1.069% | 2.093% | 3.603%
| PDM3 |  | 0.9575% | -15.57% | -10.49% | -1.889% | 10.76% | -7.19% | 9.012% | 1.051% | 1.927% | 3.344% | 0.1931% | 1.722% | 7.682% | 0.6779% | -5.273% | -1.149% | -0.9628% | -15.11% | 10.92% | -1.216%
| Z-Score (Absolute) |  | 0.13 | 4.13 | 2.82 | 0.6 | 2.66 | 1.97 | 2.21 | 0.16 | 0.38 | 0.75 | 0.06 | 0.33 | 1.87 | 0.06 | 1.47 | 0.41 | 0.36 | 4.01 | 2.7 | 0.43
| Analytical results for Yb in OREAS 239b (Certified Value 0.87 ppm) | [25 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 1.03 | 1.06 | 1.20 | 1.21 | 1.30 | 1.32 | [18 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | [18 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [18 empty] | 2
| 1 | 1 | 0.81865 | 0.8 | 0.721 | 0.8 | 1 | 1.0812 | [18 empty] | 1
| 1 | 2 | 0.79182 | 0.8 | 0.775 | 0.8 | 0.96 | 1.0518 | [18 empty] | 39
| 1 | 3 | 0.81016 | 0.8 | 0.761 | 0.8 | 0.95 | 1.0221 | [18 empty] | 16
| 1 | 4 | 0.78626 | 0.8 | 0.798 | 0.8 | 0.98 | 1.1399 | [18 empty] | 0.86918
| 1 | 5 | 0.82848 | 0.8 | 0.761 | 0.8 | 0.95 | 1.1103 | [18 empty] | 118
| 1 | 6 | 0.85545 | 0.8 | 0.737 | 0.8 | 0.95 | 1.0512
| Mean |  | 0.81514 | 0.8 | 0.75883 | 0.8 | 0.965 | 1.0761
| Median |  | 0.81441 | 0.8 | 0.761 | 0.8 | 0.955 | 1.0665
| Std Dev. |  | 0.025371 | 1.2162e-16 | 0.02725 | 1.2162e-16 | 0.020736 | 0.043342
| Rel.Std.Dev. |  | 3.112% | 1.52e-14% | 3.591% | 1.52e-14% | 2.149% | 4.028%
| PDM3 |  | -6.217% | -7.959% | -12.7% | -7.959% | 11.02% | 23.81%
| Z-Score (Absolute) |  | 0.18 | 0.18 | 1.17 | 0.18 | 3.77 | 6.43
| Analytical results for Zn in OREAS 239b (Certified Value 86 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 1
| 1 | 1 | 79.1 | 90.09 | 86.1 | 81.275 | 88 | 90 | 90 | 87 | 85 | 85 | 87 | 86 | 85.1 | 88.5266 | 86 | 87.69 | 78 | 91.79 | 83.3 | 75 | 100.3632 | 92 | 82.595 | 84.0772 | 1
| 1 | 2 | 80.3 | 88.131 | 85.7 | 83.86 | 84 | 86 | 91 | 87 | 85 | 83 | 86 | 85 | 86 | 88.4282 | 88 | 81.43 | 76 | 92.34 | 85.2 | 75 | 99.6576 | 93 | 77.2894 | 84.891 | 40
| 1 | 3 | 80.5 | 89.936 | 85.6 | 81.205 | 83 | 85 | 90 | 84 | 85 | 85 | 86 | 84 | 85.8 | 86.6027 | 87 | 82.67 | 77 | 91.21 | 86.5 | 76 | 98.6496 | 93 | 84.2638 | 84.0285 | 16
| 1 | 4 | 80.3 | 87.362 | 89.2 | 84.37 | 88 | 86 | 91 | 89 | 85 | 81 | 87 | 86 | 84.5 | 86.3268 | 89 | 77.19 | 81 | 90.53 | 85.2 | 75 | 101.1696 | 95 | 77.3464 | 85.9043 | 85.7743
| 1 | 5 | 79 | 88.194 | 84.9 | 81.25 | 86 | 87 | 90 | 85 | 84 | 82 | 88 | 86 | 83.7 | 86.7129 | 88 | 82.48 | 73 | 89.38 | 82.3 | 77 | 99.5344 | 94 | 77.4285 | 86.5227 | 119
| 1 | 6 | 79.6 | 90.211 | 87.4 | 82.7 | 85 | 87 | 93 | 85 | 85 | 83 | 87 | 86 | 84.5 | 86.9583 | 89 | 79.93 | 74 | 88.85 | 84.7 | 77 | 98.9856 | 95 | 76.6276 | 85.7185
| Mean |  | 79.8 | 88.9873 | 86.4833 | 82.4433 | 85.6667 | 86.8333 | 90.8333 | 86.1667 | 84.8333 | 83.1667 | 86.8333 | 85.5 | 84.9333 | 87.2593 | 87.8333 | 81.8983 | 76.5 | 90.6833 | 84.5333 | 75.8333 | 99.7267 | 93.6667 | 79.2584 | 85.1904
| Median |  | 79.95 | 89.065 | 85.9 | 81.9875 | 85.5 | 86.5 | 90.5 | 86 | 85 | 83 | 87 | 86 | 84.8 | 86.8356 | 88 | 81.955 | 76.5 | 90.87 | 84.95 | 75.5 | 99.596 | 93.5 | 77.3875 | 85.3047
| Std Dev. |  | 0.65727 | 1.2343 | 1.5664 | 1.4218 | 2.0656 | 1.7224 | 1.169 | 1.8348 | 0.40825 | 1.6021 | 0.75277 | 0.83666 | 0.87331 | 0.96564 | 1.169 | 3.4854 | 2.881 | 1.366 | 1.5029 | 0.98319 | 0.92068 | 1.2111 | 3.286 | 1.0239
| Rel.Std.Dev. |  | 0.8236% | 1.387% | 1.811% | 1.725% | 2.411% | 1.984% | 1.287% | 2.129% | 0.4812% | 1.926% | 0.8669% | 0.9785% | 1.028% | 1.107% | 1.331% | 4.256% | 3.766% | 1.506% | 1.778% | 1.297% | 0.9232% | 1.293% | 4.146% | 1.202%
| PDM3 |  | -6.965% | 3.746% | 0.8267% | -3.883% | -0.1254% | 1.235% | 5.898% | 0.4575% | -1.097% | -3.04% | 1.235% | -0.3197% | -0.9804% | 1.731% | 2.401% | -4.519% | -10.81% | 5.723% | -1.447% | -11.59% | 16.27% | 9.201% | -7.596% | -0.6807%
| Z-Score (Absolute) |  | 1.67 | 0.98 | 0.26 | 0.91 | 0.02 | 0.36 | 1.52 | 0.17 | 0.22 | 0.7 | 0.36 | 0.02 | 0.19 | 0.48 | 0.65 | 1.06 | 2.62 | 1.47 | 0.3 | 2.82 | 4.09 | 2.34 | 1.83 | 0.11
| Analytical results for Zr in OREAS 239b (Certified Value 21.9 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 20.8 | 32.76 | 12.6 | 23.6506 | 22.5 | 11.5 | 28.4 | 20.3 | 21.7 | 22.6 | 22.4 | 21.9 | 24.3439 | 14 | 17.6 | 22.8 | 23.1 | 15.93 | 94.6292 | 19.8121 |  |  |  |  | 1
| 1 | 2 | 21.51 | 33.2 | 7.3 | 23.7535 | 21.8 | 13.1 | 26.9 | 20.5 | 22.6 | 20.3 | 23 | 22.6 | 24.5226 | 16 | 17.5 | 23.9 | 23.3 | 15.67 | 94.4934 | 19.5389 |  |  |  |  | 41
| 1 | 3 | 21.43 | 32.72 | 7.4 | 23.3799 | 21.6 | 12.7 | 27.8 | 20.1 | 21.2 | 22.1 | 22.4 | 21.5 | 24.8016 | 15 | 17.9 | 22.6 | 23.2 | 15.81 | 93.6803 | 18.0074 |  |  |  |  | 16
| 1 | 4 | 21.47 | 32 | 7.1 | 23.6168 | 22.4 | 11.9 | 26.3 | 20.6 | 21.7 | 21.4 | 22 | 22.3 | 24.3059 | 14 | 18.4 | 23 | 23.5 | 16.26 | 94.626 | 21.6831 |  |  |  |  | 21.8632
| 1 | 5 | 21.2 | 32.17 | 6.6 | 23.5947 | 21.5 | 12.9 | 28.2 | 20.5 | 21.3 | 21.3 | 22.1 | 22.2 | 24.3164 | 14 | 18.8 | 23.9 | 24.1 | 16.06 | 94.9897 | 20.8047 |  |  |  |  | 120
| 1 | 6 | 20.86 | 31.98 | 11.1 | 23.6552 | 21.4 | 12.2 | 27.6 | 20.6 | 22 | 21.6 | 22.5 | 21.4 | 24.6036 | 15 | 18.9 | 24.5 | 22.7 | 16.02 | 94.0658 | 21.4798
| Mean |  | 21.2117 | 32.4717 | 8.6833 | 23.6084 | 21.8667 | 12.3833 | 27.5333 | 20.4333 | 21.75 | 21.55 | 22.4 | 21.9833 | 24.4823 | 14.6667 | 18.1833 | 23.45 | 23.3167 | 15.9583 | 94.4141 | 20.221
| Median |  | 21.315 | 32.445 | 7.35 | 23.6337 | 21.7 | 12.45 | 27.7 | 20.5 | 21.7 | 21.5 | 22.4 | 22.05 | 24.4332 | 14.5 | 18.15 | 23.45 | 23.25 | 15.975 | 94.5597 | 20.3084
| Std Dev. |  | 0.31518 | 0.49608 | 2.5135 | 0.1245 | 0.47188 | 0.62102 | 0.79917 | 0.19664 | 0.50892 | 0.78166 | 0.35214 | 0.47081 | 0.19803 | 0.8165 | 0.6047 | 0.75565 | 0.46655 | 0.20547 | 0.46659 | 1.3853
| Rel.Std.Dev. |  | 1.486% | 1.528% | 28.95% | 0.5274% | 2.158% | 5.015% | 2.903% | 0.9623% | 2.34% | 3.627% | 1.572% | 2.142% | 0.8089% | 5.567% | 3.326% | 3.222% | 2.001% | 1.288% | 0.4942% | 6.851%
| PDM3 |  | -2.98% | 48.52% | -60.28% | 7.982% | 0.01572% | -43.36% | 25.93% | -6.54% | -0.5179% | -1.433% | 2.455% | 0.5493% | 11.98% | -32.92% | -16.83% | 7.258% | 6.648% | -27.01% | 331.8% | -7.511%
| Z-Score (Absolute) |  | 0.23 | 4.18 | 5.14 | 0.71 | 0.02 | 3.69 | 2.24 | 0.54 | 0.02 | 0.1 | 0.23 | 0.07 | 1.05 | 2.8 | 1.42 | 0.64 | 0.59 | 2.29 | 28.45 | 0.62
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1124: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1142: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1160: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1178: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1196: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1214: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1232: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 239b (Indicative Value 14.43 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.42 | 1
| 1 | 2 | 14.44 | 16
| Mean |  | 14.43 | 16
| Median |  | 14.43 | 14.43
| Std Dev. |  | 0.014142 | 22
| Rel.Std.Dev. |  | 0.09801%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 239b (Indicative Value 1.21 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.21 | 1
| 1 | 2 | 1.21 | 17
| Mean |  | 1.21 | 16
| Median |  | 1.21 | 1.21
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 239b (Indicative Value 5.83 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.85 | 1
| 1 | 2 | 5.81 | 18
| Mean |  | 5.83 | 16
| Median |  | 5.83 | 5.83
| Std Dev. |  | 0.028284 | 24
| Rel.Std.Dev. |  | 0.4852%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 239b (Indicative Value 3.3 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.3 | 1
| 1 | 2 | 3.3 | 19
| Mean |  | 3.3 | 16
| Median |  | 3.3 | 3.3
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 239b (Indicative Value 2.88 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.88 | 1
| 1 | 2 | 2.88 | 16
| Mean |  | 2.88 | 16
| Median |  | 2.88 | 2.88
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 239b (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 17
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 239b (Indicative Value 0.92 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.93 | 1
| 1 | 2 | 0.91 | 18
| Mean |  | 0.92 | 16
| Median |  | 0.92 | 0.92
| Std Dev. |  | 0.014142 | 24
| Rel.Std.Dev. |  | 1.537%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 239b (Indicative Value 0.151 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.152 | 1
| 1 | 2 | 0.15 | 19
| Mean |  | 0.151 | 16
| Median |  | 0.151 | 0.151
| Std Dev. |  | 0.0014142 | 25
| Rel.Std.Dev. |  | 0.9366%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 239b (Indicative Value 67.8 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.77 | 1
| 1 | 2 | 67.82 | 16
| Mean |  | 67.795 | 16
| Median |  | 67.795 | 67.795
| Std Dev. |  | 0.035355 | 22
| Rel.Std.Dev. |  | 0.05215%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 239b (Indicative Value 0.513 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.517 | 1
| 1 | 2 | 0.509 | 17
| Mean |  | 0.513 | 16
| Median |  | 0.513 | 0.513
| Std Dev. |  | 0.0056569 | 23
| Rel.Std.Dev. |  | 1.103%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 239b (Indicative Value 0.81 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.81 | 1
| 1 | 2 | 0.81 | 18
| Mean |  | 0.81 | 16
| Median |  | 0.81 | 0.81
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 239b (Indicative Value 2.5 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.49 | 1
| 1 | 2 | 2.5 | 21
| Mean |  | 2.495 | 16
| Median |  | 2.495 | 2.495
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 0.2834%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 239b (Indicative Value 0.11 wt.%) |  |  |  | Indicative
C | Round | Replicate | BV Geo |  | 1
| No | No | Perth |  | wt.%
|  |  | IRC |  | 3
3
| 1 | 1 | 0.12 |  | 1
| 1 | 2 | 0.1 |  | 23
| Mean |  | 0.11 |  | 16
| Median |  | 0.11 |  | 0.11
| Std Dev. |  | 0.014142 |  | 29
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 239b (Indicative Value 0.196 wt.%) |  |  |  | Indicative
S | Round | Replicate | Inspectorate | BV Geo | 1
| No | No | Gandhidham | Perth | wt.%
|  |  | IRC | IRC | 3
3
| 1 | 1 | 0.2 | 0.19 | 1
| 1 | 2 | 0.2 | 0.18 | 23
| 1 | 3 | 0.21 |  | 16
| 1 | 4 | 0.21 |  | 0.19583
| 1 | 5 | 0.21 |  | 29
| 1 | 6 | 0.21
| Mean |  | 0.20667 | 0.185
| Median |  | 0.21 | 0.185
| Std Dev. |  | 0.005164 | 0.0070711
| Rel.Std.Dev. |  | 2.499% | 3.822%
| PDM3 |  | 5.532% | -5.532%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 239b (Indicative Value 0.35 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.4 | 25
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 239b (Indicative Value 848 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 845 | 1
| 1 | 2 | 850 | 26
| Mean |  | 847.5 | 16
| Median |  | 847.5 | 847.5
| Std Dev. |  | 3.5355 | 32
| Rel.Std.Dev. |  | 0.4172%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 239b (Indicative Value 700 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 695 | 1
| 1 | 2 | 705 | 27
| Mean |  | 700 | 16
| Median |  | 700 | 700
| Std Dev. |  | 7.0711 | 33
| Rel.Std.Dev. |  | 1.01%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 239b (Indicative Value 2.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.6 | 28
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 239b (Indicative Value 0.43 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.42 | 1
| 1 | 2 | 0.44 | 29
| Mean |  | 0.43 | 16
| Median |  | 0.43 | 0.43
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 239b (Indicative Value 0.3 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.3 | 30
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 239b (Indicative Value 80 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 80.3 | 1
| 1 | 2 | 80.4 | 31
| Mean |  | 80.35 | 16
| Median |  | 80.35 | 80.35
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.088%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 239b (Indicative Value 18.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.9 | 1
| 1 | 2 | 18.3 | 32
| Mean |  | 18.1 | 16
| Median |  | 18.1 | 18.1
| Std Dev. |  | 0.28284 | 38
| Rel.Std.Dev. |  | 1.563%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 239b (Indicative Value 151 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 150 | 1
| 1 | 2 | 151 | 33
| Mean |  | 150.5 | 16
| Median |  | 150.5 | 150.5
| Std Dev. |  | 0.70711 | 39
| Rel.Std.Dev. |  | 0.4698%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 239b (Indicative Value 9.45 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.45 | 1
| 1 | 2 | 9.44 | 34
| Mean |  | 9.445 | 16
| Median |  | 9.445 | 9.445
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.07487%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 239b (Indicative Value 33 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 34 | 1
| 1 | 2 | 32 | 35
| Mean |  | 33 | 16
| Median |  | 33 | 33
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 239b (Indicative Value 5.78 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.77 | 1
| 1 | 2 | 5.78 | 9
| Mean |  | 5.775 | 16
| Median |  | 5.775 | 5.775
| Std Dev. |  | 0.0070711 | 42
| Rel.Std.Dev. |  | 0.1224%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 239b (Indicative Value 3.39 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.42 | 1
| 1 | 2 | 3.36 | 10
| Mean |  | 3.39 | 16
| Median |  | 3.39 | 3.39
| Std Dev. |  | 0.042426 | 43
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 239b (Indicative Value 1.36 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.36 | 1
| 1 | 2 | 1.35 | 11
| Mean |  | 1.355 | 16
| Median |  | 1.355 | 1.355
| Std Dev. |  | 0.0070711 | 44
| Rel.Std.Dev. |  | 0.5219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 239b (Indicative Value 19.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.2 | 1
| 1 | 2 | 19.4 | 39
| Mean |  | 19.3 | 16
| Median |  | 19.3 | 19.3
| Std Dev. |  | 0.14142 | 45
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 239b (Indicative Value 6.07 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.1 | 1
| 1 | 2 | 6.03 | 12
| Mean |  | 6.065 | 16
| Median |  | 6.065 | 6.065
| Std Dev. |  | 0.049497 | 46
| Rel.Std.Dev. |  | 0.8161%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 239b (Indicative Value 1.65 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.55 | 1
| 1 | 2 | 1.75 | 5
| Mean |  | 1.65 | 16
| Median |  | 1.65 | 1.65
| Std Dev. |  | 0.14142 | 47
| Rel.Std.Dev. |  | 8.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 239b (Indicative Value 6.62 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.65 | 1
| 1 | 2 | 6.59 | 25
| Mean |  | 6.62 | 16
| Median |  | 6.62 | 6.62
| Std Dev. |  | 0.042426 | 31
| Rel.Std.Dev. |  | 0.6409%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 239b (Indicative Value 1.18 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.19 | 1
| 1 | 2 | 1.17 | 14
| Mean |  | 1.18 | 16
| Median |  | 1.18 | 1.18
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 1.198%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 239b (Indicative Value < 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.05 | 1
| 1 | 2 | < 0.05 | 27
| Mean |  |  | 16
| Median |  |  | < 0.05
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 239b (Indicative Value 40.6 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.6 | 1
| 1 | 2 | 40.6 | 28
| Mean |  | 40.6 | 16
| Median |  | 40.6 | 40.6
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 239b (Indicative Value 0.48 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.48 | 1
| 1 | 2 | 0.47 | 29
| Mean |  | 0.475 | 16
| Median |  | 0.475 | 0.475
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 1.489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 239b (Indicative Value 0.043 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0424 | 1
| 1 | 2 | 0.043 | 30
| Mean |  | 0.0427 | 16
| Median |  | 0.0427 | 0.0427
| Std Dev. |  | 0.00042426 | 36
| Rel.Std.Dev. |  | 0.9936%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 239b (Indicative Value 1.4 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.4 | 31
| Mean |  | 1.4 | 16
| Median |  | 1.4 | 1.4
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 239b (Indicative Value 15.9 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.8 | 1
| 1 | 2 | 16 | 32
| Mean |  | 15.9 | 16
| Median |  | 15.9 | 15.9
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 0.8894%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 239b (Indicative Value 36.6 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.8 | 1
| 1 | 2 | 36.3 | 9
| Mean |  | 36.55 | 16
| Median |  | 36.55 | 36.55
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 0.9673%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 239b (Indicative Value 69 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 68 | 34
| Mean |  | 69 | 16
| Median |  | 69 | 69
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 239b (Indicative Value 25.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25 | 1
| 1 | 2 | 26 | 35
| Mean |  | 25.5 | 16
| Median |  | 25.5 | 25.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 2.773%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 239b (Indicative Value 9.9 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.79 | 1
| 1 | 2 | 10 | 11
| Mean |  | 9.895 | 16
| Median |  | 9.895 | 9.895
| Std Dev. |  | 0.14849 | 42
| Rel.Std.Dev. |  | 1.501%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 239b (Indicative Value 162 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 161 | 1
| 1 | 2 | 162 | 37
| Mean |  | 161.5 | 16
| Median |  | 161.5 | 161.5
| Std Dev. |  | 0.70711 | 43
| Rel.Std.Dev. |  | 0.4378%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 239b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 239b (Indicative Value 779 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 786 | 1
| 1 | 2 | 772 | 39
| Mean |  | 779 | 16
| Median |  | 779 | 779
| Std Dev. |  | 9.8995 | 45
| Rel.Std.Dev. |  | 1.271%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 239b (Indicative Value 15.4 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.1 | 1
| 1 | 2 | 15.7 | 40
| Mean |  | 15.4 | 16
| Median |  | 15.4 | 15.4
| Std Dev. |  | 0.42426 | 46
| Rel.Std.Dev. |  | 2.755%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 239b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 239b (Indicative Value 7.32 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.18 | 1
| 1 | 2 | 7.45 | 10
| Mean |  | 7.315 | 16
| Median |  | 7.315 | 7.315
| Std Dev. |  | 0.19092 | 31
| Rel.Std.Dev. |  | 2.61%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 239b (Indicative Value 3.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 3.8 | 26
| Mean |  | 3.8 | 16
| Median |  | 3.8 | 3.8
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 239b (Indicative Value 120 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 119 | 1
| 1 | 2 | 121 | 27
| Mean |  | 120 | 16
| Median |  | 120 | 120
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 1.179%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 239b (Indicative Value 1.24 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.19 | 1
| 1 | 2 | 1.29 | 28
| Mean |  | 1.24 | 16
| Median |  | 1.24 | 1.24
| Std Dev. |  | 0.070711 | 34
| Rel.Std.Dev. |  | 5.702%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 239b (Indicative Value 0.97 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.97 | 1
| 1 | 2 | 0.96 | 11
| Mean |  | 0.965 | 16
| Median |  | 0.965 | 0.965
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 239b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 5
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 239b (Indicative Value 15.2 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.2 | 1
| 1 | 2 | 15.2 | 31
| Mean |  | 15.2 | 16
| Median |  | 15.2 | 15.2
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 239b (Indicative Value 0.496 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.493 | 1
| 1 | 2 | 0.499 | 32
| Mean |  | 0.496 | 16
| Median |  | 0.496 | 0.496
| Std Dev. |  | 0.0042426 | 38
| Rel.Std.Dev. |  | 0.8554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 239b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 33
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 239b (Indicative Value 0.51 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.52 | 1
| 1 | 2 | 0.5 | 13
| Mean |  | 0.51 | 16
| Median |  | 0.51 | 0.51
| Std Dev. |  | 0.014142 | 40
| Rel.Std.Dev. |  | 2.773%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 239b (Indicative Value 3.16 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.11 | 1
| 1 | 2 | 3.2 | 35
| Mean |  | 3.155 | 16
| Median |  | 3.155 | 3.155
| Std Dev. |  | 0.06364 | 41
| Rel.Std.Dev. |  | 2.017%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 239b (Indicative Value 113 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 111 | 1
| 1 | 2 | 114 | 36
| Mean |  | 112.5 | 16
| Median |  | 112.5 | 112.5
| Std Dev. |  | 2.1213 | 42
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 239b (Indicative Value 4 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 4 | 37
| Mean |  | 4 | 16
| Median |  | 4 | 4
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 239b (Indicative Value 30.8 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.9 | 1
| 1 | 2 | 30.6 | 38
| Mean |  | 30.75 | 16
| Median |  | 30.75 | 30.75
| Std Dev. |  | 0.21213 | 44
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 239b (Indicative Value 3.22 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.28 | 1
| 1 | 2 | 3.16 | 39
| Mean |  | 3.22 | 16
| Median |  | 3.22 | 3.22
| Std Dev. |  | 0.084853 | 45
| Rel.Std.Dev. |  | 2.635%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 239b (Indicative Value 100 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 95 | 1
| 1 | 2 | 105 | 40
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 7.0711 | 46
| Rel.Std.Dev. |  | 7.071%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 239b (Indicative Value 235 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 233 | 1
| 1 | 2 | 236 | 41
| Mean |  | 234.5 | 16
| Median |  | 234.5 | 234.5
| Std Dev. |  | 2.1213 | 47
| Rel.Std.Dev. |  | 0.9046%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.