## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 712
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Oxidising Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 10.7091 | 10.5991 | 10.8191 | 10.6609 | 10.7573
As, Arsenic (ppm) | < 100 | IND | IND | IND | IND
Bi, Bismuth (ppm) | < 100 | IND | IND | IND | IND
CaO, Calcium oxide (wt.%) | 4.312 | 4.2614 | 4.3625 | 4.2922 | 4.3317
CeO2, Cerium(IV) oxide (ppm) | 573.2492 | 445.6631 | 700.8354 | IND | IND
Co, Cobalt (ppm) | < 100 | IND | IND | IND | IND
Cr, Chromium (ppm) | 146.1003 | 124.3861 | 167.8146 | 136.2629 | 155.9377
Cu, Copper (ppm) | < 50 | IND | IND | IND | IND
Fe, Iron (wt.%) | 14.5188 | 14.4023 | 14.6353 | 14.4521 | 14.5855
HfO2, Hafnium dioxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 0.75315 | 0.74087 | 0.76543 | 0.74464 | 0.76167
La2O3, Lanthanum(III) oxide (ppm) | 341.3451 | 225.4714 | 457.2188 | IND | IND
MgO, Magnesium oxide (wt.%) | 0.97603 | 0.95735 | 0.99472 | 0.96227 | 0.9898
Mn, Manganese (wt.%) | 1.2923 | 1.2728 | 1.3119 | 1.2815 | 1.3032
Na2O, Sodium oxide (wt.%) | 0.8312 | 0.79898 | 0.86342 | 0.81306 | 0.84935
Nb, Niobium (ppm) | 115.5005 | 90.5866 | 140.4144 | IND | IND
Ni, Nickel (ppm) | < 50 | IND | IND | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.18206 | 0.1757 | 0.18842 | 0.17802 | 0.1861
SiO2, Silicon dioxide (wt.%) | 47.9683 | 47.5538 | 48.3828 | 47.8403 | 48.0962
Sr, Strontium (ppm) | 309.5523 | 214.8886 | 404.2161 | 298.567 | 320.5376
TiO2, Titanium dioxide (wt.%) | 12.0381 | 11.9467 | 12.1296 | 11.9924 | 12.0838
Y2O3, Yttrium(III) oxide (ppm) | 175.4976 | 145.414 | 205.5812 | IND | IND
Zn, Zinc (ppm) | 70.7401 | 56.7775 | 84.7028 | IND | IND
ZrO2, Zirconium dioxide (wt.%) | 0.15328 | 0.14119 | 0.16537 | 0.15038 | 0.15618
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | -0.61608 | -0.69001 | -0.54215 | -0.64617 | -0.58599
4-Acid Digestion
Al, Aluminium (wt.%) | 5.5361 | 5.3648 | 5.7074 | 5.4233 | 5.6489
Ba, Barium (ppm) | 164.8815 | 158.1336 | 171.6295 | 161.5882 | 168.1749
Be, Beryllium (ppm) | 1.0188 | 0.92225 | 1.1153 | 0.966 | 1.0715
Bi, Bismuth (ppm) | 0.24636 | 0.22424 | 0.26849 | 0.22293 | 0.2698
Ca, Calcium (wt.%) | 3.0485 | 2.9672 | 3.1297 | 3.0005 | 3.0965
Cd, Cadmium (ppm) | 0.29368 | 0.26695 | 0.32042 | 0.27362 | 0.31374
Ce, Cerium (ppm) | 468.004 | 443.8026 | 492.2055 | 457.895 | 478.1131
Co, Cobalt (ppm) | 11.8497 | 11.1001 | 12.5993 | 11.5508 | 12.1486
Cr, Chromium (ppm) | 132.2942 | 124.8595 | 139.7289 | 128.7592 | 135.8293
Cs, Caesium (ppm) | 1.5176 | 1.428 | 1.6071 | 1.4597 | 1.5754
Dy, Dysprosium (ppm) | 17.9696 | 16.9433 | 18.9959 | 17.4156 | 18.5236
Er, Erbium (ppm) | 17.1517 | 15.9394 | 18.3639 | 16.5054 | 17.798
Eu, Europium (ppm) | 3.1671 | 2.9648 | 3.3693 | 3.0393 | 3.2948
Fe, Iron (wt.%) | 14.2755 | 13.8372 | 14.7139 | 14.0525 | 14.4986
Ga, Gallium (ppm) | 12.1442 | 11.0432 | 13.2453 | 11.7722 | 12.5162
Gd, Gadolinium (ppm) | 16.8963 | 15.9868 | 17.8058 | 16.5221 | 17.2705
Ho, Holmium (ppm) | 4.651 | 4.3738 | 4.9283 | 4.476 | 4.8261
In, Indium (ppm) | 0.093255 | 0.084747 | 0.10176 | 0.086813 | 0.099697
K, Potassium (wt.%) | 0.62437 | 0.60178 | 0.64696 | 0.61158 | 0.63716
La, Lanthanum (ppm) | 249.5199 | 237.6533 | 261.3866 | 243.5174 | 255.5225
Li, Lithium (ppm) | 21.1738 | 19.81 | 22.5376 | 20.3639 | 21.9837
Lu, Lutetium (ppm) | 3.3668 | 3.107 | 3.6265 | 3.2219 | 3.5117
Mg, Magnesium (wt.%) | 0.55585 | 0.52904 | 0.58266 | 0.54528 | 0.56642
Mn, Manganese (wt.%) | 1.2442 | 1.1916 | 1.2967 | 1.2259 | 1.2624
Mo, Molybdenum (ppm) | 3.2223 | 2.832 | 3.6127 | 3.1006 | 3.3441
Na, Sodium (wt.%) | 0.61059 | 0.59222 | 0.62897 | 0.59929 | 0.62189
Nb, Niobium (ppm) | 105.1542 | 98.6275 | 111.6808 | 100.1279 | 110.1804
Nd, Neodymium (ppm) | 158.1134 | 149.3734 | 166.8535 | 155.0217 | 161.2052
Ni, Nickel (ppm) | 13.4217 | 12.2859 | 14.5575 | 12.7939 | 14.0495
P, Phosphorus (wt.%) | 0.075706 | 0.071376 | 0.080036 | 0.073101 | 0.078312
Pb, Lead (ppm) | 10.195 | 9.2913 | 11.0987 | 9.6606 | 10.7294
Pr, Praseodymium (ppm) | 47.675 | 44.8413 | 50.5088 | 46.4813 | 48.8688
Rb, Rubidium (ppm) | 32.4438 | 30.3869 | 34.5008 | 31.0778 | 33.8098
S, Sulphur (wt.%) | < 0.01 | IND | IND | IND | IND
Sb, Antimony (ppm) | 0.51973 | 0.44717 | 0.59228 | 0.47214 | 0.56731
Sc, Scandium (ppm) | 37.0017 | 34.8056 | 39.1978 | 35.8574 | 38.1459
Sm, Samarium (ppm) | 23.4807 | 22.1312 | 24.8301 | 22.7379 | 24.2234
Sn, Tin (ppm) | 3.9475 | 3.6248 | 4.2702 | 3.757 | 4.1379
Sr, Strontium (ppm) | 361.7672 | 346.8849 | 376.6495 | 353.4885 | 370.046
Ta, Tantalum (ppm) | 10.2398 | 9.6033 | 10.8763 | 9.8558 | 10.6238
Tb, Terbium (ppm) | 2.4885 | 2.3265 | 2.6505 | 2.4013 | 2.5757
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 72.1164 | 66.9003 | 77.3325 | 69.3364 | 74.8964
Ti, Titanium (wt.%) | 7.0642 | 6.8226 | 7.3058 | 6.959 | 7.1693
Tl, Thallium (ppm) | 0.18591 | 0.16887 | 0.20294 | 0.16574 | 0.20607
Tm, Thulium (ppm) | 2.9987 | 2.7458 | 3.2515 | 2.8763 | 3.1211
U, Uranium (ppm) | 7.659 | 7.1572 | 8.1608 | 7.3539 | 7.9641
V, Vanadium (ppm) | 110.8166 | 103.9703 | 117.663 | 107.7477 | 113.8855
W, Tungsten (ppm) | 3.366 | 3.0719 | 3.6601 | 3.231 | 3.5011
Y, Yttrium (ppm) | 127.9255 | 121.7082 | 134.1429 | 125.0096 | 130.8414
Yb, Ytterbium (ppm) | 21.8428 | 20.5063 | 23.1794 | 21.1439 | 22.5417
Zn, Zinc (ppm) | 70.1359 | 66.5827 | 73.6891 | 66.2747 | 73.9971
Gas / Liquid Pycnometry
SG, Specific Gravity (Unity) | 3.3458 | 3.2689 | 3.4228 | 3.3083 | 3.3833
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 712
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Oxidising Fusion XRF
BaO | ppm | 245.3053 | Ho | ppm | < 90 | Sm | ppm | < 170
Cd | ppm | < 100 | In | ppm | < 100 | Sn | ppm | < 50
Cl | ppm | 63.3333 | Lu | ppm | < 90 | Ta | ppm | 10.9089
Cs | ppm | < 90 | Mo | ppm | < 50 | Tb | ppm | < 90
Dy | ppm | 87.1308 | Nd | ppm | 158.4677 | Te | ppm | < 100
Er | ppm | 80.1632 | Pb | ppm | < 50 | Th | ppm | 70.1783
Eu | ppm | 86.3632 | Pr | ppm | 81.3883 | Tl | ppm | < 100
F | ppm | < 5000 | Rb | ppm | 61.6385 | Tm | ppm | < 90
Ga | ppm | < 100 | S | wt.% | 0.0033664 | U | ppm | 15.8918
Gd | ppm | < 90 | Sb | ppm | < 50 | V | ppm | 135.2788
Ge | ppm | < 100 | Sc | ppm | < 100 | W | ppm | 10.3509
Hg | ppm | < 100 | Se | ppm | < 100 | Yb | ppm | < 90
Thermogravimetry
H2O- | wt.% | 0.16385
4-Acid Digestion
Ag | ppm | 0.08 | Ge | ppm | 0.79423 | Re | ppm | 0.0030333
As | ppm | 2.8376 | Hf | ppm | 1.9357 | Se | ppm | 1.5411
Cu | ppm | 7.1546 | Hg | ppm | < 0.01 | Zr | wt.% | 0.005747
Borate / Peroxide Fusion ICP
Ce | ppm | 456.6414 | Lu | ppm | 3.2799 | Ti | wt.% | 7.1233
Dy | ppm | 17.5413 | Mn | wt.% | 1.285 | Tm | ppm | 2.8608
Er | ppm | 17.2387 | Nd | ppm | 150.4872 | U | ppm | 6.232
Eu | ppm | 3.2444 | Pr | ppm | 47.7147 | W | ppm | 3.3506
Gd | ppm | 18.4567 | Sm | ppm | 23.5918 | Yb | ppm | 21.4658
Ho | ppm | 4.647 | Tb | ppm | 2.6371 | Zr | wt.% | 0.11632
La | ppm | 246.7128 | Th | ppm | 63.7668
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 712
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Oxidising Fusion XRF
Al2O3, wt.% | 10.7091 | 0.14665 | 10.4158 | 11.0024 | 10.2691 | 11.149 | 1.369% | 2.739% | 4.108% | 10.1736 | 11.2445
As, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Bi, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
CaO, wt.% | 4.312 | 0.06976 | 4.1724 | 4.4515 | 4.1027 | 4.5212 | 1.618% | 3.236% | 4.853% | 4.0964 | 4.5276
CeO2, ppm | 573.2492 | 78.4356 | 416.3781 | 730.1204 | 337.9425 | 808.556 | 13.68% | 27.37% | 41.05% | 544.5868 | 601.9117
Co, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cr, ppm | 146.1003 | 27.9301 | 90.2401 | 201.9605 | 62.31 | 229.8906 | 19.12% | 38.23% | 57.35% | 138.7953 | 153.4053
Cu, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Fe, wt.% | 14.5188 | 0.14087 | 14.2371 | 14.8006 | 14.0962 | 14.9414 | 0.9703% | 1.941% | 2.911% | 13.7929 | 15.2448
HfO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 0.75315 | 0.011495 | 0.73016 | 0.77614 | 0.71867 | 0.78764 | 1.526% | 3.053% | 4.579% | 0.71549 | 0.79081
La2O3, ppm | 341.3451 | 50.9749 | 239.3953 | 443.2949 | 188.4204 | 494.2697 | 14.93% | 29.87% | 44.8% | 324.2778 | 358.4123
MgO, wt.% | 0.97603 | 0.028687 | 0.91866 | 1.0334 | 0.88997 | 1.0621 | 2.939% | 5.878% | 8.818% | 0.92723 | 1.0248
Mn, wt.% | 1.2923 | 0.024574 | 1.2432 | 1.3415 | 1.2186 | 1.3661 | 1.902% | 3.803% | 5.705% | 1.2277 | 1.3569
Na2O, wt.% | 0.8312 | 0.047657 | 0.73589 | 0.92652 | 0.68823 | 0.97417 | 5.733% | 11.47% | 17.2% | 0.78964 | 0.87276
Nb, ppm | 115.5005 | 20.3061 | 74.8882 | 156.1128 | 54.5821 | 176.4189 | 17.58% | 35.16% | 52.74% | 109.7255 | 121.2755
Ni, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
P2O5, wt.% | 0.18206 | 0.004872 | 0.17231 | 0.1918 | 0.16744 | 0.19667 | 2.676% | 5.352% | 8.028% | 0.17296 | 0.19116
SiO2, wt.% | 47.9683 | 0.48868 | 46.9909 | 48.9457 | 46.5022 | 49.4343 | 1.019% | 2.038% | 3.056% | 45.5699 | 50.3667
Sr, ppm | 309.5523 | 69.0529 | 171.4464 | 447.6582 | 102.3935 | 516.7111 | 22.31% | 44.61% | 66.92% | 294.0747 | 325.0299
TiO2, wt.% | 12.0381 | 0.11102 | 11.8161 | 12.2602 | 11.7051 | 12.3712 | 0.9222% | 1.844% | 2.767% | 11.4362 | 12.64
Y2O3, ppm | 175.4976 | 25.5009 | 124.4958 | 226.4995 | 98.9949 | 252.0004 | 14.53% | 29.06% | 43.59% | 166.7227 | 184.2725
Zn, ppm | 70.7401 | 12.1802 | 46.3798 | 95.1005 | 34.1996 | 107.2807 | 17.22% | 34.44% | 51.65% | 67.2031 | 74.2771
ZrO2, wt.% | 0.15328 | 0.0076227 | 0.13804 | 0.16853 | 0.13042 | 0.17615 | 4.973% | 9.946% | 14.92% | 0.14562 | 0.16095
Thermogravimetry
LOI1000, wt.% | -0.61608 | 0.097192 | -0.81046 | -0.42169 | -0.90765 | -0.3245 | 15.78% | 31.55% | 47.33% | -0.64688 | -0.58527
4-Acid Digestion
Al, wt.% | 5.5361 | 0.18899 | 5.1581 | 5.9141 | 4.9691 | 6.1031 | 3.414% | 6.827% | 10.24% | 5.2593 | 5.8129
Ba, ppm | 164.8815 | 8.1071 | 148.6673 | 181.0958 | 140.5601 | 189.2029 | 4.917% | 9.834% | 14.75% | 156.6375 | 173.1256
Be, ppm | 1.0188 | 0.099956 | 0.81884 | 1.2187 | 0.71888 | 1.3186 | 9.812% | 19.62% | 29.43% | 0.96782 | 1.0697
Bi, ppm | 0.24636 | 0.013678 | 0.219 | 0.27372 | 0.20533 | 0.2874 | 5.552% | 11.1% | 16.66% | 0.23404 | 0.25868
Ca, wt.% | 3.0485 | 0.060965 | 2.9265 | 3.1704 | 2.8656 | 3.2314 | 2% | 4% | 6% | 2.896 | 3.2009
Cd, ppm | 0.29368 | 0.014433 | 0.26481 | 0.32255 | 0.25038 | 0.33698 | 4.915% | 9.829% | 14.74% | 0.279 | 0.30837
Ce, ppm | 468.004 | 17.0821 | 433.8398 | 502.1683 | 416.7577 | 519.2504 | 3.65% | 7.3% | 10.95% | 444.6038 | 491.4042
Co, ppm | 11.8497 | 0.87642 | 10.0969 | 13.6025 | 9.2204 | 14.4789 | 7.396% | 14.79% | 22.19% | 11.2572 | 12.4422
Cr, ppm | 132.2942 | 10.7724 | 110.7495 | 153.8389 | 99.9771 | 164.6113 | 8.143% | 16.29% | 24.43% | 125.6795 | 138.9089
Cs, ppm | 1.5176 | 0.089894 | 1.3378 | 1.6973 | 1.2479 | 1.7872 | 5.924% | 11.85% | 17.77% | 1.4417 | 1.5934
Dy, ppm | 17.9696 | 1.0292 | 15.9111 | 20.0281 | 14.8819 | 21.0573 | 5.728% | 11.46% | 17.18% | 17.0711 | 18.8681
Er, ppm | 17.1517 | 1.2738 | 14.604 | 19.6993 | 13.3302 | 20.9732 | 7.427% | 14.85% | 22.28% | 16.2941 | 18.0093
Eu, ppm | 3.1671 | 0.11484 | 2.9374 | 3.3967 | 2.8225 | 3.5116 | 3.626% | 7.252% | 10.88% | 3.0087 | 3.3254
Fe, wt.% | 14.2755 | 0.40888 | 13.4578 | 15.0933 | 13.0489 | 15.5022 | 2.864% | 5.728% | 8.593% | 13.5618 | 14.9893
Ga, ppm | 12.1442 | 1.3017 | 9.5407 | 14.7477 | 8.239 | 16.0495 | 10.72% | 21.44% | 32.16% | 11.537 | 12.7514
Gd, ppm | 16.8963 | 0.61289 | 15.6705 | 18.1221 | 15.0577 | 18.735 | 3.627% | 7.255% | 10.88% | 16.0515 | 17.7411
Ho, ppm | 4.651 | 0.24931 | 4.1524 | 5.1496 | 3.9031 | 5.399 | 5.36% | 10.72% | 16.08% | 4.4185 | 4.8836
In, ppm | 0.093255 | 0.0082536 | 0.076748 | 0.10976 | 0.068494 | 0.11802 | 8.851% | 17.7% | 26.55% | 0.088592 | 0.097918
K, wt.% | 0.62437 | 0.019983 | 0.5844 | 0.66433 | 0.56442 | 0.68432 | 3.201% | 6.401% | 9.602% | 0.59315 | 0.65559
La, ppm | 249.5199 | 8.3584 | 232.8031 | 266.2368 | 224.4447 | 274.5952 | 3.35% | 6.7% | 10.05% | 237.044 | 261.9959
Li, ppm | 21.1738 | 1.2459 | 18.682 | 23.6657 | 17.436 | 24.9116 | 5.884% | 11.77% | 17.65% | 20.1151 | 22.2325
Lu, ppm | 3.3668 | 0.23105 | 2.9047 | 3.8289 | 2.6736 | 4.0599 | 6.863% | 13.73% | 20.59% | 3.1984 | 3.5351
Mg, wt.% | 0.55585 | 0.037974 | 0.4799 | 0.6318 | 0.44193 | 0.66977 | 6.832% | 13.66% | 20.5% | 0.52806 | 0.58364
Mn, wt.% | 1.2442 | 0.078855 | 1.0865 | 1.4019 | 1.0076 | 1.4808 | 6.338% | 12.68% | 19.01% | 1.182 | 1.3064
Mo, ppm | 3.2223 | 0.44561 | 2.3311 | 4.1136 | 1.8855 | 4.5592 | 13.83% | 27.66% | 41.49% | 3.0612 | 3.3835
Na, wt.% | 0.61059 | 0.017058 | 0.57647 | 0.64471 | 0.55942 | 0.66177 | 2.794% | 5.587% | 8.381% | 0.58006 | 0.64112
Nb, ppm | 105.1542 | 4.9735 | 95.2072 | 115.1011 | 90.2337 | 120.0746 | 4.73% | 9.459% | 14.19% | 99.8964 | 110.4119
Nd, ppm | 158.1134 | 5.6657 | 146.7821 | 169.4448 | 141.1164 | 175.1105 | 3.583% | 7.167% | 10.75% | 150.2078 | 166.0191
Ni, ppm | 13.4217 | 0.92485 | 11.572 | 15.2714 | 10.6471 | 16.1962 | 6.891% | 13.78% | 20.67% | 12.7506 | 14.0928
P, wt.% | 0.075706 | 0.0044685 | 0.066769 | 0.084643 | 0.062301 | 0.089111 | 5.902% | 11.8% | 17.71% | 0.071921 | 0.079491
Pb, ppm | 10.195 | 0.84752 | 8.5 | 11.89 | 7.6524 | 12.7375 | 8.313% | 16.63% | 24.94% | 9.6852 | 10.7047
Pr, ppm | 47.675 | 1.858 | 43.9591 | 51.391 | 42.1011 | 53.249 | 3.897% | 7.794% | 11.69% | 45.2913 | 50.0588
Rb, ppm | 32.4438 | 2.3982 | 27.6474 | 37.2402 | 25.2493 | 39.6383 | 7.392% | 14.78% | 22.18% | 30.8216 | 34.066
S, wt.% | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sb, ppm | 0.51973 | 0.052787 | 0.41415 | 0.6253 | 0.36137 | 0.67809 | 10.16% | 20.31% | 30.47% | 0.49374 | 0.54571
Sc, ppm | 37.0017 | 2.3797 | 32.2423 | 41.761 | 29.8627 | 44.1407 | 6.431% | 12.86% | 19.29% | 35.1516 | 38.8518
Sm, ppm | 23.4807 | 0.79976 | 21.8811 | 25.0802 | 21.0814 | 25.88 | 3.406% | 6.812% | 10.22% | 22.3066 | 24.6547
Sn, ppm | 3.9475 | 0.35665 | 3.2342 | 4.6608 | 2.8775 | 5.0174 | 9.035% | 18.07% | 27.1% | 3.7501 | 4.1448
Sr, ppm | 361.7672 | 18.1711 | 325.4249 | 398.1095 | 307.2538 | 416.2806 | 5.023% | 10.05% | 15.07% | 343.6789 | 379.8556
Ta, ppm | 10.2398 | 0.54978 | 9.1402 | 11.3394 | 8.5905 | 11.8891 | 5.369% | 10.74% | 16.11% | 9.7278 | 10.7518
Tb, ppm | 2.4885 | 0.090982 | 2.3066 | 2.6705 | 2.2156 | 2.7615 | 3.656% | 7.312% | 10.97% | 2.3641 | 2.6129
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 72.1164 | 5.2454 | 61.6257 | 82.6071 | 56.3803 | 87.8525 | 7.273% | 14.55% | 21.82% | 68.5106 | 75.7222
Ti, wt.% | 7.0642 | 0.2811 | 6.502 | 7.6264 | 6.2208 | 7.9075 | 3.979% | 7.959% | 11.94% | 6.7109 | 7.4174
Tl, ppm | 0.18591 | 0.014255 | 0.15739 | 0.21442 | 0.14314 | 0.22867 | 7.668% | 15.34% | 23% | 0.17661 | 0.1952
Tm, ppm | 2.9987 | 0.22676 | 2.5452 | 3.4522 | 2.3184 | 3.679 | 7.562% | 15.12% | 22.69% | 2.8487 | 3.1486
U, ppm | 7.659 | 0.4221 | 6.8148 | 8.5032 | 6.3927 | 8.9253 | 5.511% | 11.02% | 16.53% | 7.2761 | 8.042
V, ppm | 110.8166 | 8.7622 | 93.2922 | 128.3411 | 84.5299 | 137.1033 | 7.907% | 15.81% | 23.72% | 105.2758 | 116.3575
W, ppm | 3.366 | 0.26026 | 2.8455 | 3.8865 | 2.5852 | 4.1468 | 7.732% | 15.46% | 23.2% | 3.1977 | 3.5343
Y, ppm | 127.9255 | 6.9074 | 114.1107 | 141.7403 | 107.2033 | 148.6477 | 5.4% | 10.8% | 16.2% | 121.5293 | 134.3218
Yb, ppm | 21.8428 | 1.2962 | 19.2504 | 24.4353 | 17.9541 | 25.7315 | 5.934% | 11.87% | 17.8% | 20.7507 | 22.935
Zn, ppm | 70.1359 | 4.4983 | 61.1393 | 79.1325 | 56.641 | 83.6307 | 6.414% | 12.83% | 19.24% | 66.6291 | 73.6427
Gas / Liquid Pycnometry
SG, Unity | 3.3458 | 0.089511 | 3.1668 | 3.5248 | 3.0773 | 3.6144 | 2.675% | 5.351% | 8.026% | 3.1785 | 3.5131
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 712
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
GASPYC | gas expansion pycnometry
LIQPYC | liquid pycnometry
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 712
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek, Perth, WA, Australia
Lucid Laboratories Private Limited, Hyderabad, Telangana, India
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan
UIS Analytical Services, Centurion, South Africa

## Oxidising Fusion XRF
| Analytical results for Al2O3 in OREAS 712 (Certified Value 10.71 wt.%) | [21 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 10.5 | 10.59 | 10.65 | 10.8 | 10.8099 | 10.8 | 10.12 | 10.9 | 10.55 | 10.87 | 10.61 | 10.83 | 10.5458 | 10.7417 | 10.6 | 10.7607 | 10.95 | 10.7 | 10.519 |  | 1
| 1 | 2 | 10.51 | 10.61 | 10.75 | 10.7 | 10.8017 | 10.5 | 10.15 | 10.9 | 10.45 | 10.82 | 10.62 | 10.83 | 10.5565 | 10.6569 | 10.6 | 10.7434 | 11.05 | 10.7 | 10.773 |  | #N/A
| 1 | 3 | 10.5 | 10.59 | 10.7 | 10.75 | 10.8216 | 10.5 | 10.1 | 10.8 | 10.42 | 10.97 | 10.61 | 10.79 | 10.5545 | 10.6978 | 10.6 | 10.7703 | 11.05 | 10.7 | 10.575 |  | 16
| 1 | 4 | 10.51 | 10.6 | 10.6 | 10.8 | 10.8581 | 10.7 | 10.12 | 10.8 | 10.65 | 11.03 | 10.61 | 10.8 | 10.5645 | 10.7573 | 10.6 | 10.7694 | 11 | 10.65 | 10.571 |  | 10.7091
| 1 | 5 | 10.48 | 10.58 | 10.7 | 10.7 | 10.8489 | 10.7 | 10.11 | 10.9 | 10.7 | 10.9 | 10.61 | 10.82 | 10.5692 | 10.7216 | 10.6 | 10.7793 | 11 | 10.7 | 10.567 |  | 7
| 1 | 6 | 10.57 | 10.57 | 10.7 | 10.8 | 10.7682 | 10.6 | 10.2 | 10.9 | 10.73 | 10.99 | 10.61 | 10.81 | 10.5342 | 10.7174 | 10.7 | 10.7734 | 11.1 | 10.65 | 10.614
| Mean |  | 10.5117 | 10.59 | 10.6833 | 10.7583 | 10.818 | 10.6333 | 10.1333 | 10.8667 | 10.5833 | 10.93 | 10.6117 | 10.8133 | 10.5541 | 10.7155 | 10.6167 | 10.7661 | 11.025 | 10.6833 | 10.6032
| Median |  | 10.505 | 10.59 | 10.7 | 10.775 | 10.8157 | 10.65 | 10.12 | 10.9 | 10.6 | 10.935 | 10.61 | 10.815 | 10.5555 | 10.7195 | 10.6 | 10.7698 | 11.025 | 10.7 | 10.573
| Std Dev. |  | 0.030605 | 0.014142 | 0.05164 | 0.04916 | 0.032818 | 0.12111 | 0.036697 | 0.05164 | 0.13049 | 0.07975 | 0.0040825 | 0.01633 | 0.012724 | 0.03528 | 0.040825 | 0.012667 | 0.05244 | 0.02582 | 0.088522
| Rel.Std.Dev. |  | 0.2912% | 0.1335% | 0.4834% | 0.4569% | 0.3034% | 1.139% | 0.3621% | 0.4752% | 1.233% | 0.7296% | 0.03847% | 0.151% | 0.1206% | 0.3292% | 0.3845% | 0.1177% | 0.4757% | 0.2417% | 0.8349%
| PDM3 |  | -1.843% | -1.112% | -0.2405% | 0.4599% | 1.017% | -0.7074% | -5.376% | 1.471% | -1.174% | 2.063% | -0.9097% | 0.9734% | -1.447% | 0.05944% | -0.863% | 0.5322% | 2.95% | -0.2405% | -0.9891%
| Z-Score (Absolute) |  | 1.24 | 0.67 | 0 | 0.54 | 0.97 | 0.36 | 3.97 | 1.32 | 0.72 | 1.78 | 0.52 | 0.94 | 0.93 | 0.23 | 0.48 | 0.6 | 2.47 | 0 | 0.58
| Analytical results for As in OREAS 712 (Certified Value < 100 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 23 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 38 | < 80 | < 100 | 100 | [13 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 38 | < 80 | < 100 | < 100 | [13 empty] | 17
| 1 | 3 | < 100 | < 100 | < 100 | < 38 | < 80 | < 100 | 100 | [13 empty] | 16
| 1 | 4 | < 100 | 100 | < 100 | < 38 | < 80 | < 100 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 38 | < 80 | < 100 | 100 | [13 empty] | 8
| 1 | 6 | < 100 | 100 | < 100 | < 38 | < 80 | < 100 | < 100
| Mean |  |  | 100 |  |  |  |  | 100
| Median |  |  | 100 |  |  |  |  | 100
| Std Dev. |  |  | 0 |  |  |  |  | 0
| Rel.Std.Dev. |  |  | 0% |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 0.93 | 0 | 1.73 | 0.67 | 0 | 1.39
| Analytical results for BaO in OREAS 712 (Indicative Value 245 ppm) | [21 empty] | Indicative
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 200 | 400 | 200 | 100 | < 100 | 340 | 106.1467 | 400 | 200 | 263.511 | 300 | 100 | 620 | [7 empty] | 1
| 1 | 2 | 200 | 500 | 200 | 200 | < 100 | 310 | 151.06 | 340 | 100 | 254.4727 | 200 | 200 | 590 | [7 empty] | 1
| 1 | 3 | 400 | 400 | 100 | 200 | < 100 | 270 | 154.4467 | 343.275 | 300 | 286.6681 | 200 | 200 | 660 | [7 empty] | 16
| 1 | 4 | 200 | 500 | 200 | 200 | < 100 | 290 | 130.46 | 380 | 200 | 282.6486 | 300 | 200 | 630 | [7 empty] | 245.3053
| 1 | 5 | < 100 | 300 | 200 | 100 | < 100 | 320 | 147.7267 | 335 | 100 | 284.6278 | 400 | 300 | 630 | [7 empty] | 7
| 1 | 6 | 200 | 300 | 200 | 100 | < 100 | 280 | 158.4 | 325 | 200 | 296.7039 | 200 | 200 | 670
| Mean |  | 240 | 400 | 183.3333 | 150 |  | 301.6667 | 141.3733 | 353.8792 | 183.3333 | 278.1054 | 266.6667 | 200 | 633.3333
| Median |  | 200 | 400 | 200 | 150 |  | 300 | 149.3933 | 341.6375 | 200 | 283.6382 | 250 | 200 | 630
| Std Dev. |  | 89.4427 | 89.4427 | 40.8248 | 54.7723 |  | 26.3944 | 19.7739 | 29.3426 | 75.2773 | 15.8361 | 81.6497 | 63.2456 | 28.7518
| Rel.Std.Dev. |  | 37.27% | 22.36% | 22.27% | 36.51% |  | 8.75% | 13.99% | 8.292% | 41.06% | 5.694% | 30.62% | 31.62% | 4.54%
| PDM3 |  | -2.163% | 63.06% | -25.26% | -38.85% |  | 22.98% | -42.37% | 44.26% | -25.26% | 13.37% | 8.708% | -18.47% | 158.2%
| Z-Score (Absolute) |  | 0 | 1.93 | 0.25 | 0.59 | 1.59 | 0.94 | 0.67 | 1.47 | 0.25 | 0.7 | 0.59 | 0.08 | 4.28
| Analytical results for Bi in OREAS 712 (Certified Value < 100 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 23 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 100 | < 100 | < 100 | < 90 | < 90 | < 100 | < 100 | [13 empty] | 1
| 1 | 2 | 100 | < 100 | < 100 | < 90 | < 90 | < 100 | < 100 | [13 empty] | #N/A
| 1 | 3 | 100 | < 100 | < 100 | < 90 | < 90 | < 100 | < 100 | [13 empty] | 16
| 1 | 4 | 100 | < 100 | < 100 | < 90 | < 90 | < 100 | < 100 | [13 empty] | < 100
| 1 | 5 | 100 | < 100 | < 100 | < 90 | < 90 | < 100 | < 100 | [13 empty] | 9
| 1 | 6 | 100 | < 100 | < 100 | < 90 | < 90 | < 100 | < 100
| Mean |  | 100
| Median |  | 100
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 712 (Certified Value 4.31 wt.%) | [21 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 4.38 | 4.28 | 4.3 | 4.39 | 4.3597 | 4.48 | 5.67 | 4.18 | 4.36 | 4.24 | 4.32 | 4.27 | 4.2895 | 4.2586 | 4.17 | 4.2712 | 4.35 | 4.38 | 4.123 |  | 1
| 1 | 2 | 4.4 | 4.27 | 4.33 | 4.29 | 4.3542 | 4.41 | 5.6 | 4.25 | 4.35 | 4.27 | 4.32 | 4.29 | 4.2879 | 4.2714 | 4.16 | 4.2724 | 4.39 | 4.38 | 4.106 |  | #N/A
| 1 | 3 | 4.4 | 4.28 | 4.32 | 4.33 | 4.3725 | 4.34 | 5.41 | 4.2 | 4.32 | 4.23 | 4.32 | 4.29 | 4.2925 | 4.2586 | 4.16 | 4.2739 | 4.42 | 4.37 | 4.124 |  | 16
| 1 | 4 | 4.38 | 4.27 | 4.29 | 4.37 | 4.3765 | 4.4 | 5.6 | 4.22 | 4.41 | 4.27 | 4.32 | 4.28 | 4.2839 | 4.2656 | 4.17 | 4.2774 | 4.4 | 4.4 | 4.135 |  | 4.312
| 1 | 5 | 4.38 | 4.29 | 4.34 | 4.34 | 4.3699 | 4.43 | 5.54 | 4.22 | 4.41 | 4.3 | 4.32 | 4.29 | 4.2822 | 4.2752 | 4.15 | 4.2737 | 4.41 | 4.39 | 4.12 |  | 10
| 1 | 6 | 4.41 | 4.29 | 4.32 | 4.32 | 4.3446 | 4.34 | 5.6 | 4.2 | 4.36 | 4.27 | 4.33 | 4.28 | 4.2749 | 4.2656 | 4.15 | 4.2685 | 4.42 | 4.37 | 4.101
| Mean |  | 4.3917 | 4.28 | 4.3167 | 4.34 | 4.3629 | 4.4 | 5.57 | 4.2117 | 4.3683 | 4.2633 | 4.3217 | 4.2833 | 4.2852 | 4.2658 | 4.16 | 4.2729 | 4.3983 | 4.3817 | 4.1182
| Median |  | 4.39 | 4.28 | 4.32 | 4.335 | 4.3648 | 4.405 | 5.6 | 4.21 | 4.36 | 4.27 | 4.32 | 4.285 | 4.2859 | 4.2656 | 4.16 | 4.2731 | 4.405 | 4.38 | 4.1215
| Std Dev. |  | 0.013292 | 0.0089443 | 0.018619 | 0.035777 | 0.012219 | 0.054037 | 0.088544 | 0.024014 | 0.035449 | 0.025033 | 0.0040825 | 0.008165 | 0.0062712 | 0.0066886 | 0.0089443 | 0.0029909 | 0.026394 | 0.01169 | 0.012545
| Rel.Std.Dev. |  | 0.3027% | 0.209% | 0.4313% | 0.8244% | 0.2801% | 1.228% | 1.59% | 0.5702% | 0.8115% | 0.5872% | 0.09447% | 0.1906% | 0.1463% | 0.1568% | 0.215% | 0.07% | 0.6001% | 0.2668% | 0.3046%
| PDM3 |  | 1.848% | -0.7414% | 0.109% | 0.6501% | 1.182% | 2.042% | 29.18% | -2.326% | 1.307% | -1.128% | 0.2249% | -0.6641% | -0.6216% | -1.07% | -3.524% | -0.9068% | 2.003% | 1.616% | -4.494%
| Z-Score (Absolute) |  | 0.98 | 0.48 | 0 | 0.3 | 0.6 | 1.09 | 16.36 | 1.37 | 0.67 | 0.7 | 0.07 | 0.44 | 0.41 | 0.66 | 2.04 | 0.57 | 1.07 | 0.85 | 2.59
| Analytical results for Cd in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Cd | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 2
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for CeO2 in OREAS 712 (Certified Value 573 ppm) | [21 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 05 | 13 | 15 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 600 | 600 | < 90 | 550 | 451 | 630 | 500 | 600 | 600 | [11 empty] | 1
| 1 | 2 | 600 | 300 | < 90 | 640 | 396 | 730 | 500 | 600 | 600 | [11 empty] | #N/A
| 1 | 3 | 600 | 500 | < 90 | 600 | 454.2175 | 630 | 600 | 600 | 600 | [11 empty] | 16
| 1 | 4 | 500 | 600 | < 90 | 650 | 484 | 730 | 600 | 600 | 500 | [11 empty] | 573.2492
| 1 | 5 | 600 | 500 | < 90 | 630 | 440 | 730 | 500 | 600 | 500 | [11 empty] | 11
| 1 | 6 | 600 | 700 | < 90 | 610 | 445.5 | 520 | 500 | 600 | 600
| Mean |  | 583.3333 | 533.3333 |  | 613.3333 | 445.1196 | 661.6667 | 533.3333 | 600 | 566.6667
| Median |  | 600 | 550 |  | 620 | 448.25 | 680 | 500 | 600 | 600
| Std Dev. |  | 40.8248 | 136.626 |  | 36.1478 | 28.5241 | 84.951 | 51.6398 | 0 | 51.6398
| Rel.Std.Dev. |  | 6.999% | 25.62% |  | 5.894% | 6.408% | 12.84% | 9.682% | 0% | 9.113%
| PDM3 |  | 1.759% | -6.963% |  | 6.992% | -22.35% | 15.42% | -6.963% | 4.667% | -1.148%
| Z-Score (Absolute) |  | 0.34 | 0.67 | 10.55 | 0.94 | 2.46 | 1.97 | 0.67 | 0.67 | 0
| Analytical results for Cl in OREAS 712 (Indicative Value 63 ppm) | [21 empty] | Indicative
Cl | Round | Replicate | Lab | [19 empty] | 1
| No | No | 13 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | 60 | [19 empty] | 1
| 1 | 2 | 70 | [19 empty] | 3
| 1 | 3 | 60 | [19 empty] | 16
| 1 | 4 | 70 | [19 empty] | 63.3333
| 1 | 5 | 60 | [19 empty] | 9
| 1 | 6 | 60
| Mean |  | 63.3333
| Median |  | 60
| Std Dev. |  | 5.164
| Rel.Std.Dev. |  | 8.154%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 712 (Certified Value < 100 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 17 | 21 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | [12 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | [12 empty] | 12
| 1 | 6 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cr in OREAS 712 (Certified Value 146 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  |  | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 34 | 137 | 200 | 200 | 258.3 | < 70 | < 70 | 210 | 270 | 140 | 144 | 150.5303 | 147.109 | 140 | 143.1403 | 200 | 200 | 171 |  |  | 1
| 1 | 2 | < 34 | 68 | 200 | 200 | 239.96 | < 70 | < 70 | 210 | 270 | 140 | 164 | 150.5303 | 136.846 | 140 | 138.3305 | 200 | 100 | 171 |  |  | #N/A
| 1 | 3 | < 34 | 68 | 200 | 200 | 268.15 | < 70 | < 70 | 210 | 270 | 140 | 144 | 157.3726 | 142.97 | 140 | 139.7114 | 200 | 100 | 171 |  |  | 16
| 1 | 4 | < 34 | 68 | 200 | 200 | 264.66 | < 70 | < 70 | 270 | 270 | 140 | 144 | 148.2495 | 153.951 | 140 | 141.778 | 200 | 100 | 130 |  |  | 146.1003
| 1 | 5 | < 34 | 137 | 100 | 100 | 265 | < 70 | < 70 | 210 | 270 | 140 | 144 | 145.9688 | 136.846 | 140 | 141.0737 | 200 | 200 | 144 |  |  | 13
| 1 | 6 | < 34 | 137 | 100 | 200 | 254.19 | < 70 | < 70 | 210 | 270 | 140 | 137 | 152.811 | 143.688 | 140 | 137.6361 | 100 | 200 | 151
| Mean |  |  | 102.5 | 166.6667 | 183.3333 | 258.3767 |  |  | 220 | 270 | 140 | 146.1667 | 150.9104 | 143.5683 | 140 | 140.2783 | 183.3333 | 150 | 156.3333
| Median |  |  | 102.5 | 200 | 200 | 261.48 |  |  | 210 | 270 | 140 | 144 | 150.5303 | 143.329 | 140 | 140.3925 | 200 | 150 | 161
| Std Dev. |  |  | 37.7929 | 51.6398 | 40.8248 | 10.3496 |  |  | 24.4949 | 0 | 0 | 9.1742 | 3.9284 | 6.4975 | 0 | 2.1057 | 40.8248 | 54.7723 | 17.4318
| Rel.Std.Dev. |  |  | 36.87% | 30.98% | 22.27% | 4.006% |  |  | 11.13% | 0% | 0% | 6.277% | 2.603% | 4.526% | 0% | 1.501% | 22.27% | 36.51% | 11.15%
| PDM3 |  |  | -29.84% | 14.08% | 25.48% | 76.85% |  |  | 50.58% | 84.8% | -4.175% | 0.04541% | 3.292% | -1.733% | -4.175% | -3.985% | 25.48% | 2.669% | 7.004%
| Z-Score (Absolute) |  | 3.14 | 0.99 | NA | NA | 2.93 | 2.71 | 2.71 | 1.88 | 3.32 | 0.13 | 0.1 | 0.23 | 0.04 | 0.13 | 0.04 | NA | NA | 0.36
Results from laboratories 3, 4, 21 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 08 | 17 | [18 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | < 90 | [18 empty] | 1
| 1 | 2 | < 100 | < 90 | [18 empty] | 4
| 1 | 3 | < 100 | < 90 | [18 empty] | 16
| 1 | 4 | < 100 | < 90 | [18 empty] | < 90
| 1 | 5 | < 100 | < 90 | [18 empty] | 10
| 1 | 6 | < 100 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Cu in OREAS 712 (Certified Value < 50 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 14 | 15 | 17 | 21 | 23 | 24 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 50 | < 50 | 300 | < 40 | < 40 | < 160 | 19.9712 | < 80 | < 50 | 60 | < 48 | [9 empty] | 1
| 1 | 2 | < 50 | < 50 | 200 | < 40 | < 40 | < 160 | 31.954 | < 80 | < 50 | < 50 | < 48 | [9 empty] | 18
| 1 | 3 | < 50 | < 50 | 100 | < 40 | < 40 | < 160 | 35.7685 | < 80 | < 50 | < 50 | < 48 | [9 empty] | 16
| 1 | 4 | < 50 | < 50 | 200 | < 40 | < 40 | < 160 | 31.954 | < 80 | < 50 | < 50 | < 48 | [9 empty] | < 50
| 1 | 5 | < 50 | < 50 | 400 | < 40 | < 40 | < 160 | 31.954 | < 80 | < 50 | < 50 | < 48 | [9 empty] | 14
| 1 | 6 | 50 | < 50 | 300 | < 40 | < 40 | < 160 | 31.954 | < 80 | < 50 | < 50 | < 48
| Mean |  | 50 |  | 250 |  |  |  | 30.5926 |  |  | 60
| Median |  | 50 |  | 250 |  |  |  | 31.954 |  |  | 60
| Std Dev. |  |  |  | 104.8809 |  |  |  | 5.4225
| Rel.Std.Dev. |  |  |  | 41.95% |  |  |  | 17.72%
PDM3
| Z-Score (Absolute) |  | 0 | 0.54 | 28.63 | 1.19 | 1.19 | 6.58 | 0.18 | 1.4 | 0.54 | 0.22 | 0.67
| Analytical results for Dy in OREAS 712 (Indicative Value 87 ppm) | [21 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 17 | 22 | [18 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | 90 | [18 empty] | 1
| 1 | 2 | < 170 | 90 | [18 empty] | 5
| 1 | 3 | < 170 | 90 | [18 empty] | 16
| 1 | 4 | < 170 | 90 | [18 empty] | 87.1308
| 1 | 5 | < 170 | 90 | [18 empty] | 11
| 1 | 6 | < 170 | 90
| Mean |  |  | 90
| Median |  |  | 90
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 3.293%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Er in OREAS 712 (Indicative Value 80 ppm) | [21 empty] | Indicative
Er | Round | Replicate | Lab | [19 empty] | 1
| No | No | 22 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | 90 | [19 empty] | 1
| 1 | 2 | 90 | [19 empty] | 6
| 1 | 3 | < 90 | [19 empty] | 16
| 1 | 4 | 90 | [19 empty] | 80.1632
| 1 | 5 | 90 | [19 empty] | 12
| 1 | 6 | 90
| Mean |  | 90
| Median |  | 90
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 12.27%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 712 (Indicative Value 86 ppm) | [21 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 17 | 22 | [18 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | 90 | [18 empty] | 1
| 1 | 2 | < 170 | 90 | [18 empty] | 7
| 1 | 3 | < 170 | 90 | [18 empty] | 16
| 1 | 4 | < 170 | 90 | [18 empty] | 86.3632
| 1 | 5 | < 170 | 90 | [18 empty] | 13
| 1 | 6 | < 170 | 90
| Mean |  |  | 90
| Median |  |  | 90
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 4.211%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for F in OREAS 712 (Indicative Value < 5000 ppm) | [21 empty] | Indicative
F | Round | Replicate | Lab | [19 empty] | 1
| No | No | 17 | [19 empty] | ppm
|  |  | OxBF*XRF | [19 empty] | 0
0
| 1 | 1 | < 5000 | [19 empty] | 1
| 1 | 2 | < 5000 | [19 empty] | 8
| 1 | 3 | < 5000 | [19 empty] | 16
| 1 | 4 | < 5000 | [19 empty] | < 5000
| 1 | 5 | < 5000 | [19 empty] | 14
| 1 | 6 | < 5000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 712 (Certified Value 14.52 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 14.493 | 14.5905 | 14.25 | 14.45 | 14.6299 | 14.69 | 14.765 | 14.22 | 14.835 | 14.667 | 14.5 | 14.423 | 14.4944 | 14.3983 | 14.646 | 14.4817 | 14.7 | 14.62 | 14.45 | 14.604 | 1
| 1 | 2 | 14.514 | 14.5835 | 14.4 | 14.2 | 14.6047 | 14.41 | 14.933 | 14.297 | 14.66 | 14.709 | 14.486 | 14.472 | 14.4753 | 14.4141 | 14.597 | 14.4437 | 14.85 | 14.55 | 14.55 | 14.583 | #N/A
| 1 | 3 | 14.493 | 14.5905 | 14.3 | 14.55 | 14.6209 | 14.34 | 14.863 | 14.262 | 14.639 | 14.639 | 14.479 | 14.458 | 14.5641 | 14.3929 | 14.688 | 14.4867 | 14.9 | 14.55 | 14.5 | 14.625 | 16
| 1 | 4 | 14.402 | 14.5835 | 14.2 | 14.55 | 14.5982 | 14.48 | 14.262 | 14.248 | 14.87 | 14.514 | 14.479 | 14.458 | 14.5009 | 14.4247 | 14.667 | 14.5032 | 14.85 | 14.41 | 14.6 | 14.569 | 14.5188
| 1 | 5 | 14.402 | 14.5975 | 14.35 | 14.45 | 14.6274 | 14.55 | 14.702 | 14.304 | 14.814 | 14.562 | 14.514 | 14.472 | 14.466 | 14.4437 | 14.674 | 14.5481 | 14.8 | 14.48 | 14.5 | 14.542 | 15
| 1 | 6 | 14.549 | 14.5975 | 14.35 | 14.55 | 14.5971 | 14.34 | 15.038 | 14.248 | 14.625 | 14.59 | 14.444 | 14.472 | 14.441 | 14.4078 | 14.59 | 14.5043 | 14.95 | 14.62 | 14.5 | 14.583
| Mean |  | 14.4755 | 14.5905 | 14.3083 | 14.4583 | 14.613 | 14.4683 | 14.7605 | 14.2632 | 14.7405 | 14.6135 | 14.4837 | 14.4592 | 14.4903 | 14.4136 | 14.6437 | 14.4946 | 14.8417 | 14.5383 | 14.5167 | 14.5843
| Median |  | 14.493 | 14.5905 | 14.325 | 14.5 | 14.6128 | 14.445 | 14.814 | 14.255 | 14.737 | 14.6145 | 14.4825 | 14.465 | 14.4849 | 14.411 | 14.6565 | 14.495 | 14.85 | 14.55 | 14.5 | 14.583
| Std Dev. |  | 0.0605 | 0.006261 | 0.073598 | 0.13571 | 0.014819 | 0.13586 | 0.27174 | 0.032053 | 0.11066 | 0.071746 | 0.023704 | 0.018999 | 0.042014 | 0.01855 | 0.041215 | 0.034222 | 0.08612 | 0.081833 | 0.05164 | 0.028535
| Rel.Std.Dev. |  | 0.418% | 0.04291% | 0.5144% | 0.9386% | 0.1014% | 0.939% | 1.841% | 0.2247% | 0.7507% | 0.491% | 0.1637% | 0.1314% | 0.2899% | 0.1287% | 0.2815% | 0.2361% | 0.5803% | 0.5629% | 0.3557% | 0.1957%
| PDM3 |  | -0.2984% | 0.4937% | -1.45% | -0.4166% | 0.6489% | -0.3477% | 1.665% | -1.761% | 1.527% | 0.6521% | -0.2421% | -0.4109% | -0.1965% | -0.7249% | 0.8599% | -0.1667% | 2.224% | 0.1344% | -0.01483% | 0.4512%
| Z-Score (Absolute) |  | 0.25 | 0.7 | 1.62 | 0.39 | 0.88 | 0.32 | 2.1 | 2 | 1.94 | 0.89 | 0.18 | 0.39 | 0.13 | 0.76 | 1.14 | 0.09 | 2.77 | 0.26 | 0.09 | 0.65
| Analytical results for Ga in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 08 | 17 | [18 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | < 150 | [18 empty] | 1
| 1 | 2 | < 100 | < 150 | [18 empty] | 9
| 1 | 3 | < 100 | < 150 | [18 empty] | 16
| 1 | 4 | < 100 | < 150 | [18 empty] | < 100
| 1 | 5 | < 100 | < 150 | [18 empty] | 15
| 1 | 6 | < 100 | < 150
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Gd in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 17 | 22 | [18 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | < 90 | [18 empty] | 1
| 1 | 2 | < 170 | < 90 | [18 empty] | 10
| 1 | 3 | < 170 | < 90 | [18 empty] | 16
| 1 | 4 | < 170 | < 90 | [18 empty] | < 90
| 1 | 5 | < 170 | < 90 | [18 empty] | 16
| 1 | 6 | < 170 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ge in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Ge | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 11
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 712 (Certified Value < 100 ppm) | [21 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 09 | 11 | 13 | 14 | 15 | 17 | 21 | 23 | 24 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | 100 | < 100 | < 100 | < 100 | < 100 | < 200 | 30 | 100 | < 100 | < 100 | < 100 | [9 empty] | 1
| 1 | 2 | 100 | < 100 | < 100 | < 100 | < 100 | < 200 | 30 | 200 | < 100 | < 100 | < 100 | [9 empty] | #N/A
| 1 | 3 | 100 | < 100 | < 100 | 100 | < 100 | < 200 | 34.825 | 100 | < 100 | < 100 | < 100 | [9 empty] | 16
| 1 | 4 | 100 | < 100 | < 100 | < 100 | < 100 | < 200 | 40 | 100 | < 100 | < 100 | < 100 | [9 empty] | < 100
| 1 | 5 | 100 | < 100 | < 100 | < 100 | < 100 | < 200 | 40 | 100 | < 100 | < 100 | < 100 | [9 empty] | 16
| 1 | 6 | 100 | < 100 | < 100 | < 100 | < 100 | < 200 | 35 | 100 | < 100 | < 100 | < 100
| Mean |  | 100 |  |  | 100 |  |  | 34.9708 | 116.6667
| Median |  | 100 |  |  | 100 |  |  | 34.9125 | 100
| Std Dev. |  | 0 | [5 empty] | 4.4727 | 40.8248
| Rel.Std.Dev. |  | 0% | [5 empty] | 12.79% | 34.99%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Hg in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 12
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Ho | Round | Replicate | Lab | [19 empty] | 1
| No | No | 22 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 90 | [19 empty] | 1
| 1 | 2 | < 90 | [19 empty] | 25
| 1 | 3 | < 90 | [19 empty] | 16
| 1 | 4 | < 90 | [19 empty] | < 90
| 1 | 5 | < 90 | [19 empty] | 7
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
In | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 2
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 712 (Certified Value 0.753 wt.%) | [21 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.733 | 0.76 | 0.74 | 0.77 | 0.75509 | 0.78 | 0.96 | 0.76 | 0.75 | 0.76 | 0.74 | 0.75 | 0.74323 | 0.74466 | 0.76 | 0.73904 | 0.76 | 0.77 | 0.729 |  | 1
| 1 | 2 | 0.734 | 0.75 | 0.74 | 0.75 | 0.75609 | 0.76 | 1.02 | 0.76 | 0.75 | 0.76 | 0.74 | 0.76 | 0.74524 | 0.74671 | 0.75 | 0.73757 | 0.77 | 0.78 | 0.724 |  | #N/A
| 1 | 3 | 0.734 | 0.76 | 0.76 | 0.76 | 0.75438 | 0.75 | 0.95 | 0.76 | 0.75 | 0.75 | 0.74 | 0.76 | 0.7439 | 0.74451 | 0.76 | 0.73025 | 0.76 | 0.77 | 0.728 |  | 16
| 1 | 4 | 0.729 | 0.75 | 0.73 | 0.77 | 0.75744 | 0.76 | 1.03 | 0.76 | 0.76 | 0.76 | 0.74 | 0.75 | 0.74558 | 0.74569 | 0.76 | 0.7376 | 0.76 | 0.77 | 0.724 |  | 0.75315
| 1 | 5 | 0.735 | 0.77 | 0.75 | 0.76 | 0.75618 | 0.75 | 0.94 | 0.76 | 0.76 | 0.76 | 0.74 | 0.75 | 0.74625 | 0.74774 | 0.76 | 0.74013 | 0.76 | 0.78 | 0.724 |  | 17
| 1 | 6 | 0.738 | 0.77 | 0.75 | 0.76 | 0.75433 | 0.76 | 0.93 | 0.76 | 0.75 | 0.76 | 0.74 | 0.76 | 0.74257 | 0.74569 | 0.75 | 0.73854 | 0.77 | 0.77 | 0.72
| Mean |  | 0.73383 | 0.76 | 0.745 | 0.76167 | 0.75558 | 0.76 | 0.97167 | 0.76 | 0.75333 | 0.75833 | 0.74 | 0.755 | 0.74446 | 0.74583 | 0.75667 | 0.73719 | 0.76333 | 0.77333 | 0.72483
| Median |  | 0.734 | 0.76 | 0.745 | 0.76 | 0.75559 | 0.76 | 0.955 | 0.76 | 0.75 | 0.76 | 0.74 | 0.755 | 0.74457 | 0.74569 | 0.76 | 0.73807 | 0.76 | 0.77 | 0.724
| Std Dev. |  | 0.0029269 | 0.0089443 | 0.010488 | 0.0075277 | 0.0012101 | 0.010954 | 0.042622 | 1.2162e-16 | 0.005164 | 0.0040825 | 1.2162e-16 | 0.0054772 | 0.0014452 | 0.0012302 | 0.005164 | 0.0035323 | 0.005164 | 0.005164 | 0.0032506
| Rel.Std.Dev. |  | 0.3988% | 1.177% | 1.408% | 0.9883% | 0.1602% | 1.441% | 4.387% | 1.6e-14% | 0.6855% | 0.5383% | 1.643e-14% | 0.7255% | 0.1941% | 0.1649% | 0.6825% | 0.4792% | 0.6765% | 0.6678% | 0.4485%
| PDM3 |  | -2.565% | 0.9094% | -1.082% | 1.131% | 0.3232% | 0.9094% | 29.01% | 0.9094% | 0.02422% | 0.6881% | -1.746% | 0.2455% | -1.154% | -0.9717% | 0.4668% | -2.119% | 1.352% | 2.68% | -3.76%
| Z-Score (Absolute) |  | 1.89 | 0.38 | 0.92 | 0.53 | 0 | 0.38 | 18.8 | 0.38 | 0.2 | 0.24 | 1.36 | 0.05 | 0.97 | 0.85 | 0.09 | 1.6 | 0.67 | 1.54 | 2.68
| Analytical results for La2O3 in OREAS 712 (Certified Value 341 ppm) | [21 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 05 | 13 | 15 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 300 | 300 | 461.9 | 320 | 386.25 | 400 | 200 | 400 | 400 | [11 empty] | 1
| 1 | 2 | 300 | 400 | 488.1 | 300 | 360.5 | 400 | 200 | 300 | 300 | [11 empty] | #N/A
| 1 | 3 | 400 | 300 | 498.5 | 280 | 389.443 | 300 | 300 | 300 | 400 | [11 empty] | 16
| 1 | 4 | 300 | 400 | 492.9 | 300 | 350.2 | 400 | 200 | 300 | 400 | [11 empty] | 341.3451
| 1 | 5 | 300 | 200 | 424.4 | 330 | 339.9 | 300 | 300 | 400 | 300 | [11 empty] | 18
| 1 | 6 | 300 | 400 | 479.9 | 330 | 350.2 | 300 | 300 | 400 | 400
| Mean |  | 316.6667 | 333.3333 | 474.2833 | 310 | 362.7488 | 350 | 250 | 350 | 366.6667
| Median |  | 300 | 350 | 484 | 310 | 355.35 | 350 | 250 | 350 | 400
| Std Dev. |  | 40.8248 | 81.6497 | 27.5578 | 20 | 20.5278 | 54.7723 | 54.7723 | 54.7723 | 51.6398
| Rel.Std.Dev. |  | 12.89% | 24.49% | 5.81% | 6.452% | 5.659% | 15.65% | 21.91% | 15.65% | 14.08%
| PDM3 |  | -7.23% | -2.347% | 38.95% | -9.183% | 6.27% | 2.536% | -26.76% | 2.536% | 7.418%
| Z-Score (Absolute) |  | 1.35 | 0.67 | 5.03 | 1.62 | 0.52 | 0 | 4.05 | 0 | 0.67
| Analytical results for Lu in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Lu | Round | Replicate | Lab | [19 empty] | 1
| No | No | 22 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 90 | [19 empty] | 1
| 1 | 2 | < 90 | [19 empty] | 3
| 1 | 3 | < 90 | [19 empty] | 16
| 1 | 4 | < 90 | [19 empty] | < 90
| 1 | 5 | < 90 | [19 empty] | 9
| 1 | 6 | 90
| Mean |  | 90
| Median |  | 90
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 712 (Certified Value 0.976 wt.%) | [21 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.94 | 0.93 | 0.98 | 0.98 | 0.96416 | 0.97 | 0.99 | 0.99 | 0.86 | 0.97 | 0.99 | 0.96 | 1.0255 | 0.95732 | 1 | 0.96433 | 1.02 | 0.99 | 0.864 |  | 1
| 1 | 2 | 0.93 | 0.94 | 0.99 | 0.97 | 0.9523 | 0.98 | 1 | 0.97 | 0.88 | 1 | 0.99 | 0.95 | 1.0217 | 0.94744 | 1 | 0.96036 | 1.01 | 0.98 | 0.873 |  | #N/A
| 1 | 3 | 0.94 | 0.93 | 0.99 | 0.96 | 0.95587 | 0.96 | 0.97 | 0.98 | 0.88 | 1.04 | 0.98 | 0.98 | 1.0286 | 0.94788 | 1 | 0.96582 | 1.01 | 0.97 | 0.859 |  | 16
| 1 | 4 | 0.93 | 0.94 | 0.97 | 0.98 | 0.95589 | 0.97 | 1.04 | 0.99 | 0.89 | 0.99 | 0.98 | 0.96 | 1.0237 | 0.9516 | 1 | 0.96634 | 1.03 | 0.98 | 0.86 |  | 0.97603
| 1 | 5 | 0.94 | 0.92 | 0.99 | 0.98 | 0.95342 | 0.97 | 0.91 | 0.98 | 0.88 | 1 | 0.99 | 0.98 | 1.0248 | 0.95732 | 1 | 0.96482 | 1.02 | 0.98 | 0.861 |  | 19
| 1 | 6 | 0.94 | 0.92 | 0.98 | 0.96 | 0.95898 | 0.97 | 0.9 | 0.98 | 0.89 | 1.01 | 0.99 | 0.97 | 1.0249 | 0.95992 | 1 | 0.96223 | 1.01 | 0.98 | 0.868
| Mean |  | 0.93667 | 0.93 | 0.98333 | 0.97167 | 0.95677 | 0.97 | 0.96833 | 0.98167 | 0.88 | 1.0017 | 0.98667 | 0.96667 | 1.0249 | 0.95358 | 1 | 0.96398 | 1.0167 | 0.98 | 0.86417
| Median |  | 0.94 | 0.93 | 0.985 | 0.975 | 0.95588 | 0.97 | 0.98 | 0.98 | 0.88 | 1 | 0.99 | 0.965 | 1.0249 | 0.95446 | 1 | 0.96457 | 1.015 | 0.98 | 0.8625
| Std Dev. |  | 0.005164 | 0.0089443 | 0.008165 | 0.0098319 | 0.0042961 | 0.0063246 | 0.054191 | 0.0075277 | 0.010954 | 0.023166 | 0.005164 | 0.012111 | 0.0022822 | 0.0053351 | 0 | 0.0022786 | 0.008165 | 0.0063246 | 0.0054191
| Rel.Std.Dev. |  | 0.5513% | 0.9617% | 0.8303% | 1.012% | 0.449% | 0.652% | 5.596% | 0.7668% | 1.245% | 2.313% | 0.5234% | 1.253% | 0.2227% | 0.5595% | 0% | 0.2364% | 0.8031% | 0.6454% | 0.6271%
| PDM3 |  | -4.033% | -4.716% | 0.7482% | -0.4471% | -1.973% | -0.6179% | -0.7886% | 0.5774% | -9.839% | 2.627% | 1.09% | -0.9594% | 5.003% | -2.3% | 2.456% | -1.235% | 4.163% | 0.4067% | -11.46%
| Z-Score (Absolute) |  | 1.37 | 1.64 | 0.55 | 0.07 | 0.54 | 0 | 0.07 | 0.48 | 3.7 | 1.3 | 0.68 | 0.14 | 2.25 | 0.67 | 1.23 | 0.25 | 1.92 | 0.41 | 4.35
| Analytical results for Mn in OREAS 712 (Certified Value 1.29 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 1.3344 | 1.2779 | 1.27 | 1.31 | 1.3078 | 1.293 | 1.479 | 1.363 | 1.3212 | 1.324 | 1.301 | 1.278 | 1.3108 | 1.2605 | 1.282 | 1.2783 | 1.29 | 1.28 | 1.2454 |  | 1
| 1 | 2 | 1.3437 | 1.2856 | 1.29 | 1.29 | 1.3005 | 1.262 | 1.464 | 1.371 | 1.3057 | 1.348 | 1.293 | 1.293 | 1.3084 | 1.262 | 1.275 | 1.2738 | 1.3 | 1.28 | 1.2497 |  | 21
| 1 | 3 | 1.3367 | 1.2779 | 1.28 | 1.32 | 1.3069 | 1.27 | 1.479 | 1.379 | 1.3112 | 1.317 | 1.301 | 1.286 | 1.3161 | 1.2568 | 1.289 | 1.2768 | 1.31 | 1.29 | 1.2432 |  | 16
| 1 | 4 | 1.3274 | 1.2856 | 1.27 | 1.32 | 1.3064 | 1.262 | 1.44 | 1.379 | 1.3251 | 1.317 | 1.301 | 1.286 | 1.3103 | 1.2564 | 1.289 | 1.2759 | 1.31 | 1.3 | 1.2454 |  | 1.2923
| 1 | 5 | 1.3266 | 1.2701 | 1.29 | 1.31 | 1.3139 | 1.278 | 1.456 | 1.371 | 1.3251 | 1.34 | 1.293 | 1.286 | 1.3079 | 1.2627 | 1.282 | 1.2766 | 1.3 | 1.29 | 1.2468 |  | 20
| 1 | 6 | 1.3383 | 1.2701 | 1.28 | 1.31 | 1.3137 | 1.255 | 1.425 | 1.363 | 1.3065 | 1.324 | 1.293 | 1.286 | 1.3047 | 1.2594 | 1.275 | 1.2824 | 1.3 | 1.29 | 1.2296
| Mean |  | 1.3345 | 1.2779 | 1.28 | 1.31 | 1.3082 | 1.27 | 1.4572 | 1.371 | 1.3158 | 1.3283 | 1.297 | 1.2858 | 1.3097 | 1.2596 | 1.282 | 1.2773 | 1.3017 | 1.2883 | 1.2433
| Median |  | 1.3356 | 1.2779 | 1.28 | 1.31 | 1.3074 | 1.266 | 1.46 | 1.371 | 1.3162 | 1.324 | 1.297 | 1.286 | 1.3094 | 1.26 | 1.282 | 1.2767 | 1.3 | 1.29 | 1.2454
| Std Dev. |  | 0.006584 | 0.0069319 | 0.0089443 | 0.010954 | 0.00503 | 0.013755 | 0.021591 | 0.0071554 | 0.0090752 | 0.012785 | 0.0043818 | 0.0047504 | 0.0038171 | 0.0026216 | 0.006261 | 0.0028894 | 0.0075277 | 0.0075277 | 0.007067
| Rel.Std.Dev. |  | 0.4934% | 0.5425% | 0.6988% | 0.8362% | 0.3845% | 1.083% | 1.482% | 0.5219% | 0.6897% | 0.9625% | 0.3378% | 0.3694% | 0.2914% | 0.2081% | 0.4884% | 0.2262% | 0.5783% | 0.5843% | 0.5684%
| PDM3 |  | 3.264% | -1.119% | -0.9542% | 1.367% | 1.228% | -1.728% | 12.75% | 6.087% | 1.816% | 2.786% | 0.3613% | -0.5028% | 1.345% | -2.529% | -0.7994% | -1.165% | 0.7224% | -0.3093% | -3.79%
| Z-Score (Absolute) |  | 1.35 | 0.7 | 0.62 | 0.46 | 0.4 | 0.98 | 5.81 | 2.67 | 0.67 | 1.12 | 0 | 0.42 | 0.45 | 1.36 | 0.55 | 0.72 | 0.16 | 0.32 | 1.95
| Analytical results for Mo in OREAS 712 (Indicative Value < 50 ppm) | [21 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 23 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 100 | < 50 | < 100 | 33 | < 70 | < 50 | < 50 | [13 empty] | 1
| 1 | 2 | 90 | < 50 | < 100 | 33 | < 70 | 60 | < 50 | [13 empty] | 4
| 1 | 3 | 110 | < 50 | < 100 | < 33 | < 70 | 70 | 60 | [13 empty] | 16
| 1 | 4 | 80 | < 50 | < 100 | < 33 | < 70 | < 50 | < 50 | [13 empty] | < 50
| 1 | 5 | 80 | < 50 | < 100 | < 33 | < 70 | < 50 | < 50 | [13 empty] | 10
| 1 | 6 | 90 | < 50 | < 100 | < 33 | < 70 | < 50 | < 50
| Mean |  | 91.6667 |  |  | 33 |  | 65 | 60
| Median |  | 90 |  |  | 33 |  | 65 | 60
| Std Dev. |  | 11.6905 |  |  | 0 |  | 7.0711
| Rel.Std.Dev. |  | 12.75% |  |  | 0% |  | 10.88%
PDM3
| Z-Score (Absolute) |  | 4.72 | 0.67 | 1.35 | 0.9 | 0 | 0.41 | 0.2
| Analytical results for Na2O in OREAS 712 (Certified Value 0.831 wt.%) | [21 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 24 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.8 | 0.78 | 0.80704 | 0.87 | 0.91 | 0.86 | 0.8 | 0.79 | 0.89 | 0.87753 | 0.846 | 0.8 | 0.87441 | 0.766 | [6 empty] | 1
| 1 | 2 | 0.81 | 0.77 | 0.81082 | 0.86 | 0.81 | 0.87 | 0.79 | 0.8 | 0.9 | 0.87174 | 0.845 | 0.8 | 0.86993 | 0.773 | [6 empty] | #N/A
| 1 | 3 | 0.81 | 0.78 | 0.80985 | 0.85 | 0.84 | 0.83 | 0.77 | 0.79 | 0.9 | 0.8699 | 0.86018 | 0.8 | 0.87177 | 0.762 | [6 empty] | 16
| 1 | 4 | 0.83 | 0.77 | 0.80311 | 0.84 | 0.96 | 0.85 | 0.79 | 0.79 | 0.9 | 0.87643 | 0.8675 | 0.8 | 0.87128 | 0.764 | [6 empty] | 0.8312
| 1 | 5 | 0.81 | 0.79 | 0.80587 | 0.83 | 0.72 | 0.86 | 0.8 | 0.8 | 0.9 | 0.8728 | 0.8525 | 0.8 | 0.86867 | 0.757 | [6 empty] | 21
| 1 | 6 | 0.82 | 0.79 | 0.79948 | 0.85 | 0.92 | 0.83 | 0.78 | 0.79 | 0.91 | 0.86922 | 0.879 | 0.8 | 0.86587 | 0.756
| Mean |  | 0.81333 | 0.78 | 0.80603 | 0.85 | 0.86 | 0.85 | 0.78833 | 0.79333 | 0.9 | 0.87294 | 0.85836 | 0.8 | 0.87032 | 0.763
| Median |  | 0.81 | 0.78 | 0.80645 | 0.85 | 0.875 | 0.855 | 0.79 | 0.79 | 0.9 | 0.87227 | 0.85634 | 0.8 | 0.87061 | 0.763
| Std Dev. |  | 0.010328 | 0.0089443 | 0.0042447 | 0.014142 | 0.087864 | 0.016733 | 0.01169 | 0.005164 | 0.0063246 | 0.0033986 | 0.013255 | 1.2162e-16 | 0.0029139 | 0.006261
| Rel.Std.Dev. |  | 1.27% | 1.147% | 0.5266% | 1.664% | 10.22% | 1.969% | 1.483% | 0.6509% | 0.7027% | 0.3893% | 1.544% | 1.52e-14% | 0.3348% | 0.8206%
| PDM3 |  | -2.15% | -6.16% | -3.029% | 2.261% | 3.464% | 2.261% | -5.158% | -4.556% | 8.277% | 5.021% | 3.267% | -3.754% | 4.706% | -8.205%
| Z-Score (Absolute) |  | 0.67 | 1.29 | 0.81 | 0 | 0.18 | 0 | 1.13 | 1.04 | 0.92 | 0.42 | 0.15 | NA | 0.37 | 1.6
Results from laboratory 17 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Nb in OREAS 712 (Certified Value 116 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 05 | 13 | 15 | 17 | 21 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 120 | 110 | 209.65 | 140 | 62.9151 | 70 | 100 | 90 | [12 empty] | 1
| 1 | 2 | 120 | 100 | 220.83 | 140 | 69.9056 | 140 | 120 | 100 | [12 empty] | #N/A
| 1 | 3 | 130 | 120 | 190.49 | 140 | 73.0339 | 140 | 120 | 110 | [12 empty] | 16
| 1 | 4 | 110 | 120 | 223.98 | 140 | 59.4198 | 70 | 120 | 110 | [12 empty] | 115.5005
| 1 | 5 | 110 | 110 | 179.24 | 140 | 66.4103 | 70 | 100 | 90 | [12 empty] | 22
| 1 | 6 | 120 | 120 | 199.23 | 140 | 69.9056 | 140 | 120 | 120
| Mean |  | 118.3333 | 113.3333 | 203.9033 | 140 | 66.9317 | 105 | 113.3333 | 103.3333
| Median |  | 120 | 115 | 204.44 | 140 | 68.1579 | 105 | 120 | 105
| Std Dev. |  | 7.5277 | 8.165 | 17.508 | 0 | 5.0549 | 38.3406 | 10.328 | 12.1106
| Rel.Std.Dev. |  | 6.361% | 7.204% | 8.586% | 0% | 7.552% | 36.51% | 9.113% | 11.72%
| PDM3 |  | 2.453% | -1.876% | 76.54% | 21.21% | -42.05% | -9.091% | -1.876% | -10.53%
| Z-Score (Absolute) |  | 0.36 | 0 | 6.61 | 1.93 | 3.39 | 0.62 | 0 | 0.73
| Analytical results for Nd in OREAS 712 (Indicative Value 158 ppm) | [21 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 13 | 15 | 17 | 22 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 203.1 | 163 | 73.9455 | < 170 | 170 | [15 empty] | 1
| 1 | 2 | 204.65 | 197 | 78.8752 | < 170 | 170 | [15 empty] | 5
| 1 | 3 | 210.99 | 146 | 88.2909 | < 170 | 170 | [15 empty] | 16
| 1 | 4 | 208.08 | 163 | 83.8049 | < 170 | 170 | [15 empty] | 158.4677
| 1 | 5 | 216.14 | 163 | 78.8752 | < 170 | 170 | [15 empty] | 11
| 1 | 6 | 227.28 | 180 | 88.7346 | < 170 | 170
| Mean |  | 211.7067 | 168.6667 | 82.0877 |  | 170
| Median |  | 209.535 | 163 | 81.34 |  | 170
| Std Dev. |  | 8.9411 | 17.5575 | 5.8744 |  | 0
| Rel.Std.Dev. |  | 4.223% | 10.41% | 7.156% |  | 0%
| PDM3 |  | 33.6% | 6.436% | -48.2% |  | 7.277%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.35 | 1.3 | 0.04
| Analytical results for Ni in OREAS 712 (Certified Value < 50 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 14 | 15 | 17 | 21 | 23 | 24 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | < 24 | < 39 | < 196 | 15.7171 | < 80 | < 50 | < 50 | < 47 | [9 empty] | 1
| 1 | 2 | 80 | < 50 | < 100 | < 24 | < 39 | < 196 | 19.6464 | < 80 | < 50 | < 50 | < 47 | [9 empty] | #N/A
| 1 | 3 | 90 | < 50 | < 100 | < 24 | < 39 | < 196 | 11.7289 | < 80 | < 50 | < 50 | < 47 | [9 empty] | 16
| 1 | 4 | 60 | < 50 | < 100 | 55 | < 39 | < 196 | 27.5049 | < 80 | < 50 | < 50 | < 47 | [9 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 100 | < 24 | < 39 | < 196 | 11.7878 | < 80 | 50 | < 50 | < 47 | [9 empty] | 23
| 1 | 6 | < 50 | < 50 | < 100 | < 24 | < 39 | < 196 | 31.4342 | < 80 | < 50 | < 50 | < 47
| Mean |  | 76.6667 |  |  | 55 |  |  | 19.6365 |  | 50
| Median |  | 80 |  |  | 55 |  |  | 17.6818 |  | 50
| Std Dev. |  | 15.2753 | [5 empty] | 8.2534
| Rel.Std.Dev. |  | 19.92% | [5 empty] | 42.03%
PDM3
| Z-Score (Absolute) |  | 3.25 | 0 | 3.14 | 0.76 | 0.67 | 9.21 | 0.67 | 1.8 | 0.52 | 0 | 0.18
| Analytical results for P2O5 in OREAS 712 (Certified Value 0.182 wt.%) | [21 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.161 | 0.19 | 0.18 | 0.18 | 0.18263 | 0.19 | 0.31 | 0.2 | 0.19 | 0.18 | 0.183 | 0.181 | 0.18021 | 0.18126 | 0.15 | 0.17752 | 0.18 | 0.18 | 0.174 |  | 1
| 1 | 2 | 0.165 | 0.2 | 0.18 | 0.18 | 0.18775 | 0.19 | 0.33 | 0.2 | 0.19 | 0.18 | 0.183 | 0.183 | 0.18051 | 0.18179 | 0.15 | 0.17823 | 0.18 | 0.19 | 0.176 |  | #N/A
| 1 | 3 | 0.166 | 0.19 | 0.17 | 0.18 | 0.18838 | 0.19 | 0.3 | 0.2 | 0.19 | 0.18 | 0.185 | 0.185 | 0.17915 | 0.18035 | 0.14 | 0.17894 | 0.18 | 0.18 | 0.173 |  | 16
| 1 | 4 | 0.163 | 0.2 | 0.17 | 0.19 | 0.18051 | 0.19 | 0.33 | 0.2 | 0.19 | 0.18 | 0.183 | 0.182 | 0.17921 | 0.18868 | 0.14 | 0.17643 | 0.18 | 0.18 | 0.174 |  | 0.18206
| 1 | 5 | 0.16 | 0.18 | 0.18 | 0.18 | 0.18487 | 0.2 | 0.31 | 0.2 | 0.19 | 0.18 | 0.185 | 0.184 | 0.18057 | 0.18338 | 0.15 | 0.17752 | 0.18 | 0.18 | 0.175 |  | 24
| 1 | 6 | 0.164 | 0.18 | 0.18 | 0.18 | 0.1922 | 0.19 | 0.28 | 0.2 | 0.18 | 0.18 | 0.183 | 0.183 | 0.17851 | 0.18232 | 0.15 | 0.17792 | 0.18 | 0.18 | 0.169
| Mean |  | 0.16317 | 0.19 | 0.17667 | 0.18167 | 0.18606 | 0.19167 | 0.31 | 0.2 | 0.18833 | 0.18 | 0.18367 | 0.183 | 0.17969 | 0.18296 | 0.14667 | 0.17776 | 0.18 | 0.18167 | 0.1735
| Median |  | 0.1635 | 0.19 | 0.18 | 0.18 | 0.18631 | 0.19 | 0.31 | 0.2 | 0.19 | 0.18 | 0.183 | 0.183 | 0.17971 | 0.18206 | 0.15 | 0.17772 | 0.18 | 0.18 | 0.174
| Std Dev. |  | 0.0023166 | 0.0089443 | 0.005164 | 0.0040825 | 0.0042408 | 0.0040825 | 0.018974 | 3.0405e-17 | 0.0040825 | 3.0405e-17 | 0.0010328 | 0.0014142 | 0.00085305 | 0.0029785 | 0.005164 | 0.00084068 | 3.0405e-17 | 0.0040825 | 0.002429
| Rel.Std.Dev. |  | 1.42% | 4.708% | 2.923% | 2.247% | 2.279% | 2.13% | 6.121% | 1.52e-14% | 2.168% | 1.689e-14% | 0.5623% | 0.7728% | 0.4747% | 1.628% | 3.521% | 0.4729% | 1.689e-14% | 2.247% | 1.4%
| PDM3 |  | -10.38% | 4.362% | -2.961% | -0.215% | 2.196% | 5.278% | 70.28% | 9.855% | 3.447% | -1.13% | 0.8836% | 0.5174% | -1.3% | 0.4976% | -19.44% | -2.361% | -1.13% | -0.215% | -4.701%
| Z-Score (Absolute) |  | 3.01 | 1.35 | 0.81 | 0 | 0.71 | 1.63 | 20.86 | NA | 1.08 | 0.27 | 0.33 | 0.22 | 0.32 | 0.21 | 5.69 | 0.64 | 0.27 | 0 | 1.33
Results from laboratory 9 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 712 (Indicative Value < 50 ppm) | [21 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 08 | 11 | 13 | 14 | 15 | 17 | 21 | 23 | 24 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 370 | < 50 | 70 | < 100 | < 90 | < 46 | < 190 | 139.2499 | < 90 | < 50 | 100 | < 74 | [8 empty] | 1
| 1 | 2 | 280 | < 50 | < 50 | < 100 | < 90 | < 46 | < 190 | 106.7583 | < 90 | < 50 | 80 | < 74 | [8 empty] | 6
| 1 | 3 | 370 | < 50 | < 50 | < 100 | < 90 | < 46 | < 190 | 157.0275 | < 90 | < 50 | 120 | < 74 | [8 empty] | 16
| 1 | 4 | 280 | < 50 | 110 | < 100 | < 90 | < 46 | < 190 | 139.2499 | < 90 | < 50 | 80 | < 74 | [8 empty] | < 50
| 1 | 5 | 460 | < 50 | < 50 | < 100 | < 90 | < 46 | < 190 | 102.1166 | < 90 | < 50 | 100 | < 74 | [8 empty] | 12
| 1 | 6 | 460 | < 50 | 110 | < 100 | < 90 | < 46 | < 190 | 157.8166 | < 90 | < 50 | 70 | < 74
| Mean |  | 370 |  | 96.6667 |  |  |  |  | 133.7031 |  |  | 91.6667
| Median |  | 370 |  | 110 |  |  |  |  | 139.2499 |  |  | 90
| Std Dev. |  | 80.4984 |  | 23.094 |  |  |  |  | 24.1278 |  |  | 18.3485
| Rel.Std.Dev. |  | 21.76% |  | 23.89% |  |  |  |  | 18.05% |  |  | 20.02%
PDM3
| Z-Score (Absolute) |  | 9.39 | 0.67 | 0.37 | 0.05 | 0.05 | 0.73 | 1.3 | 2.48 | 0.05 | 0.67 | 1.26 | 0.32
| Analytical results for Pr in OREAS 712 (Indicative Value 81 ppm) | [21 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 13 | 17 | 22 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | [17 empty] | 0
0
| 1 | 1 | 58 | < 170 | 80 | [17 empty] | 1
| 1 | 2 | 50 | < 170 | 80 | [17 empty] | 7
| 1 | 3 | 83 | < 170 | 170 | [17 empty] | 16
| 1 | 4 | 66 | < 170 | 80 | [17 empty] | 81.3883
| 1 | 5 | 74 | < 170 | 80 | [17 empty] | 13
| 1 | 6 | 66 | < 170 | 80
| Mean |  | 66.1667 |  | 95
| Median |  | 66 |  | 80
| Std Dev. |  | 11.6003 |  | 36.7423
| Rel.Std.Dev. |  | 17.53% |  | 38.68%
| PDM3 |  | -18.7% |  | 16.72%
| Z-Score (Absolute) |  | 0.81 | 0 | 0.67
| Analytical results for Rb in OREAS 712 (Indicative Value 62 ppm) | [21 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 23 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 50 | < 50 | < 100 | 101 | < 180 | < 50 | 50 | [13 empty] | 1
| 1 | 2 | 60 | < 50 | < 100 | 73 | < 180 | < 50 | 60 | [13 empty] | 8
| 1 | 3 | 50 | < 50 | < 100 | 64 | < 180 | < 50 | 50 | [13 empty] | 16
| 1 | 4 | 50 | < 50 | < 100 | 101 | < 180 | < 50 | 50 | [13 empty] | 61.6385
| 1 | 5 | 50 | < 50 | < 100 | 64 | < 180 | < 50 | 60 | [13 empty] | 14
| 1 | 6 | 50 | < 50 | < 100 | 73 | < 180 | < 50 | < 50
| Mean |  | 51.6667 |  |  | 79.3333 |  |  | 54
| Median |  | 50 |  |  | 73 |  |  | 50
| Std Dev. |  | 4.0825 |  |  | 17.2588 |  |  | 5.4772
| Rel.Std.Dev. |  | 7.902% |  |  | 21.75% |  |  | 10.14%
| PDM3 |  | -16.18% |  |  | 28.71% |  |  | -12.39%
| Z-Score (Absolute) |  | 0.04 | 0.67 | 0 | 0.79 | 1.12 | 0.67 | 0.02
| Analytical results for S in OREAS 712 (Indicative Value 0.003 wt.%) | [21 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 21 | 23 | 24 | [7 empty] | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 3
3
| 1 | 1 | < 0.004 | < 0.01 | < 0.01 | 0.024 | < 0.004 | < 0.002 | 0.0032 | < 0.008 | 0.0021024 | < 0.08 | < 0.01 | 0.01 | 0.0012 | [7 empty] | 1
| 1 | 2 | < 0.004 | < 0.01 | < 0.01 | 0.024 | < 0.004 | < 0.002 | 0.0024 | < 0.008 | 0.0018921 | < 0.08 | < 0.01 | < 0.01 | 0.006 | [7 empty] | 9
| 1 | 3 | < 0.004 | < 0.01 | < 0.01 | 0.02 | < 0.004 | < 0.002 | 0.0032 | < 0.008 | 0.0020919 | < 0.08 | < 0.01 | < 0.01 | 0.0024 | [7 empty] | 16
| 1 | 4 | < 0.004 | < 0.01 | < 0.01 | 0.036 | < 0.004 | < 0.002 | 0.0024 | < 0.008 | 0.0021024 | < 0.08 | < 0.01 | < 0.01 | 0.0036 | [7 empty] | 0.0033664
| 1 | 5 | < 0.004 | < 0.01 | < 0.01 | 0.028 | < 0.004 | < 0.002 | 0.0024 | < 0.008 | 0.0021024 | < 0.08 | < 0.01 | < 0.01 | 0.0024 | [7 empty] | 15
| 1 | 6 | < 0.004 | < 0.01 | < 0.01 | 0.028 | < 0.004 | < 0.002 | 0.0032 | < 0.008 | 0.0025228 | < 0.08 | < 0.01 | < 0.01 | 0.006
| Mean |  |  |  |  | 0.026667 |  |  | 0.0028 |  | 0.0021356 |  |  | 0.01 | 0.0036
| Median |  |  |  |  | 0.026 |  |  | 0.0028 |  | 0.0021024 |  |  | 0.01 | 0.003
| Std Dev. |  |  |  |  | 0.005465 |  |  | 0.00043818 |  | 0.0002071 |  |  |  | 0.002008
| Rel.Std.Dev. |  |  |  |  | 20.49% |  |  | 15.65% |  | 9.697% |  |  |  | 55.78%
| PDM3 |  |  |  |  | 692.2% |  |  | -16.82% |  | -36.56% |  |  | 197.1% | 6.941%
| Z-Score (Absolute) |  | 0.74 | 0.37 | 0.37 | 8.37 | 0.74 | 1.11 | 0.44 | 0 | 0.69 | 13.29 | 0.37 | 0.67 | 0.15
| Analytical results for Sb in OREAS 712 (Indicative Value < 50 ppm) | [21 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 08 | 17 | 21 | 23 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | < 50 | 60 | < 100 | < 80 | < 50 | < 50 | [14 empty] | 1
| 1 | 2 | < 50 | 50 | < 100 | < 80 | < 50 | < 50 | [14 empty] | 10
| 1 | 3 | 140 | < 50 | < 100 | < 80 | < 50 | < 50 | [14 empty] | 16
| 1 | 4 | < 50 | 90 | < 100 | < 80 | < 50 | < 50 | [14 empty] | < 50
| 1 | 5 | < 50 | 50 | < 100 | < 80 | < 50 | < 50 | [14 empty] | 16
| 1 | 6 | < 50 | 50 | < 100 | < 80 | < 50 | < 50
| Mean |  | 140 | 60
| Median |  | 140 | 50
| Std Dev. |  |  | 17.3205
| Rel.Std.Dev. |  |  | 28.87%
PDM3
| Z-Score (Absolute) |  | 0.09 | 0.83 | 0.52 | 0.09 | 1.33 | 1.33
| Analytical results for Sc in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Sc | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 11
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 12
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 712 (Certified Value 47.97 wt.%) | [21 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 46.81 | 47.45 | 48.1 | 48.1 | 47.2578 | 49 | 44.19 | 46.1 | 48.34 | 47.91 | 47.92 | 47.94 | 47.9999 | 48.4335 | 47.9 | 48.075 | 48.4 | 48.3 | 45.695 |  | 1
| 1 | 2 | 46.99 | 47.44 | 48.5 | 47.5 | 47.2057 | 48.3 | 44.22 | 46 | 47.91 | 48.68 | 47.94 | 48.11 | 47.9535 | 48.4261 | 47.9 | 48.0592 | 48.6 | 48.2 | 45.64 |  | #N/A
| 1 | 3 | 46.88 | 47.45 | 48.5 | 47.8 | 47.1161 | 47.9 | 44.38 | 46.1 | 47.65 | 48.4 | 47.97 | 47.99 | 47.9927 | 48.3822 | 47.7 | 48.1627 | 48.8 | 48.2 | 45.668 |  | 16
| 1 | 4 | 46.7 | 47.44 | 48 | 48.1 | 47.291 | 48.1 | 44.29 | 46.1 | 48.31 | 48.04 | 47.96 | 47.88 | 48.034 | 48.5222 | 47.9 | 48.1556 | 48.6 | 48.2 | 45.648 |  | 47.9683
| 1 | 5 | 46.68 | 47.43 | 48.6 | 47.8 | 47.2045 | 48.6 | 44.48 | 46.1 | 48.37 | 48.42 | 47.95 | 48.06 | 48.0617 | 48.532 | 47.9 | 48.1331 | 48.5 | 48.3 | 45.673 |  | 25
| 1 | 6 | 47.1 | 47.43 | 48.2 | 47.7 | 47.009 | 47.9 | 44.34 | 46.1 | 47.56 | 48.2 | 47.96 | 47.99 | 47.8728 | 48.5011 | 48 | 48.1144 | 48.9 | 48.2 | 45.79
| Mean |  | 46.86 | 47.44 | 48.3167 | 47.8333 | 47.1807 | 48.3 | 44.3167 | 46.0833 | 48.0233 | 48.275 | 47.95 | 47.995 | 47.9858 | 48.4662 | 47.8833 | 48.1167 | 48.6333 | 48.2333 | 45.6857
| Median |  | 46.845 | 47.44 | 48.35 | 47.8 | 47.2051 | 48.2 | 44.315 | 46.1 | 48.11 | 48.3 | 47.955 | 47.99 | 47.9963 | 48.4673 | 47.9 | 48.1237 | 48.6 | 48.2 | 45.6705
| Std Dev. |  | 0.16456 | 0.0089443 | 0.24833 | 0.23381 | 0.1029 | 0.43359 | 0.10708 | 0.040825 | 0.36593 | 0.28098 | 0.017889 | 0.082158 | 0.066549 | 0.060703 | 0.098319 | 0.042289 | 0.18619 | 0.05164 | 0.05468
| Rel.Std.Dev. |  | 0.3512% | 0.01885% | 0.514% | 0.4888% | 0.2181% | 0.8977% | 0.2416% | 0.08859% | 0.762% | 0.582% | 0.03731% | 0.1712% | 0.1387% | 0.1252% | 0.2053% | 0.08789% | 0.3828% | 0.1071% | 0.1197%
| PDM3 |  | -2.31% | -1.101% | 0.7263% | -0.2813% | -1.642% | 0.6915% | -7.613% | -3.93% | 0.1148% | 0.6394% | -0.03812% | 0.05569% | 0.03642% | 1.038% | -0.1771% | 0.3093% | 1.386% | 0.5525% | -4.759%
| Z-Score (Absolute) |  | 2.42 | 1.17 | 0.71 | 0.33 | 1.73 | 0.67 | 7.87 | 4.08 | 0.08 | 0.62 | 0.08 | 0.02 | 0 | 1.03 | 0.22 | 0.28 | 1.39 | 0.53 | 4.94
| Analytical results for Sm in OREAS 712 (Indicative Value < 170 ppm) | [21 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 17 | 22 | [18 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | 90 | [18 empty] | 1
| 1 | 2 | < 170 | < 90 | [18 empty] | 24
| 1 | 3 | < 170 | < 90 | [18 empty] | 16
| 1 | 4 | < 170 | < 90 | [18 empty] | < 170
| 1 | 5 | < 170 | 90 | [18 empty] | 7
| 1 | 6 | < 170 | 90
| Mean |  |  | 90
| Median |  |  | 90
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 712 (Indicative Value < 50 ppm) | [21 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 08 | 11 | 13 | 14 | 15 | 17 | 21 | 23 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 80 | < 100 | < 80 | < 39 | < 160 | 63.012 | < 80 | < 50 | 60 | [9 empty] | 1
| 1 | 2 | < 50 | < 50 | < 80 | < 100 | < 80 | < 39 | < 160 | 59.0737 | < 80 | < 50 | < 50 | [9 empty] | 2
| 1 | 3 | < 50 | < 50 | < 80 | < 100 | < 80 | < 39 | < 160 | 47.0227 | < 80 | < 50 | < 50 | [9 empty] | 16
| 1 | 4 | < 50 | < 50 | < 80 | < 100 | < 80 | < 39 | < 160 | 47.259 | < 80 | < 50 | 50 | [9 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 80 | < 100 | < 80 | < 39 | < 160 | 63.012 | < 80 | < 50 | 50 | [9 empty] | 8
| 1 | 6 | < 50 | < 50 | < 80 | < 100 | < 80 | < 39 | < 160 | 63.012 | < 80 | < 50 | 60
| Mean | [8 empty] | 57.0652 |  |  | 55
| Median | [8 empty] | 61.0429 |  |  | 55
| Std Dev. | [8 empty] | 7.8376 |  |  | 5.7735
| Rel.Std.Dev. | [8 empty] | 13.73% |  |  | 10.5%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0 | 0.5 | 0 | 0.92 | 1.85 | 0.83 | 0 | 0.67 | 0.26
| Analytical results for Sr in OREAS 712 (Certified Value 310 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 372 | 85 | 200 | 300 | 590 | 250 | 205.7613 | 359.3776 | 340 | 337.9981 | 300 | 400 | 271 | [7 empty] | 1
| 1 | 2 | 381 | 85 | 200 | 300 | 590 | 170 | 202.9427 | 355.1497 | 420 | 336.3566 | 300 | 400 | 271 | [7 empty] | #N/A
| 1 | 3 | 397 | 85 | 200 | 300 | 680 | 250 | 205.7613 | 412.2696 | 340 | 337.6749 | 300 | 400 | 279 | [7 empty] | 16
| 1 | 4 | 372 | 85 | 200 | 400 | 590 | 250 | 205.7613 | 346.6937 | 420 | 338.7458 | 300 | 400 | 288 | [7 empty] | 309.5523
| 1 | 5 | 364 | 85 | 200 | 300 | 590 | 250 | 222.6732 | 414.3413 | 340 | 339.6022 | 300 | 400 | 304 | [7 empty] | 26
| 1 | 6 | 389 | 85 | 200 | 400 | 590 | 250 | 208.58 | 350.9217 | 250 | 337.0758 | 300 | 400 | 262
| Mean |  | 379.1667 | 85 | 200 | 333.3333 | 605 | 236.6667 | 208.58 | 373.1256 | 351.6667 | 337.9089 | 300 | 400 | 279.1667
| Median |  | 376.5 | 85 | 200 | 300 | 590 | 250 | 205.7613 | 357.2636 | 340 | 337.8365 | 300 | 400 | 275
| Std Dev. |  | 12.2543 | 0 | 0 | 51.6398 | 36.7423 | 32.6599 | 7.1307 | 31.4159 | 63.3772 | 1.1605 | 0 | 0 | 14.9855
| Rel.Std.Dev. |  | 3.232% | 0% | 0% | 15.49% | 6.073% | 13.8% | 3.419% | 8.42% | 18.02% | 0.3434% | 0% | 0% | 5.368%
| PDM3 |  | 22.49% | -72.54% | -35.39% | 7.682% | 95.44% | -23.55% | -32.62% | 20.54% | 13.6% | 9.161% | -3.086% | 29.22% | -9.816%
| Z-Score (Absolute) |  | 0.47 | 2.9 | NA | NA | 3.07 | 1.13 | 1.48 | 0.4 | 0.17 | 0 | NA | NA | 0.67
Results from laboratories 3, 4, 21 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for Ta in OREAS 712 (Indicative Value 10.9 ppm) | [21 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 08 | 15 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | [17 empty] | 1
1
| 1 | 1 | < 100 | 12.284 | < 80 | [17 empty] | 1
| 1 | 2 | < 100 | 12.284 | < 80 | [17 empty] | 3
| 1 | 3 | < 100 | 12.2226 | < 80 | [17 empty] | 16
| 1 | 4 | < 100 | 8.1893 | < 80 | [17 empty] | 10.9089
| 1 | 5 | < 100 | 8.1893 | < 80 | [17 empty] | 9
| 1 | 6 | < 100 | 12.284 | < 80
| Mean |  |  | 10.9089
| Median |  |  | 12.2533
| Std Dev. |  |  | 2.1067
| Rel.Std.Dev. |  |  | 19.31%
| PDM3 |  |  | -0.000122%
| Z-Score (Absolute) |  | 0.67 | 2.24 | 0
| Analytical results for Tb in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Tb | Round | Replicate | Lab | [19 empty] | 1
| No | No | 22 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 90 | [19 empty] | 1
| 1 | 2 | < 90 | [19 empty] | 25
| 1 | 3 | < 90 | [19 empty] | 16
| 1 | 4 | < 90 | [19 empty] | < 90
| 1 | 5 | < 90 | [19 empty] | 10
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Te | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 5
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 11
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 712 (Indicative Value 70 ppm) | [21 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 11 | 13 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 74 | 260 | 79 | 70.3049 | 180 | 80 | [14 empty] | 1
| 1 | 2 | 71 | 260 | 62 | 74.699 | 180 | 60 | [14 empty] | 26
| 1 | 3 | 61 | 260 | 70 | 69.9534 | 180 | 80 | [14 empty] | 16
| 1 | 4 | 59 | 260 | 79 | 74.699 | 180 | 70 | [14 empty] | 70.1783
| 1 | 5 | 61 | 260 | 70 | 70.3049 | 180 | 70 | [14 empty] | 12
| 1 | 6 | 63 | 260 | 79 | 65.9109 | 180 | 70
| Mean |  | 64.8333 | 260 | 73.1667 | 70.9787 | 180 | 71.6667
| Median |  | 62 | 260 | 74.5 | 70.3049 | 180 | 70
| Std Dev. |  | 6.1455 | 0 | 7.0261 | 3.3264 | 0 | 7.5277
| Rel.Std.Dev. |  | 9.479% | 0% | 9.603% | 4.686% | 0% | 10.5%
| PDM3 |  | -7.616% | 270.5% | 4.258% | 1.141% | 156.5% | 2.121%
| Z-Score (Absolute) |  | 1.13 | 28.37 | 0.12 | 0.22 | 15.33 | 0.12
| Analytical results for TiO2 in OREAS 712 (Certified Value 12.04 wt.%) | [21 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 11.99 | 12.23 | 11.9 | 11.95 | 12.0801 | 12.3 | 12.24 | 12.044 | 12.1 | 11.86 | 12.01 | 11.88 | 11.9307 | 12.102 | 12.04 | 11.9072 | 11.95 | 12.05 | 12.15 | 12.01 | 1
| 1 | 2 | 12.07 | 12.22 | 12 | 11.9 | 12.0416 | 12.2 | 12.28 | 12.094 | 11.96 | 11.94 | 12.01 | 11.91 | 11.9025 | 12.119 | 11.97 | 11.9003 | 12.05 | 12.05 | 12.15 | 11.99 | #N/A
| 1 | 3 | 12.05 | 12.23 | 11.95 | 11.9 | 12.0671 | 12.1 | 12.37 | 11.994 | 11.93 | 11.93 | 11.99 | 11.9 | 11.9337 | 12.1028 | 12.05 | 11.93 | 12.1 | 12.05 | 12.15 | 12.06 | 16
| 1 | 4 | 11.97 | 12.22 | 11.85 | 11.95 | 12.0875 | 12.1 | 12.27 | 12.344 | 12.1 | 11.94 | 11.98 | 11.93 | 11.9478 | 12.1206 | 12.02 | 11.9273 | 12.1 | 11.95 | 12.15 | 12.01 | 12.0381
| 1 | 5 | 11.98 | 12.24 | 12 | 11.95 | 12.029 | 12.2 | 12.24 | 12.144 | 12.1 | 11.95 | 12.01 | 11.92 | 11.9562 | 12.1434 | 12.01 | 11.9342 | 12.05 | 12.05 | 12.2 | 12.03 | 27
| 1 | 6 | 12.06 | 12.24 | 11.95 | 11.9 | 12.0054 | 12 | 12.27 | 12.044 | 11.94 | 11.93 | 12 | 11.9 | 11.9171 | 12.1233 | 11.97 | 11.925 | 12.1 | 12.1 | 12.1 | 12.01
| Mean |  | 12.02 | 12.23 | 11.9417 | 11.925 | 12.0518 | 12.15 | 12.2783 | 12.1107 | 12.0217 | 11.925 | 12 | 11.9067 | 11.9313 | 12.1185 | 12.01 | 11.9206 | 12.0583 | 12.0417 | 12.15 | 12.0183
| Median |  | 12.02 | 12.23 | 11.95 | 11.925 | 12.0544 | 12.15 | 12.27 | 12.069 | 12.03 | 11.935 | 12.005 | 11.905 | 11.9322 | 12.1198 | 12.015 | 11.9261 | 12.075 | 12.05 | 12.15 | 12.01
| Std Dev. |  | 0.044721 | 0.0089443 | 0.058452 | 0.027386 | 0.031902 | 0.10488 | 0.047924 | 0.12517 | 0.086352 | 0.032711 | 0.012649 | 0.017512 | 0.019647 | 0.01528 | 0.034059 | 0.013635 | 0.058452 | 0.04916 | 0.031623 | 0.024014
| Rel.Std.Dev. |  | 0.3721% | 0.07313% | 0.4895% | 0.2297% | 0.2647% | 0.8632% | 0.3903% | 1.034% | 0.7183% | 0.2743% | 0.1054% | 0.1471% | 0.1647% | 0.1261% | 0.2836% | 0.1144% | 0.4847% | 0.4082% | 0.2603% | 0.1998%
| PDM3 |  | -0.1506% | 1.594% | -0.8013% | -0.9398% | 0.1135% | 0.9293% | 1.995% | 0.6025% | -0.1368% | -0.9398% | -0.3168% | -1.092% | -0.8873% | 0.6678% | -0.2337% | -0.976% | 0.1678% | 0.02936% | 0.9293% | -0.1645%
| Z-Score (Absolute) |  | 0.01 | 1.58 | 0.6 | 0.72 | 0.23 | 0.97 | 1.94 | 0.67 | 0.01 | 0.72 | 0.16 | 0.86 | 0.67 | 0.74 | 0.08 | 0.75 | 0.28 | 0.16 | 0.97 | 0.02
| Analytical results for Tl in OREAS 712 (Indicative Value < 100 ppm) | [21 empty] | Indicative
Tl | Round | Replicate | Lab | [19 empty] | 1
| No | No | 08 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 7
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Tm | Round | Replicate | Lab | [19 empty] | 1
| No | No | 22 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 90 | [19 empty] | 1
| 1 | 2 | < 90 | [19 empty] | 8
| 1 | 3 | < 90 | [19 empty] | 16
| 1 | 4 | < 90 | [19 empty] | < 90
| 1 | 5 | < 90 | [19 empty] | 14
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 712 (Indicative Value 15.9 ppm) | [21 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 11 | 13 | 15 | 17 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | [15 empty] | 1
1
| 1 | 1 | < 20 | 170 | 25 | 21.2008 | < 80 | [15 empty] | 1
| 1 | 2 | < 20 | 170 | < 17 | 16.9607 | < 80 | [15 empty] | 9
| 1 | 3 | < 20 | 170 | < 17 | 21.0948 | < 80 | [15 empty] | 16
| 1 | 4 | < 20 | 170 | 25 | 12.7205 | < 80 | [15 empty] | 15.8918
| 1 | 5 | < 20 | 170 | < 17 | 16.9607 | < 80 | [15 empty] | 15
| 1 | 6 | < 20 | 170 | < 17 | 16.9607 | < 80
| Mean |  |  | 170 | 25 | 17.6497
| Median |  |  | 170 | 25 | 16.9607
| Std Dev. |  |  | 0 | 0 | 3.1686
| Rel.Std.Dev. |  |  | 0% | 0% | 17.95%
| PDM3 |  |  | 969.7% | 57.31% | 11.06%
| Z-Score (Absolute) |  | 0.67 | 13.39 | 0.31 | 0 | 2.18
| Analytical results for V in OREAS 712 (Indicative Value 135 ppm) | [21 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 168 | 100 | 200 | 143.23 | < 100 | < 60 | < 17 | 110 | 123 | 217.716 | 56.0161 | 110 | 180.4648 | < 100 | 100 | 84 |  |  |  |  | 1
| 1 | 2 | 174 | 100 | 200 | 129.62 | < 100 | < 60 | < 17 | 170 | 95 | 199.1187 | 58.8169 | 110 | 162.4099 | < 100 | 100 | 90 |  |  |  |  | 10
| 1 | 3 | 146 | 100 | 200 | 146.99 | < 100 | < 60 | < 17 | 170 | 101 | 225.1662 | 61.3097 | 170 | 172.6044 | < 100 | 100 | 78 |  |  |  |  | 16
| 1 | 4 | 162 | 100 | 300 | 142.95 | < 100 | < 60 | < 17 | 110 | 106 | 212.1518 | 64.4186 | 110 | 185.5623 | < 100 | 100 | 84 |  |  |  |  | 135.2788
| 1 | 5 | 140 | 100 | 200 | 134.16 | < 100 | < 60 | < 17 | 170 | 106 | 208.3427 | 72.821 | 110 | 186.5871 | < 100 | 100 | 90 |  |  |  |  | 16
| 1 | 6 | 157 | 100 | 200 | 162.89 | < 100 | < 60 | < 17 | 170 | 112 | 210.3406 | 61.6177 | 110 | 184.6705 | < 100 | 100 | 84
| Mean |  | 157.8333 | 100 | 216.6667 | 143.3067 |  |  |  | 150 | 107.1667 | 212.1393 | 62.5 | 120 | 178.7165 |  | 100 | 85
| Median |  | 159.5 | 100 | 200 | 143.09 |  |  |  | 170 | 106 | 211.2462 | 61.4637 | 110 | 182.5677 |  | 100 | 84
| Std Dev. |  | 12.9679 | 0 | 40.8248 | 11.5602 |  |  |  | 30.9839 | 9.6212 | 8.8089 | 5.7966 | 24.4949 | 9.493 |  | 0 | 4.5166
| Rel.Std.Dev. |  | 8.216% | 0% | 18.84% | 8.067% |  |  |  | 20.66% | 8.978% | 4.152% | 9.275% | 20.41% | 5.312% |  | 0% | 5.314%
| PDM3 |  | 16.67% | -26.08% | 60.16% | 5.934% |  |  |  | 10.88% | -20.78% | 56.82% | -53.8% | -11.29% | 32.11% |  | -26.08% | -37.17%
| Z-Score (Absolute) |  | 0.67 | NA | NA | 0.48 | 0.77 | 1.06 | 1.32 | 0.56 | 0 | 1.4 | 0.6 | 0.19 | 0.95 | 0.77 | NA | 0.3
Results from laboratories 3, 4 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 712 (Indicative Value 10.4 ppm) | [21 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 13 | 15 | 17 | 21 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | 10 | < 10 | < 80 | < 100 | 87 | 15.8604 | < 80 | < 10 | 10 | [11 empty] | 1
| 1 | 2 | 10 | < 10 | < 80 | < 100 | 95 | 11.8953 | < 80 | 10 | 10 | [11 empty] | 11
| 1 | 3 | 20 | < 10 | < 80 | < 100 | 71 | 3.9453 | < 80 | 10 | 20 | [11 empty] | 16
| 1 | 4 | < 10 | < 10 | < 80 | < 100 | 79 | 3.9651 | < 80 | < 10 | < 10 | [11 empty] | 10.3509
| 1 | 5 | 10 | < 10 | < 80 | < 100 | 71 | 11.8953 | < 80 | 10 | 20 | [11 empty] | 17
| 1 | 6 | 10 | < 10 | < 80 | < 100 | 79 | 15.8604 | < 80 | 10 | < 10
| Mean |  | 12 |  |  |  | 80.3333 | 10.5703 |  | 10 | 15
| Median |  | 10 |  |  |  | 79 | 11.8953 |  | 10 | 15
| Std Dev. |  | 4.4721 |  |  |  | 9.3524 | 5.4222 |  | 0 | 5.7735
| Rel.Std.Dev. |  | 37.27% |  |  |  | 11.64% | 51.3% |  | 0% | 38.49%
| PDM3 |  | 15.93% |  |  |  | 676.1% | 2.12% |  | -3.39% | 44.91%
| Z-Score (Absolute) |  | 0.08 | 0.67 | 2.83 | 3.88 | 6.97 | 0.11 | 2.83 | 0.34 | 0
| Analytical results for Y2O3 in OREAS 712 (Certified Value 175 ppm) | [21 empty] | Certified
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 11 | 13 | 15 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | 160 | 180 | 200.2 | 300 | 120 | 88 | 200 | 140 | 250 | 180 | [10 empty] | 1
| 1 | 2 | 160 | 210 | 201.6 | 200 | 170 | 93.5 | 200 | 160 | 250 | 180 | [10 empty] | #N/A
| 1 | 3 | 170 | 190 | 206.3 | 300 | 70 | 93.0325 | 200 | 160 | 250 | 190 | [10 empty] | 16
| 1 | 4 | 150 | 190 | 194.5 | 200 | 120 | 82.5 | < 200 | 160 | 250 | 170 | [10 empty] | 175.4976
| 1 | 5 | 160 | 180 | 204.2 | 300 | 160 | 88 | 200 | 150 | 250 | 190 | [10 empty] | 28
| 1 | 6 | 160 | 210 | 198.1 | 200 | 110 | 88 | < 200 | 170 | 250 | 180
| Mean |  | 160 | 193.3333 | 200.8167 | 250 | 125 | 88.8388 | 200 | 156.6667 | 250 | 181.6667
| Median |  | 160 | 190 | 200.9 | 250 | 120 | 88 | 200 | 160 | 250 | 180
| Std Dev. |  | 6.3246 | 13.6626 | 4.2376 | 54.7723 | 36.1939 | 4.0399 | 0 | 10.328 | 0 | 7.5277
| Rel.Std.Dev. |  | 3.953% | 7.067% | 2.11% | 21.91% | 28.96% | 4.547% | 0% | 6.592% | 0% | 4.144%
| PDM3 |  | -8.831% | 10.16% | 14.43% | 42.45% | -28.77% | -49.38% | 13.96% | -10.73% | 42.45% | 3.515%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 0.86 | NA | 1.05 | 1.97 | 0 | 0.25 | 2.21 | 0.38
Results from laboratory 11 were removed due to their 100 ppm reading resolution.
| Analytical results for Yb in OREAS 712 (Indicative Value < 90 ppm) | [21 empty] | Indicative
Yb | Round | Replicate | Lab | [19 empty] | 1
| No | No | 22 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | < 90 | [19 empty] | 1
| 1 | 2 | < 90 | [19 empty] | 12
| 1 | 3 | < 90 | [19 empty] | 16
| 1 | 4 | < 90 | [19 empty] | < 90
| 1 | 5 | < 90 | [19 empty] | 18
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 712 (Certified Value 71 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 80 | 60 | 80 | < 100 | < 40 | < 160 | 66.838 | < 80 | 69.6091 | < 50 | 80 | 88 | [8 empty] | 1
| 1 | 2 | 80 | < 50 | 70 | < 100 | < 40 | < 160 | 66.838 | < 80 | 67.5107 | < 50 | 90 | 80 | [8 empty] | #N/A
| 1 | 3 | 80 | < 50 | 70 | < 100 | < 40 | < 160 | 62.3474 | 80 | 68.398 | < 50 | 100 | 72 | [8 empty] | 16
| 1 | 4 | 80 | < 50 | 70 | < 100 | < 40 | < 160 | 66.838 | < 80 | 67.7954 | < 50 | 90 | 72 | [8 empty] | 70.7401
| 1 | 5 | 80 | < 50 | 70 | < 100 | < 40 | < 160 | 71.0154 | 80 | 68.1498 | < 50 | 80 | 72 | [8 empty] | 29
| 1 | 6 | 80 | < 50 | 80 | < 100 | < 40 | < 160 | 71.0154 | 80 | 66.4013 | < 50 | 60 | 80
| Mean |  | 80 | 60 | 73.3333 |  |  |  | 67.482 | 80 | 67.9774 |  | 83.3333 | 77.3333
| Median |  | 80 | 60 | 70 |  |  |  | 66.838 | 80 | 67.9726 |  | 85 | 76
| Std Dev. |  | 0 |  | 5.164 |  |  |  | 3.2428 | 0 | 1.0587 |  | 13.6626 | 6.532
| Rel.Std.Dev. |  | 0% |  | 7.042% |  |  |  | 4.805% | 0% | 1.557% |  | 16.4% | 8.447%
| PDM3 |  | 13.09% | -15.18% | 3.666% |  |  |  | -4.606% | 13.09% | -3.905% |  | 17.8% | 9.32%
| Z-Score (Absolute) |  | 0.67 | NA | NA | 0.92 | 2.49 | 0.67 | 0 | 0.38 | 0.03 | 2.23 | NA | 0.53
Results from laboratories 3, 4 and 23 were removed due to their 10 ppm reading resolution.
| Analytical results for ZrO2 in OREAS 712 (Certified Value 0.153 wt.%) | [21 empty] | Certified
ZrO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 08 | 09 | 11 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  |  |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 0.15 | 0.162 | 0.162 | 0.14537 | 0.162 | 0.14 | 0.15 | 0.15 | 0.15433 | 0.14397 | 0.16 | 0.14891 | 0.162 | 0.149 | 0.135 | 0.158 |  |  |  |  | 1
| 1 | 2 | 0.14 | 0.162 | 0.149 | 0.14485 | 0.162 | 0.15 | 0.16 | 0.16 | 0.15333 | 0.16047 | 0.16 | 0.14923 | 0.162 | 0.149 | 0.135 | 0.153 |  |  |  |  | #N/A
| 1 | 3 | 0.15 | 0.162 | 0.149 | 0.14653 | 0.149 | 0.14 | 0.15 | 0.15 | 0.15433 | 0.15669 | 0.17 | 0.14913 | 0.162 | 0.149 | 0.135 | 0.16 |  |  |  |  | 16
| 1 | 4 | 0.14 | 0.162 | 0.162 | 0.15145 | 0.135 | 0.15 | 0.15 | 0.15 | 0.15733 | 0.15947 | 0.15 | 0.14924 | 0.162 | 0.149 | 0.135 | 0.156 |  |  |  |  | 0.15328
| 1 | 5 | 0.16 | 0.162 | 0.162 | 0.14724 | 0.162 | 0.14 | 0.16 | 0.15 | 0.15533 | 0.15297 | 0.15 | 0.1493 | 0.162 | 0.135 | 0.149 | 0.157 |  |  |  |  | 30
| 1 | 6 | 0.16 | 0.162 | 0.162 | 0.15183 | 0.135 | 0.14 | 0.16 | 0.15 | 0.153 | 0.15397 | 0.15 | 0.14961 | 0.162 | 0.149 | 0.135 | 0.16
| Mean |  | 0.15 | 0.162 | 0.15767 | 0.14788 | 0.15083 | 0.14333 | 0.155 | 0.15167 | 0.15461 | 0.15459 | 0.15667 | 0.14924 | 0.162 | 0.14667 | 0.13733 | 0.15733
| Median |  | 0.15 | 0.162 | 0.162 | 0.14688 | 0.1555 | 0.14 | 0.155 | 0.15 | 0.15433 | 0.15533 | 0.155 | 0.14923 | 0.162 | 0.149 | 0.135 | 0.1575
| Std Dev. |  | 0.0089443 | 0 | 0.0067132 | 0.0030351 | 0.013258 | 0.005164 | 0.0054772 | 0.0040825 | 0.0015694 | 0.0059749 | 0.008165 | 0.00022715 | 0 | 0.0057155 | 0.0057155 | 0.0026583
| Rel.Std.Dev. |  | 5.963% | 0% | 4.258% | 2.052% | 8.79% | 3.603% | 3.534% | 2.692% | 1.015% | 3.865% | 5.212% | 0.1522% | 0% | 3.897% | 4.162% | 1.69%
| PDM3 |  | -2.142% | 5.686% | 2.859% | -3.526% | -1.599% | -6.492% | 1.12% | -1.055% | 0.8659% | 0.8533% | 2.207% | -2.641% | 5.686% | -4.317% | -10.41% | 2.642%
| Z-Score (Absolute) |  | 0.52 | 1.49 | 0.74 | 0.87 | 0.38 | 1.63 | 0.31 | 0.24 | 0.25 | 0.24 | 0.59 | 0.65 | 1.49 | 1.13 | 2.63 | 0.7
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C558: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C722: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C758: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C776: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C812: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C830: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C867: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C885: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C903: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C921: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C939: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C957: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C975: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C993: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1012: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1086: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 712 (Indicative Value 0.164 wt.%) | [21 empty] | Indicative
H2O- | Round | Replicate | Lab | [19 empty] | 1
| No | No | 20 | [19 empty] | wt.%
|  |  | BF*XRF | [19 empty] | 3
3
| 1 | 1 | 0.1744 | [19 empty] | 1
| 1 | 2 | 0.1791 | [19 empty] | 14
| 1 | 3 | 0.1232 | [19 empty] | 16
| 1 | 4 | 0.1746 | [19 empty] | 0.16385
| 1 | 5 | 0.1606 | [19 empty] | 20
| 1 | 6 | 0.1712
| Mean |  | 0.16385
| Median |  | 0.1728
| Std Dev. |  | 0.020864
| Rel.Std.Dev. |  | 12.73%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 712 (Certified Value -0.616 wt.%) | [21 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*Furnace | LOI*TGA | OxBF*XRF | BF*XRF | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | BF*XRF | LOI*Furnace | LOI*TGA | LOI*TGA | 3
3
| 1 | 1 | -0.01 | < 0.01 | -0.79 | -0.67 | -0.6685 | -0.73 | -0.49 | < 0.01 | -0.54 | -0.52717 | < 0.01 | -0.68 | < 0.01 | -0.43895 | -0.6 | < 0.5 | -0.5767 | -0.63 | -0.7 | -0.9 | 1
| 1 | 2 | -0.01 | < 0.01 | -0.49 | -0.6 | -0.6481 | -0.7 | -0.48 | < 0.01 | -0.56 | -0.53815 | < 0.01 | -0.67 | < 0.01 | -0.42946 | -0.57 | < 0.5 | -0.5821 | -0.56 | -0.7 | -0.93 | #N/A
| 1 | 3 | -0.01 | < 0.01 | -0.82 | -0.69 | -0.6682 | -0.79 | -0.5 | < 0.01 | -0.54 | -0.51222 | < 0.01 | -0.68 | < 0.01 | -0.46948 | -0.57 | < 0.5 | -0.5853 | -0.64 | -0.7 | -0.94 | 16
| 1 | 4 | -0.01 | < 0.01 | -0.79 | -0.67 | -0.6765 | -0.77 | -0.53 | < 0.01 | -0.54 | -0.52508 | < 0.01 | -0.68 | < 0.01 | -0.42411 | -0.58 | < 0.5 | -0.5867 | -0.59 | -0.69 | -0.79 | -0.61608
| 1 | 5 | -0.01 | < 0.01 | -0.72 | -0.31 | -0.6767 | -0.8 | -0.53 | < 0.01 | -0.53 | -0.53714 | < 0.01 | -0.69 | < 0.01 | -0.51902 | -0.56 | < 0.5 | -0.5752 | -0.6 | -0.7 | -0.91 | 32
| 1 | 6 | -0.01 | < 0.01 | -0.81 | -0.62 | -0.6662 | -0.79 | -0.49 | < 0.01 | -0.58 | -0.55635 | < 0.01 | -0.68 | < 0.01 | -0.50921 | -0.56 | < 0.5 | -0.5776 | -0.65 | -0.69 | -0.87
| Mean |  | -0.01 |  | -0.73667 | -0.59333 | -0.66737 | -0.76333 | -0.50333 |  | -0.54833 | -0.53268 |  | -0.68 |  | -0.46504 | -0.57333 |  | -0.5806 | -0.61167 | -0.69667 | -0.89
| Median |  | -0.01 |  | -0.79 | -0.645 | -0.66835 | -0.78 | -0.495 |  | -0.54 | -0.53215 |  | -0.68 |  | -0.45422 | -0.57 |  | -0.57985 | -0.615 | -0.7 | -0.905
| Std Dev. |  | 0 |  | 0.1258 | 0.14292 | 0.010441 | 0.039833 | 0.021602 |  | 0.018348 | 0.014945 |  | 0.0063246 |  | 0.041246 | 0.015055 |  | 0.0047942 | 0.034303 | 0.005164 | 0.054772
| Rel.Std.Dev. |  | 0% |  | -17.08% | -24.09% | -1.565% | -5.218% | -4.292% |  | -3.346% | -2.806% |  | -0.9301% |  | -8.869% | -2.626% |  | -0.8257% | -5.608% | -0.7412% | -6.154%
| PDM3 |  | -98.38% |  | 19.57% | -3.692% | 8.325% | 23.9% | -18.3% |  | -11% | -13.54% |  | 10.38% |  | -24.52% | -6.938% |  | -5.759% | -0.7161% | 13.08% | 44.46%
| Z-Score (Absolute) |  | 2.91 | 2.99 | 0.93 | 0.17 | 0.56 | 1.07 | 0.3 | 2.99 | 0.07 | 0.15 | 2.99 | 0.63 | 2.99 | 0.51 | 0.07 | 4.29 | 0.1 | 0.27 | 0.72 | 1.74
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 712 (Indicative Value 0.08 ppm) | [20 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 3
3
| 1 | 1 | < 0.3 | 0.07 | 0.11 | < 0.1 | 0.26 | 0.5 | < 0.1 | < 0.1 | < 1 | < 0.5 | 0.08 | < 0.1 | 0.12 | 0.04 | 0.04 | < 0.1 |  |  |  | 1
| 1 | 2 | < 0.3 | 0.05 | 0.06 | < 0.1 | 0.26 | 0.5 | < 0.1 | 0.1 | < 1 | < 0.5 | 2.26 | < 0.1 | 0.12 | 0.04 | 0.02 | < 0.1 |  |  |  | 16
| 1 | 3 | < 0.3 | 0.11 | 0.08 | < 0.1 | 0.27 | 0.8 | < 0.1 | 0.1 | < 1 | < 0.5 | < 0.05 | < 0.1 | 0.09 | 0.06 | 0.16 | < 0.1 |  |  |  | 16
| 1 | 4 | < 0.3 | 0.07 | 0.06 | < 0.1 | 0.26 | 0.5 | < 0.1 | < 0.1 | < 1 | < 0.5 | < 0.05 | < 0.1 | 0.09 | < 0.01 | 0.07 | < 0.1 |  |  |  | 0.08
| 1 | 5 | < 0.3 | 0.09 | 0.12 | < 0.1 | 0.26 | 0.8 | < 0.1 | < 0.1 | < 1 | < 0.5 | 0.13 | < 0.1 | 0.1 | 0.02 | 0.09 | < 0.1 |  |  |  | 22
| 1 | 6 | < 0.3 | 0.07 | 0.07 | < 0.1 | 0.26 | 0.6 | < 0.1 | < 0.1 | < 1 | < 0.5 | 0.06 | < 0.1 | 0.11 | < 0.01 | 0.01 | < 0.1
| Mean |  |  | 0.076667 | 0.083333 |  | 0.26167 | 0.61667 |  | 0.1 |  |  | 0.6325 |  | 0.105 | 0.04 | 0.065
| Median |  |  | 0.07 | 0.075 |  | 0.26 | 0.55 |  | 0.1 |  |  | 0.105 |  | 0.105 | 0.04 | 0.055
| Std Dev. |  |  | 0.020656 | 0.02582 |  | 0.0040825 | 0.1472 |  | 0 |  |  | 1.0854 |  | 0.013784 | 0.01633 | 0.055408
| Rel.Std.Dev. |  |  | 26.94% | 30.98% |  | 1.56% | 23.87% |  | 0% |  |  | 171.6% |  | 13.13% | 40.82% | 85.24%
| PDM3 |  |  | -4.167% | 4.167% |  | 227.1% | 670.8% |  | 25% |  |  | 690.6% |  | 31.25% | -50% | -18.75%
| Z-Score (Absolute) |  | 1.57 | 0.07 | 0.07 | 0.67 | 4.08 | 12.06 | 0.67 | 0.3 | 9.44 | 3.82 | 7.87 | 0.67 | 0.56 | 1.16 | 0.34 | 0.67
| Analytical results for Al in OREAS 712 (Certified Value 5.54 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 5.2377 | 5.7795 | 5.33 | 5.22 | 5.44 | 5.66 | 6.09 | 5.67 | 5.75 | 5.3806 | 5.6 | 5.5 | 5.7888 | 5.54 | 5.55 | 5.47 | 5.44 | 5.5912 |  | 1
| 1 | 2 | 5.2655 | 5.7321 | 5.53 | 5.15 | 5.45 | 5.72 | 5.5 | 5.54 | 5.91 | 5.3153 | 5.69 | 5.55 | 5.9274 | 5.79 | 5.51 | 5.47 | 5.38 | 5.5094 |  | #N/A
| 1 | 3 | 5.2935 | 5.7243 | 5.49 | 5.29 | 5.49 | 5.66 | 5.14 | 5.83 | 5.63 | 5.3368 | 5.55 | 5.73 | 5.7233 | 5.47 | 5.42 | 5.49 | 5.35 | 5.5799 |  | 16
| 1 | 4 | 5.232 | 5.799 | 5.4 | 5.14 | 5.42 | 5.7 | 5.46 | 5.5 | 5.75 | 5.4285 | 5.59 | 5.61 | 5.7649 | 5.65 | 5.34 | 5.34 | 5.45 | 5.5264 |  | 5.5361
| 1 | 5 | 5.1851 | 5.7828 | 5.48 | 5.35 | 5.4 | 5.71 | 5.42 | 5.68 | 5.78 | 5.306 | 5.61 | 5.56 | 5.8848 | 5.67 | 5.73 | 5.41 | 5.46 | 5.5791 |  | 34
| 1 | 6 | 5.2517 | 5.7264 | 5.48 | 5.36 | 5.42 | 5.64 | 6.14 | 5.58 | 5.73 | 5.3836 | 5.67 | 5.57 | 5.761 | 5.75 | 5.34 | 5.45 | 5.45 | 5.541
| Mean |  | 5.2443 | 5.7573 | 5.4517 | 5.2517 | 5.4367 | 5.6817 | 5.625 | 5.6333 | 5.7583 | 5.3584 | 5.6183 | 5.5867 | 5.8084 | 5.645 | 5.4817 | 5.4383 | 5.4217 | 5.5545
| Median |  | 5.2447 | 5.7558 | 5.48 | 5.255 | 5.43 | 5.68 | 5.48 | 5.625 | 5.75 | 5.3587 | 5.605 | 5.565 | 5.7768 | 5.66 | 5.465 | 5.46 | 5.445 | 5.5601
| Std Dev. |  | 0.036399 | 0.03337 | 0.073052 | 0.096626 | 0.031411 | 0.032506 | 0.40039 | 0.11961 | 0.090425 | 0.047141 | 0.052313 | 0.078655 | 0.079703 | 0.12194 | 0.14905 | 0.055287 | 0.04535 | 0.033476
| Rel.Std.Dev. |  | 0.6941% | 0.5796% | 1.34% | 1.84% | 0.5778% | 0.5721% | 7.118% | 2.123% | 1.57% | 0.8798% | 0.9311% | 1.408% | 1.372% | 2.16% | 2.719% | 1.017% | 0.8365% | 0.6027%
| PDM3 |  | -5.272% | 3.996% | -1.525% | -5.138% | -1.796% | 2.629% | 1.606% | 1.756% | 4.014% | -3.209% | 1.485% | 0.9133% | 4.918% | 1.967% | -0.9833% | -1.766% | -2.067% | 0.3323%
| Z-Score (Absolute) |  | 1.75 | 1 | 0.64 | 1.71 | 0.72 | 0.6 | 0.29 | 0.34 | 1.01 | 1.14 | 0.26 | 0.09 | 1.28 | 0.4 | 0.48 | 0.71 | 0.8 | 0.09
| Analytical results for As in OREAS 712 (Indicative Value 2.84 ppm) | [20 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 20.6 | 17.56 | 2.5 | 2.2 | < 1 | 12.49 | 0.6 | 2 | < 5 | < 10 | 5 | 11 | 6.1 | 2.4 | 2.2 | 2.6 | 7.4731 |  |  | 1
| 1 | 2 | 16.7 | 17.46 | 1.9 | 2.1 | < 1 | 12.33 | < 0.1 | 2 | < 5 | < 10 | 4 | 11 | 5.9 | 2.4 | 2.1 | 2.3 | 1.4168 |  |  | 17
| 1 | 3 | 18.6 | 18.11 | 1.9 | 2.1 | < 1 | 12.72 | 0.3 | 3 | < 5 | < 10 | 4 | 11 | 6.2 | 2.1 | 2.1 | 2.4 | 6.3625 |  |  | 16
| 1 | 4 | 19.5 | 17.77 | 1.8 | 1.6 | < 1 | 12.27 | 1.3 | 2 | < 5 | < 10 | 4 | 11 | 6 | 2 | 2 | 2.4 | 6.6915 |  |  | 2.8376
| 1 | 5 | 19 | 16.72 | 2.4 | 2.1 | < 1 | 12.79 | 0.1 | 2 | < 5 | < 10 | 4 | 12 | 6.3 | 2.3 | 2.4 | 2.1 | 2.5065 |  |  | 23
| 1 | 6 | 20.6 | 17.31 | 2.3 | 2.2 | < 1 | 12.92 | 0.5 | 2 | < 5 | < 10 | 4 | 12 | 6.3 | 2.3 | 2.1 | 2.2 | 1.9034
| Mean |  | 19.1667 | 17.4883 | 2.1333 | 2.05 |  | 12.5867 | 0.56 | 2.1667 |  |  | 4.1667 | 11.3333 | 6.1333 | 2.25 | 2.15 | 2.3333 | 4.3923
| Median |  | 19.25 | 17.51 | 2.1 | 2.1 |  | 12.605 | 0.5 | 2 |  |  | 4 | 11 | 6.15 | 2.3 | 2.1 | 2.35 | 4.4345
| Std Dev. |  | 1.4597 | 0.46774 | 0.30111 | 0.22583 |  | 0.26296 | 0.45607 | 0.40825 |  |  | 0.40825 | 0.5164 | 0.1633 | 0.16432 | 0.13784 | 0.17512 | 2.73
| Rel.Std.Dev. |  | 7.616% | 2.675% | 14.11% | 11.02% |  | 2.089% | 81.44% | 18.84% |  |  | 9.798% | 4.556% | 2.662% | 7.303% | 6.411% | 7.505% | 62.15%
| PDM3 |  | 575.5% | 516.3% | -24.82% | -27.75% |  | 343.6% | -80.26% | -23.64% |  |  | 46.84% | 299.4% | 116.1% | -20.71% | -24.23% | -17.77% | 54.79%
| Z-Score (Absolute) |  | 5.94 | 5.34 | 0.13 | 0.16 | 0.71 | 3.59 | 0.72 | 0.12 | 0 | 0.89 | 0.59 | 3.15 | 1.29 | 0.09 | 0.12 | 0.06 | 0.67
| Analytical results for Ba in OREAS 712 (Certified Value 165 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 151 | 171.9055 | 170 | 170 | 172.9 | 162 | 171 | 175 | 161 | 160.8389 | 161 | 162 | 175.4668 | 156 | 152.7345 | 180 | 170 | 170 | 169.9362 | 1
| 1 | 2 | 151 | 173.1535 | 170 | 170 | 172.6 | 162 | 168 | 170 | 165 | 157.9587 | 164 | 164 | 179.8118 | 162 | 156.6811 | 180 | 170 | 170 | 168.3507 | #N/A
| 1 | 3 | 151 | 171.999 | 170 | 170 | 177.7 | 158 | 159 | 176 | 158 | 159.8129 | 162 | 165 | 174.5538 | 153 | 159.7936 | 170 | 170 | 170 | 172.9056 | 16
| 1 | 4 | 152 | 172.873 | 170 | 170 | 169.4 | 162 | 176 | 172 | 158 | 160.8087 | 162 | 167 | 174.9278 | 158 | 159.3173 | 170 | 170 | 170 | 168.2947 | 164.8815
| 1 | 5 | 149 | 173.823 | 170 | 170 | 171.7 | 160 | 164 | 171 | 158 | 157.6943 | 162 | 168 | 176.0388 | 154 | 154.7247 | 190 | 170 | 170 | 177.042 | 35
| 1 | 6 | 150 | 173.1535 | 170 | 180 | 177.2 | 155 | 176 | 175 | 165 | 161.411 | 162 | 169 | 173.7838 | 157 | 156.4695 | 170 | 170 | 170 | 173.2841
| Mean |  | 150.6667 | 172.8179 | 170 | 171.6667 | 173.5833 | 159.8333 | 169 | 173.1667 | 160.8333 | 159.7541 | 162.1667 | 165.8333 | 175.7638 | 156.6667 | 156.6201 | 176.6667 | 170 | 170 | 171.6355
| Median |  | 151 | 173.0132 | 170 | 170 | 172.75 | 161 | 169.5 | 173.5 | 159.5 | 160.3108 | 162 | 166 | 175.1973 | 156.5 | 156.5753 | 175 | 170 | 170 | 171.4209
| Std Dev. |  | 1.0328 | 0.74048 | 0 | 4.0825 | 3.2406 | 2.8577 | 6.7528 | 2.4833 | 3.4303 | 1.5814 | 0.98319 | 2.6394 | 2.1276 | 3.2042 | 2.6856 | 8.165 | 0 | 0 | 3.4176
| Rel.Std.Dev. |  | 0.6855% | 0.4285% | 0% | 2.378% | 1.867% | 1.788% | 3.996% | 1.434% | 2.133% | 0.9899% | 0.6063% | 1.592% | 1.21% | 2.045% | 1.715% | 4.622% | 0% | 0% | 1.991%
| PDM3 |  | -8.621% | 4.813% | 3.104% | 4.115% | 5.278% | -3.062% | 2.498% | 5.025% | -2.455% | -3.11% | -1.647% | 0.5773% | 6.6% | -4.982% | -5.011% | 7.148% | 3.104% | 3.104% | 4.096%
| Z-Score (Absolute) |  | 1.22 | 0.81 | NA | NA | 0.88 | 0.38 | 0.46 | 0.84 | 0.29 | 0.39 | 0.17 | 0.17 | 1.08 | 0.67 | 0.68 | NA | NA | NA | 0.7
Results from laboratories 3, 4, 21, 22 and 23 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 712 (Certified Value 1.02 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.71 | 1.07 | 0.94 | < 2 | 0.91 | 1.1 | 1.1 | 1 | 1.0374 | 1 | 0.94 | 1.2333 | < 1 | 0.99806 | 1.22 | 0.6 | 1.35 | 0.8841 |  | 1
| 1 | 2 | 1.65 | 1.13 | 0.94 | < 2 | 0.9 | 1 | 1 | 0.9 | 0.90081 | 1 | 0.98 | 1.1892 | < 1 | 1.0289 | 1.2 | 0.57 | 1.36 | 0.8445 |  | #N/A
| 1 | 3 | 1.66 | 1.11 | 0.94 | < 2 | 0.89 | 1 | 1 | 0.9 | 1.016 | 1 | 0.96 | 1.1874 | < 1 | 1.0517 | 1.18 | 0.58 | 1.46 | 0.8476 |  | 16
| 1 | 4 | 1.66 | 1.08 | 0.93 | < 2 | 0.91 | 1 | 1.1 | 1 | 0.94784 | 1.1 | 0.94 | 1.0409 | < 1 | 1.0401 | 1.13 | 0.59 | 1.36 | 0.8952 |  | 1.0188
| 1 | 5 | 1.69 | 1.09 | 0.98 | < 2 | 0.94 | 1.1 | 1.1 | 1.2 | 1.0185 | 1.1 | 0.96 | 1.0176 | < 1 | 1.0044 | 1.26 | 0.59 | 1.43 | 0.8317 |  | 36
| 1 | 6 | 1.64 | 1.12 | 1 | < 2 | 0.91 | 1 | 1.1 | 0.9 | 1.0071 | 1 | 1.03 | 1.1045 | < 1 | 0.95918 | 1.18 | 0.59 | 1.37 | 0.9069
| Mean |  | 1.6683 | 1.1 | 0.955 |  | 0.91 | 1.0333 | 1.0667 | 0.98333 | 0.98794 | 1.0333 | 0.96833 | 1.1288 |  | 1.0137 | 1.195 | 0.58667 | 1.3883 | 0.86833
| Median |  | 1.66 | 1.1 | 0.94 |  | 0.91 | 1 | 1.1 | 0.95 | 1.0116 | 1 | 0.96 | 1.1459 |  | 1.0167 | 1.19 | 0.59 | 1.365 | 0.86585
| Std Dev. |  | 0.026394 | 0.023664 | 0.028107 |  | 0.016733 | 0.05164 | 0.05164 | 0.1169 | 0.052415 | 0.05164 | 0.033714 | 0.087978 |  | 0.033703 | 0.043704 | 0.010328 | 0.04535 | 0.030976
| Rel.Std.Dev. |  | 1.582% | 2.151% | 2.943% |  | 1.839% | 4.997% | 4.841% | 11.89% | 5.305% | 4.997% | 3.482% | 7.794% |  | 3.325% | 3.657% | 1.76% | 3.267% | 3.567%
| PDM3 |  | 63.76% | 7.975% | -6.258% |  | -10.68% | 1.431% | 4.703% | -3.477% | -3.024% | 1.431% | -4.949% | 10.8% |  | -0.493% | 17.3% | -42.41% | 36.28% | -14.77%
| Z-Score (Absolute) |  | 5.83 | 0.82 | 0.46 | 0.06 | 0.85 | 0.23 | 0.53 | 0.21 | 0.17 | 0.23 | 0.34 | 1.08 | 4.47 | 0.06 | 1.66 | 3.71 | 3.36 | 1.22
| Analytical results for Bi in OREAS 712 (Certified Value 0.25 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 0.24 | 0.21 | 0.24 | < 0.5 | 1.42 | 0.27 | 0.23 | 0.3 | < 3.5 | 0.3 | 0.31 | 0.2 | 0.24631 | 0.26 | 0.25 | 0.24 |  |  |  | 1
| 1 | 2 | 0.25 | 0.26 | 0.24 | < 0.5 | 1.5 | 0.25 | 0.23 | 0.3 | < 3.5 | 0.3 | 0.26 | 0.2 | 0.24788 | 0.26 | 0.23 | 0.26 | < 0.1 |  |  | #N/A
| 1 | 3 | 0.25 | 0.24 | 0.25 | < 0.5 | 1.44 | 0.24 | 0.25 | 0.3 | < 3.5 | 0.3 | 0.25 | 0.2 | 0.24181 | 0.25 | 0.24 | 0.23 | 2.4617 |  |  | 16
| 1 | 4 | 0.34 | 0.24 | 0.24 | < 0.5 | 1.47 | 0.28 | 0.23 | 0.3 | < 3.5 | 0.3 | 0.24 | 0.2 | 0.25205 | 0.24 | 0.23 | 0.26 | < 0.1 |  |  | 0.24636
| 1 | 5 | 0.23 | 0.25 | 0.24 | < 0.5 | 1.56 | 0.25 | 0.24 | 0.3 | < 3.5 | 0.3 | 0.24 | 0.1 | 0.26325 | 0.27 | 0.26 | 0.25 | 1.708 |  |  | 37
| 1 | 6 | 0.23 | 0.26 | 0.26 | < 0.5 | 1.55 | 0.28 | 0.23 | 0.3 | < 3.5 | 0.3 | 0.24 | 0.2 | 0.2544 | 0.25 | 0.22 | 0.25 | 1.8707
| Mean |  | 0.25667 | 0.24333 | 0.245 |  | 1.49 | 0.26167 | 0.235 | 0.3 |  | 0.3 | 0.25667 | 0.18333 | 0.25095 | 0.255 | 0.23833 | 0.24833 | 2.0135
| Median |  | 0.245 | 0.245 | 0.24 |  | 1.485 | 0.26 | 0.23 | 0.3 |  | 0.3 | 0.245 | 0.2 | 0.24996 | 0.255 | 0.235 | 0.25 | 1.8707
| Std Dev. |  | 0.041793 | 0.018619 | 0.0083666 |  | 0.057271 | 0.017224 | 0.0083666 | 0 |  | 0 | 0.027325 | 0.040825 | 0.0074686 | 0.010488 | 0.01472 | 0.01169 | 0.39661
| Rel.Std.Dev. |  | 16.28% | 7.652% | 3.415% |  | 3.844% | 6.582% | 3.56% | 0% |  | 0% | 10.65% | 22.27% | 2.976% | 4.113% | 6.176% | 4.708% | 19.7%
| PDM3 |  | 4.183% | -1.229% | -0.5527% |  | 504.8% | 6.212% | -4.612% | 21.77% |  | 21.77% | 4.183% | -25.58% | 1.862% | 3.506% | -3.259% | 0.8003% | 717.3%
| Z-Score (Absolute) |  | 0.11 | 0.79 | 0.67 | 0.34 | 83.28 | 0.45 | 1.35 | NA | 100.81 | 3.03 | 0.11 | NA | 0.27 | 0 | 1.12 | 0.45 | 65.62
Results from laboratories 13 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 712 (Certified Value 3.05 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 3.0915 | 3.0947 | 2.99 | 3.05 | 3.26 | 2.97 | 3 | 3.06 | 3.19 | 2.9305 | 3.03 | 2.978 | 3.2462 | 3.08 | 2.99 | 3.11 | 3.05 | 2.9284 |  | 1
| 1 | 2 | 3.0757 | 3.0906 | 3.11 | 2.96 | 3.22 | 3.03 | 3.18 | 3.02 | 3.26 | 2.8463 | 3.07 | 2.99 | 3.2945 | 3.04 | 3.04 | 3.14 | 3.05 | 2.8707 |  | #N/A
| 1 | 3 | 3.0893 | 3.088 | 3.09 | 3.07 | 3.27 | 2.9 | 3.05 | 3.19 | 3.13 | 2.8733 | 3.01 | 3.07 | 3.2291 | 3.09 | 2.97 | 3.06 | 3.05 | 2.8971 |  | 16
| 1 | 4 | 3.0858 | 3.0889 | 3.04 | 2.97 | 3.28 | 2.9 | 2.97 | 3.01 | 3.19 | 2.8818 | 3.04 | 2.997 | 3.2502 | 3.08 | 2.9 | 3 | 3.05 | 2.8987 |  | 3.0485
| 1 | 5 | 3.0598 | 3.0895 | 3.09 | 3.11 | 3.21 | 3 | 3.2 | 3.07 | 3.2 | 2.8508 | 3.05 | 2.973 | 3.2484 | 3.09 | 3.18 | 3.04 | 3.12 | 2.9057 |  | 38
| 1 | 6 | 3.0668 | 3.0909 | 3.13 | 3.1 | 3.21 | 3 | 3.03 | 3.06 | 3.18 | 2.8459 | 3.07 | 2.986 | 3.2342 | 3.1 | 2.92 | 3.05 | 3.1 | 2.9204
| Mean |  | 3.0782 | 3.0904 | 3.075 | 3.0433 | 3.2417 | 2.9667 | 3.0717 | 3.0683 | 3.1917 | 2.8714 | 3.045 | 2.999 | 3.2504 | 3.08 | 3 | 3.0667 | 3.07 | 2.9035
| Median |  | 3.0808 | 3.09 | 3.09 | 3.06 | 3.24 | 2.985 | 3.04 | 3.06 | 3.19 | 2.8621 | 3.045 | 2.988 | 3.2473 | 3.085 | 2.98 | 3.055 | 3.05 | 2.9022
| Std Dev. |  | 0.012904 | 0.0023492 | 0.051284 | 0.064395 | 0.031885 | 0.055015 | 0.095795 | 0.064317 | 0.041673 | 0.032577 | 0.023452 | 0.035811 | 0.023163 | 0.020976 | 0.10139 | 0.050465 | 0.031623 | 0.02025
| Rel.Std.Dev. |  | 0.4192% | 0.07601% | 1.668% | 2.116% | 0.9836% | 1.854% | 3.119% | 2.096% | 1.306% | 1.135% | 0.7702% | 1.194% | 0.7126% | 0.681% | 3.38% | 1.646% | 1.03% | 0.6974%
| PDM3 |  | 0.9741% | 1.377% | 0.8708% | -0.168% | 6.338% | -2.683% | 0.7614% | 0.6521% | 4.698% | -5.807% | -0.1133% | -1.622% | 6.625% | 1.035% | -1.589% | 0.5974% | 0.7067% | -4.755%
| Z-Score (Absolute) |  | 0.24 | 0.57 | 0.16 | 0.7 | 4.65 | 2.76 | 0.07 | 0.02 | 3.3 | 5.33 | 0.65 | 1.89 | 4.89 | 0.29 | 1.87 | 0.07 | 0.02 | 4.47
| Analytical results for Cd in OREAS 712 (Certified Value 0.29 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 0.29 | 0.29 | 0.3 | 0.43 | 0.32 | < 0.3 | 0.29 | 0.3 | 0.24072 | < 0.5 | 0.15 | 0.3 | 0.31551 | 0.32 | 0.31 | 0.28 | < 0.1 |  |  | 1
| 1 | 2 | 0.31 | 0.27 | 0.28 | 0.45 | 0.31 | 1 | 0.3 | 0.3 | 0.18954 | < 0.5 | 0.16 | 0.3 | 0.30013 | 0.3 | 0.28 | 0.28 | < 0.1 |  |  | 2
| 1 | 3 | 0.32 | 0.25 | 0.29 | 0.42 | 0.31 | < 0.3 | 0.3 | 0.3 | 0.20582 | < 0.5 | 0.16 | 0.3 | 0.3129 | 0.28 | 0.31 | 0.27 | < 0.1 |  |  | 16
| 1 | 4 | 0.29 | 0.28 | 0.28 | 0.5 | 0.32 | < 0.3 | 0.3 | 0.4 | 0.2068 | < 0.5 | 0.14 | 0.3 | 0.27933 | 0.29 | 0.28 | 0.29 | < 0.1 |  |  | 0.29368
| 1 | 5 | 0.3 | 0.28 | 0.29 | 0.49 | 0.3 | < 0.3 | 0.3 | 0.3 | 0.20392 | < 0.5 | 0.15 | 0.3 | 0.2693 | 0.31 | 0.28 | 0.29 | < 0.1 |  |  | 39
| 1 | 6 | 0.3 | 0.28 | 0.29 | 0.44 | 0.31 | < 0.3 | 0.29 | 0.5 | 0.20741 | < 0.5 | 0.15 | 0.3 | 0.30161 | 0.28 | 0.27 | 0.29 | < 0.1
| Mean |  | 0.30167 | 0.275 | 0.28833 | 0.455 | 0.31167 | 1 | 0.29667 | 0.35 | 0.20903 |  | 0.15167 | 0.3 | 0.29646 | 0.29667 | 0.28833 | 0.28333
| Median |  | 0.3 | 0.28 | 0.29 | 0.445 | 0.31 | 1 | 0.3 | 0.3 | 0.20631 |  | 0.15 | 0.3 | 0.30087 | 0.295 | 0.28 | 0.285
| Std Dev. |  | 0.01169 | 0.013784 | 0.0075277 | 0.032711 | 0.0075277 |  | 0.005164 | 0.083666 | 0.016899 |  | 0.0075277 | 0 | 0.018462 | 0.01633 | 0.017224 | 0.008165
| Rel.Std.Dev. |  | 3.875% | 5.012% | 2.611% | 7.189% | 2.415% |  | 1.741% | 23.9% | 8.084% |  | 4.963% | 0% | 6.227% | 5.504% | 5.974% | 2.882%
| PDM3 |  | 2.719% | -6.361% | -1.821% | 54.93% | 6.124% | 240.5% | 1.017% | 19.18% | -28.82% |  | -48.36% | 2.152% | 0.9474% | 1.017% | -1.821% | -3.523%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0 | 8.43 | 1.18 | 0.17 | 0.42 | NA | 4.01 | 1.94 | 6.91 | NA | 0.41 | 0.42 | 0 | 0.25 | 12.05
Results from laboratories 13 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 712 (Certified Value 468 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 401.1 | 737.7124 | 496 | 468 | 463 | 466 | 446.9 | 453 | 481 | 467.0913 | 485 | 449.9301 | 529 | 460 | 480 | 476.5281 |  |  |  | 1
| 1 | 2 | 399.2 | 735.5213 | 506 | 451 | 451 | 475 | 463.2 | 463 | 480 | 464.8779 | 485 | 450.2979 | 499 | 462 | 504 | 462.2981 |  |  |  | 20
| 1 | 3 | 399.4 | 735.0579 | 469 | 474 | 457 | 443 | 447.6 | 450 | 481 | 476.326 | 461 | 448.7889 | 466 | 469 | 475 | 442.6465 |  |  |  | 16
| 1 | 4 | 402.4 | 733.5722 | 459 | 460 | 456 | 453 | 452 | 452 | 475 | 493.1027 | 476 | 447.9697 | 486 | 454 | 501 | 465.4829 |  |  |  | 468.004
| 1 | 5 | 394.9 | 737.0915 | 495 | 473 | 464 | 468 | 448.5 | 457 | 479 | 453.5074 | 477 | 455.229 | 515 | 459 | 502 | 435.9435 |  |  |  | 40
| 1 | 6 | 393 | 736.8226 | 484 | 467 | 456 | 470 | 452.3 | 458 | 482 | 442.4525 | 498 | 459.4592 | 474 | 456 | 486 | 456.907
| Mean |  | 398.3333 | 735.963 | 484.8333 | 465.5 | 457.8333 | 462.5 | 451.75 | 455.5 | 479.6667 | 466.2263 | 480.3333 | 451.9458 | 494.8333 | 460 | 491.3333 | 456.6344
| Median |  | 399.3 | 736.172 | 489.5 | 467.5 | 456.5 | 467 | 450.25 | 455 | 480.5 | 465.9846 | 481 | 450.114 | 492.5 | 459.5 | 493.5 | 459.6025
| Std Dev. |  | 3.6418 | 1.5364 | 17.8596 | 8.6891 | 4.8751 | 12.0457 | 6.0461 | 4.7645 | 2.5033 | 17.6288 | 12.3234 | 4.4673 | 24.2274 | 5.2536 | 12.5804 | 15.0332
| Rel.Std.Dev. |  | 0.9143% | 0.2088% | 3.684% | 1.867% | 1.065% | 2.604% | 1.338% | 1.046% | 0.5219% | 3.781% | 2.566% | 0.9885% | 4.896% | 1.142% | 2.56% | 3.292%
| PDM3 |  | -14.89% | 57.26% | 3.596% | -0.535% | -2.173% | -1.176% | -3.473% | -2.672% | 2.492% | -0.3799% | 2.634% | -3.431% | 5.733% | -1.71% | 4.985% | -2.429%
| Z-Score (Absolute) |  | 3.64 | 15.09 | 1.16 | 0.08 | 0.34 | 0.08 | 0.68 | 0.47 | 0.87 | 0.12 | 0.91 | 0.67 | 1.71 | 0.22 | 1.52 | 0.41
| Analytical results for Co in OREAS 712 (Certified Value 11.8 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 60.3 | 11.8795 | 10.5 | 11 | 133.1 | 12.59 | 12 | 12 | 12 | 11.1676 | 45 | 11.9 | 13.0157 | 10.4 | 11.1898 | 13.8 | 11.9 | 11.6 | 13.1987 | 1
| 1 | 2 | 60.3 | 11.7777 | 10.6 | 10.5 | 130.5 | 12.08 | 12 | 12 | 12 | 11.2759 | 45 | 12.4 | 12.6238 | 10.2 | 11.4233 | 13.6 | 12 | 11.8 | 12.924 | #N/A
| 1 | 3 | 60.5 | 11.8187 | 10.4 | 11.2 | 133.3 | 12.37 | 12 | 12 | 11 | 11.3329 | 45 | 12.1 | 13.0973 | 11.3 | 11.2635 | 12.4 | 12 | 11.6 | 12.7806 | 16
| 1 | 4 | 60.8 | 11.8146 | 10.4 | 10.7 | 136.9 | 11.89 | 12 | 12 | 12 | 11.7973 | 40 | 12.2 | 13.1166 | 11 | 11.4263 | 12.9 | 11.8 | 12 | 12.7488 | 11.8497
| 1 | 5 | 59.7 | 11.781 | 10.8 | 10.9 | 133.9 | 12.34 | 12 | 12 | 12 | 10.9769 | 40 | 12.4 | 12.4033 | 10.5 | 11.7809 | 13.5 | 11.9 | 11.9 | 13.5877 | 41
| 1 | 6 | 60.3 | 11.9034 | 10.4 | 10.9 | 132.3 | 11.99 | 12 | 12 | 12 | 11.2122 | 40 | 12.8 | 12.5145 | 11.4 | 11.5474 | 12.9 | 12 | 11.8 | 13.3358
| Mean |  | 60.3167 | 11.8292 | 10.5167 | 10.8667 | 133.3333 | 12.21 | 12 | 12 | 11.8333 | 11.2938 | 42.5 | 12.3 | 12.7952 | 10.8 | 11.4385 | 13.1833 | 11.9333 | 11.7833 | 13.0959
| Median |  | 60.3 | 11.8166 | 10.45 | 10.9 | 133.2 | 12.21 | 12 | 12 | 12 | 11.2441 | 42.5 | 12.3 | 12.8197 | 10.75 | 11.4248 | 13.2 | 11.95 | 11.8 | 13.0614
| Std Dev. |  | 0.36009 | 0.051648 | 0.16021 | 0.24221 | 2.1068 | 0.26631 | 0 | 0 | 0.40825 | 0.27508 | 2.7386 | 0.30984 | 0.31777 | 0.502 | 0.21068 | 0.53448 | 0.08165 | 0.16021 | 0.33449
| Rel.Std.Dev. |  | 0.597% | 0.4366% | 1.523% | 2.229% | 1.58% | 2.181% | 0% | 0% | 3.45% | 2.436% | 6.444% | 2.519% | 2.484% | 4.648% | 1.842% | 4.054% | 0.6842% | 1.36% | 2.554%
| PDM3 |  | 409% | -0.1733% | -11.25% | -8.296% | 1025% | 3.041% | 1.268% | 1.268% | -0.138% | -4.691% | 258.7% | 3.8% | 7.979% | -8.858% | -3.47% | 11.25% | 0.7059% | -0.56% | 10.52%
| Z-Score (Absolute) |  | 37.85 | 0.08 | 1.11 | 0.83 | 94.98 | 0.22 | NA | NA | NA | 0.5 | NA | 0.29 | 0.67 | 0.89 | 0.39 | 0.98 | 0 | 0.12 | 0.91
Results from laboratories 10, 12, 13 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 712 (Certified Value 132 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
0
| 1 | 1 | 72.7 | 147.871 | 128 | 131 | 115 | 126 | 91 | 122 | 145 | 132.1093 | 144 | 115 | 149.2464 | 141 | 129 | 134 | 135.6888 |  |  | 1
| 1 | 2 | 72.7 | 146.74 | 134 | 129 | 115.3 | 123 | 108 | 124 | 123 | 131.3796 | 144 | 118 | 149.6355 | 139 | 128 | 131 | 130.9927 |  |  | #N/A
| 1 | 3 | 73.9 | 146.223 | 131 | 131 | 114.3 | 124 | 115 | 124 | 142 | 133.0556 | 139 | 117 | 150.1869 | 134 | 127 | 133 | 122.8181 |  |  | 16
| 1 | 4 | 72.1 | 146.777 | 132 | 128 | 115.8 | 124 | 95 | 127 | 152 | 133.3124 | 147 | 116 | 150.1794 | 133 | 123 | 132 | 91.1878 |  |  | 132.2942
| 1 | 5 | 72.5 | 147.882 | 131 | 132 | 113.8 | 125 | 107 | 135 | 137 | 131.8359 | 150 | 112 | 152.9493 | 147 | 125 | 136 | 140.0913 |  |  | 42
| 1 | 6 | 71.9 | 148.1915 | 132 | 133 | 113.8 | 125 | 99 | 123 | 134 | 129.7423 | 139 | 114 | 151.2036 | 134 | 125 | 131 | 133.7077
| Mean |  | 72.6333 | 147.2807 | 131.3333 | 130.6667 | 114.6667 | 124.5 | 102.5 | 125.8333 | 138.8333 | 131.9059 | 143.8333 | 115.3333 | 150.5669 | 138 | 126.1667 | 132.8333 | 125.7477
| Median |  | 72.6 | 147.324 | 131.5 | 131 | 114.65 | 124.5 | 103 | 124 | 139.5 | 131.9726 | 144 | 115.5 | 150.1832 | 136.5 | 126 | 132.5 | 132.3502
| Std Dev. |  | 0.70048 | 0.80055 | 1.9664 | 1.8619 | 0.82865 | 1.0488 | 9.0277 | 4.7924 | 9.9883 | 1.2887 | 4.3551 | 2.1602 | 1.3407 | 5.4406 | 2.2286 | 1.9408 | 17.8796
| Rel.Std.Dev. |  | 0.9644% | 0.5436% | 1.497% | 1.425% | 0.7227% | 0.8424% | 8.808% | 3.808% | 7.194% | 0.977% | 3.028% | 1.873% | 0.8904% | 3.942% | 1.766% | 1.461% | 14.22%
| PDM3 |  | -45.1% | 11.33% | -0.7263% | -1.23% | -13.32% | -5.892% | -22.52% | -4.884% | 4.943% | -0.2936% | 8.722% | -12.82% | 13.81% | 4.313% | -4.632% | 0.4075% | -4.948%
| Z-Score (Absolute) |  | 5.34 | 1.53 | 0.06 | 0 | 1.47 | 0.57 | 2.59 | 0.44 | 0.75 | 0.11 | 1.21 | 1.41 | 1.83 | 0.67 | 0.41 | 0.2 | 0.45
| Analytical results for Cs in OREAS 712 (Certified Value 1.52 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.6 | 1.3447 | 1.46 | 1.42 | 1.43 | 1.46 | 1.6 | 1.5 | 1.6 | 1.53 | 1.5248 | 1.5 | 1.4883 | 1.74 | 1.45 | 1.58 | 1.5458 |  |  | 1
| 1 | 2 | 1.6 | 1.3484 | 1.49 | 1.4 | 1.38 | 1.49 | 1.6 | 1.5 | 1.6 | 1.53 | 1.5509 | 1.6 | 1.5308 | 1.7 | 1.42 | 1.52 | 1.5315 |  |  | 4
| 1 | 3 | 1.6 | 1.3701 | 1.4 | 1.43 | 1.41 | 1.47 | 1.6 | 1.5 | 1.6 | 1.58 | 1.6053 | 1.6 | 1.5417 | 1.61 | 1.46 | 1.51 | 1.5247 |  |  | 16
| 1 | 4 | 1.5 | 1.4332 | 1.41 | 1.4 | 1.49 | 1.53 | 1.6 | 1.5 | 1.6 | 1.57 | 1.5795 | 1.7 | 1.5322 | 1.64 | 1.41 | 1.56 | 1.5865 |  |  | 1.5176
| 1 | 5 | 1.5 | 1.2999 | 1.49 | 1.44 | 1.47 | 1.56 | 1.6 | 1.4 | 1.6 | 1.58 | 1.4838 | 1.7 | 1.5081 | 1.75 | 1.42 | 1.52 | 1.5094 |  |  | 43
| 1 | 6 | 1.5 | 1.3444 | 1.45 | 1.41 | 1.4 | 1.78 | 1.6 | 1.4 | 1.6 | 1.55 | 1.5014 | 1.9 | 1.5084 | 1.62 | 1.42 | 1.58 | 1.5349
| Mean |  | 1.55 | 1.3568 | 1.45 | 1.4167 | 1.43 | 1.5483 | 1.6 | 1.4667 | 1.6 | 1.5567 | 1.5409 | 1.6667 | 1.5183 | 1.6767 | 1.43 | 1.545 | 1.5388
| Median |  | 1.55 | 1.3466 | 1.455 | 1.415 | 1.42 | 1.51 | 1.6 | 1.5 | 1.6 | 1.56 | 1.5378 | 1.65 | 1.5196 | 1.67 | 1.42 | 1.54 | 1.5332
| Std Dev. |  | 0.054772 | 0.043854 | 0.038471 | 0.01633 | 0.042426 | 0.11957 | 2.4324e-16 | 0.05164 | 2.4324e-16 | 0.023381 | 0.046514 | 0.13663 | 0.020006 | 0.061536 | 0.02 | 0.032094 | 0.026291
| Rel.Std.Dev. |  | 3.534% | 3.232% | 2.653% | 1.153% | 2.967% | 7.722% | 1.52e-14% | 3.521% | 1.52e-14% | 1.502% | 3.019% | 8.198% | 1.318% | 3.67% | 1.399% | 2.077% | 1.709%
| PDM3 |  | 2.138% | -10.59% | -4.452% | -6.648% | -5.77% | 2.028% | 5.433% | -3.353% | 5.433% | 2.577% | 1.542% | 9.826% | 0.04659% | 10.48% | -5.77% | 1.808% | 1.4%
| Z-Score (Absolute) |  | 0.1 | 2.1 | 1.04 | 1.42 | 1.27 | 0.08 | 0.67 | 0.85 | 0.67 | 0.18 | 0 | 1.44 | 0.26 | 1.55 | 1.27 | 0.05 | 0.02
| Analytical results for Cu in OREAS 712 (Indicative Value 7.15 ppm) | [20 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
2
| 1 | 1 | 2.5 | 5.9 | 31 | 19.14 | 6.45 | 2 | 13 | 23 | 6.4528 | < 5 | 73 | 5.3585 | 6.4 | 6.1861 | 15.8 | 6.2 | 7.1 | 11.6797 |  | 1
| 1 | 2 | < 0.5 | 5.5 | 28.1 | 19.72 | 6.44 | 4 | 13 | 24 | 6.2575 | 5 | 88 | 5.9511 | 7.5 | 6.4047 | 14.5 | 5.6 | 5.2 | 16.2967 |  | 18
| 1 | 3 | 1.6 | 6.2 | 29.5 | 19.5 | 6.47 | 5 | 15 | 18 | 6.8584 | < 5 | 77 | 5.3856 | 7.9 | 6.6215 | 13.2 | 6.5 | 5.2 | 12.2703 |  | 16
| 1 | 4 | < 0.5 | 5.8 | 31.2 | 19.99 | 6.58 | 2 | 15 | 24 | 6.7822 | < 5 | 79 | 5.3059 | 6.9 | 6.4539 | 13.4 | 5.4 | 8.1 | 10.2405 |  | 7.1546
| 1 | 5 | 3.1 | 6.3 | 27.5 | 19.3 | 6.43 | 6 | 12 | 18 | 6.8848 | < 5 | 80 | 6.3119 | 6.9 | 6.3656 | 14.3 | 5.5 | 8.3 | 11.5354 |  | 24
| 1 | 6 | 3.3 | 6.3 | 27.5 | 19.83 | 6.55 | 1 | 12 | 21 | 6.6453 | < 5 | 87 | 6.1666 | 7.1 | 6.2821 | 14 | 4.4 | 4.3 | 10.8914
| Mean |  | 2.625 | 6 | 29.1333 | 19.58 | 6.4867 | 3.3333 | 13.3333 | 21.3333 | 6.6469 | 5 | 80.6667 | 5.7466 | 7.1167 | 6.3856 | 14.2 | 5.6 | 6.3667 | 12.1523
| Median |  | 2.8 | 6.05 | 28.8 | 19.61 | 6.46 | 3 | 13 | 22 | 6.7138 | 5 | 79.5 | 5.6684 | 7 | 6.3852 | 14.15 | 5.55 | 6.15 | 11.6075
| Std Dev. |  | 0.76322 | 0.32249 | 1.6908 | 0.32515 | 0.062823 | 1.9664 | 1.3663 | 2.8048 | 0.24856 |  | 5.8195 | 0.45009 | 0.52313 | 0.14939 | 0.93167 | 0.72938 | 1.6896 | 2.1466
| Rel.Std.Dev. |  | 29.07% | 5.375% | 5.804% | 1.661% | 0.9685% | 58.99% | 10.25% | 13.15% | 3.74% |  | 7.214% | 7.832% | 7.351% | 2.339% | 6.561% | 13.02% | 26.54% | 17.66%
| PDM3 |  | -63.31% | -16.14% | 307.2% | 173.7% | -9.336% | -53.41% | 86.36% | 198.2% | -7.097% | -30.11% | 1027% | -19.68% | -0.53% | -10.75% | 98.47% | -21.73% | -11.01% | 69.85%
| Z-Score (Absolute) |  | 0.93 | 0.11 | 4.42 | 2.55 | 0.02 | 0.63 | 1.33 | 2.89 | 0.02 | 0.72 | 14.52 | 0.16 | 0.11 | 0.04 | 1.5 | 0.19 | 0.04 | 1.09
| Analytical results for Dy in OREAS 712 (Certified Value 18 ppm) | [20 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 19.48 | 19.4138 | 17.59 | 17.8 | 16.6 | 18.3 | 17.64 | 16.6356 | 18.7104 | 18.2405 | [9 empty] | 1
| 1 | 2 | 20.24 | 19.3833 | 17.57 | 18.4 | 17.6 | 18.4 | 17.9 | 16.3911 | 19.086 | 17.0254 | [9 empty] | 5
| 1 | 3 | 19.65 | 19.7752 | 17.63 | 17.7 | 16.5 | 18.2 | 18.03 | 16.9095 | 18.8183 | 16.582 | [9 empty] | 16
| 1 | 4 | 19.36 | 19.2088 | 17.63 | 17.7 | 16.8 | 17.9 | 17.62 | 17.2977 | 19.3318 | 16.8479 | [9 empty] | 17.9696
| 1 | 5 | 18.49 | 19.0952 | 17.66 | 18 | 16.8 | 18.3 | 17.49 | 16.172 | 19.146 | 16.7231 | [9 empty] | 44
| 1 | 6 | 18.7 | 19.2743 | 17.18 | 18.3 | 16.9 | 18.3 | 17.63 | 15.8029 | 19.0013 | 17.3141
| Mean |  | 19.32 | 19.3584 | 17.5433 | 17.9833 | 16.8667 | 18.2333 | 17.7183 | 16.5348 | 19.0157 | 17.1222
| Median |  | 19.42 | 19.3288 | 17.61 | 17.9 | 16.8 | 18.3 | 17.635 | 16.5133 | 19.0437 | 16.9367
| Std Dev. |  | 0.64122 | 0.23501 | 0.18085 | 0.30605 | 0.38816 | 0.17512 | 0.20292 | 0.53301 | 0.2255 | 0.60384
| Rel.Std.Dev. |  | 3.319% | 1.214% | 1.031% | 1.702% | 2.301% | 0.9604% | 1.145% | 3.224% | 1.186% | 3.527%
| PDM3 |  | 7.515% | 7.729% | -2.372% | 0.07641% | -6.138% | 1.468% | -1.398% | -7.985% | 5.821% | -4.716%
| Z-Score (Absolute) |  | 1.16 | 1.19 | 0.24 | 0.1 | 0.77 | 0.3 | 0.1 | 1.04 | 0.92 | 0.57
| Analytical results for Er in OREAS 712 (Certified Value 17.2 ppm) | [20 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 19.26 | 18.5207 | 16.1 | 16.4 | 15.8 | 16.6 | 17.4 | 15.8724 | 18.6612 | 16.8934 | [9 empty] | 1
| 1 | 2 | 19.63 | 18.6469 | 16.22 | 18.3 | 16.6 | 16.4 | 17.73 | 15.8918 | 19.0493 | 15.7328 | [9 empty] | 6
| 1 | 3 | 19.37 | 18.9434 | 16.35 | 15.9 | 16.6 | 16.7 | 17.8 | 16.2247 | 18.8747 | 15.1493 | [9 empty] | 16
| 1 | 4 | 19 | 18.339 | 16.23 | 16.1 | 16.2 | 16.4 | 17.45 | 16.5994 | 18.7872 | 15.4903 | [9 empty] | 17.1517
| 1 | 5 | 18.25 | 18.6418 | 16.03 | 17.4 | 16.3 | 16.9 | 17.25 | 15.592 | 18.5119 | 15.6455 | [9 empty] | 45
| 1 | 6 | 18.29 | 18.8851 | 15.81 | 16.6 | 16.1 | 17.1 | 17.21 | 15.0107 | 19.2053 | 16.1526
| Mean |  | 18.9667 | 18.6628 | 16.1233 | 16.7833 | 16.2667 | 16.6833 | 17.4733 | 15.8652 | 18.8483 | 15.844
| Median |  | 19.13 | 18.6443 | 16.16 | 16.5 | 16.25 | 16.65 | 17.425 | 15.8821 | 18.8309 | 15.6892
| Std Dev. |  | 0.57643 | 0.22533 | 0.18949 | 0.90646 | 0.30768 | 0.27869 | 0.24402 | 0.54293 | 0.25317 | 0.60925
| Rel.Std.Dev. |  | 3.039% | 1.207% | 1.175% | 5.401% | 1.891% | 1.67% | 1.397% | 3.422% | 1.343% | 3.845%
| PDM3 |  | 10.58% | 8.81% | -5.996% | -2.148% | -5.16% | -2.731% | 1.875% | -7.501% | 9.892% | -7.624%
| Z-Score (Absolute) |  | 1.87 | 1.62 | 0.51 | 0.04 | 0.39 | 0.04 | 0.62 | 0.73 | 1.77 | 0.75
| Analytical results for Eu in OREAS 712 (Certified Value 3.17 ppm) | [20 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 3.24 | 3.656 | 3.52 | 3.18 | 3.3 | 3.1 | 3.08 | 3.0777 | 3.1514 | 3.3404 | [9 empty] | 1
| 1 | 2 | 3.31 | 3.6253 | 3.28 | 3.63 | 3.2 | 3.5 | 3.09 | 3.0009 | 3.2863 | 3.0801 | [9 empty] | 7
| 1 | 3 | 3.25 | 3.7255 | 3.32 | 3.04 | 3.2 | 3.25 | 3.07 | 3.1038 | 3.1488 | 3.0098 | [9 empty] | 16
| 1 | 4 | 3.26 | 3.6229 | 3.26 | 3.16 | 3.2 | 3.15 | 3.03 | 3.2364 | 3.2636 | 3.0339 | [9 empty] | 3.1671
| 1 | 5 | 3.06 | 3.7597 | 3.23 | 3.16 | 3.3 | 3.25 | 2.99 | 2.9861 | 3.1719 | 3.1501 | [9 empty] | 46
| 1 | 6 | 3.05 | 3.744 | 3.27 | 3.1 | 3.2 | 3.25 | 3.01 | 2.9013 | 3.1721 | 3.1662
| Mean |  | 3.195 | 3.6889 | 3.3133 | 3.2117 | 3.2333 | 3.25 | 3.045 | 3.051 | 3.199 | 3.1301
| Median |  | 3.245 | 3.6907 | 3.275 | 3.16 | 3.2 | 3.25 | 3.05 | 3.0393 | 3.172 | 3.1151
| Std Dev. |  | 0.11113 | 0.061451 | 0.10539 | 0.21132 | 0.05164 | 0.13784 | 0.040866 | 0.11578 | 0.060069 | 0.12013
| Rel.Std.Dev. |  | 3.478% | 1.666% | 3.181% | 6.58% | 1.597% | 4.241% | 1.342% | 3.795% | 1.878% | 3.838%
| PDM3 |  | 0.8825% | 16.48% | 4.619% | 1.409% | 2.093% | 2.619% | -3.854% | -3.663% | 1.009% | -1.167%
| Z-Score (Absolute) |  | 0.12 | 5.44 | 1.21 | 0.07 | 0.31 | 0.5 | 1.8 | 1.74 | 0.07 | 0.85
| Analytical results for Fe in OREAS 712 (Certified Value 14.28 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 14.3997 | 14.4943 | 13.6 | 13.8 | 15.64 | 14.32 | 14 | 14.5 | 14.52 | 14.4428 | 14.1 | 14.47 | 15.0066 | 11.73 | 13.95 | 14.15 | 13.8 | 15.4013 |  | 1
| 1 | 2 | 14.3434 | 14.2646 | 14.05 | 13.55 | 15.88 | 14.33 | 14.4 | 14.2 | 14.88 | 14.2496 | 14.3 | 14.5 | 15.4849 | 11.96 | 13.9 | 14.05 | 13.75 | 14.8975 |  | #N/A
| 1 | 3 | 14.3563 | 14.3304 | 13.85 | 14.1 | 15.58 | 14.24 | 13.7 | 14.7 | 14.29 | 14.4483 | 14 | 14.3 | 14.8499 | 11.88 | 13.6 | 14.1 | 13.85 | 14.9583 |  | 16
| 1 | 4 | 14.3164 | 14.4763 | 13.85 | 13.6 | 15.59 | 14.2 | 13.7 | 14.5 | 14.54 | 14.3994 | 14.1 | 14.48 | 15.1153 | 11.75 | 13.45 | 13.75 | 13.9 | 15.1117 |  | 14.2755
| 1 | 5 | 14.2574 | 14.5163 | 13.8 | 14.05 | 15.65 | 14.28 | 14.5 | 14.3 | 14.58 | 14.2821 | 14.2 | 14.47 | 15.0848 | 11.69 | 14.35 | 13.9 | 13.9 | 15.9927 |  | 47
| 1 | 6 | 14.2478 | 14.3197 | 14 | 14.1 | 15.64 | 14.28 | 14 | 14.4 | 14.4 | 14.3236 | 14.3 | 14.5 | 14.9904 | 11.63 | 13.4 | 14.05 | 14 | 14.4414
| Mean |  | 14.3202 | 14.4003 | 13.8583 | 13.8667 | 15.6633 | 14.275 | 14.05 | 14.4333 | 14.535 | 14.3576 | 14.1667 | 14.4533 | 15.0887 | 11.7733 | 13.775 | 14 | 13.8667 | 15.1338
| Median |  | 14.3299 | 14.4034 | 13.85 | 13.925 | 15.64 | 14.28 | 14 | 14.45 | 14.53 | 14.3615 | 14.15 | 14.475 | 15.0457 | 11.74 | 13.75 | 14.05 | 13.875 | 15.035
| Std Dev. |  | 0.058927 | 0.10756 | 0.15943 | 0.25232 | 0.11003 | 0.048888 | 0.33912 | 0.17512 | 0.19977 | 0.084558 | 0.12111 | 0.076333 | 0.21499 | 0.1234 | 0.36159 | 0.14832 | 0.08756 | 0.52456
| Rel.Std.Dev. |  | 0.4115% | 0.7469% | 1.15% | 1.82% | 0.7025% | 0.3425% | 2.414% | 1.213% | 1.374% | 0.5889% | 0.8549% | 0.5281% | 1.425% | 1.048% | 2.625% | 1.059% | 0.6314% | 3.466%
| PDM3 |  | 0.3126% | 0.8738% | -2.923% | -2.864% | 9.721% | -0.003841% | -1.58% | 1.105% | 1.817% | 0.575% | -0.7627% | 1.245% | 5.696% | -17.53% | -3.506% | -1.93% | -2.864% | 6.012%
| Z-Score (Absolute) |  | 0.06 | 0.25 | 1.09 | 1.07 | 3.38 | 0.06 | 0.61 | 0.34 | 0.59 | 0.15 | 0.32 | 0.39 | 1.96 | 6.24 | 1.29 | 0.74 | 1.07 | 2.07
| Analytical results for Ga in OREAS 712 (Certified Value 12.1 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
1
| 1 | 1 | 3.16 | 14.668 | 11.2 | 12.05 | 13.78 | 11.6 | 11.1 | 11.8 | 13.4 | 17.42 | 22.7073 | 24 | 11.5986 | 14.4 | 11.55 | 12.1 | 11.7429 |  |  | 1
| 1 | 2 | 3.42 | 15.094 | 11.4 | 11.7 | 14.06 | 11.4 | 10.5 | 12.2 | 13.8 | 17.71 | 22.8122 | 25 | 11.4124 | 13.5 | 11.55 | 11.85 | 10.5964 |  |  | 9
| 1 | 3 | 3.16 | 14.619 | 11 | 12.15 | 14 | 10.4 | 11.1 | 11.7 | 13.6 | 17.91 | 23.3587 | 25 | 11.193 | 12.85 | 11.7 | 11.6 | 10.8498 |  |  | 16
| 1 | 4 | 3.13 | 14.276 | 10.9 | 11.95 | 14.13 | 11.7 | 11 | 11.4 | 13.4 | 17.78 | 23.7485 | 25 | 11.0764 | 13 | 11.6 | 12.05 | 10.3391 |  |  | 12.1442
| 1 | 5 | 3.99 | 14.464 | 11.5 | 12.05 | 13.8 | 11.4 | 11.4 | 11.1 | 13.6 | 17.73 | 21.6702 | 25 | 10.8474 | 14.05 | 11.55 | 12.2 | 10.856 |  |  | 48
| 1 | 6 | 3.15 | 14.807 | 11 | 12 | 13.85 | 11.9 | 10.1 | 10.8 | 13.6 | 17.73 | 21.922 | 26 | 10.8538 | 12.85 | 11.65 | 11.35 | 10.0362
| Mean |  | 3.335 | 14.6547 | 11.1667 | 11.9833 | 13.9367 | 11.4 | 10.8667 | 11.5 | 13.5667 | 17.7133 | 22.7031 | 25 | 11.1636 | 13.4417 | 11.6 | 11.8583 | 10.7367
| Median |  | 3.16 | 14.6435 | 11.1 | 12.025 | 13.925 | 11.5 | 11.05 | 11.55 | 13.6 | 17.73 | 22.7597 | 25 | 11.1347 | 13.25 | 11.575 | 11.95 | 10.7231
| Std Dev. |  | 0.33875 | 0.28193 | 0.24221 | 0.15384 | 0.14652 | 0.52536 | 0.4761 | 0.50596 | 0.15055 | 0.16108 | 0.80146 | 0.63246 | 0.30186 | 0.66137 | 0.063246 | 0.32774 | 0.58445
| Rel.Std.Dev. |  | 10.16% | 1.924% | 2.169% | 1.284% | 1.051% | 4.608% | 4.381% | 4.4% | 1.11% | 0.9094% | 3.53% | 2.53% | 2.704% | 4.92% | 0.5452% | 2.764% | 5.443%
| PDM3 |  | -72.54% | 20.67% | -8.05% | -1.325% | 14.76% | -6.128% | -10.52% | -5.305% | 11.71% | 45.86% | 86.95% | 105.9% | -8.075% | 10.68% | -4.481% | -2.354% | -11.59%
| Z-Score (Absolute) |  | 5.12 | 1.68 | 0.42 | 0.08 | 1.25 | 0.28 | 0.6 | 0.22 | 1.03 | 3.52 | 6.52 | 7.9 | 0.42 | 0.95 | 0.16 | 0 | 0.67
| Analytical results for Gd in OREAS 712 (Certified Value 16.9 ppm) | [20 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 19.39 | 44.3132 | 17.32 | 16.9 | 15.6 | 16.8 | 16.9 | 17.9863 | 16.7411 | 16.0374 | [9 empty] | 1
| 1 | 2 | 19.82 | 44.3969 | 16.65 | 17.8 | 16 | 17.2 | 16.8 | 17.8466 | 17.1229 | 13.8359 | [9 empty] | 10
| 1 | 3 | 19.42 | 44.6625 | 16.77 | 16.5 | 16.1 | 17 | 17 | 18.1987 | 16.9025 | 14.114 | [9 empty] | 16
| 1 | 4 | 19.2 | 43.9421 | 16.32 | 17 | 15.8 | 17 | 16.6 | 18.6269 | 16.9849 | 13.6398 | [9 empty] | 16.8963
| 1 | 5 | 18.32 | 44.0788 | 16.99 | 16.7 | 16.3 | 17 | 16.6 | 17.4555 | 16.9721 | 14.2367 | [9 empty] | 49
| 1 | 6 | 18.19 | 44.2027 | 16.5 | 17.2 | 15.8 | 17 | 16.6 | 17.0883 | 16.9691 | 13.3047
| Mean |  | 19.0567 | 44.266 | 16.7583 | 17.0167 | 15.9333 | 17 | 16.75 | 17.867 | 16.9488 | 14.1948
| Median |  | 19.295 | 44.258 | 16.71 | 16.95 | 15.9 | 17 | 16.7 | 17.9165 | 16.9706 | 13.9749
| Std Dev. |  | 0.65424 | 0.2532 | 0.35774 | 0.4535 | 0.25033 | 0.12649 | 0.17607 | 0.54367 | 0.12478 | 0.96256
| Rel.Std.Dev. |  | 3.433% | 0.572% | 2.135% | 2.665% | 1.571% | 0.7441% | 1.051% | 3.043% | 0.7362% | 6.781%
| PDM3 |  | 12.79% | 162% | -0.8166% | 0.7123% | -5.699% | 0.6137% | -0.8659% | 5.745% | 0.3106% | -15.99%
| Z-Score (Absolute) |  | 2.51 | 32.95 | 0.26 | 0.05 | 1.26 | 0.03 | 0.27 | 1.08 | 0.03 | 3.36
| Analytical results for Ge in OREAS 712 (Indicative Value 0.79 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.59 | 0.84273 | 0.35 | 0.54 | 1.06 | 0.2 | 5.7 | < 0.5 | 1 | 1.7316 | 13.2 | 0.56 | 0.48 | 0.45 | [5 empty] | 1
| 1 | 2 | 1.71 | 0.83718 | 0.35 | 0.48 | 1.06 | 0.2 | 5.6 | 1.1 | 0.9 | 1.6394 | 13.2 | 0.58 | 0.49 | 0.5 | [5 empty] | 11
| 1 | 3 | 1.49 | 0.80805 | 0.35 | 0.53 | 1.09 | 0.2 | 5.9 | 1 | 0.9 | 1.8273 | 14 | 0.51 | 0.46 | 0.46 | [5 empty] | 16
| 1 | 4 | 1.21 | 0.82477 | 0.33 | 0.52 | 0.99 | 0.2 | 5.7 | 0.7 | 0.9 | 1.7665 | 13.4 | 0.55 | 0.48 | 0.49 | [5 empty] | 0.79423
| 1 | 5 | 1.76 | 0.80904 | 0.37 | 0.53 | 1.07 | 0.2 | 6 | 0.9 | 0.9 | 1.5495 | 13.4 | 0.58 | 0.49 | 0.47 | [5 empty] | 22
| 1 | 6 | 1.6 | 0.83948 | 0.36 | 0.5 | 1.02 | < 0.1 | 6 | < 0.5 | 0.9 | 1.5487 | 14.7 | 0.55 | 0.47 | 0.48
| Mean |  | 1.56 | 0.82688 | 0.35167 | 0.51667 | 1.0483 | 0.2 | 5.8167 | 0.925 | 0.91667 | 1.6772 | 13.65 | 0.555 | 0.47833 | 0.475
| Median |  | 1.595 | 0.83098 | 0.35 | 0.525 | 1.06 | 0.2 | 5.8 | 0.95 | 0.9 | 1.6855 | 13.4 | 0.555 | 0.48 | 0.475
| Std Dev. |  | 0.19616 | 0.01545 | 0.013292 | 0.022509 | 0.03656 | 0 | 0.17224 | 0.17078 | 0.040825 | 0.11637 | 0.59245 | 0.025884 | 0.01169 | 0.018708
| Rel.Std.Dev. |  | 12.57% | 1.868% | 3.78% | 4.357% | 3.487% | 0% | 2.961% | 18.46% | 4.454% | 6.939% | 4.34% | 4.664% | 2.444% | 3.939%
| PDM3 |  | 96.42% | 4.111% | -55.72% | -34.95% | 31.99% | -74.82% | 632.4% | 16.47% | 15.42% | 111.2% | 1619% | -30.12% | -39.77% | -40.19%
| Z-Score (Absolute) |  | 1.87 | 0.15 | 0.97 | 0.58 | 0.67 | 1.38 | 11.88 | 0.15 | 0.36 | 2.15 | 30.3 | 0.49 | 0.67 | 0.68
| Analytical results for Hf in OREAS 712 (Indicative Value 1.94 ppm) | [20 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.22 | 1.4 | 1.9 | 3.67 | < 0.1 | 4.4 | 4.4 | 1.8 | 2.6 | 1.6213 | 3 | 28.2089 | 1.6 | 1.5 | 1.5 |  |  |  |  | 1
| 1 | 2 | 1.26 | 1.5 | 1.7 | 3.52 | 0.1 | 4.3 | 3.8 | 1 | 2.63 | 1.6021 | 3 | 28.8621 | 1.3 | 1.6 | 1.3 |  |  |  |  | 17
| 1 | 3 | 1.32 | 1.5 | 2.2 | 3.63 | 0.6 | 4.6 | 2.5 | 1.4 | 2.62 | 1.7097 | 2.6 | 28.9382 | 1.4 | 1.9 | 1.7 |  |  |  |  | 16
| 1 | 4 | 1.27 | 1.6 | 2 | 3.69 | 0.1 | 4.5 | 2.1 | 0.8 | 2.69 | 1.8544 | 3.3 | 29.8014 | 1.4 | 1.6 | 1.7 |  |  |  |  | 1.9357
| 1 | 5 | 1.13 | 1.6 | 1.8 | 3.63 | 0.2 | 4.4 | 4.8 | 0.8 | 2.47 | 1.9085 | 3 | 29.4064 | 1.4 | 2 | 1.5 |  |  |  |  | 23
| 1 | 6 | 1.12 | 1.6 | 2.6 | 3.69 | < 0.1 | 4.5 | 2.3 | 1 | 2.6 | 1.5724 | 3.5 | 29.0333 | 1.3 | 1.4 | 1.7
| Mean |  | 1.22 | 1.5333 | 2.0333 | 3.6383 | 0.25 | 4.45 | 3.3167 | 1.1333 | 2.6017 | 1.7114 | 3.0667 | 29.0418 | 1.4 | 1.6667 | 1.5667
| Median |  | 1.24 | 1.55 | 1.95 | 3.65 | 0.15 | 4.45 | 3.15 | 1 | 2.61 | 1.6655 | 3 | 28.9858 | 1.4 | 1.6 | 1.6
| Std Dev. |  | 0.08025 | 0.08165 | 0.3266 | 0.064005 | 0.23805 | 0.10488 | 1.1652 | 0.39328 | 0.072503 | 0.1405 | 0.30768 | 0.53792 | 0.10954 | 0.23381 | 0.1633
| Rel.Std.Dev. |  | 6.578% | 5.325% | 16.06% | 1.759% | 95.22% | 2.357% | 35.13% | 34.7% | 2.787% | 8.21% | 10.03% | 1.852% | 7.825% | 14.03% | 10.42%
| PDM3 |  | -36.97% | -20.79% | 5.043% | 87.96% | -87.08% | 129.9% | 71.34% | -41.45% | 34.4% | -11.59% | 58.42% | 1400% | -27.68% | -13.9% | -19.07%
| Z-Score (Absolute) |  | 0.57 | 0.21 | 0.38 | 2.25 | 1.78 | 3.19 | 1.87 | 0.67 | 1.04 | 0 | 1.58 | 31.88 | 0.36 | 0.05 | 0.17
| Analytical results for Hg in OREAS 712 (Indicative Value < 0.01 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 14 | [17 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [17 empty] | 3
3
| 1 | 1 | < 0.01 | < 1 | [17 empty] | 1
| 1 | 2 | < 0.01 | < 1 | [17 empty] | 12
| 1 | 3 | < 0.01 | < 1 | [17 empty] | 16
| 1 | 4 | < 0.01 | < 1 | [17 empty] | < 0.01
| 1 | 5 | < 0.01 | < 1 | [17 empty] | 24
| 1 | 6 | < 0.01 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ho in OREAS 712 (Certified Value 4.65 ppm) | [20 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 4.56 | 4.9239 | 4.75 | 4.4 | 4.7 | 4.66 | 4.54 | 4.3421 | 4.9239 | 4.7459 | [9 empty] | 1
| 1 | 2 | 4.66 | 5.0746 | 4.67 | 5.1 | 4.8 | 4.66 | 4.61 | 4.2946 | 5.0811 | 4.3618 | [9 empty] | 25
| 1 | 3 | 4.63 | 5.085 | 4.75 | 4.7 | 4.6 | 4.66 | 4.61 | 4.4366 | 4.9275 | 4.3205 | [9 empty] | 16
| 1 | 4 | 4.5 | 5.0182 | 4.77 | 4.4 | 4.5 | 4.68 | 4.53 | 4.5302 | 5.1657 | 4.3842 | [9 empty] | 4.651
| 1 | 5 | 4.33 | 4.8837 | 4.59 | 4.8 | 4.6 | 4.72 | 4.47 | 4.2135 | 4.9979 | 4.3947 | [9 empty] | 50
| 1 | 6 | 4.34 | 5.1232 | 4.62 | 4.4 | 4.7 | 4.74 | 4.51 | 4.1122 | 4.9936 | 4.4666
| Mean |  | 4.5033 | 5.0181 | 4.6917 | 4.6333 | 4.65 | 4.6867 | 4.545 | 4.3215 | 5.015 | 4.4456
| Median |  | 4.53 | 5.0464 | 4.71 | 4.55 | 4.65 | 4.67 | 4.535 | 4.3184 | 4.9957 | 4.3895
| Std Dev. |  | 0.1418 | 0.095561 | 0.076004 | 0.28752 | 0.10488 | 0.035024 | 0.055767 | 0.15064 | 0.093567 | 0.1547
| Rel.Std.Dev. |  | 3.149% | 1.904% | 1.62% | 6.205% | 2.256% | 0.7473% | 1.227% | 3.486% | 1.866% | 3.48%
| PDM3 |  | -3.175% | 7.892% | 0.8739% | -0.3803% | -0.02193% | 0.7664% | -2.279% | -7.084% | 7.825% | -4.416%
| Z-Score (Absolute) |  | 0.79 | 2.16 | 0.29 | 0.05 | 0.05 | 0.26 | 0.55 | 1.84 | 2.14 | 1.13
| Analytical results for In in OREAS 712 (Certified Value 0.093 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.106 | 0.085 | 0.098 | 0.12 | < 0.1 | 0.09 | 0.1 | 0.1 | 0.095 | 0.08 | 0.08781 | 0.102 | 0.087 | 0.093 | [5 empty] | 1
| 1 | 2 | 0.108 | 0.095 | 0.089 | 0.12 | 0.1 | 0.09 | 0.07 | 0.1 | 0.1 | 0.09 | 0.08698 | 0.099 | 0.089 | 0.095 | [5 empty] | 2
| 1 | 3 | 0.111 | 0.09 | 0.095 | 0.12 | < 0.1 | 0.09 | 0.07 | 0.1 | 0.103 | 0.09 | 0.08436 | 0.086 | 0.088 | 0.091 | [5 empty] | 16
| 1 | 4 | 0.121 | 0.091 | 0.091 | 0.12 | < 0.1 | 0.09 | 0.11 | 0.1 | 0.102 | 0.09 | 0.09123 | 0.091 | 0.088 | 0.096 | [5 empty] | 0.093255
| 1 | 5 | 0.104 | 0.088 | 0.096 | 0.12 | 0.1 | 0.09 | 0.08 | 0.1 | 0.101 | 0.08 | 0.08984 | 0.099 | 0.084 | 0.097 | [5 empty] | 51
| 1 | 6 | 0.109 | 0.102 | 0.095 | 0.12 | 0.1 | 0.09 | 0.08 | 0.1 | 0.095 | 0.09 | 0.08833 | 0.092 | 0.078 | 0.095
| Mean |  | 0.10983 | 0.091833 | 0.094 | 0.12 | 0.1 | 0.09 | 0.085 | 0.1 | 0.099333 | 0.086667 | 0.088092 | 0.094833 | 0.085667 | 0.0945
| Median |  | 0.1085 | 0.0905 | 0.095 | 0.12 | 0.1 | 0.09 | 0.08 | 0.1 | 0.1005 | 0.09 | 0.08807 | 0.0955 | 0.0875 | 0.095
| Std Dev. |  | 0.0059805 | 0.0059805 | 0.0033466 | 0 | 1.6997e-17 | 1.5202e-17 | 0.016432 | 1.5202e-17 | 0.0035024 | 0.005164 | 0.002375 | 0.0061128 | 0.0041312 | 0.0021679
| Rel.Std.Dev. |  | 5.445% | 6.512% | 3.56% | 0% | 1.7e-14% | 1.689e-14% | 19.33% | 1.52e-14% | 3.526% | 5.958% | 2.696% | 6.446% | 4.822% | 2.294%
| PDM3 |  | 17.78% | -1.524% | 0.799% | 28.68% | 7.233% | -3.49% | -8.852% | 7.233% | 6.518% | -7.065% | -5.537% | 1.693% | -8.137% | 1.335%
| Z-Score (Absolute) |  | 1.81 | 0.25 | 0 | 2.97 | NA | 0.46 | 1.03 | 0.68 | 0.61 | 0.84 | 0.67 | 0.1 | 0.95 | 0.06
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 712 (Certified Value 0.624 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.6323 | 0.63526 | 0.6 | 0.61 | 0.67 | 0.64 | 0.62 | 0.658 | 0.62 | 0.60595 | 0.62 | 0.63139 | 0.55 | 0.62 | 0.62 | 0.61 | 0.6273 |  |  | 1
| 1 | 2 | 0.6325 | 0.63445 | 0.63 | 0.58 | 0.66 | 0.65 | 0.63 | 0.642 | 0.63 | 0.59944 | 0.62 | 0.64593 | 0.56 | 0.62 | 0.62 | 0.6 | 0.6171 |  |  | #N/A
| 1 | 3 | 0.6359 | 0.63569 | 0.61 | 0.6 | 0.65 | 0.61 | 0.59 | 0.676 | 0.59 | 0.59904 | 0.64 | 0.63153 | 0.56 | 0.6 | 0.62 | 0.6 | 0.6262 |  |  | 16
| 1 | 4 | 0.6396 | 0.63268 | 0.61 | 0.59 | 0.68 | 0.62 | 0.61 | 0.641 | 0.61 | 0.61272 | 0.63 | 0.63114 | 0.56 | 0.6 | 0.62 | 0.61 | 0.6211 |  |  | 0.62437
| 1 | 5 | 0.6265 | 0.63987 | 0.62 | 0.6 | 0.66 | 0.64 | 0.62 | 0.661 | 0.62 | 0.59622 | 0.62 | 0.6273 | 0.56 | 0.64 | 0.62 | 0.61 | 0.6275 |  |  | 52
| 1 | 6 | 0.6317 | 0.63261 | 0.63 | 0.61 | 0.68 | 0.63 | 0.63 | 0.645 | 0.61 | 0.60314 | 0.62 | 0.66274 | 0.57 | 0.6 | 0.62 | 0.61 | 0.6207
| Mean |  | 0.63308 | 0.63509 | 0.61667 | 0.59833 | 0.66667 | 0.63167 | 0.61667 | 0.65383 | 0.61333 | 0.60275 | 0.625 | 0.63834 | 0.56 | 0.61333 | 0.62 | 0.60667 | 0.62332
| Median |  | 0.6324 | 0.63486 | 0.615 | 0.6 | 0.665 | 0.635 | 0.62 | 0.6515 | 0.615 | 0.60129 | 0.62 | 0.63146 | 0.56 | 0.61 | 0.62 | 0.61 | 0.62365
| Std Dev. |  | 0.0043956 | 0.0026685 | 0.012111 | 0.01169 | 0.012111 | 0.01472 | 0.015055 | 0.013732 | 0.013663 | 0.0059514 | 0.0083666 | 0.013575 | 0.0063246 | 0.01633 | 0 | 0.005164 | 0.0042916
| Rel.Std.Dev. |  | 0.6943% | 0.4202% | 1.964% | 1.954% | 1.817% | 2.33% | 2.441% | 2.1% | 2.228% | 0.9874% | 1.339% | 2.127% | 1.129% | 2.662% | 0% | 0.8512% | 0.6885%
| PDM3 |  | 1.396% | 1.718% | -1.233% | -4.17% | 6.775% | 1.169% | -1.233% | 4.719% | -1.767% | -3.462% | 0.1014% | 2.238% | -10.31% | -1.767% | -0.6994% | -2.835% | -0.1682%
| Z-Score (Absolute) |  | 0.67 | 0.78 | 0.17 | 1.12 | 2.41 | 0.6 | 0.17 | 1.74 | 0.34 | 0.89 | 0.26 | 0.95 | 3.09 | 0.34 | 0 | 0.69 | 0.17
| Analytical results for La in OREAS 712 (Certified Value 250 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
0
| 1 | 1 | 225 | 258.25 | 240 | 243 | 231.27 | 252 | 247.2 | 235.4807 | 249 | 251 | 253.0357 | 272 | 246.8206 | 255 | 242 | 253 | 261.0998 |  |  | 1
| 1 | 2 | 224.5 | 259.4 | 259 | 237 | 225.11 | 264 | 256.5 | 232.7852 | 257 | 255 | 251.8666 | 295 | 247.7139 | 247 | 245 | 264 | 250.9485 |  |  | 26
| 1 | 3 | 224.4 | 258.73 | 242 | 244 | 228.85 | 249 | 248.2 | 236.475 | 250 | 264 | 259.0072 | 279 | 249.8309 | 238 | 245 | 254 | 242.4528 |  |  | 16
| 1 | 4 | 224.6 | 260.7 | 241 | 234 | 231.95 | 247 | 250.6 | 239.6371 | 250 | 253 | 267.302 | 285 | 247.9178 | 240 | 237 | 264 | 260.0188 |  |  | 249.5199
| 1 | 5 | 222.3 | 260.72 | 250 | 242 | 229.81 | 260 | 249.5 | 232.7204 | 252 | 255 | 245.9678 | 294 | 243.7402 | 253 | 240 | 259 | 242.7125 |  |  | 53
| 1 | 6 | 222.1 | 259.16 | 243 | 253 | 230.91 | 257 | 251.4 | 241.1014 | 253 | 253 | 240.2707 | 300 | 246.1524 | 237 | 242 | 259 | 251.2578
| Mean |  | 223.8167 | 259.4933 | 245.8333 | 242.1667 | 229.65 | 254.8333 | 250.5667 | 236.3666 | 251.8333 | 255.1667 | 252.9083 | 287.5 | 247.0293 | 245 | 241.8333 | 258.8333 | 251.415
| Median |  | 224.45 | 259.28 | 242.5 | 242.5 | 230.36 | 254.5 | 250.05 | 235.9778 | 251 | 254 | 252.4512 | 289.5 | 247.2672 | 243.5 | 242 | 259 | 251.1031
| Std Dev. |  | 1.2703 | 1.0207 | 7.3598 | 6.5549 | 2.4811 | 6.6156 | 3.2855 | 3.465 | 2.9269 | 4.5789 | 9.5221 | 10.7098 | 2.0366 | 7.823 | 3.0605 | 4.7081 | 8.0513
| Rel.Std.Dev. |  | 0.5676% | 0.3933% | 2.994% | 2.707% | 1.08% | 2.596% | 1.311% | 1.466% | 1.162% | 1.794% | 3.765% | 3.725% | 0.8244% | 3.193% | 1.266% | 1.819% | 3.202%
| PDM3 |  | -10.3% | 3.997% | -1.477% | -2.947% | -7.963% | 2.129% | 0.4195% | -5.271% | 0.9271% | 2.263% | 1.358% | 15.22% | -0.9982% | -1.811% | -3.081% | 3.733% | 0.7595%
| Z-Score (Absolute) |  | 3.24 | 1.08 | 0.57 | 1.02 | 2.53 | 0.52 | 0 | 1.72 | 0.15 | 0.56 | 0.28 | 4.47 | 0.43 | 0.67 | 1.06 | 1 | 0.1
| Analytical results for Li in OREAS 712 (Certified Value 21.2 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 20 | 16.151 | 22.7 | 22.2 | 20.47 | 21.75 | 22 | 19.7 | 21 | 20.2088 | 20.5 | 20.6 | 19.41 | 20.8 | 20.0568 | 22.1 | 23.4 | 22.7 | 30.1065 | 1
| 1 | 2 | 20 | 15.4736 | 23.8 | 21.9 | 21.82 | 21.32 | 20 | 19.2 | 21 | 18.9381 | 21 | 19.9 | 21.46 | 20.6 | 19.9566 | 22.4 | 23.7 | 22.4 | 28.6201 | #N/A
| 1 | 3 | 20 | 16.2444 | 23.1 | 22.7 | 20.22 | 21.08 | 20 | 20.6 | 20 | 20.502 | 20.5 | 21.8 | 19.56 | 20.8 | 21.0774 | 21.8 | 23.3 | 22.7 | 29.3746 | 16
| 1 | 4 | 20 | 15.7023 | 22.7 | 22.1 | 21.83 | 20.98 | 21 | 20.5 | 21 | 20.5415 | 20.5 | 22.1 | 18.27 | 22.5 | 20.647 | 21.4 | 22.4 | 22.8 | 28.4606 | 21.1738
| 1 | 5 | 20 | 15.9208 | 23.5 | 22.5 | 20.43 | 21.46 | 21 | 20.1 | 20 | 19.9805 | 21 | 22.2 | 20.45 | 20.6 | 19.9864 | 23.3 | 22.6 | 23.1 | 29.6798 | 54
| 1 | 6 | 20 | 15.9337 | 23.5 | 22 | 20.57 | 21.17 | 22 | 20 | 21 | 18.7874 | 20.5 | 21.8 | 18.67 | 20.8 | 19.9271 | 21.1 | 22.9 | 22.8 | 27.9656
| Mean |  | 20 | 15.9043 | 23.2167 | 22.2333 | 20.89 | 21.2933 | 21 | 20.0167 | 20.6667 | 19.8264 | 20.6667 | 21.4 | 19.6367 | 21.0167 | 20.2752 | 22.0167 | 23.05 | 22.75 | 29.0345
| Median |  | 20 | 15.9273 | 23.3 | 22.15 | 20.52 | 21.245 | 21 | 20.05 | 21 | 20.0946 | 20.5 | 21.8 | 19.485 | 20.8 | 20.0216 | 21.95 | 23.1 | 22.75 | 28.9974
| Std Dev. |  | 0 | 0.28424 | 0.45789 | 0.30768 | 0.73318 | 0.2814 | 0.89443 | 0.51929 | 0.5164 | 0.77544 | 0.2582 | 0.93167 | 1.1701 | 0.73326 | 0.47655 | 0.78337 | 0.501 | 0.22583 | 0.81546
| Rel.Std.Dev. |  | 0% | 1.787% | 1.972% | 1.384% | 3.51% | 1.322% | 4.259% | 2.594% | 2.499% | 3.911% | 1.249% | 4.354% | 5.959% | 3.489% | 2.35% | 3.558% | 2.174% | 0.9927% | 2.809%
| PDM3 |  | -5.544% | -24.89% | 9.648% | 5.004% | -1.34% | 0.5645% | -0.8209% | -5.465% | -2.395% | -6.364% | -2.395% | 1.068% | -7.26% | -0.7422% | -4.244% | 3.981% | 8.861% | 7.444% | 37.12%
| Z-Score (Absolute) |  | 0.67 | 3.44 | 1.49 | 0.83 | 0.07 | 0.2 | 0 | 0.66 | 0.22 | 0.79 | 0.22 | 0.27 | 0.92 | 0.01 | 0.49 | 0.69 | 1.38 | 1.18 | 5.42
| Analytical results for Lu in OREAS 712 (Certified Value 3.37 ppm) | [20 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 3.66 | 3.444 | 3.59 | 3.1 | 3.26 | 3.56 | 3.36 | 2.9882 | 3.6217 | 2.3175 | [9 empty] | 1
| 1 | 2 | 3.74 | 3.3883 | 3.48 | 3.2 | 3.45 | 3.4 | 3.28 | 2.9537 | 3.6193 | 2.0801 | [9 empty] | 3
| 1 | 3 | 3.7 | 3.4707 | 3.5 | 3.1 | 3.32 | 3.38 | 3.33 | 3.0999 | 3.6285 | 2.0809 | [9 empty] | 16
| 1 | 4 | 3.63 | 3.4541 | 3.55 | 3.1 | 3.26 | 3.34 | 3.04 | 3.1376 | 3.7155 | 2.0209 | [9 empty] | 3.3668
| 1 | 5 | 3.45 | 3.3517 | 3.6 | 3.2 | 3.39 | 3.42 | 3.08 | 2.9439 | 3.6455 | 2.1301 | [9 empty] | 55
| 1 | 6 | 3.48 | 3.5435 | 3.59 | 3 | 3.31 | 3.46 | 3.1 | 2.8087 | 3.5804 | 1.9918
| Mean |  | 3.61 | 3.442 | 3.5517 | 3.1167 | 3.3317 | 3.4267 | 3.1983 | 2.9887 | 3.6351 | 2.1036
| Median |  | 3.645 | 3.449 | 3.57 | 3.1 | 3.315 | 3.41 | 3.19 | 2.971 | 3.6251 | 2.0805
| Std Dev. |  | 0.11866 | 0.066801 | 0.051153 | 0.075277 | 0.075211 | 0.076594 | 0.14063 | 0.11845 | 0.044819 | 0.11564
| Rel.Std.Dev. |  | 3.287% | 1.941% | 1.44% | 2.415% | 2.257% | 2.235% | 4.397% | 3.963% | 1.233% | 5.497%
| PDM3 |  | 7.225% | 2.236% | 5.492% | -7.428% | -1.042% | 1.779% | -5.003% | -11.23% | 7.971% | -37.52%
| Z-Score (Absolute) |  | 0.76 | 0.21 | 0.57 | 0.86 | 0.16 | 0.16 | 0.59 | 1.28 | 0.84 | 4.18
| Analytical results for Mg in OREAS 712 (Certified Value 0.556 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.5654 | 0.61898 | 0.5 | 0.5 | 0.56 | 0.57 | 0.54 | 0.59 | 0.59 | 0.53319 | 0.565 | 0.63488 | 0.52 | 0.52 | 0.53 | 0.52 | 0.5772 |  |  | 1
| 1 | 2 | 0.5648 | 0.61405 | 0.53 | 0.49 | 0.56 | 0.56 | 0.57 | 0.576 | 0.61 | 0.53478 | 0.567 | 0.65402 | 0.52 | 0.53 | 0.53 | 0.52 | 0.5687 |  |  | #N/A
| 1 | 3 | 0.5654 | 0.61554 | 0.52 | 0.51 | 0.57 | 0.56 | 0.54 | 0.605 | 0.58 | 0.54236 | 0.58 | 0.6346 | 0.51 | 0.51 | 0.52 | 0.52 | 0.575 |  |  | 16
| 1 | 4 | 0.5652 | 0.61821 | 0.51 | 0.49 | 0.55 | 0.55 | 0.52 | 0.574 | 0.59 | 0.55702 | 0.568 | 0.63658 | 0.53 | 0.5 | 0.5 | 0.53 | 0.5715 |  |  | 0.55585
| 1 | 5 | 0.5607 | 0.61366 | 0.52 | 0.52 | 0.55 | 0.57 | 0.58 | 0.574 | 0.6 | 0.52819 | 0.566 | 0.63774 | 0.53 | 0.55 | 0.51 | 0.53 | 0.5751 |  |  | 56
| 1 | 6 | 0.5617 | 0.61141 | 0.53 | 0.51 | 0.57 | 0.56 | 0.54 | 0.579 | 0.59 | 0.53218 | 0.566 | 0.6358 | 0.53 | 0.51 | 0.51 | 0.52 | 0.573
| Mean |  | 0.56387 | 0.61531 | 0.51833 | 0.50333 | 0.56 | 0.56167 | 0.54833 | 0.583 | 0.59333 | 0.53795 | 0.56867 | 0.63894 | 0.52333 | 0.52 | 0.51667 | 0.52333 | 0.57342
| Median |  | 0.565 | 0.61479 | 0.52 | 0.505 | 0.56 | 0.56 | 0.54 | 0.5775 | 0.59 | 0.53398 | 0.5665 | 0.63619 | 0.525 | 0.515 | 0.515 | 0.52 | 0.574
| Std Dev. |  | 0.0021011 | 0.0028793 | 0.01169 | 0.012111 | 0.0089443 | 0.0075277 | 0.022286 | 0.012329 | 0.010328 | 0.010435 | 0.0056451 | 0.0074782 | 0.008165 | 0.017889 | 0.012111 | 0.005164 | 0.0030248
| Rel.Std.Dev. |  | 0.3726% | 0.4679% | 2.255% | 2.406% | 1.597% | 1.34% | 4.064% | 2.115% | 1.741% | 1.94% | 0.9927% | 1.17% | 1.56% | 3.44% | 2.344% | 0.9867% | 0.5275%
| PDM3 |  | 1.442% | 10.7% | -6.75% | -9.448% | 0.7463% | 1.046% | -1.353% | 4.884% | 6.743% | -3.22% | 2.305% | 14.95% | -5.85% | -6.45% | -7.05% | -5.85% | 3.16%
| Z-Score (Absolute) |  | 0.08 | 1.12 | 0.84 | 1.15 | 0 | 0.03 | 0.24 | 0.47 | 0.67 | 0.45 | 0.18 | 1.6 | 0.74 | 0.81 | 0.88 | 0.74 | 0.27
| Analytical results for Mn in OREAS 712 (Certified Value 1.24 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.1057 | 1.3107 | 1.16 | 1.2 | > 1 | 1.3 | > 1 | 1.24 | 1.3784 | 1.176 | 1.28 | 1.3067 | 1.363 | > 1 | 1.2 | 1.215 | 1.155 | 1.2106 |  | 1
| 1 | 2 | 1.1023 | 1.3093 | 1.23 | 1.17 | > 1 | 1.3 | > 1 | 1.24 | 1.4192 | 1.1669 | 1.29 | 1.3102 | 1.4008 | > 1 | 1.215 | 1.21 | 1.165 | 1.2213 |  | 21
| 1 | 3 | 1.1038 | 1.3023 | 1.215 | 1.21 | > 1 | 1.3 | > 1 | 1.29 | 1.3556 | 1.1795 | 1.26 | 1.3097 | 1.3658 | > 1 | 1.175 | 1.2 | 1.165 | 1.2239 |  | 16
| 1 | 4 | 1.1045 | 1.3111 | 1.2 | 1.18 | > 1 | 1.3 | > 1 | 1.22 | 1.383 | 1.1895 | 1.27 | 1.3179 | 1.3708 | > 1 | 1.16 | 1.16 | 1.175 | 1.2103 |  | 1.2442
| 1 | 5 | 1.0964 | 1.3027 | 1.215 | 1.23 | > 1 | 1.3 | > 1 | 1.24 | 1.3876 | 1.1646 | 1.28 | 1.3049 | 1.3627 | > 1 | 1.265 | 1.18 | 1.175 | 1.2536 |  | 57
| 1 | 6 | 1.0951 | 1.3067 | 1.22 | 1.22 | > 1 | 1.3 | > 1 | 1.22 | 1.377 | 1.1726 | 1.29 | 1.3151 | 1.3635 | > 1 | 1.17 | 1.18 | 1.165 | 1.2091
| Mean |  | 1.1013 | 1.3072 | 1.2067 | 1.2017 |  | 1.3 |  | 1.2417 | 1.3835 | 1.1748 | 1.2783 | 1.3108 | 1.3711 |  | 1.1975 | 1.1908 | 1.1667 | 1.2215
| Median |  | 1.1031 | 1.308 | 1.215 | 1.205 |  | 1.3 |  | 1.24 | 1.3807 | 1.1743 | 1.28 | 1.3099 | 1.3647 |  | 1.1875 | 1.19 | 1.165 | 1.2159
| Std Dev. |  | 0.0044565 | 0.0039118 | 0.024833 | 0.023166 |  | 0 |  | 0.025626 | 0.02068 | 0.0090527 | 0.01169 | 0.0049412 | 0.014861 |  | 0.038827 | 0.021075 | 0.0075277 | 0.016936
| Rel.Std.Dev. |  | 0.4047% | 0.2993% | 2.058% | 1.928% |  | 0% |  | 2.064% | 1.495% | 0.7705% | 0.9145% | 0.377% | 1.084% |  | 3.242% | 1.77% | 0.6452% | 1.387%
| PDM3 |  | -11.48% | 5.061% | -3.016% | -3.417% |  | 4.486% |  | -0.2024% | 11.19% | -5.573% | 2.745% | 5.35% | 10.2% |  | -3.752% | -4.288% | -6.23% | -1.826%
| Z-Score (Absolute) |  | 1.48 | 1.05 | 0.18 | 0.24 | NA | 0.97 | NA | 0.25 | 1.99 | 0.57 | 0.7 | 1.1 | 1.84 | NA | 0.29 | 0.38 | 0.67 | 0
| Analytical results for Mo in OREAS 712 (Certified Value 3.22 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
2
| 1 | 1 | 4.3 | 3.28 | 3.17 | 1.17 | 5.22 | < 1 | 3.5 | 3 | < 10 | 3.5 | 3.4 | 3.0066 | 1.3 | 3.2835 | 3.64 | 3.16 | 3.04 | 2.4477 |  | 1
| 1 | 2 | 4 | 3.46 | 3.01 | 1.14 | 5.14 | < 1 | 3.5 | 3 | < 10 | 3 | 3.3 | 2.942 | 1.4 | 3.2861 | 3.55 | 3.28 | 3.01 | 2.2569 |  | 4
| 1 | 3 | 3.7 | 3.19 | 3.18 | 1.2 | 5.13 | < 1 | 3.5 | 3 | < 10 | 3 | 3.5 | 3.5148 | 1.5 | 3.2876 | 3.23 | 3.28 | 2.95 | 2.5767 |  | 16
| 1 | 4 | 4.3 | 3.31 | 3.11 | 1.3 | 5.17 | < 1 | 3.4 | 3 | < 10 | 2 | 3.5 | 3.5496 | 1.7 | 3.2595 | 3.35 | 3.16 | 3.05 | 2.1752 |  | 3.2223
| 1 | 5 | 3.8 | 3.53 | 3.22 | 1.22 | 5.12 | < 1 | 3.5 | 3 | < 10 | 2 | 3.5 | 2.9937 | 1.2 | 3.2088 | 3.61 | 3.18 | 3.09 | 2.3886 |  | 58
| 1 | 6 | 3.9 | 3.32 | 3.17 | 1.18 | 5.31 | < 1 | 3.5 | 4 | < 10 | 2.5 | 3.4 | 3.3331 | 1.8 | 3.2238 | 3.31 | 3.25 | 3.01 | 2.5147
| Mean |  | 4 | 3.3483 | 3.1433 | 1.2017 | 5.1817 |  | 3.4833 | 3.1667 |  | 2.6667 | 3.4333 | 3.2233 | 1.4833 | 3.2582 | 3.4483 | 3.2183 | 3.025 | 2.3933
| Median |  | 3.95 | 3.315 | 3.17 | 1.19 | 5.155 |  | 3.5 | 3 |  | 2.75 | 3.45 | 3.1699 | 1.45 | 3.2715 | 3.45 | 3.215 | 3.025 | 2.4181
| Std Dev. |  | 0.25298 | 0.12449 | 0.074207 | 0.055287 | 0.072503 |  | 0.040825 | 0.40825 |  | 0.60553 | 0.08165 | 0.27653 | 0.23166 | 0.034375 | 0.17302 | 0.058109 | 0.047223 | 0.15332
| Rel.Std.Dev. |  | 6.325% | 3.718% | 2.361% | 4.601% | 1.399% |  | 1.172% | 12.89% |  | 22.71% | 2.378% | 8.579% | 15.62% | 1.055% | 5.018% | 1.806% | 1.561% | 6.406%
| PDM3 |  | 24.13% | 3.91% | -2.452% | -62.71% | 60.8% |  | 8.099% | -1.728% |  | -17.24% | 6.548% | 0.02936% | -53.97% | 1.113% | 7.013% | -0.1245% | -6.124% | -25.73%
| Z-Score (Absolute) |  | 2.01 | 0.32 | 0.21 | 5.24 | 5.08 | 7.06 | 0.67 | NA | 4.61 | 1.44 | 0.54 | 0 | 4.51 | 0.09 | 0.58 | 0.01 | 0.51 | 2.15
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 712 (Certified Value 0.611 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.5391 | 0.57562 | 0.61 | 0.61 | 0.6 | 0.64 | 0.58 | 0.655 | 0.66 | 0.61696 | 0.69 | 0.62 | 0.55 | 0.62 | 0.63 | 0.62 | 0.6121 |  |  | 1
| 1 | 2 | 0.534 | 0.57702 | 0.63 | 0.6 | 0.61 | 0.66 | 0.62 | 0.657 | 0.67 | 0.60624 | 0.7 | 0.615 | 0.57 | 0.62 | 0.63 | 0.61 | 0.5983 |  |  | #N/A
| 1 | 3 | 0.5374 | 0.57421 | 0.62 | 0.61 | 0.6 | 0.63 | 0.6 | 0.684 | 0.64 | 0.59977 | 0.68 | 0.636 | 0.55 | 0.6 | 0.62 | 0.61 | 0.6075 |  |  | 16
| 1 | 4 | 0.5358 | 0.57817 | 0.61 | 0.6 | 0.59 | 0.62 | 0.58 | 0.651 | 0.66 | 0.60919 | 0.69 | 0.622 | 0.56 | 0.6 | 0.61 | 0.62 | 0.6063 |  |  | 0.61059
| 1 | 5 | 0.5309 | 0.5738 | 0.62 | 0.62 | 0.6 | 0.64 | 0.62 | 0.675 | 0.66 | 0.60519 | 0.69 | 0.617 | 0.55 | 0.64 | 0.61 | 0.62 | 0.6052 |  |  | 59
| 1 | 6 | 0.534 | 0.57286 | 0.62 | 0.62 | 0.61 | 0.63 | 0.6 | 0.659 | 0.65 | 0.61064 | 0.69 | 0.617 | 0.57 | 0.6 | 0.62 | 0.62 | 0.6065
| Mean |  | 0.5352 | 0.57528 | 0.61833 | 0.61 | 0.60167 | 0.63667 | 0.6 | 0.6635 | 0.65667 | 0.608 | 0.69 | 0.62117 | 0.55833 | 0.61333 | 0.62 | 0.61667 | 0.60598
| Median |  | 0.5349 | 0.57492 | 0.62 | 0.61 | 0.6 | 0.635 | 0.6 | 0.658 | 0.66 | 0.60771 | 0.69 | 0.6185 | 0.555 | 0.61 | 0.62 | 0.62 | 0.6064
| Std Dev. |  | 0.0028907 | 0.0020365 | 0.0075277 | 0.0089443 | 0.0075277 | 0.013663 | 0.017889 | 0.012988 | 0.010328 | 0.0057847 | 0.0063246 | 0.007679 | 0.0098319 | 0.01633 | 0.0089443 | 0.005164 | 0.0044661
| Rel.Std.Dev. |  | 0.5401% | 0.354% | 1.217% | 1.466% | 1.251% | 2.146% | 2.981% | 1.958% | 1.573% | 0.9514% | 0.9166% | 1.236% | 1.761% | 2.662% | 1.443% | 0.8374% | 0.737%
| PDM3 |  | -12.35% | -5.783% | 1.268% | -0.09682% | -1.462% | 4.271% | -1.735% | 8.665% | 7.546% | -0.4247% | 13.01% | 1.732% | -8.559% | 0.4491% | 1.541% | 0.995% | -0.7547%
| Z-Score (Absolute) |  | 4.52 | 2.2 | 0.29 | 0.19 | 0.67 | 1.35 | 0.77 | 2.9 | 2.5 | 0.31 | 4.43 | 0.45 | 3.18 | 0 | 0.39 | 0.19 | 0.42
| Analytical results for Nb in OREAS 712 (Certified Value 105 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  |  | 0
0
| 1 | 1 | 123.15 | 99.1 | 107 | 21.3 | < 0.1 | 98.9 | 107.7 | 101 | 114 | 96.6 | 108.9287 | 112 | 107.5 | 108.5 | 24.5772 |  |  |  |  | 1
| 1 | 2 | 127.93 | 102 | 105.5 | 22 | 0.1 | 96.9 | 103.9 | 71 | 105 | 98.2 | 109.4232 | 107.5 | 107.5 | 105.5 | 16.2515 |  |  |  |  | #N/A
| 1 | 3 | 128.16 | 98.2 | 111 | 20.2 | 1 | 99.4 | 100.4 | 76 | 109 | 94.2 | 111.53 | 101.5 | 108 | 102.5 | 19.3012 |  |  |  |  | 16
| 1 | 4 | 139.17 | 97.3 | 107 | 20.7 | 0.1 | 99.7 | 108.6 | 56 | 113 | 102.7 | 110.1921 | 104.5 | 108 | 110.5 | 13.1436 |  |  |  |  | 105.1542
| 1 | 5 | 121.95 | 102 | 109.5 | 22.1 | 0.7 | 99.6 | 103.4 | 63.5 | 111 | 95.5 | 109.4496 | 109 | 106 | 111.5 | 21.0017 |  |  |  |  | 60
| 1 | 6 | 123.22 | 100 | 107.5 | 21.2 | 0.1 | 101 | 102.2 | 78.5 | 112 | 103.7 | 109.0258 | 102.5 | 107 | 108.5 | 22.4316
| Mean |  | 127.2633 | 99.7667 | 107.9167 | 21.25 | 0.4 | 99.25 | 104.3667 | 74.3333 | 110.6667 | 98.4833 | 109.7582 | 106.1667 | 107.3333 | 107.8333 | 19.4511
| Median |  | 125.575 | 99.55 | 107.25 | 21.25 | 0.1 | 99.5 | 103.65 | 73.5 | 111.5 | 97.4 | 109.4364 | 106 | 107.5 | 108.5 | 20.1515
| Std Dev. |  | 6.3956 | 1.95 | 1.9854 | 0.73417 | 0.42426 | 1.3457 | 3.1816 | 15.465 | 3.266 | 3.8953 | 0.97565 | 4.0456 | 0.75277 | 3.3267 | 4.1859
| Rel.Std.Dev. |  | 5.025% | 1.955% | 1.84% | 3.455% | 106.1% | 1.356% | 3.048% | 20.8% | 2.951% | 3.955% | 0.8889% | 3.811% | 0.7013% | 3.085% | 21.52%
| PDM3 |  | 21.03% | -5.123% | 2.627% | -79.79% | -99.62% | -5.615% | -0.7489% | -29.31% | 5.242% | -6.344% | 4.378% | 0.9629% | 2.072% | 2.548% | -81.5%
| Z-Score (Absolute) |  | 2.86 | 0.58 | 0.44 | 10.4 | 13.01 | 0.64 | 0 | 3.76 | 0.79 | 0.74 | 0.67 | 0.23 | 0.37 | 0.43 | 10.62
| Analytical results for Nd in OREAS 712 (Certified Value 158 ppm) | [20 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 0
0
| 1 | 1 | 193.8 | 168.2987 | 160 | 164 | 153.6 | 152 | 155 | 154.0518 | 159.5613 | 171.1659 | [9 empty] | 1
| 1 | 2 | 198.27 | 168.3476 | 151 | 158 | 157.9 | 156 | 155 | 154.2095 | 161.0519 | 155.3226 | [9 empty] | 5
| 1 | 3 | 193.98 | 169.1935 | 153 | 152 | 155.4 | 154 | 153 | 158.7413 | 160.0966 | 154.8978 | [9 empty] | 16
| 1 | 4 | 192.63 | 169.5594 | 153 | 158 | 151.8 | 153 | 157 | 166.9545 | 160.9576 | 157.1217 | [9 empty] | 158.1134
| 1 | 5 | 181.44 | 168.1888 | 155 | 154 | 153.1 | 155 | 161 | 150.0037 | 160.3223 | 160.2917 | [9 empty] | 61
| 1 | 6 | 181.08 | 168.7205 | 150 | 163 | 151.6 | 155 | 161 | 156.4625 | 160.5343 | 162.6692
| Mean |  | 190.2 | 168.7181 | 153.6667 | 158.1667 | 153.9 | 154.1667 | 157 | 156.7372 | 160.4207 | 160.2448
| Median |  | 193.215 | 168.5341 | 153 | 158 | 153.35 | 154.5 | 156 | 155.336 | 160.4283 | 158.7067
| Std Dev. |  | 7.186 | 0.55265 | 3.559 | 4.7504 | 2.395 | 1.472 | 3.3466 | 5.7866 | 0.55729 | 6.1274
| Rel.Std.Dev. |  | 3.778% | 0.3276% | 2.316% | 3.003% | 1.556% | 0.9548% | 2.132% | 3.692% | 0.3474% | 3.824%
| PDM3 |  | 20.29% | 6.707% | -2.812% | 0.03367% | -2.665% | -2.496% | -0.7042% | -0.8704% | 1.459% | 1.348%
| Z-Score (Absolute) |  | 7.03 | 2.4 | 0.84 | 0.13 | 0.79 | 0.74 | 0.13 | 0.18 | 0.61 | 0.57
| Analytical results for Ni in OREAS 712 (Certified Value 13.4 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 12 | 13.85 | 13 | 13.5 | 14.94 | 14.99 | 15 | 14 | 15 | 13.4242 | 15 | 14 | 9.1155 | 13.1 | 14.0472 | 15.5 | 14.1 | 13.2 | 13.8459 | 1
| 1 | 2 | 11.7 | 14.02 | 12.6 | 12.8 | 13.62 | 13.96 | 14 | 14 | < 10 | 13.2843 | 15 | 14 | 9.4164 | 13.8 | 13.6109 | 14.8 | 14.3 | 13 | 11.2468 | #N/A
| 1 | 3 | 11.5 | 13.39 | 13.9 | 13.6 | 13.39 | 14.37 | 13 | 14 | < 10 | 13.1002 | 15 | 15 | 9.6681 | 14.2 | 13.1726 | 13.6 | 14.4 | 13 | 12.0681 | 16
| 1 | 4 | 11.6 | 14.18 | 12.6 | 13 | 12.23 | 14.11 | 14 | 14 | 14 | 13.429 | 15 | 15 | 8.853 | 14.1 | 13.288 | 14.2 | 13.8 | 13.2 | 11.9312 | 13.4217
| 1 | 5 | 11.8 | 13.73 | 13.3 | 13.3 | 13.46 | 14.91 | 15 | 14 | 12 | 13.1257 | 15 | 15 | 8.8941 | 14.4 | 13.8503 | 15.2 | 14 | 13.6 | 12.2919 | 62
| 1 | 6 | 11.3 | 13.92 | 12.3 | 13.2 | 13.25 | 13.84 | 15 | 14 | < 10 | 13.4118 | 15 | 15 | 8.5712 | 14.6 | 13.3575 | 14 | 14.2 | 13 | 11.9468
| Mean |  | 11.65 | 13.8483 | 12.95 | 13.2333 | 13.4817 | 14.3633 | 14.3333 | 14 | 13.6667 | 13.2959 | 15 | 14.6667 | 9.0864 | 14.0333 | 13.5544 | 14.55 | 14.1333 | 13.1667 | 12.2218
| Median |  | 11.65 | 13.885 | 12.8 | 13.25 | 13.425 | 14.24 | 14.5 | 14 | 14 | 13.348 | 15 | 15 | 9.0048 | 14.15 | 13.4842 | 14.5 | 14.15 | 13.1 | 12.0074
| Std Dev. |  | 0.2429 | 0.27154 | 0.58224 | 0.30111 | 0.86892 | 0.48833 | 0.8165 | 0 | 1.5275 | 0.15166 | 0 | 0.5164 | 0.40126 | 0.53166 | 0.34326 | 0.73689 | 0.21602 | 0.23381 | 0.86913
| Rel.Std.Dev. |  | 2.085% | 1.961% | 4.496% | 2.275% | 6.445% | 3.4% | 5.696% | 0% | 11.18% | 1.141% | 0% | 3.521% | 4.416% | 3.789% | 2.532% | 5.065% | 1.528% | 1.776% | 7.111%
| PDM3 |  | -13.2% | 3.179% | -3.514% | -1.403% | 0.4468% | 7.016% | 6.792% | 4.309% | 1.825% | -0.9375% | 11.76% | 9.276% | -32.3% | 4.557% | 0.9886% | 8.407% | 5.302% | -1.9% | -8.94%
| Z-Score (Absolute) |  | 2.12 | 0.56 | 0.54 | 0.19 | 0.11 | 1.19 | NA | NA | NA | 0.11 | NA | NA | 5.25 | 0.79 | 0.2 | 1.42 | 0.91 | 0.27 | 1.43
Results from laboratories 10, 12, 13, 16 and 17 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 712 (Certified Value 0.076 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.0724 | 0.075 | 0.082 | 0.09 | 0.0634 | 0.025 | 0.081 | 0.08 | 0.076658 | 0.075 | 0.077 | 0.065 | 0.077 | 0.079 | 0.076 | 0.0591 |  |  |  | 1
| 1 | 2 | 0.0721 | 0.08 | 0.078 | 0.09 | 0.0632 | 0.035 | 0.08 | 0.08 | 0.074752 | 0.075 | 0.075 | 0.07 | 0.078 | 0.079 | 0.077 | 0.0718 |  |  |  | #N/A
| 1 | 3 | 0.0711 | 0.077 | 0.079 | 0.09 | 0.0613 | 0.037 | 0.08 | 0.08 | 0.076076 | 0.075 | 0.077 | 0.067 | 0.077 | 0.078 | 0.078 | 0.0721 |  |  |  | 16
| 1 | 4 | 0.0728 | 0.076 | 0.079 | 0.09 | 0.0643 | 0.027 | 0.08 | 0.08 | 0.074464 | 0.07 | 0.077 | 0.068 | 0.074 | 0.076 | 0.079 | 0.0655 |  |  |  | 0.075706
| 1 | 5 | 0.0709 | 0.079 | 0.08 | 0.09 | 0.0624 | 0.044 | 0.081 | 0.08 | 0.076597 | 0.07 | 0.075 | 0.065 | 0.083 | 0.077 | 0.078 | 0.0742 |  |  |  | 63
| 1 | 6 | 0.0717 | 0.079 | 0.084 | 0.09 | 0.0645 | 0.023 | 0.081 | 0.08 | 0.074169 | 0.075 | 0.077 | 0.069 | 0.074 | 0.078 | 0.076 | 0.0617
| Mean |  | 0.071833 | 0.077667 | 0.080333 | 0.09 | 0.063183 | 0.031833 | 0.0805 | 0.08 | 0.075453 | 0.073333 | 0.076333 | 0.067333 | 0.077167 | 0.077833 | 0.077333 | 0.0674
| Median |  | 0.0719 | 0.078 | 0.0795 | 0.09 | 0.0633 | 0.031 | 0.0805 | 0.08 | 0.075414 | 0.075 | 0.077 | 0.0675 | 0.077 | 0.078 | 0.0775 | 0.06865
| Std Dev. |  | 0.00074207 | 0.0019664 | 0.0022509 | 1.5202e-17 | 0.001199 | 0.0081588 | 0.00054772 | 0 | 0.0011196 | 0.002582 | 0.0010328 | 0.0020656 | 0.0033116 | 0.001169 | 0.0012111 | 0.0062077
| Rel.Std.Dev. |  | 1.033% | 2.532% | 2.802% | 1.689e-14% | 1.898% | 25.63% | 0.6804% | 0% | 1.484% | 3.521% | 1.353% | 3.068% | 4.291% | 1.502% | 1.566% | 9.21%
| PDM3 |  | -5.116% | 2.59% | 6.112% | 18.88% | -16.54% | -57.95% | 6.332% | 5.672% | -0.3348% | -3.134% | 0.8285% | -11.06% | 1.929% | 2.81% | 2.149% | -10.97%
| Z-Score (Absolute) |  | 0.95 | 0.18 | 0.69 | 2.55 | 2.61 | 8.65 | 0.72 | 0.63 | 0.25 | 0.66 | 0.08 | 1.81 | 0.08 | 0.21 | 0.11 | 1.8
| Analytical results for Pb in OREAS 712 (Certified Value 10.2 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 1
2
| 1 | 1 | < 3 | 9.2 | 9.9 | < 2 | 12.43 | 14 | 11 | 10.3525 | 11 | 9.9 | 17.69 | 10.8 | 9.8904 | 10.9 | 9.8 | 9.5 | 16.4838 |  |  | 1
| 1 | 2 | < 3 | 9.8 | 9.6 | < 2 | 11.75 | 14 | 10 | 10.2173 | 13 | 10 | 16.28 | 10.8 | 9.9113 | 10.3 | 9.9 | 9.9 | 16.9504 |  |  | 6
| 1 | 3 | < 3 | 9.3 | 10 | < 2 | 12 | 15 | 11 | 10.653 | 11 | 9.2 | 17.41 | 12.2 | 10.0256 | 9.8 | 10.2 | 9.1 | 18.2995 |  |  | 16
| 1 | 4 | < 3 | 9.8 | 9.7 | < 2 | 11.86 | 13 | 10 | 10.8185 | 11 | 9 | 16.9 | 10.7 | 10.0469 | 10 | 9.7 | 10 | 18.6259 |  |  | 10.195
| 1 | 5 | < 3 | 9.7 | 10 | < 2 | 12.3 | 15 | 10 | 10.4131 | 11 | 8.6 | 17.44 | 9.7 | 10.2346 | 10.9 | 9.8 | 10.2 | 17.6217 |  |  | 64
| 1 | 6 | < 3 | 9.6 | 10 | < 2 | 11.79 | 13 | 10 | 10.4027 | 11 | 8.7 | 16.26 | 11 | 10.1035 | 10 | 9.6 | 9.7 | 16.0675
| Mean |  |  | 9.5667 | 9.8667 |  | 12.0217 | 14 | 10.3333 | 10.4762 | 11.3333 | 9.2333 | 16.9967 | 10.8667 | 10.0354 | 10.3167 | 9.8333 | 9.7333 | 17.3415
| Median |  |  | 9.65 | 9.95 |  | 11.93 | 14 | 10 | 10.4079 | 11 | 9.1 | 17.155 | 10.8 | 10.0363 | 10.15 | 9.8 | 9.8 | 17.286
| Std Dev. |  |  | 0.2582 | 0.17512 |  | 0.28224 | 0.89443 | 0.5164 | 0.21918 | 0.8165 | 0.59554 | 0.61867 | 0.79917 | 0.12727 | 0.47924 | 0.20656 | 0.39328 | 1.0156
| Rel.Std.Dev. |  |  | 2.699% | 1.775% |  | 2.348% | 6.389% | 4.997% | 2.092% | 7.204% | 6.45% | 3.64% | 7.354% | 1.268% | 4.645% | 2.101% | 4.041% | 5.857%
| PDM3 |  |  | -6.163% | -3.22% |  | 17.92% | 37.32% | 1.357% | 2.758% | 11.17% | -9.433% | 66.72% | 6.588% | -1.565% | 1.193% | -3.547% | -4.528% | 70.1%
| Z-Score (Absolute) |  | 9.17 | 0.42 | 0.09 | 9.71 | 2.25 | NA | NA | 0.57 | NA | 0.78 | 7.65 | 0.99 | 0.09 | 0.4 | 0.13 | 0.24 | 8.02
Results from laboratories 10, 12 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 712 (Certified Value 47.7 ppm) | [20 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 50.84 | 48.9458 | 44.95 | 48.5 | 45.54 | 46.6 | 57.83 | 48.2966 | 47.1303 | 50.7345 | [9 empty] | 1
| 1 | 2 | 51.5 | 49.1694 | 43.66 | 50.5 | 47.58 | 48.5 | 57.66 | 48.5577 | 47.7757 | 48.3551 | [9 empty] | 7
| 1 | 3 | 50.26 | 49.3963 | 44.81 | 45.9 | 46.53 | 46.9 | 58.31 | 49.4883 | 46.3623 | 46.8292 | [9 empty] | 16
| 1 | 4 | 50.34 | 49.3163 | 44.61 | 47.3 | 45.87 | 46.6 | 58.19 | 51.3193 | 46.1086 | 49.2056 | [9 empty] | 47.675
| 1 | 5 | 47.18 | 49.1174 | 45.48 | 48.3 | 46.04 | 47 | 57.39 | 47.1566 | 47.3302 | 47.3506 | [9 empty] | 65
| 1 | 6 | 47.75 | 49.0838 | 44.06 | 49.3 | 46.57 | 47.5 | 57.52 | 45.7493 | 47.2366 | 47.9668
| Mean |  | 49.645 | 49.1715 | 44.595 | 48.3 | 46.355 | 47.1833 | 57.8167 | 48.428 | 46.9906 | 48.407
| Median |  | 50.3 | 49.1434 | 44.71 | 48.4 | 46.285 | 46.95 | 57.745 | 48.4272 | 47.1835 | 48.1609
| Std Dev. |  | 1.7549 | 0.16313 | 0.65056 | 1.59 | 0.7178 | 0.72503 | 0.36811 | 1.9136 | 0.62999 | 1.4041
| Rel.Std.Dev. |  | 3.535% | 0.3318% | 1.459% | 3.292% | 1.548% | 1.537% | 0.6367% | 3.951% | 1.341% | 2.901%
| PDM3 |  | 4.132% | 3.139% | -6.46% | 1.311% | -2.769% | -1.031% | 21.27% | 1.579% | -1.436% | 1.535%
| Z-Score (Absolute) |  | 0.71 | 0.45 | 2.06 | 0.03 | 1.09 | 0.64 | 5.18 | 0.04 | 0.75 | 0.03
| Analytical results for Rb in OREAS 712 (Certified Value 32.4 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 30.2 | 30.1872 | 32.3 | 32.3 | 32.25 | 36.3 | 30.6 | 31 | 35.6 | 33 | 33.5761 | 30.4 | 32.5638 | 37.9 | 31.9 | 33.9 | 28.9916 |  |  | 1
| 1 | 2 | 31.3 | 30.6951 | 31.6 | 34.6 | 30.69 | 29 | 29.6 | 32.5 | 36.4 | 32.9 | 32.6636 | 32.8 | 32.1601 | 36.3 | 32.3 | 34.3 | 29.3227 |  |  | 8
| 1 | 3 | 31.4 | 30.143 | 28.8 | 33.4 | 32.24 | 28.6 | 31.9 | 31.4 | 35.2 | 33.5 | 34.515 | 29.5 | 31.8068 | 34.3 | 31.8 | 34.1 | 28.2176 |  |  | 16
| 1 | 4 | 30.3 | 31.0179 | 29.6 | 32 | 32.13 | 33.1 | 30 | 32.1 | 35.6 | 33.9 | 34.8124 | 33.1 | 32.5157 | 35.7 | 31.4 | 35.8 | 29.7279 |  |  | 32.4438
| 1 | 5 | 30.2 | 30.8846 | 32.9 | 33 | 30.83 | 29 | 32 | 30.4 | 36 | 33.6 | 32.1581 | 35.9 | 30.1131 | 37.8 | 31.9 | 36.2 | 28.7249 |  |  | 66
| 1 | 6 | 30.3 | 30.4822 | 30.3 | 32.8 | 30.37 | 38.5 | 30.1 | 30.4 | 36 | 33.1 | 35.6826 | 37.8 | 32.8964 | 34.5 | 31.7 | 35.8 | 28.9244
| Mean |  | 30.6167 | 30.5683 | 30.9167 | 33.0167 | 31.4183 | 32.4167 | 30.7 | 31.3 | 35.8 | 33.3333 | 33.9013 | 33.25 | 32.0093 | 36.0833 | 31.8333 | 35.0167 | 28.9848
| Median |  | 30.3 | 30.5887 | 30.95 | 32.9 | 31.48 | 31.05 | 30.35 | 31.2 | 35.8 | 33.3 | 34.0456 | 32.95 | 32.3379 | 36 | 31.85 | 35.05 | 28.958
| Std Dev. |  | 0.57067 | 0.36106 | 1.6068 | 0.92177 | 0.87737 | 4.2537 | 1.0198 | 0.87178 | 0.41952 | 0.39328 | 1.3458 | 3.1665 | 1.001 | 1.5575 | 0.29439 | 1.0226 | 0.51502
| Rel.Std.Dev. |  | 1.864% | 1.181% | 5.197% | 2.792% | 2.793% | 13.12% | 3.322% | 2.785% | 1.172% | 1.18% | 3.97% | 9.523% | 3.127% | 4.316% | 0.9248% | 2.92% | 1.777%
| PDM3 |  | -5.632% | -5.781% | -4.707% | 1.766% | -3.161% | -0.08365% | -5.375% | -3.525% | 10.34% | 2.742% | 4.492% | 2.485% | -1.339% | 11.22% | -1.882% | 7.93% | -10.66%
| Z-Score (Absolute) |  | 0.72 | 0.74 | 0.56 | 0.52 | 0.3 | 0.21 | 0.67 | 0.37 | 1.95 | 0.68 | 0.97 | 0.64 | 0 | 2.1 | 0.09 | 1.55 | 1.56
| Analytical results for Re in OREAS 712 (Indicative Value 0.003 ppm) | [20 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.017 | 0.002 | 0.004 | 0.006 | 0.001 | < 0.05 | < 0.02 | < 0.1 | < 0.1 | 0.008 | 0.004 | 0.002 | 0.002 | [6 empty] | 1
| 1 | 2 | 0.018 | < 0.002 | 0.004 | 0.006 | < 0.001 | < 0.05 | < 0.02 | < 0.1 | < 0.1 | 0.012 | 0.002 | 0.002 | 0.002 | [6 empty] | 16
| 1 | 3 | 0.017 | 0.002 | 0.004 | 0.006 | 0.001 | < 0.05 | < 0.02 | < 0.1 | < 0.1 | 0.009 | 0.002 | 0.002 | 0.003 | [6 empty] | 16
| 1 | 4 | 0.015 | 0.002 | 0.004 | 0.006 | 0.001 | < 0.05 | < 0.02 | < 0.1 | < 0.1 | 0.011 | 0.003 | 0.002 | 0.003 | [6 empty] | 0.0030333
| 1 | 5 | 0.016 | 0.003 | 0.005 | 0.006 | < 0.001 | < 0.05 | < 0.02 | < 0.1 | < 0.1 | 0.011 | 0.003 | 0.002 | 0.004 | [6 empty] | 22
| 1 | 6 | 0.015 | 0.003 | 0.005 | 0.006 | 0.001 | < 0.05 | < 0.02 | < 0.1 | < 0.1 | 0.007 | 0.003 | 0.002 | 0.002
| Mean |  | 0.016333 | 0.0024 | 0.0043333 | 0.006 | 0.001 |  |  |  |  | 0.0096667 | 0.0028333 | 0.002 | 0.0026667
| Median |  | 0.0165 | 0.002 | 0.004 | 0.006 | 0.001 |  |  |  |  | 0.01 | 0.003 | 0.002 | 0.0025
| Std Dev. |  | 0.0012111 | 0.00054772 | 0.0005164 | 9.5015e-19 | 0 |  |  |  |  | 0.0019664 | 0.00075277 | 0 | 0.0008165
| Rel.Std.Dev. |  | 7.415% | 22.82% | 11.92% | 1.584e-14% | 0% |  |  |  |  | 20.34% | 26.57% | 0% | 30.62%
| PDM3 |  | 438.5% | -20.88% | 42.86% | 97.8% | -67.03% |  |  |  |  | 218.7% | -6.593% | -34.07% | -12.09%
| Z-Score (Absolute) |  | 1.74 | 0.65 | 0.28 | 0 | 0.87 | 3.2 | 0.67 | 7.42 | 7.42 | 0.62 | 0.53 | 0.67 | 0.56
| Analytical results for S in OREAS 712 (Certified Value < 0.01 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 17 | 19 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | < 0.003 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.002 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | [5 empty] | 1
| 1 | 2 | < 0.003 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.002 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | [5 empty] | 9
| 1 | 3 | < 0.003 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.004 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | [5 empty] | 16
| 1 | 4 | < 0.003 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.006 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | [5 empty] | < 0.01
| 1 | 5 | < 0.003 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.004 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | [5 empty] | 67
| 1 | 6 | < 0.003 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.004 | < 0.05 | < 0.01 | < 0.01 | < 0.01
| Mean |  |  |  | 0.01 | [6 empty] | 0.0036667
| Median |  |  |  | 0.01 | [6 empty] | 0.004
| Std Dev. |  |  |  | 0 | [6 empty] | 0.0015055
| Rel.Std.Dev. |  |  |  | 0% | [6 empty] | 41.06%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Sb in OREAS 712 (Certified Value 0.52 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.47 | 0.83 | 0.5 | < 1 | 8.64 | < 0.1 | 0.6 | 0.5 | < 2.5 | 0.6 | 0.66 | 0.9 | 0.52375 | 0.58 | 0.51 | 0.43 | 7.057 |  |  | 1
| 1 | 2 | 0.49 | 0.53 | 0.5 | < 1 | 8.45 | < 0.1 | 0.6 | < 0.5 | < 2.5 | 0.4 | 0.56 | 0.8 | 0.51997 | 0.51 | 0.52 | 0.4 | 8.6159 |  |  | 10
| 1 | 3 | 0.5 | 0.52 | 0.53 | < 1 | 8.36 | < 0.1 | 0.6 | < 0.5 | < 2.5 | 0.6 | 0.64 | 0.9 | 0.52418 | 0.52 | 0.52 | 0.44 | 8.5299 |  |  | 16
| 1 | 4 | 0.48 | 0.55 | 0.51 | < 1 | 8.54 | 0.1 | 0.6 | < 0.5 | < 2.5 | 0.4 | 0.57 | 0.9 | 0.52081 | 0.56 | 0.52 | 0.41 | 6.9041 |  |  | 0.51973
| 1 | 5 | 0.47 | 0.56 | 0.53 | < 1 | 8.87 | 0.1 | 0.6 | 0.5 | < 2.5 | 0.4 | 0.57 | 0.8 | 0.53058 | 0.59 | 0.5 | 0.33 | 7.354 |  |  | 68
| 1 | 6 | 0.47 | 0.63 | 0.52 | < 1 | 8.35 | 0.2 | 0.6 | < 0.5 | < 2.5 | 0.4 | 0.56 | 0.7 | 0.5336 | 0.51 | 0.52 | 0.45 | 6.5312
| Mean |  | 0.48 | 0.60333 | 0.515 |  | 8.535 | 0.13333 | 0.6 | 0.5 |  | 0.46667 | 0.59333 | 0.83333 | 0.52548 | 0.545 | 0.515 | 0.41 | 7.4987
| Median |  | 0.475 | 0.555 | 0.515 |  | 8.495 | 0.1 | 0.6 | 0.5 |  | 0.4 | 0.57 | 0.85 | 0.52397 | 0.54 | 0.52 | 0.42 | 7.2055
| Std Dev. |  | 0.012649 | 0.11759 | 0.013784 |  | 0.19766 | 0.057735 | 0 | 0 |  | 0.10328 | 0.044572 | 0.08165 | 0.0054552 | 0.036194 | 0.0083666 | 0.043359 | 0.87375
| Rel.Std.Dev. |  | 2.635% | 19.49% | 2.677% |  | 2.316% | 43.3% | 0% | 0% |  | 22.13% | 7.512% | 9.798% | 1.038% | 6.641% | 1.625% | 10.58% | 11.65%
| PDM3 |  | -7.644% | 16.09% | -0.9095% |  | 1542% | -74.35% | 15.45% | -3.796% |  | -10.21% | 14.16% | 60.34% | 1.107% | 4.863% | -0.9095% | -21.11% | 1343%
| Z-Score (Absolute) |  | 0.45 | 0.77 | 0.1 | 0.25 | 79.6 | 4.31 | NA | NA | 7.2 | NA | 0.67 | NA | 0 | 0.19 | 0.1 | 1.15 | 69.3
Results from laboratories 12, 13, 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 712 (Certified Value 37 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 33.41 | 38.43 | 34.9 | 38.4 | 39.4 | 37 | 36 | 37 | 35.2467 | 36 | 36.8 | 39.7105 | 39.3 | 34.4693 | 42.7 | 37.4 | 36.9 | 38.3144 |  | 1
| 1 | 2 | 33.22 | 36.99 | 33.5 | 36.6 | 39.19 | 34 | 37 | 38 | 36.6044 | 36 | 37 | 40.0447 | 40.8 | 32.641 | 40.2 | 36.7 | 38.3 | 38.465 |  | 11
| 1 | 3 | 33.17 | 37.06 | 32.9 | 38.7 | 39.49 | 32 | 37 | 36 | 34.4739 | 34 | 38.2 | 40.9938 | 38.4 | 34.4155 | 37.8 | 37.5 | 36.7 | 36.8587 |  | 16
| 1 | 4 | 33.26 | 37.75 | 32.7 | 37.7 | 39.99 | 34 | 38 | 37 | 35.5588 | 36 | 37.3 | 42.423 | 40.1 | 33.7411 | 38.8 | 37.1 | 39 | 39.7143 |  | 37.0017
| 1 | 5 | 33.03 | 37.75 | 35.5 | 38.2 | 39.89 | 33 | 37 | 38 | 34.5478 | 36 | 36.7 | 38.497 | 40.2 | 31.4953 | 41.5 | 37 | 39.6 | 37.2187 |  | 69
| 1 | 6 | 33.12 | 37.09 | 34.3 | 38.1 | 38.76 | 37 | 38 | 37 | 37.3567 | 36 | 36.6 | 38.2257 | 40.7 | 33.0393 | 39.1 | 37.2 | 37.9 | 39.1253
| Mean |  | 33.2017 | 37.5117 | 33.9667 | 37.95 | 39.4533 | 34.5 | 37.1667 | 37.1667 | 35.6314 | 35.6667 | 37.1 | 39.9824 | 39.9167 | 33.3003 | 40.0167 | 37.15 | 38.0667 | 38.2827
| Median |  | 33.195 | 37.42 | 33.9 | 38.15 | 39.445 | 34 | 37 | 37 | 35.4027 | 36 | 36.9 | 39.8776 | 40.15 | 33.3902 | 39.65 | 37.15 | 38.1 | 38.3897
| Std Dev. |  | 0.12983 | 0.56761 | 1.1219 | 0.73959 | 0.45452 | 2.0736 | 0.75277 | 0.75277 | 1.1481 | 0.8165 | 0.5933 | 1.5716 | 0.91524 | 1.1459 | 1.8258 | 0.2881 | 1.1431 | 1.0912
| Rel.Std.Dev. |  | 0.391% | 1.513% | 3.303% | 1.949% | 1.152% | 6.011% | 2.025% | 2.025% | 3.222% | 2.289% | 1.599% | 3.931% | 2.293% | 3.441% | 4.563% | 0.7755% | 3.003% | 2.85%
| PDM3 |  | -10.27% | 1.378% | -8.202% | 2.563% | 6.626% | -6.761% | 0.4459% | 0.4459% | -3.703% | -3.608% | 0.2657% | 8.056% | 7.878% | -10% | 8.148% | 0.4009% | 2.878% | 3.462%
| Z-Score (Absolute) |  | 1.76 | 0.15 | 1.42 | 0.35 | 1.02 | 1.18 | 0 | 0 | 0.68 | 0.67 | 0.03 | 1.25 | 1.22 | 1.72 | 1.27 | 0.01 | 0.4 | 0.5
| Analytical results for Se in OREAS 712 (Indicative Value 1.54 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
2
| 1 | 1 | 2.04 | 3.2969 | 1 | 2 | 1.64 | 0.6 | < 1 | < 5 | < 5 | 5 | 10 | 5.2749 | 1 | 1 | 2 | 2.1699 |  |  |  | 1
| 1 | 2 | 2.15 | 3.3747 | < 1 | 1 | 1.67 | < 0.1 | < 1 | < 5 | < 5 | 5 | 9 | 5.5006 | 1 | < 1 | 1 | 2.8287 |  |  |  | 12
| 1 | 3 | 2.18 | 3.4014 | 1 | 1 | 1.72 | < 0.1 | < 1 | < 5 | < 5 | 5 | 10 | 5.4042 | 1 | 1 | 1 | 2.3819 |  |  |  | 16
| 1 | 4 | 1.94 | 3.3573 | 1 | 1 | 1.73 | 0.8 | < 1 | < 5 | < 5 | 5 | 10 | 5.4474 | 1 | 1 | 1 | 2.2792 |  |  |  | 1.5411
| 1 | 5 | 2.05 | 3.3319 | 1 | 2 | 1.68 | < 0.1 | < 1 | < 5 | < 5 | 5 | 10 | 4.9205 | 1 | 1 | 1 | 1.4496 |  |  |  | 23
| 1 | 6 | 1.95 | 3.3477 | 1 | 1 | 1.57 | 0.7 | < 1 | < 5 | < 5 | 5 | 9 | 5.0445 | 1 | 1 | 1 | 1.725
| Mean |  | 2.0517 | 3.3517 | 1 | 1.3333 | 1.6683 | 0.7 |  |  |  | 5 | 9.6667 | 5.2654 | 1 | 1 | 1.1667 | 2.139
| Median |  | 2.045 | 3.3525 | 1 | 1 | 1.675 | 0.7 |  |  |  | 5 | 10 | 5.3396 | 1 | 1 | 1 | 2.2245
| Std Dev. |  | 0.099079 | 0.035905 | 0 | 0.5164 | 0.058452 | 0.1 |  |  |  | 0 | 0.5164 | 0.23475 | 0 | 0 | 0.40825 | 0.4903
| Rel.Std.Dev. |  | 4.829% | 1.071% | 0% | 38.73% | 3.504% | 14.29% |  |  |  | 0% | 5.342% | 4.458% | 0% | 0% | 34.99% | 22.92%
| PDM3 |  | 33.13% | 117.5% | -35.11% | -13.48% | 8.258% | -54.58% |  |  |  | 224.4% | 527.3% | 241.7% | -35.11% | -35.11% | -24.3% | 38.8%
| Z-Score (Absolute) |  | 0.14 | 1.12 | 0.71 | 0.39 | 0.14 | 1.11 | 1.02 | 0.48 | 0.48 | 2.35 | 5.84 | 2.55 | 0.64 | 0.71 | 0.52 | 0.21
| Analytical results for Sm in OREAS 712 (Certified Value 23.5 ppm) | [20 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 24.37 | 25.0955 | 24.28 | 25.1 | 22.2 | 23.5 | 24.3 | 23.1476 | 23.1944 | 24.3401 | [9 empty] | 1
| 1 | 2 | 25.01 | 25.1416 | 23 | 24.6 | 22.6 | 24.2 | 24.1 | 23.1332 | 23.3256 | 21.9631 | [9 empty] | 24
| 1 | 3 | 24.66 | 25.4945 | 23.2 | 23.5 | 23 | 23.6 | 24.3 | 23.5819 | 23.1583 | 22.115 | [9 empty] | 16
| 1 | 4 | 24.4 | 25.3443 | 24 | 23.3 | 22.4 | 23.7 | 24.1 | 24.143 | 23.5662 | 21.9036 | [9 empty] | 23.4807
| 1 | 5 | 23.02 | 25.1381 | 24.1 | 23.5 | 22 | 24 | 23.9 | 22.7853 | 23.1924 | 22.7502 | [9 empty] | 70
| 1 | 6 | 22.93 | 25.5346 | 23.16 | 24.5 | 22.5 | 24 | 23.5 | 22.4605 | 23.8409 | 22.8251
| Mean |  | 24.065 | 25.2914 | 23.6233 | 24.0833 | 22.45 | 23.8333 | 24.0333 | 23.2086 | 23.3796 | 22.6495
| Median |  | 24.385 | 25.243 | 23.6 | 24 | 22.45 | 23.85 | 24.1 | 23.1404 | 23.26 | 22.4326
| Std Dev. |  | 0.87551 | 0.19364 | 0.56262 | 0.74409 | 0.34496 | 0.27325 | 0.30111 | 0.59342 | 0.27159 | 0.91786
| Rel.Std.Dev. |  | 3.638% | 0.7656% | 2.382% | 3.09% | 1.537% | 1.147% | 1.253% | 2.557% | 1.162% | 4.052%
| PDM3 |  | 2.489% | 7.712% | 0.6076% | 2.567% | -4.389% | 1.502% | 2.354% | -1.159% | -0.4303% | -3.54%
| Z-Score (Absolute) |  | 0.65 | 3 | 0.2 | 0.68 | 2.45 | 0.2 | 0.58 | 1 | 0.67 | 2.07
| Analytical results for Sn in OREAS 712 (Certified Value 3.95 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 4.02 | 4 | 4.1 | < 5 | 4.72 | < 1 | 4.2 | 4 | 4.007 | 4 | 3.6 | 2.7155 | 3.6 | 3.9682 | 4.7 | 3.8 | 3.7 | 3.3284 |  | 1
| 1 | 2 | 4.13 | 4.2 | 3.9 | < 5 | 4.49 | < 1 | 4 | 4 | 3.8077 | 4 | 3.9 | 2.6516 | 3.5 | 4.0529 | 4.4 | 3.8 | 3.7 | 2.5958 |  | 2
| 1 | 3 | 4.18 | 3.9 | 4.1 | < 5 | 4.76 | < 1 | 4.1 | 3 | 3.8845 | 4 | 3.8 | 7.7308 | 3.6 | 4.0598 | 4.2 | 3.9 | 3.6 | 3.1762 |  | 16
| 1 | 4 | 4.35 | 3.9 | 4 | < 5 | 4.66 | < 1 | 4.1 | 4 | 4.1201 | 4 | 3.9 | 3.1887 | 3.8 | 4.1165 | 4.3 | 3.8 | 3.7 | 3.289 |  | 3.9475
| 1 | 5 | 3.99 | 4.1 | 4.1 | < 5 | 4.53 | < 1 | 4.1 | 3 | 3.739 | 4 | 3.6 | 2.3267 | 3.5 | 4.0857 | 4.6 | 3.8 | 3.5 | 3.2957 |  | 71
| 1 | 6 | 3.99 | 3.9 | 4.1 | < 5 | 4.64 | < 1 | 4 | 4 | 4.0921 | 4 | 3.6 | 2.2437 | 3.7 | 4.054 | 4.3 | 3.7 | 3.7 | 3.0398
| Mean |  | 4.11 | 4 | 4.05 |  | 4.6333 |  | 4.0833 | 3.6667 | 3.9417 | 4 | 3.7333 | 3.4762 | 3.6167 | 4.0562 | 4.4167 | 3.8 | 3.65 | 3.1208
| Median |  | 4.075 | 3.95 | 4.1 |  | 4.65 |  | 4.1 | 4 | 3.9457 | 4 | 3.7 | 2.6835 | 3.6 | 4.0569 | 4.35 | 3.8 | 3.7 | 3.2326
| Std Dev. |  | 0.14128 | 0.12649 | 0.083666 |  | 0.10539 |  | 0.075277 | 0.5164 | 0.15557 | 0 | 0.15055 | 2.1111 | 0.1169 | 0.049566 | 0.19408 | 0.063246 | 0.083666 | 0.27829
| Rel.Std.Dev. |  | 3.437% | 3.162% | 2.066% |  | 2.275% |  | 1.844% | 14.08% | 3.947% | 0% | 4.033% | 60.73% | 3.232% | 1.222% | 4.394% | 1.664% | 2.292% | 8.917%
| PDM3 |  | 4.117% | 1.331% | 2.597% |  | 17.37% |  | 3.442% | -7.113% | -0.1453% | 1.331% | -5.425% | -11.94% | -8.38% | 2.754% | 11.89% | -3.736% | -7.536% | -20.94%
| Z-Score (Absolute) |  | 0.7 | 0.38 | 0.53 | 4.02 | 2.24 | 9.88 | 0.62 | NA | 0.21 | NA | 0.4 | 1.16 | 0.75 | 0.54 | 1.6 | 0.21 | 0.65 | 2.2
Results from laboratories 13 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 712 (Certified Value 362 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 358 | 371.456 | 376 | 387 | 337.5 | 372 | 354 | 370 | 335.8 | 340.3007 | 359 | 363 | 349.471 | 354 | 345.7758 | 385 | 390 | 389 | 490.3224 | 1
| 1 | 2 | 355 | 369.8435 | 393 | 377 | 331.8 | 361 | 362 | 355 | 346.5 | 337.3562 | 368 | 363 | 356.676 | 365 | 337.2581 | 388 | 389 | 380 | 504.1827 | #N/A
| 1 | 3 | 357 | 370.406 | 384 | 388 | 335.1 | 365 | 341 | 364 | 335.5 | 336.2129 | 358 | 374 | 367.2479 | 353 | 349.3697 | 377 | 384 | 379 | 487.5822 | 16
| 1 | 4 | 357 | 374.092 | 376 | 377 | 330.2 | 357 | 353 | 361 | 344.4 | 346.12 | 358 | 364 | 350.4001 | 355 | 342.0429 | 371 | 374 | 387 | 470.5044 | 361.7672
| 1 | 5 | 354 | 373.0755 | 391 | 393 | 337.4 | 366 | 355 | 362 | 339 | 322.2279 | 361 | 361 | 343.7998 | 355 | 326.3782 | 406 | 376 | 390 | 485.8779 | 72
| 1 | 6 | 354 | 368.943 | 388 | 396 | 335 | 354 | 356 | 358 | 337.3 | 344.2223 | 363 | 361 | 333.5926 | 358 | 343.0923 | 370 | 383 | 388 | 489.6695
| Mean |  | 355.8333 | 371.3027 | 384.6667 | 386.3333 | 334.5 | 362.5 | 353.5 | 361.6667 | 339.75 | 337.74 | 361.1667 | 364.3333 | 350.1979 | 356.6667 | 340.6528 | 382.8333 | 382.6667 | 385.5 | 488.0232
| Median |  | 356 | 370.931 | 386 | 387.5 | 335.05 | 363 | 354.5 | 361.5 | 338.15 | 338.8285 | 360 | 363 | 349.9355 | 355 | 342.5676 | 381 | 383.5 | 387.5 | 488.6259
| Std Dev. |  | 1.7224 | 1.9722 | 7.3666 | 7.9415 | 2.9597 | 6.5345 | 6.892 | 5.164 | 4.635 | 8.5077 | 3.8687 | 4.8854 | 11.401 | 4.4121 | 8.0672 | 13.4672 | 6.5625 | 4.7645 | 10.778
| Rel.Std.Dev. |  | 0.484% | 0.5312% | 1.915% | 2.056% | 0.8848% | 1.803% | 1.95% | 1.428% | 1.364% | 2.519% | 1.071% | 1.341% | 3.256% | 1.237% | 2.368% | 3.518% | 1.715% | 1.236% | 2.209%
| PDM3 |  | -1.64% | 2.636% | 6.33% | 6.791% | -7.537% | 0.2026% | -2.285% | -0.0278% | -6.086% | -6.642% | -0.166% | 0.7093% | -3.198% | -1.41% | -5.836% | 5.823% | 5.777% | 6.56% | 34.9%
| Z-Score (Absolute) |  | 0.19 | 0.31 | 0.74 | 0.79 | 0.87 | 0.03 | 0.26 | 0 | 0.7 | 0.77 | 0.02 | 0.09 | 0.37 | 0.16 | 0.67 | 0.68 | 0.67 | 0.77 | 4.06
| Analytical results for Ta in OREAS 712 (Certified Value 10.2 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 13.15 | 10.15 | 10.55 | 2.16 | < 0.1 | 9.66 | 9.6 | 7.6 | < 20 | 9.5 | 10.3279 | 11.85 | 10.2 | 9.99 | [5 empty] | 1
| 1 | 2 | 14.37 | 10.8 | 10.45 | 2.18 | < 0.1 | 9.59 | 9.6 | 5.6 | < 20 | 9.9 | 10.513 | 11 | 10.3 | 10.6 | [5 empty] | 3
| 1 | 3 | 14.23 | 10.8 | 10.85 | 2.29 | < 0.1 | 9.85 | 9.5 | 7 | < 20 | 9.4 | 10.2545 | 10.45 | 10.25 | 10.15 | [5 empty] | 16
| 1 | 4 | 13.99 | 10.55 | 10.65 | 2.25 | < 0.1 | 9.44 | 9.5 | 4.4 | < 20 | 10.1 | 10.733 | 10.45 | 10.1 | 10.45 | [5 empty] | 10.2398
| 1 | 5 | 13.22 | 11 | 10.7 | 2.13 | < 0.1 | 9.84 | 9.4 | 5.1 | < 20 | 9.1 | 10.422 | 11.35 | 10.2 | 10.45 | [5 empty] | 73
| 1 | 6 | 13.38 | 10.65 | 11.1 | 2.15 | < 0.1 | 9.67 | 9.8 | 7 | < 20 | 10.1 | 10.4089 | 10.55 | 10.15 | 10
| Mean |  | 13.7233 | 10.6583 | 10.7167 | 2.1933 |  | 9.675 | 9.5667 | 6.1167 |  | 9.6833 | 10.4432 | 10.9417 | 10.2 | 10.2733
| Median |  | 13.685 | 10.725 | 10.675 | 2.17 |  | 9.665 | 9.55 | 6.3 |  | 9.7 | 10.4154 | 10.775 | 10.2 | 10.3
| Std Dev. |  | 0.53776 | 0.29226 | 0.23166 | 0.062823 |  | 0.15527 | 0.13663 | 1.2656 |  | 0.41191 | 0.16694 | 0.57133 | 0.070711 | 0.26051
| Rel.Std.Dev. |  | 3.919% | 2.742% | 2.162% | 2.864% |  | 1.605% | 1.428% | 20.69% |  | 4.254% | 1.599% | 5.222% | 0.6932% | 2.536%
| PDM3 |  | 34.02% | 4.087% | 4.657% | -78.58% |  | -5.516% | -6.574% | -40.27% |  | -5.434% | 1.987% | 6.854% | -0.3887% | 0.3275%
| Z-Score (Absolute) |  | 4.48 | 0.69 | 0.76 | 9.77 | 12.42 | 0.53 | 0.66 | 4.92 | 0.12 | 0.51 | 0.42 | 1.04 | 0.12 | 0.21
| Analytical results for Tb in OREAS 712 (Certified Value 2.49 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | < 0.2 | 3.957 | 2.54 | 2.5 | 2.54 | 2.42 | 2.57 | 2.4396 | 2.5904 | 2.214 | [9 empty] | 1
| 1 | 2 | < 0.2 | 3.9806 | 2.41 | 2.6 | 2.61 | 2.44 | 2.52 | 2.4322 | 2.6154 | 1.9599 | [9 empty] | 25
| 1 | 3 | < 0.2 | 4.0035 | 2.36 | 2.3 | 2.54 | 2.42 | 2.55 | 2.506 | 2.5607 | 1.9753 | [9 empty] | 16
| 1 | 4 | < 0.2 | 4.0566 | 2.38 | 2.5 | 2.41 | 2.4 | 2.52 | 2.5281 | 2.6839 | 1.8745 | [9 empty] | 2.4885
| 1 | 5 | < 0.2 | 3.9679 | 2.4 | 2.5 | 2.52 | 2.46 | 2.48 | 2.3856 | 2.6155 | 2.0053 | [9 empty] | 74
| 1 | 6 | < 0.2 | 4.1448 | 2.4 | 2.4 | 2.61 | 2.44 | 2.48 | 2.3146 | 2.6259 | 1.8624
| Mean |  |  | 4.0184 | 2.415 | 2.4667 | 2.5383 | 2.43 | 2.52 | 2.4343 | 2.6153 | 1.9819
| Median |  |  | 3.992 | 2.4 | 2.5 | 2.54 | 2.43 | 2.52 | 2.4359 | 2.6154 | 1.9676
| Std Dev. |  |  | 0.07129 | 0.063797 | 0.10328 | 0.073598 | 0.020976 | 0.036332 | 0.078343 | 0.041026 | 0.12709
| Rel.Std.Dev. |  |  | 1.774% | 2.642% | 4.187% | 2.899% | 0.8632% | 1.442% | 3.218% | 1.569% | 6.413%
| PDM3 |  |  | 61.48% | -2.954% | -0.8781% | 2.002% | -2.352% | 1.265% | -2.177% | 5.094% | -20.36%
| Z-Score (Absolute) |  | 20.15 | 13.44 | 0.3 | 0.14 | 0.75 | 0.18 | 0.6 | 0.14 | 1.41 | 4.02
| Analytical results for Te in OREAS 712 (Certified Value < 0.05 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | < 0.03 | < 0.05 | < 0.05 | 0.15 | < 0.1 | < 0.1 | < 0.5 | < 5 | < 0.2 | 0.09 | < 0.1 | 0.01815 | < 0.05 | < 0.05 | < 0.05 |  |  |  |  | 1
| 1 | 2 | 0.05 | < 0.05 | < 0.05 | 0.15 | < 0.1 | < 0.1 | < 0.5 | < 5 | < 0.2 | < 0.05 | 0.2 | 0.02485 | 0.06 | < 0.05 | < 0.05 |  |  |  |  | 5
| 1 | 3 | < 0.03 | < 0.05 | < 0.05 | 0.14 | < 0.1 | < 0.1 | < 0.5 | < 5 | < 0.2 | < 0.05 | < 0.1 | 0.02258 | 0.07 | < 0.05 | < 0.05 |  |  |  |  | 16
| 1 | 4 | < 0.03 | < 0.05 | < 0.05 | 0.15 | < 0.1 | < 0.1 | < 0.5 | < 5 | < 0.2 | < 0.05 | < 0.1 | 0.02266 | < 0.05 | < 0.05 | < 0.05 |  |  |  |  | < 0.05
| 1 | 5 | < 0.03 | < 0.05 | < 0.05 | 0.15 | < 0.1 | < 0.1 | < 0.5 | < 5 | < 0.2 | < 0.05 | 0.1 | 0.02776 | < 0.05 | < 0.05 | < 0.05 |  |  |  |  | 75
| 1 | 6 | < 0.03 | < 0.05 | < 0.05 | 0.15 | < 0.1 | < 0.1 | < 0.5 | < 5 | < 0.2 | < 0.05 | < 0.1 | 0.03007 | 0.06 | < 0.05 | < 0.05
| Mean |  | 0.05 |  |  | 0.14833 | [5 empty] | 0.09 | 0.15 | 0.024345 | 0.063333
| Median |  | 0.05 |  |  | 0.15 | [5 empty] | 0.09 | 0.15 | 0.023755 | 0.06
| Std Dev. |  |  |  |  | 0.0040825 | [6 empty] | 0.070711 | 0.0042196 | 0.0057735
| Rel.Std.Dev. |  |  |  |  | 2.752% | [6 empty] | 47.14% | 17.33% | 9.116%
PDM3
| Z-Score (Absolute) |  | 0.82 | 0.67 | 0.67 | 3.66 | 0.21 | 0.21 | 7.24 | 86.4 | 1.96 | 0.29 | 1.38 | 0.7 | 0 | 0.67 | 0.67
| Analytical results for Th in OREAS 712 (Certified Value 72 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 77.18 | 76.3 | 71.1 | 67 | 72 | 74.4 | 70.3 | 69.6 | 65 | 69.3358 | 79 | 71.1498 | 85.5 | 69.9 | 70.1 | 76.2788 |  |  |  | 1
| 1 | 2 | 78.28 | 81.6 | 68.8 | 67 | 69.5 | 74.1 | 71.1 | 72 | 65.87 | 68.3707 | 81.3 | 72.3529 | 79.5 | 71.6 | 77 | 62.7569 |  |  |  | 26
| 1 | 3 | 76.5 | 77.4 | 72.6 | 67 | 62.9 | 80.3 | 69.3 | 70.2 | 64.4 | 70.6829 | 79.9 | 70.4421 | 74 | 73 | 67.5 | 63.1495 |  |  |  | 16
| 1 | 4 | 74.13 | 78.5 | 69 | 69 | 66.6 | 74.7 | 68.9 | 69.5 | 67.09 | 73.0034 | 82.1 | 71.669 | 77.5 | 69.4 | 75.9 | 61.3909 |  |  |  | 72.1164
| 1 | 5 | 72.39 | 79.5 | 69.9 | 68 | 70.8 | 74.2 | 69.5 | 71.2 | 67.35 | 66.4855 | 81.6 | 72.0064 | 84.5 | 69.9 | 74.8 | 72.1098 |  |  |  | 76
| 1 | 6 | 72.78 | 80.4 | 73.2 | 66 | 72 | 75.1 | 70.6 | 72.3 | 70 | 65.5756 | 84.1 | 71.9847 | 76 | 69.6 | 72.9 | 59.2605
| Mean |  | 75.21 | 78.95 | 70.7667 | 67.3333 | 68.9667 | 75.4667 | 69.95 | 70.8 | 66.6183 | 68.909 | 81.3333 | 71.6008 | 79.5 | 70.5667 | 73.0333 | 65.8244
| Median |  | 75.315 | 79 | 70.5 | 67 | 70.15 | 74.55 | 69.9 | 70.7 | 66.48 | 68.8532 | 81.45 | 71.8269 | 78.5 | 69.9 | 73.85 | 62.9532
| Std Dev. |  | 2.4492 | 1.9522 | 1.8511 | 1.0328 | 3.5904 | 2.3956 | 0.84794 | 1.2116 | 2.0146 | 2.7338 | 1.7784 | 0.69686 | 4.6368 | 1.4264 | 3.6418 | 6.7543
| Rel.Std.Dev. |  | 3.257% | 2.473% | 2.616% | 1.534% | 5.206% | 3.174% | 1.212% | 1.711% | 3.024% | 3.967% | 2.187% | 0.9733% | 5.832% | 2.021% | 4.986% | 10.26%
| PDM3 |  | 4.29% | 9.476% | -1.872% | -6.632% | -4.368% | 4.646% | -3.004% | -1.825% | -7.624% | -4.448% | 12.78% | -0.7149% | 10.24% | -2.149% | 1.271% | -8.725%
| Z-Score (Absolute) |  | 1.05 | 1.93 | 0 | 0.82 | 0.43 | 1.11 | 0.2 | 0 | 0.99 | 0.44 | 2.5 | 0.19 | 2.06 | 0.05 | 0.53 | 1.17
| Analytical results for Ti in OREAS 712 (Certified Value 7.06 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.626 | 7.2569 | 6.98 | 7.3 | 6.71 | 7.19 | 6.75 | 7.26 | 6.9789 | 6.7381 | 7.22 | 6.876 | 7.3777 | > 1 | 7.0117 | 7.41 | 7.51 | 7.36 | 6.5301 | 1
| 1 | 2 | 5.5118 | 7.251 | 7.01 | 7.14 | 6.64 | 7 | 6.55 | 7.12 | 7.1844 | 6.7252 | 7.07 | 6.895 | 7.3948 | > 1 | 7.0178 | 7.44 | 7.43 | 7.16 | 6.581 | 20
| 1 | 3 | 5.5427 | 7.2569 | 6.96 | 7.39 | 6.73 | 6.99 | 6.85 | 7.19 | 6.8943 | 6.6409 | 7.16 | 7.078 | 7.3089 | > 1 | 7.0323 | 7.36 | 7.39 | 7.28 | 6.6804 | 16
| 1 | 4 | 5.7563 | 7.251 | 7.03 | 7.17 | 6.59 | 7.05 | 6.94 | 7.07 | 6.9813 | 6.9107 | 7.16 | 6.919 | 7.3638 | > 1 | 6.9561 | 7.18 | 7.39 | 7.25 | 6.4416 | 7.0642
| 1 | 5 | 5.4921 | 7.2628 | 6.97 | 7.42 | 6.69 | 7.06 | 6.92 | 7.26 | 7.0619 | 6.5272 | 6.72 | 6.904 | 7.3266 | > 1 | 7.1242 | 7.79 | 7.37 | 7.34 | 6.8377 | 77
| 1 | 6 | 5.4736 | 7.2628 | 7 | 7.47 | 6.69 | 6.89 | 6.42 | 7.17 | 6.9938 | 6.754 | 7.35 | 6.918 | 7.313 | > 1 | 7.0481 | 7.16 | 7.37 | 7.4 | 6.3638
| Mean |  | 5.5671 | 7.2569 | 6.9917 | 7.315 | 6.675 | 7.03 | 6.7383 | 7.1783 | 7.0158 | 6.716 | 7.1133 | 6.9317 | 7.3475 |  | 7.0317 | 7.39 | 7.41 | 7.2983 | 6.5724
| Median |  | 5.5273 | 7.2569 | 6.99 | 7.345 | 6.69 | 7.025 | 6.8 | 7.18 | 6.9876 | 6.7317 | 7.16 | 6.911 | 7.3452 |  | 7.0251 | 7.385 | 7.39 | 7.31 | 6.5556
| Std Dev. |  | 0.10704 | 0.0053073 | 0.026394 | 0.13605 | 0.051284 | 0.098995 | 0.21123 | 0.075741 | 0.098335 | 0.12745 | 0.2137 | 0.073443 | 0.036144 |  | 0.055027 | 0.22839 | 0.053666 | 0.086814 | 0.17002
| Rel.Std.Dev. |  | 1.923% | 0.07313% | 0.3775% | 1.86% | 0.7683% | 1.408% | 3.135% | 1.055% | 1.402% | 1.898% | 3.004% | 1.06% | 0.4919% |  | 0.7826% | 3.09% | 0.7242% | 1.19% | 2.587%
| PDM3 |  | -21.19% | 2.729% | -1.026% | 3.551% | -5.509% | -0.4835% | -4.612% | 1.616% | -0.685% | -4.928% | 0.6962% | -1.876% | 4.011% |  | -0.4593% | 4.613% | 4.896% | 3.315% | -6.961%
| Z-Score (Absolute) |  | 3.58 | 0.55 | 0.1 | 0.69 | 0.87 | 0 | 0.72 | 0.36 | 0.04 | 0.77 | 0.2 | 0.24 | 0.77 | NA | 0 | 0.88 | 0.93 | 0.65 | 1.12
| Analytical results for Tl in OREAS 712 (Certified Value 0.19 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.17 | 0.17 | 0.19 | 0.2 | 0.18 | 0.18 | < 0.2 | 0.2 | 0.24 | 0.1 | 0.18468 | 0.19 | 0.18 | 0.18 | 0.2933 |  |  |  |  | 1
| 1 | 2 | 0.183 | 0.17 | 0.19 | 0.2 | 0.18 | 0.18 | < 0.2 | 0.2 | 0.23 | 0.2 | 0.18464 | 0.19 | 0.17 | 0.18 | 0.24 |  |  |  |  | 7
| 1 | 3 | 0.179 | 0.16 | 0.19 | 0.2 | 0.18 | 0.18 | < 0.2 | 0.2 | 0.23 | 0.2 | 0.18794 | 0.19 | 0.18 | 0.17 | 0.3171 |  |  |  |  | 16
| 1 | 4 | 0.17 | 0.18 | 0.19 | 0.2 | 0.18 | 0.17 | < 0.2 | 0.2 | 0.21 | 0.2 | 0.18507 | 0.19 | 0.18 | 0.19 | 0.2418 |  |  |  |  | 0.18591
| 1 | 5 | 0.162 | 0.18 | 0.19 | 0.2 | 0.19 | 0.17 | < 0.2 | 0.2 | 0.2 | 0.2 | 0.19239 | 0.2 | 0.18 | 0.2 | 0.2692 |  |  |  |  | 78
| 1 | 6 | 0.165 | 0.18 | 0.18 | 0.2 | 0.18 | 0.16 | < 0.2 | 0.2 | 0.22 | 0.2 | 0.18805 | 0.19 | 0.17 | 0.18 | 0.2281
| Mean |  | 0.1715 | 0.17333 | 0.18833 | 0.2 | 0.18167 | 0.17333 |  | 0.2 | 0.22167 | 0.18333 | 0.18713 | 0.19167 | 0.17667 | 0.18333 | 0.26492
| Median |  | 0.17 | 0.175 | 0.19 | 0.2 | 0.18 | 0.175 |  | 0.2 | 0.225 | 0.2 | 0.18651 | 0.19 | 0.18 | 0.18 | 0.2555
| Std Dev. |  | 0.0080685 | 0.008165 | 0.0040825 | 3.0405e-17 | 0.0040825 | 0.008165 |  | 3.0405e-17 | 0.01472 | 0.040825 | 0.0030205 | 0.0040825 | 0.005164 | 0.010328 | 0.034807
| Rel.Std.Dev. |  | 4.705% | 4.711% | 2.168% | 1.52e-14% | 2.247% | 4.711% |  | 1.52e-14% | 6.64% | 22.27% | 1.614% | 2.13% | 2.923% | 5.633% | 13.14%
| PDM3 |  | -7.749% | -6.763% | 1.306% | 7.581% | -2.28% | -6.763% |  | 7.581% | 19.24% | -1.384% | 0.6577% | 3.099% | -4.97% | -1.384% | 42.5%
| Z-Score (Absolute) |  | 0.8 | 0.67 | 0.34 | 1.12 | 0.11 | 0.67 | 5.62 | NA | 2.58 | NA | 0.26 | 0.56 | 0.45 | 0 | 5.5
Results from laboratories 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 712 (Certified Value 3 ppm) | [20 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 3.38 | 3.147 | 3.06 | 2.9 | 2.9 | 2.84 | 2.87 | 2.7468 | 3.2244 | 2.9167 | [9 empty] | 1
| 1 | 2 | 3.41 | 3.2022 | 2.91 | 2.9 | 3 | 2.94 | 2.85 | 2.7107 | 3.2899 | 2.7456 | [9 empty] | 8
| 1 | 3 | 3.39 | 3.3078 | 2.95 | 2.9 | 2.9 | 2.86 | 2.92 | 2.844 | 3.2322 | 2.6254 | [9 empty] | 16
| 1 | 4 | 3.3 | 3.2713 | 2.94 | 2.9 | 2.9 | 2.84 | 2.84 | 2.9125 | 3.2751 | 2.7381 | [9 empty] | 2.9987
| 1 | 5 | 3.18 | 3.1478 | 2.94 | 2.9 | 3 | 2.92 | 2.86 | 2.6684 | 3.1831 | 2.72 | [9 empty] | 79
| 1 | 6 | 3.19 | 3.2991 | 2.87 | 2.8 | 2.9 | 2.92 | 2.81 | 2.6307 | 3.2892 | 2.8187
| Mean |  | 3.3083 | 3.2292 | 2.945 | 2.8833 | 2.9333 | 2.8867 | 2.8583 | 2.7522 | 3.249 | 2.7608
| Median |  | 3.34 | 3.2368 | 2.94 | 2.9 | 2.9 | 2.89 | 2.855 | 2.7288 | 3.2537 | 2.7419
| Std Dev. |  | 0.10265 | 0.073391 | 0.063482 | 0.040825 | 0.05164 | 0.045019 | 0.03656 | 0.10735 | 0.042904 | 0.098388
| Rel.Std.Dev. |  | 3.103% | 2.273% | 2.156% | 1.416% | 1.76% | 1.56% | 1.279% | 3.901% | 1.321% | 3.564%
| PDM3 |  | 10.33% | 7.687% | -1.79% | -3.847% | -2.179% | -3.735% | -4.68% | -8.22% | 8.347% | -7.934%
| Z-Score (Absolute) |  | 1.66 | 1.33 | 0.12 | NA | NA | 0.12 | 0.24 | 0.69 | 1.41 | 0.66
Results from laboratories 10 and 13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 712 (Certified Value 7.66 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 7.2 | 5.5439 | 7.4 | 7.7 | 7.44 | 7.4 | 8.33 | 7.1 | < 50 | 7.5 | 7.8 | 6.3761 | 8.2 | 9.2788 | 8.5 | 7.5 | 7.6 | 7.675 |  | 1
| 1 | 2 | 7.33 | 5.2818 | 8 | 7.5 | 7.25 | 7.5 | 7.92 | 7.5 | < 50 | 7.4 | 7.46 | 6.4352 | 8.4 | 9.4014 | 8 | 7.7 | 8.4 | 7.2457 |  | 9
| 1 | 3 | 7.31 | 5.683 | 7.6 | 7.8 | 7.19 | 7.6 | 8.71 | 7.4 | < 50 | 7.6 | 7.43 | 6.6381 | 7.9 | 9.1517 | 7.8 | 7.8 | 7.2 | 7.2922 |  | 16
| 1 | 4 | 7.11 | 5.4335 | 7.6 | 7.6 | 7.39 | 7.3 | 8.06 | 7.1 | < 50 | 7.3 | 7.53 | 6.8343 | 8.6 | 9.4686 | 7.9 | 7.5 | 8 | 7.5071 |  | 7.659
| 1 | 5 | 6.8 | 5.2742 | 7.8 | 7.6 | 7.26 | 7.4 | 8.5 | 7.1 | < 50 | 7.4 | 7.85 | 6.2417 | 8.4 | 9.3792 | 8.4 | 7.6 | 7.9 | 7.3232 |  | 80
| 1 | 6 | 6.82 | 5.6241 | 7.9 | 7.9 | 7.35 | 7.4 | 8.14 | 7.4 | < 50 | 7.6 | 8.01 | 6.0734 | 8.8 | 9.3879 | 7.9 | 7.5 | 7.8 | 7.424
| Mean |  | 7.095 | 5.4734 | 7.7167 | 7.6833 | 7.3133 | 7.4333 | 8.2767 | 7.2667 |  | 7.4667 | 7.68 | 6.4331 | 8.3833 | 9.3446 | 8.0833 | 7.6 | 7.8167 | 7.4112
| Median |  | 7.155 | 5.4887 | 7.7 | 7.65 | 7.305 | 7.4 | 8.235 | 7.25 |  | 7.45 | 7.665 | 6.4057 | 8.4 | 9.3835 | 7.95 | 7.55 | 7.85 | 7.3736
| Std Dev. |  | 0.23467 | 0.17295 | 0.22286 | 0.1472 | 0.095219 | 0.10328 | 0.29439 | 0.18619 |  | 0.12111 | 0.239 | 0.27284 | 0.31252 | 0.11242 | 0.29269 | 0.12649 | 0.40208 | 0.16015
| Rel.Std.Dev. |  | 3.308% | 3.16% | 2.888% | 1.916% | 1.302% | 1.389% | 3.557% | 2.562% |  | 1.622% | 3.112% | 4.241% | 3.728% | 1.203% | 3.621% | 1.664% | 5.144% | 2.161%
| PDM3 |  | -7.364% | -28.54% | 0.7527% | 0.3175% | -4.513% | -2.947% | 8.064% | -5.123% |  | -2.511% | 0.274% | -16.01% | 9.457% | 22.01% | 5.54% | -0.7705% | 2.058% | -3.236%
| Z-Score (Absolute) |  | 1.05 | 4.17 | 0.15 | 0.08 | 0.63 | 0.4 | 1.23 | 0.72 | 33.45 | 0.33 | 0.08 | 2.33 | 1.43 | 3.28 | 0.85 | 0.08 | 0.34 | 0.44
| Analytical results for V in OREAS 712 (Certified Value 111 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 104 | 106 | 112 | 109 | 111 | 6 | 120 | 112 | 106.1531 | 125 | 163 | 118.7366 | 99 | 119.6119 | 102 | 106 | 97 | 128.0153 |  | 1
| 1 | 2 | 103 | 113 | 109 | 105.4 | 109 | 10 | 123 | 113 | 105.0601 | 120 | 164 | 118.9027 | 101 | 119.0774 | 103 | 107 | 102 | 129.3597 |  | 10
| 1 | 3 | 102 | 111 | 112 | 108.2 | 99 | 14 | 124 | 111 | 106.6665 | 125 | 168 | 122.5547 | 108 | 119.8588 | 100 | 106 | 100 | 130.0339 |  | 16
| 1 | 4 | 102 | 111 | 109 | 102.3 | 109 | 9 | 123 | 111 | 107.4734 | 115 | 165 | 120.3943 | 107 | 120.4853 | 98 | 101 | 103 | 125.1692 |  | 110.8166
| 1 | 5 | 103 | 110 | 115 | 103.5 | 99 | 12 | 119 | 111 | 105.0232 | 115 | 162 | 116.1531 | 100 | 121.8869 | 107 | 104 | 101 | 131.8906 |  | 81
| 1 | 6 | 102 | 112 | 113 | 106.8 | 100 | 7 | 121 | 113 | 105.7809 | 120 | 165 | 113.9962 | 110 | 120.6453 | 99 | 104 | 100 | 128.2683
| Mean |  | 102.6667 | 110.5 | 111.6667 | 105.8667 | 104.5 | 9.6667 | 121.6667 | 111.8333 | 106.0262 | 120 | 164.5 | 118.4563 | 104.1667 | 120.2609 | 101.5 | 104.6667 | 100.5 | 128.7895
| Median |  | 102.5 | 111 | 112 | 106.1 | 104.5 | 9.5 | 122 | 111.5 | 105.967 | 120 | 164.5 | 118.8197 | 104 | 120.172 | 101 | 105 | 100.5 | 128.814
| Std Dev. |  | 0.8165 | 2.429 | 2.3381 | 2.6334 | 5.7184 | 3.0111 | 1.9664 | 0.98319 | 0.9508 | 4.4721 | 2.0736 | 3.0341 | 4.7081 | 0.9824 | 3.2711 | 2.1602 | 2.0736 | 2.2569
| Rel.Std.Dev. |  | 0.7953% | 2.198% | 2.094% | 2.487% | 5.472% | 31.15% | 1.616% | 0.8792% | 0.8968% | 3.727% | 1.261% | 2.561% | 4.52% | 0.8169% | 3.223% | 2.064% | 2.063% | 1.752%
| PDM3 |  | -7.354% | -0.2857% | 0.7671% | -4.467% | -5.7% | -91.28% | 9.791% | 0.9175% | -4.323% | 8.287% | 48.44% | 6.894% | -6.001% | 8.522% | -8.407% | -5.55% | -9.31% | 16.22%
| Z-Score (Absolute) |  | 0.61 | 0.24 | 0.37 | 0.26 | 0.41 | 10.76 | 1.46 | 0.39 | 0.24 | 1.28 | 6.14 | 1.11 | 0.45 | 1.31 | 0.74 | 0.39 | 0.85 | 2.24
| Analytical results for W in OREAS 712 (Certified Value 3.37 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 3.41 | 3.7 | 3.2 | < 5 | 7.75 | < 0.1 | 3.5 | 3 | < 10 | 2.5 | 3 | 1 | 3.7371 | 3.7 | 3.4 | 3.2 | 3.1927 |  |  | 1
| 1 | 2 | 3.5 | 3.6 | 3.1 | < 5 | 7.8 | < 0.1 | 3.4 | 3 | < 10 | 1.5 | 3.1 | 1.1 | 3.7753 | 3.5 | 3.4 | 3.3 | 3.0714 |  |  | 11
| 1 | 3 | 3.42 | 3.6 | 3.3 | < 5 | 7.67 | < 0.1 | 3.7 | 3 | < 10 | 2 | 2.6 | 1.3 | 3.7642 | 3.3 | 3.5 | 3.2 | 3.2132 |  |  | 16
| 1 | 4 | 3.36 | 3.6 | 3.2 | < 5 | 7.81 | < 0.1 | 3.5 | 3 | < 10 | 1 | 3.8 | 1.3 | 3.8352 | 3.3 | 3.4 | 3.2 | 3.0472 |  |  | 3.366
| 1 | 5 | 3.2 | 3.7 | 3.3 | < 5 | 7.57 | < 0.1 | 3.6 | 3 | < 10 | 1 | 2.9 | 1.6 | 3.7414 | 3.6 | 3.4 | 3.3 | 3.3227 |  |  | 82
| 1 | 6 | 3.18 | 3.6 | 3.3 | < 5 | 7.64 | < 0.1 | 3.7 | 3 | < 10 | 1.5 | 2.8 | 1.4 | 3.6818 | 3.4 | 3.3 | 3.2 | 3.0281
| Mean |  | 3.345 | 3.6333 | 3.2333 |  | 7.7067 |  | 3.5667 | 3 |  | 1.5833 | 3.0333 | 1.2833 | 3.7559 | 3.4667 | 3.4 | 3.2333 | 3.1459
| Median |  | 3.385 | 3.6 | 3.25 |  | 7.71 |  | 3.55 | 3 |  | 1.5 | 2.95 | 1.3 | 3.7528 | 3.45 | 3.4 | 3.2 | 3.1321
| Std Dev. |  | 0.12834 | 0.05164 | 0.08165 |  | 0.095638 |  | 0.12111 | 0 |  | 0.58452 | 0.41312 | 0.2137 | 0.050572 | 0.1633 | 0.063246 | 0.05164 | 0.11588
| Rel.Std.Dev. |  | 3.837% | 1.421% | 2.525% |  | 1.241% |  | 3.395% | 0% |  | 36.92% | 13.62% | 16.65% | 1.346% | 4.711% | 1.86% | 1.597% | 3.684%
| PDM3 |  | -0.6241% | 7.942% | -3.942% |  | 129% |  | 5.961% | -10.87% |  | -52.96% | -9.883% | -61.87% | 11.58% | 2.99% | 1.01% | -3.942% | -6.54%
| Z-Score (Absolute) |  | 0.12 | 0.75 | 0.12 | 1.71 | 9.58 | 7.03 | 0.6 | NA | 3.71 | 3.7 | 0.55 | 4.35 | 1.01 | 0.39 | 0.24 | 0.12 | 0.31
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 712 (Certified Value 128 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 121.83 | 122.0559 | 132.5 | 125 | 128 | 120 | 138 | 124 | 122.7682 | 126 | 133 | 129.9292 | 138.1 | 146.3309 | 146 | 123 | 130 | 124.888 |  | 1
| 1 | 2 | 122.12 | 121.9002 | 137.5 | 124 | 118 | 130 | 131 | 124.3 | 120.7692 | 127 | 134 | 130.896 | 140 | 147.2309 | 137 | 125.5 | 130 | 121.6214 |  | #N/A
| 1 | 3 | 121.83 | 122.0648 | 130.5 | 128 | 119 | 122 | 134 | 121.4 | 120.6068 | 125 | 139 | 133.1748 | 139.4 | 152.5345 | 129.5 | 126.5 | 129 | 116.3437 |  | 16
| 1 | 4 | 122.02 | 121.7456 | 130.5 | 124.5 | 118 | 118 | 133 | 120.1 | 123.1033 | 125 | 137 | 137.5944 | 144.4 | 146.4173 | 134.5 | 123.5 | 136.5 | 124.3067 |  | 127.9255
| 1 | 5 | 120.97 | 121.1378 | 138.5 | 129.5 | 120 | 129 | 134 | 122.5 | 119.9949 | 127 | 134 | 129.4942 | 146.1 | 143.2051 | 141 | 123.5 | 135.5 | 120.0876 |  | 83
| 1 | 6 | 120.68 | 121.7589 | 134 | 126.5 | 119 | 124 | 134 | 122.3 | 122.235 | 127 | 133 | 128.9849 | 150 | 145.2022 | 129.5 | 123.5 | 134 | 122.2924
| Mean |  | 121.575 | 121.7772 | 133.9167 | 126.25 | 120.3333 | 123.8333 | 134 | 122.4333 | 121.5796 | 126.1667 | 135 | 131.6789 | 143 | 146.8202 | 136.25 | 124.25 | 132.5 | 121.59
| Median |  | 121.83 | 121.8296 | 133.25 | 125.75 | 119 | 123 | 134 | 122.4 | 121.5021 | 126.5 | 134 | 130.4126 | 142.2 | 146.3741 | 135.75 | 123.5 | 132 | 121.9569
| Std Dev. |  | 0.59869 | 0.34224 | 3.4412 | 2.1622 | 3.8297 | 4.8339 | 2.2804 | 1.5795 | 1.2867 | 0.98319 | 2.4495 | 3.2538 | 4.6161 | 3.1276 | 6.5173 | 1.4053 | 3.2249 | 3.1155
| Rel.Std.Dev. |  | 0.4924% | 0.281% | 2.57% | 1.713% | 3.183% | 3.904% | 1.702% | 1.29% | 1.058% | 0.7793% | 1.814% | 2.471% | 3.228% | 2.13% | 4.783% | 1.131% | 2.434% | 2.562%
| PDM3 |  | -4.964% | -4.806% | 4.683% | -1.31% | -5.935% | -3.199% | 4.748% | -4.293% | -4.961% | -1.375% | 5.53% | 2.934% | 11.78% | 14.77% | 6.507% | -2.873% | 3.576% | -4.953%
| Z-Score (Absolute) |  | 0.62 | 0.59 | 1.03 | 0.01 | 0.78 | 0.32 | 1.04 | 0.5 | 0.62 | 0.01 | 1.17 | 0.73 | 2.24 | 2.75 | 1.34 | 0.26 | 0.84 | 0.62
| Analytical results for Yb in OREAS 712 (Certified Value 21.8 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 21.8 | 23.4934 | 20.59 | 22.3 | 21.4 | 21.4 | 21.8 | 20.3924 | 23.9591 | 16.1117 | [9 empty] | 1
| 1 | 2 | 22.35 | 23.7682 | 20.12 | 23.9 | 22.9 | 21.2 | 21.7 | 20.1254 | 24.268 | 14.6513 | [9 empty] | 12
| 1 | 3 | 22.13 | 23.4703 | 20.33 | 22.6 | 21.3 | 21.2 | 22 | 20.8755 | 23.6557 | 14.5324 | [9 empty] | 16
| 1 | 4 | 21.52 | 23.6137 | 20.33 | 22.5 | 21.8 | 21.2 | 21.2 | 21.4036 | 24.1461 | 14.0087 | [9 empty] | 21.8428
| 1 | 5 | 20.72 | 23.201 | 20.72 | 22.6 | 20.8 | 21.1 | 21.1 | 19.8472 | 23.9074 | 14.7559 | [9 empty] | 84
| 1 | 6 | 20.81 | 23.2676 | 19.9 | 23.4 | 21.5 | 21.8 | 21.4 | 19.474 | 23.9837 | 13.8238
| Mean |  | 21.555 | 23.469 | 20.3317 | 22.8833 | 21.6167 | 21.3167 | 21.5333 | 20.353 | 23.9866 | 14.6473
| Median |  | 21.66 | 23.4818 | 20.33 | 22.6 | 21.45 | 21.2 | 21.55 | 20.2589 | 23.9714 | 14.5919
| Std Dev. |  | 0.67471 | 0.2114 | 0.29956 | 0.62423 | 0.70828 | 0.25626 | 0.3559 | 0.70121 | 0.21031 | 0.8069
| Rel.Std.Dev. |  | 3.13% | 0.9008% | 1.473% | 2.728% | 3.277% | 1.202% | 1.653% | 3.445% | 0.8768% | 5.509%
| PDM3 |  | -1.318% | 7.445% | -6.918% | 4.764% | -1.035% | -2.409% | -1.417% | -6.82% | 9.815% | -32.94%
| Z-Score (Absolute) |  | 0.01 | 1.08 | 0.68 | 0.75 | 0.04 | 0.13 | 0.01 | 0.67 | 1.37 | 3.87
| Analytical results for Zn in OREAS 712 (Certified Value 70 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 56 | 79.49 | 70 | 71 | 77.87 | 75 | 70 | 70 | 210 | 65.808 | 70 | 63 | 68.3276 | 61 | 64.2285 | 71 | 74 | 65 | 100.5429 | 1
| 1 | 2 | 56 | 80.05 | 75 | 69 | 74.67 | 74 | 68 | 69 | 332 | 67.0009 | 70 | 69 | 70.3243 | 62 | 64.8311 | 72 | 74 | 66 | 109.0483 | #N/A
| 1 | 3 | 56 | 79.72 | 73 | 70 | 74.51 | 73 | 70 | 73 | 229 | 67.8924 | 70 | 71 | 68.8267 | 63 | 66.9427 | 69 | 73 | 66 | 103.3166 | 16
| 1 | 4 | 56 | 79.14 | 74 | 68 | 74.38 | 74 | 66 | 69 | 321 | 68.4457 | 70 | 69 | 70.585 | 61 | 65.6175 | 68 | 71 | 67 | 91.6529 | 70.1359
| 1 | 5 | 55 | 79.63 | 74 | 72 | 74.27 | 77 | 72 | 70 | 177 | 65.7441 | 70 | 71 | 66.9507 | 60 | 65.4515 | 75 | 71 | 67 | 117.1261 | 85
| 1 | 6 | 56 | 79.46 | 74 | 73 | 76.23 | 74 | 70 | 69 | 244 | 66.1993 | 70 | 63 | 68.1429 | 64 | 63.306 | 69 | 72 | 66 | 96.6806
| Mean |  | 55.8333 | 79.5817 | 73.3333 | 70.5 | 75.3217 | 74.5 | 69.3333 | 70 | 252.1667 | 66.8484 | 70 | 67.6667 | 68.8595 | 61.8333 | 65.0629 | 70.6667 | 72.5 | 66.1667 | 103.0612
| Median |  | 56 | 79.56 | 74 | 70.5 | 74.59 | 74 | 70 | 69.5 | 236.5 | 66.6001 | 70 | 69 | 68.5772 | 61.5 | 65.1413 | 70 | 72.5 | 66 | 101.9297
| Std Dev. |  | 0.40825 | 0.30301 | 1.7512 | 1.8708 | 1.4419 | 1.3784 | 2.0656 | 1.5492 | 61.8754 | 1.1303 | 0 | 3.7238 | 1.3831 | 1.472 | 1.2513 | 2.582 | 1.3784 | 0.75277 | 9.0657
| Rel.Std.Dev. |  | 0.7312% | 0.3808% | 2.388% | 2.654% | 1.914% | 1.85% | 2.979% | 2.213% | 24.54% | 1.691% | 0% | 5.503% | 2.009% | 2.381% | 1.923% | 3.654% | 1.901% | 1.138% | 8.796%
| PDM3 |  | -20.39% | 13.47% | 4.559% | 0.5192% | 7.394% | 6.222% | -1.144% | -0.1937% | 259.5% | -4.687% | -0.1937% | -3.521% | -1.82% | -11.84% | -7.233% | 0.7568% | 3.371% | -5.659% | 46.95%
| Z-Score (Absolute) |  | 2.87 | 1.94 | 0.67 | 0.1 | 1.08 | 0.91 | 0.13 | 0 | 36.85 | 0.64 | 0 | 0.47 | 0.23 | 1.65 | 1 | 0.13 | 0.51 | 0.78 | 6.69
| Analytical results for Zr in OREAS 712 (Indicative Value 0.006 wt.%) | [20 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.00928 | 0.0018948 | 0.00492 | 0.00522 | 0.004133 | 0.0054 | 0.0002 | 0.0125 | 0.0126 | 0.0060515 | 0.0057 | 0.0075 | 0.0058475 | 0.01801 | 0.11158 | 0.00476 | 0.0075 | 0.0138 | 0.0053403 | 1
| 1 | 2 | 0.00944 | 0.0018992 | 0.00516 | 0.00515 | 0.004171 | 0.0054 | 0.0004 | 0.0118 | 0.0063 | 0.005297 | 0.0055 | 0.0071 | 0.005666 | 0.01877 | 0.11505 | 0.00453 | 0.0078 | 0.0113 | 0.0052237 | 18
| 1 | 3 | 0.00864 | 0.0018856 | 0.00509 | 0.00491 | 0.004086 | 0.0055 | 0.0018 | 0.0122 | 0.0134 | 0.006027 | 0.0062 | 0.0075 | 0.0059773 | 0.01852 | 0.11743 | 0.00571 | 0.005 | 0.00461 | 0.0053248 | 16
| 1 | 4 | 0.00925 | 0.001896 | 0.0054 | 0.00489 | 0.004093 | 0.0052 | 0.0004 | 0.0119 | 0.0115 | 0.0057607 | 0.0038 | 0.0066 | 0.0060022 | 0.01904 | 0.11536 | 0.00436 | 0.0077 | 0.0075 | 0.0042151 | 0.005747
| 1 | 5 | 0.00918 | 0.0019516 | 0.00512 | 0.00505 | 0.004119 | 0.0054 | 0.0007 | 0.0117 | 0.0136 | 0.0057381 | 0.0046 | 0.0072 | 0.0060972 | 0.01837 | 0.11322 | 0.00458 | 0.00637 | 0.00496 | 0.0051957 | 24
| 1 | 6 | 0.00881 | 0.0018496 | 0.00495 | 0.00529 | 0.004121 | 0.005 | 0.0002 | 0.0117 | 0.0068 | 0.0056923 | 0.0061 | 0.0068 | 0.005252 | 0.01948 | 0.11401 | 0.00422 | 0.0074 | 0.0122 | 0.0050627
| Mean |  | 0.0091 | 0.0018961 | 0.0051067 | 0.005085 | 0.0041205 | 0.0053167 | 0.00061667 | 0.011967 | 0.0107 | 0.0057611 | 0.0053167 | 0.0071167 | 0.0058071 | 0.018698 | 0.11444 | 0.0046933 | 0.0069617 | 0.0090617 | 0.0050604
| Median |  | 0.009215 | 0.0018954 | 0.005105 | 0.0051 | 0.00412 | 0.0054 | 0.0004 | 0.01185 | 0.01205 | 0.0057494 | 0.0056 | 0.00715 | 0.0059124 | 0.018645 | 0.11453 | 0.004555 | 0.00745 | 0.0094 | 0.0052097
| Std Dev. |  | 0.00030744 | 3.2754e-05 | 0.00017247 | 0.00016392 | 3.0514e-05 | 0.00018348 | 0.000608 | 0.00032042 | 0.0033021 | 0.00027413 | 0.00093684 | 0.0003656 | 0.00031011 | 0.0005189 | 0.0019985 | 0.00053155 | 0.0010889 | 0.0039087 | 0.00042607
| Rel.Std.Dev. |  | 3.378% | 1.727% | 3.377% | 3.224% | 0.7405% | 3.451% | 98.59% | 2.678% | 30.86% | 4.758% | 17.62% | 5.137% | 5.34% | 2.775% | 1.746% | 11.33% | 15.64% | 43.13% | 8.42%
| PDM3 |  | 58.34% | -67.01% | -11.14% | -11.52% | -28.3% | -7.489% | -89.27% | 108.2% | 86.18% | 0.2445% | -7.489% | 23.83% | 1.044% | 225.4% | 1891% | -18.33% | 21.13% | 57.68% | -11.95%
| Z-Score (Absolute) |  | 1.66 | 1.92 | 0.33 | 0.34 | 0.82 | 0.22 | 2.56 | 3.09 | 2.46 | 0 | 0.22 | 0.67 | 0.02 | 6.44 | 54.06 | 0.53 | 0.6 | 1.64 | 0.35
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Pycnometry
| Analytical results for SG in OREAS 712 (Certified Value 3.35 Unity) | [18 empty] | Certified
SG | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | Unity
|  |  | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | LIQPYC | 2
3
| 1 | 1 | 3.3 | 3.3104 | 3.4 | 3.47 | 3.4149 | 3.31 | 3.19 | 3.37 | 3.36 | 3.4 | 3.31 | 3.38 | 3.327 | 3.3471 | 3.44 | 3.26 | 3.1881 | 1
| 1 | 2 | 3.36 | 3.3342 | 3.52 | 3.45 | 3.4154 | 3.3 | 3.2 | 3.38 | 3.37 | 3.39 | 3.33 | 3.38 | 3.557 | 3.3807 | 3.46 | 3.27 | 3.2145 | #N/A
| 1 | 3 | 3.36 | 3.2537 | 3.36 | 3.47 | 3.3721 | 3.26 | 3.2 | 3.37 | 3.38 | 3.42 | 3.33 | 3.37 | 3.398 | 3.2583 | 3.41 | 3.17 | 3.1881 | 16
| 1 | 4 | 3.27 | 3.2932 | 3.53 | 3.48 | 3.3488 | 3.31 | 3.18 | 3.43 | 3.36 | 3.41 | 3.34 | 3.38 | 3.265 | 3.3222 | 3.45 | 3.3 | 3.274 | 3.3458
| 1 | 5 | 3.26 | 3.2838 | 3.54 | 3.48 | 3.363 | 3.32 | 3.18 | 3.39 | 3.37 | 3.39 | 3.37 | 3.37 | 3.129 | 3.3404 | 3.43 | 3.28 | 3.2218 | 87
| 1 | 6 | 3.33 | 3.2893 | 3.48 | 3.48 | 3.3997 | 3.26 | 3.19 | 3.38 | 3.37 | 3.42 | 3.37 | 3.38 | 3.352 | 3.2989 | 3.43 | 3.34 | 3.1829
| Mean |  | 3.3133 | 3.2941 | 3.4717 | 3.4717 | 3.3857 | 3.2933 | 3.19 | 3.3867 | 3.3683 | 3.405 | 3.3417 | 3.3767 | 3.338 | 3.3246 | 3.4367 | 3.27 | 3.2116
| Median |  | 3.315 | 3.2912 | 3.5 | 3.475 | 3.3859 | 3.305 | 3.19 | 3.38 | 3.37 | 3.405 | 3.335 | 3.38 | 3.3395 | 3.3313 | 3.435 | 3.275 | 3.2013
| Std Dev. |  | 0.043665 | 0.026966 | 0.074944 | 0.01169 | 0.028244 | 0.026583 | 0.0089443 | 0.022509 | 0.0075277 | 0.013784 | 0.024014 | 0.005164 | 0.14205 | 0.04235 | 0.017512 | 0.056569 | 0.034451
| Rel.Std.Dev. |  | 1.318% | 0.8186% | 2.159% | 0.3367% | 0.8342% | 0.8072% | 0.2804% | 0.6646% | 0.2235% | 0.4048% | 0.7186% | 0.1529% | 4.255% | 1.274% | 0.5096% | 1.73% | 1.073%
| PDM3 |  | -0.9709% | -1.546% | 3.761% | 3.761% | 1.191% | -1.569% | -4.657% | 1.221% | 0.6729% | 1.769% | -0.1241% | 0.922% | -0.2337% | -0.6345% | 2.715% | -2.266% | -4.013%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 1.84 | 1.84 | 0.62 | 0.69 | 2.15 | 0.64 | 0.38 | 0.9 | 0 | 0.5 | 0.05 | 0.24 | 1.35 | 1.02 | 1.84
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ce in OREAS 712 (Indicative Value 457 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 0
0
| 1 | 1 | 466.3681 | 1
| 1 | 2 | 453.7701 | 20
| 1 | 3 | 449.2905 | 16
| 1 | 4 | 442.8577 | 456.6414
| 1 | 5 | 483.5105 | 26
| 1 | 6 | 444.0512
| Mean |  | 456.6414
| Median |  | 451.5303
| Std Dev. |  | 15.6674
| Rel.Std.Dev. |  | 3.431%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 712 (Indicative Value 17.5 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 1
1
| 1 | 1 | 17.7549 | 1
| 1 | 2 | 18.2113 | 5
| 1 | 3 | 17.7434 | 16
| 1 | 4 | 17.4645 | 17.5413
| 1 | 5 | 16.9199 | 27
| 1 | 6 | 17.1537
| Mean |  | 17.5413
| Median |  | 17.6039
| Std Dev. |  | 0.46425
| Rel.Std.Dev. |  | 2.647%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 712 (Indicative Value 17.2 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 1
1
| 1 | 1 | 17.227 | 1
| 1 | 2 | 17.5217 | 6
| 1 | 3 | 17.2229 | 16
| 1 | 4 | 17.1294 | 17.2387
| 1 | 5 | 17.3596 | 28
| 1 | 6 | 16.9714
| Mean |  | 17.2387
| Median |  | 17.225
| Std Dev. |  | 0.18894
| Rel.Std.Dev. |  | 1.096%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 712 (Indicative Value 3.24 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 3.3891 | 1
| 1 | 2 | 3.4581 | 7
| 1 | 3 | 3.0525 | 16
| 1 | 4 | 3.0053 | 3.2444
| 1 | 5 | 3.5135 | 29
| 1 | 6 | 3.0479
| Mean |  | 3.2444
| Median |  | 3.2208
| Std Dev. |  | 0.23308
| Rel.Std.Dev. |  | 7.184%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 712 (Indicative Value 18.5 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 1
1
| 1 | 1 | 19.3182 | 1
| 1 | 2 | 17.9615 | 10
| 1 | 3 | 18.4366 | 16
| 1 | 4 | 18.4101 | 18.4567
| 1 | 5 | 18.3458 | 30
| 1 | 6 | 18.2683
| Mean |  | 18.4567
| Median |  | 18.378
| Std Dev. |  | 0.45556
| Rel.Std.Dev. |  | 2.468%
| PDM3 |  | -2.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 712 (Indicative Value 4.65 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 4.7188 | 1
| 1 | 2 | 4.707 | 25
| 1 | 3 | 4.7134 | 16
| 1 | 4 | 4.6114 | 4.647
| 1 | 5 | 4.4942 | 31
| 1 | 6 | 4.6371
| Mean |  | 4.647
| Median |  | 4.6721
| Std Dev. |  | 0.087043
| Rel.Std.Dev. |  | 1.873%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 712 (Indicative Value 247 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 0
0
| 1 | 1 | 252.0196 | 1
| 1 | 2 | 255.2133 | 26
| 1 | 3 | 245.4525 | 16
| 1 | 4 | 242.0756 | 246.7128
| 1 | 5 | 241.9829 | 32
| 1 | 6 | 243.5333
| Mean |  | 246.7128
| Median |  | 244.4929
| Std Dev. |  | 5.5856
| Rel.Std.Dev. |  | 2.264%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 712 (Indicative Value 3.28 ppm) |  |  | Indicative
Lu | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 3.321 | 1
| 1 | 2 | 3.3635 | 3
| 1 | 3 | 3.3966 | 16
| 1 | 4 | 3.301 | 3.2799
| 1 | 5 | 3.0198 | 26
| 1 | 6 | 3.2773
| Mean |  | 3.2799
| Median |  | 3.311
| Std Dev. |  | 0.13446
| Rel.Std.Dev. |  | 4.1%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 712 (Indicative Value 1.29 wt.%) |  |  | Indicative
Mn | Round | Replicate | Lab | 1
| No | No | 10 | wt.%
|  |  | BF*OES | 2
2
| 1 | 1 | 1.28 | 1
| 1 | 2 | 1.29 | 21
| 1 | 3 | 1.29 | 16
| 1 | 4 | 1.28 | 1.285
| 1 | 5 | 1.27 | 27
| 1 | 6 | 1.3
| Mean |  | 1.285
| Median |  | 1.285
| Std Dev. |  | 0.010488
| Rel.Std.Dev. |  | 0.8162%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 712 (Indicative Value 150 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 0
0
| 1 | 1 | 153.762 | 1
| 1 | 2 | 154.0564 | 5
| 1 | 3 | 147.1766 | 16
| 1 | 4 | 145.0271 | 150.4872
| 1 | 5 | 156.5803 | 28
| 1 | 6 | 146.3206
| Mean |  | 150.4872
| Median |  | 150.4693
| Std Dev. |  | 4.8729
| Rel.Std.Dev. |  | 3.238%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 712 (Indicative Value 47.7 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 1
1
| 1 | 1 | 48.4768 | 1
| 1 | 2 | 50.2359 | 7
| 1 | 3 | 44.9524 | 16
| 1 | 4 | 47.8295 | 47.7147
| 1 | 5 | 48.6082 | 29
| 1 | 6 | 46.1854
| Mean |  | 47.7147
| Median |  | 48.1531
| Std Dev. |  | 1.8824
| Rel.Std.Dev. |  | 3.945%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 712 (Indicative Value 23.6 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 1
1
| 1 | 1 | 23.5398 | 1
| 1 | 2 | 24.4824 | 24
| 1 | 3 | 22.4428 | 16
| 1 | 4 | 23.9144 | 23.5918
| 1 | 5 | 24.084 | 30
| 1 | 6 | 23.0873
| Mean |  | 23.5918
| Median |  | 23.7271
| Std Dev. |  | 0.73714
| Rel.Std.Dev. |  | 3.125%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 712 (Indicative Value 2.64 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 2.6643 | 1
| 1 | 2 | 2.727 | 25
| 1 | 3 | 2.6022 | 16
| 1 | 4 | 2.5932 | 2.6371
| 1 | 5 | 2.6653 | 31
| 1 | 6 | 2.5709
| Mean |  | 2.6371
| Median |  | 2.6332
| Std Dev. |  | 0.058553
| Rel.Std.Dev. |  | 2.22%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 712 (Indicative Value 64 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 0
0
| 1 | 1 | 64.6995 | 1
| 1 | 2 | 64.4521 | 26
| 1 | 3 | 63.9229 | 16
| 1 | 4 | 62.6701 | 63.7668
| 1 | 5 | 65.2112 | 32
| 1 | 6 | 61.6452
| Mean |  | 63.7668
| Median |  | 64.1875
| Std Dev. |  | 1.3531
| Rel.Std.Dev. |  | 2.122%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 712 (Indicative Value 7.12 wt.%) |  |  | Indicative
Ti | Round | Replicate | Lab | 1
| No | No | 13 | wt.%
|  |  | PF*OES | 2
2
| 1 | 1 | 7.15 | 1
| 1 | 2 | 7.27 | 20
| 1 | 3 | 7.22 | 16
| 1 | 4 | 7.12 | 7.1233
| 1 | 5 | 6.99 | 26
| 1 | 6 | 6.99
| Mean |  | 7.1233
| Median |  | 7.135
| Std Dev. |  | 0.11587
| Rel.Std.Dev. |  | 1.627%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 712 (Indicative Value 2.86 ppm) |  |  | Indicative
Tm | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 2.9677 | 1
| 1 | 2 | 2.9813 | 8
| 1 | 3 | 2.9768 | 16
| 1 | 4 | 2.9092 | 2.8608
| 1 | 5 | 2.3935 | 27
| 1 | 6 | 2.9362
| Mean |  | 2.8608
| Median |  | 2.9519
| Std Dev. |  | 0.23054
| Rel.Std.Dev. |  | 8.059%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 712 (Indicative Value 6.23 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 6.546 | 1
| 1 | 2 | 5.7021 | 9
| 1 | 3 | 6.1821 | 16
| 1 | 4 | 6.3277 | 6.232
| 1 | 5 | 6.3725 | 28
| 1 | 6 | 6.2618
| Mean |  | 6.232
| Median |  | 6.2948
| Std Dev. |  | 0.28695
| Rel.Std.Dev. |  | 4.604%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 712 (Indicative Value 3.35 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 2
2
| 1 | 1 | 3.7941 | 1
| 1 | 2 | 3.386 | 11
| 1 | 3 | 3.335 | 16
| 1 | 4 | 3.1127 | 3.3506
| 1 | 5 | 3.2745 | 29
| 1 | 6 | 3.2011
| Mean |  | 3.3506
| Median |  | 3.3048
| Std Dev. |  | 0.23786
| Rel.Std.Dev. |  | 7.099%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 712 (Indicative Value 21.5 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 18 | ppm
|  |  | PF*MS | 1
1
| 1 | 1 | 21.5022 | 1
| 1 | 2 | 21.7751 | 12
| 1 | 3 | 21.7256 | 16
| 1 | 4 | 21.3467 | 21.4658
| 1 | 5 | 21.0105 | 30
| 1 | 6 | 21.4345
| Mean |  | 21.4658
| Median |  | 21.4683
| Std Dev. |  | 0.2782
| Rel.Std.Dev. |  | 1.296%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 712 (Indicative Value 0.116 wt.%) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 13 | wt.%
|  |  | PF*MS | 3
3
| 1 | 1 | 0.1144 | 1
| 1 | 2 | 0.1162 | 18
| 1 | 3 | 0.1174 | 16
| 1 | 4 | 0.1167 | 0.11632
| 1 | 5 | 0.1172 | 31
| 1 | 6 | 0.116
| Mean |  | 0.11632
| Median |  | 0.11645
| Std Dev. |  | 0.0010852
| Rel.Std.Dev. |  | 0.933%
| PDM3 |  | -2.776e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.