## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 780
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Oxidising Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 13.6091 | 13.4601 | 13.758 | 13.5373 | 13.6809
As, Arsenic (ppm) | 444.1005 | 360.7474 | 527.4536 | IND | IND
BaO, Barium oxide (ppm) | 997.4881 | 908.8659 | 1086.1102 | 958.8894 | 1036.0867
Ca, Calcium (wt.%) | 5.6683 | 5.5675 | 5.7692 | 5.6166 | 5.7201
Co, Cobalt (ppm) | < 100 | IND | IND | IND | IND
Cr, Chromium (ppm) | 114.1236 | 72.8143 | 155.4329 | 100.1335 | 128.1137
Cu, Copper (ppm) | 57.2731 | 39.6305 | 74.9158 | IND | IND
F, Fluorine (wt.%) | 5.1467 | 4.9373 | 5.3561 | 4.9725 | 5.3209
Fe, Iron (wt.%) | 2.4984 | 2.4626 | 2.5342 | 2.4749 | 2.5219
HfO2, Hafnium dioxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 2.6107 | 2.5686 | 2.6528 | 2.5902 | 2.6312
MgO, Magnesium oxide (wt.%) | 0.54593 | 0.52248 | 0.56938 | 0.53363 | 0.55824
Nb, Niobium (ppm) | < 50 | IND | IND | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.063304 | 0.05489 | 0.071718 | 0.061144 | 0.065464
Pb, Lead (ppm) | < 50 | IND | IND | IND | IND
S, Sulphur (ppm) | 96.8498 | 81.7272 | 111.9724 | 85.7581 | 107.9415
SiO2, Silicon dioxide (wt.%) | 63.8312 | 63.3541 | 64.3084 | 63.6307 | 64.0317
TiO2, Titanium dioxide (wt.%) | 0.68751 | 0.66676 | 0.70826 | 0.67717 | 0.69785
W, Tungsten (ppm) | 81.3326 | 58.7826 | 103.8826 | IND | IND
Zr, Zirconium (ppm) | 200.4677 | 133.8625 | 267.0728 | 190.9084 | 210.0269
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 5.7561 | 5.4461 | 6.0662 | 5.6624 | 5.8498
Ion Selective Electrode
F, Fluorine (wt.%) | 4.3167 | 3.9392 | 4.6942 | 4.1102 | 4.5232
4-Acid Digestion
Ag, Silver (ppm) | 0.34479 | 0.30214 | 0.38744 | 0.30846 | 0.38112
Al, Aluminium (wt.%) | 6.7315 | 6.471 | 6.992 | 6.5797 | 6.8832
As, Arsenic (ppm) | 428.1965 | 413.175 | 443.2181 | 420.7877 | 435.6053
Ba, Barium (ppm) | 877.9652 | 850.0881 | 905.8424 | 865.1529 | 890.7776
Be, Beryllium (ppm) | 2.6025 | 2.4752 | 2.7298 | 2.5019 | 2.7032
Bi, Bismuth (ppm) | 116.0927 | 110.8531 | 121.3324 | 112.8157 | 119.3698
Ca, Calcium (wt.%) | 5.5864 | 5.4076 | 5.7652 | 5.4988 | 5.674
Ce, Cerium (ppm) | 87.902 | 82.4264 | 93.3777 | 85.4994 | 90.3046
Co, Cobalt (ppm) | 1.8211 | 1.7009 | 1.9413 | 1.7099 | 1.9324
Cr, Chromium (ppm) | 88.1297 | 80.3408 | 95.9186 | 83.776 | 92.4834
Cs, Caesium (ppm) | 9.7143 | 9.2553 | 10.1732 | 9.3655 | 10.063
Cu, Copper (ppm) | 52.314 | 49.5246 | 55.1033 | 50.7757 | 53.8523
Dy, Dysprosium (ppm) | 3.4002 | 3.0411 | 3.7592 | 3.2093 | 3.591
Er, Erbium (ppm) | 1.6277 | 1.3992 | 1.8562 | 1.5487 | 1.7068
Eu, Europium (ppm) | 1.2006 | 1.1218 | 1.2795 | 1.1578 | 1.2435
Fe, Iron (wt.%) | 2.4203 | 2.3411 | 2.4996 | 2.3813 | 2.4594
Ga, Gallium (ppm) | 17.8532 | 16.9962 | 18.7101 | 17.319 | 18.3873
Gd, Gadolinium (ppm) | 4.8822 | 4.4067 | 5.3577 | 4.6273 | 5.1371
Hf, Hafnium (ppm) | 3.6981 | 3.4671 | 3.929 | 3.562 | 3.8341
Ho, Holmium (ppm) | 0.59266 | 0.50036 | 0.68496 | 0.56001 | 0.62531
In, Indium (ppm) | 0.081568 | 0.072003 | 0.091134 | 0.075031 | 0.088105
K, Potassium (wt.%) | 2.1957 | 2.1302 | 2.2612 | 2.167 | 2.2244
La, Lanthanum (ppm) | 43.7113 | 41.2757 | 46.147 | 42.4741 | 44.9485
Li, Lithium (ppm) | 45.2898 | 43.5999 | 46.9796 | 44.5466 | 46.0329
Lu, Lutetium (ppm) | 0.23256 | 0.19283 | 0.27228 | 0.21775 | 0.24736
Mg, Magnesium (wt.%) | 0.26515 | 0.24468 | 0.28562 | 0.25756 | 0.27274
Mn, Manganese (wt.%) | 0.0038318 | 0.0036394 | 0.0040243 | 0.0036809 | 0.0039828
Mo, Molybdenum (ppm) | 3.7641 | 3.3781 | 4.1501 | 3.5435 | 3.9847
Na, Sodium (wt.%) | 0.13628 | 0.12773 | 0.14483 | 0.13136 | 0.1412
Nb, Niobium (ppm) | 11.6579 | 10.8066 | 12.5091 | 11.2132 | 12.1026
Nd, Neodymium (ppm) | 35.7422 | 33.5498 | 37.9347 | 34.7593 | 36.7251
Ni, Nickel (ppm) | 12.8342 | 12.1267 | 13.5418 | 12.2701 | 13.3984
P, Phosphorus (wt.%) | 0.028171 | 0.026956 | 0.029385 | 0.027577 | 0.028764
Pb, Lead (ppm) | 16.6372 | 15.7535 | 17.521 | 16.1743 | 17.1002
Pr, Praseodymium (ppm) | 9.8199 | 9.1531 | 10.4866 | 9.5721 | 10.0676
Rb, Rubidium (ppm) | 121.5177 | 114.4302 | 128.6053 | 117.3963 | 125.6391
S, Sulphur (wt.%) | 0.012163 | 0.011051 | 0.013275 | IND | IND
Sb, Antimony (ppm) | 5.578 | 5.2562 | 5.8998 | 5.3697 | 5.7862
Sc, Scandium (ppm) | 13.3377 | 12.5852 | 14.0902 | 12.9305 | 13.7449
Se, Selenium (ppm) | 0.91698 | 0.73785 | 1.0961 | 0.83601 | 0.99795
Sm, Samarium (ppm) | 6.2457 | 5.8134 | 6.6779 | 5.9988 | 6.4925
Sn, Tin (ppm) | 14.8346 | 14.0947 | 15.5745 | 14.3711 | 15.2982
Sr, Strontium (ppm) | 62.9178 | 60.2082 | 65.6274 | 61.1602 | 64.6754
Ta, Tantalum (ppm) | 0.88834 | 0.80717 | 0.96952 | 0.84409 | 0.9326
Tb, Terbium (ppm) | 0.66301 | 0.5703 | 0.75571 | 0.63507 | 0.69094
Te, Tellurium (ppm) | 0.23328 | 0.19652 | 0.27004 | 0.20496 | 0.2616
Th, Thorium (ppm) | 15.2727 | 14.4478 | 16.0976 | 14.8619 | 15.6834
Ti, Titanium (wt.%) | 0.34676 | 0.31902 | 0.3745 | 0.33549 | 0.35802
Tl, Thallium (ppm) | 0.68671 | 0.64584 | 0.72758 | 0.6581 | 0.71533
Tm, Thulium (ppm) | 0.21886 | 0.17818 | 0.25955 | 0.20193 | 0.23579
U, Uranium (ppm) | 2.4915 | 2.3137 | 2.6693 | 2.382 | 2.601
V, Vanadium (ppm) | 91.0556 | 87.278 | 94.8332 | 88.8803 | 93.2309
W, Tungsten (ppm) | 75.4289 | 71.7273 | 79.1306 | 73.4378 | 77.4201
Y, Yttrium (ppm) | 25.044 | 23.1233 | 26.9646 | 24.2095 | 25.8784
Yb, Ytterbium (ppm) | 1.618 | 1.4429 | 1.793 | 1.5372 | 1.6988
Zn, Zinc (ppm) | 28.9962 | 27.0972 | 30.8953 | 27.721 | 30.2715
Zr, Zirconium (ppm) | 129.5533 | 123.5784 | 135.5282 | 126.347 | 132.7596
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 780
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Oxidising Fusion XRF
Bi | ppm | 114.0917 | In | ppm | < 100 | Sm | ppm | < 170
Cd | ppm | < 100 | La2O3 | ppm | 89.4444 | Sn | ppm | 49.3239
CeO2 | ppm | 133.8255 | Mn | wt.% | 0.006534 | Sr | ppm | 80.4564
Cl | ppm | 44.1667 | Mo | ppm | < 50 | Ta | ppm | < 80
Cs | ppm | < 100 | Na2O | wt.% | 0.19438 | Te | ppm | < 100
Dy | ppm | < 170 | Nd | ppm | 61.4426 | Th | ppm | 68.8403
Eu | ppm | < 170 | Ni | ppm | < 50 | Tl | ppm | < 100
Ga | ppm | < 100 | Pr | ppm | 42.7633 | U | ppm | < 17
Gd | ppm | < 170 | Rb | ppm | 138.1468 | V | ppm | 100.6026
Ge | ppm | < 100 | Sb | ppm | 128.6658 | Y2O3 | ppm | 94.5833
H2O- | wt.% | 0.32302 | Sc | ppm | < 100 | Zn | ppm | 34.0651
Hg | ppm | < 100 | Se | ppm | < 100
Thermogravimetry
H2O- | wt.% | 0.32302
4-Acid Digestion
B | ppm | 23.1017 | Ge | ppm | 0.2195 | Re | ppm | 0.0024062
Cd | ppm | 0.041745 | Hg | ppm | 0.055
Borate / Peroxide Fusion ICP
Al | wt.% | 7.5135 | Gd | ppm | 5.6984 | Sr | ppm | 64.567
Ba | ppm | 991.07 | Ho | ppm | 0.95733 | Tb | ppm | 0.81625
Bi | ppm | 95.9809 | La | ppm | 42.3842 | Th | ppm | 12.6574
Ca | wt.% | 5.9437 | Li | ppm | 44.3917 | Ti | wt.% | 0.44477
Ce | ppm | 89.7254 | Lu | ppm | 0.37201 | Tl | ppm | 0.48572
Cs | ppm | 10.2984 | Mg | wt.% | 0.3596 | Tm | ppm | 0.39164
Dy | ppm | 4.6496 | Nd | ppm | 34.1847 | V | ppm | 96.3155
Er | ppm | 2.7516 | Pr | ppm | 9.2794 | W | ppm | 68.2666
Eu | ppm | 1.2825 | Si | wt.% | 30.7864 | Yb | ppm | 2.4747
Fe | wt.% | 2.3257 | Sm | ppm | 6.2081 | Zn | ppm | 43.2722
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 780
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Oxidising Fusion XRF
Al2O3, wt.% | 13.6091 | 0.15098 | 13.3071 | 13.911 | 13.1561 | 14.062 | 1.109% | 2.219% | 3.328% | 12.9286 | 14.2895
As, ppm | 444.1005 | 101.9456 | 240.2093 | 647.9917 | 138.2637 | 749.9373 | 22.96% | 45.91% | 68.87% | 421.8955 | 466.3056
BaO, ppm | 997.4881 | 56.0357 | 885.4166 | 1109.5595 | 829.3809 | 1165.5952 | 5.618% | 11.24% | 16.85% | 947.6137 | 1047.3625
Ca, wt.% | 5.6683 | 0.085971 | 5.4964 | 5.8403 | 5.4104 | 5.9262 | 1.517% | 3.033% | 4.55% | 5.3849 | 5.9517
Co, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cr, ppm | 114.1236 | 23.4674 | 67.1888 | 161.0584 | 43.7214 | 184.5258 | 20.56% | 41.13% | 61.69% | 108.4174 | 119.8298
Cu, ppm | 57.2731 | 12.1037 | 33.0658 | 81.4805 | 20.9621 | 93.5841 | 21.13% | 42.27% | 63.4% | 54.4095 | 60.1368
F, wt.% | 5.1467 | 0.16994 | 4.8069 | 5.4866 | 4.6369 | 5.6565 | 3.302% | 6.604% | 9.906% | 4.8894 | 5.4041
Fe, wt.% | 2.4984 | 0.052123 | 2.3942 | 2.6027 | 2.342 | 2.6548 | 2.086% | 4.173% | 6.259% | 2.3735 | 2.6233
HfO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 2.6107 | 0.068225 | 2.4742 | 2.7471 | 2.406 | 2.8154 | 2.613% | 5.227% | 7.84% | 2.4801 | 2.7412
MgO, wt.% | 0.54593 | 0.033555 | 0.47882 | 0.61304 | 0.44527 | 0.6466 | 6.146% | 12.29% | 18.44% | 0.51864 | 0.57323
Nb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
P2O5, wt.% | 0.063304 | 0.0065802 | 0.050143 | 0.076464 | 0.043563 | 0.083045 | 10.39% | 20.79% | 31.18% | 0.060139 | 0.066469
Pb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, ppm | 96.8498 | 23.0448 | 50.7601 | 142.9394 | 27.7153 | 165.9843 | 23.79% | 47.59% | 71.38% | 92.0073 | 101.6923
SiO2, wt.% | 63.8312 | 0.6143 | 62.6026 | 65.0598 | 61.9883 | 65.6741 | 0.9624% | 1.925% | 2.887% | 60.6396 | 67.0228
TiO2, wt.% | 0.68751 | 0.027914 | 0.63168 | 0.74334 | 0.60377 | 0.77125 | 4.06% | 8.12% | 12.18% | 0.65313 | 0.72189
W, ppm | 81.3326 | 13.7353 | 53.8621 | 108.8031 | 40.1268 | 122.5384 | 16.89% | 33.78% | 50.66% | 77.266 | 85.3993
Zr, ppm | 200.4677 | 34.8466 | 130.7744 | 270.1609 | 95.9278 | 305.0075 | 17.38% | 34.77% | 52.15% | 190.4443 | 210.491
Thermogravimetry
LOI1000, wt.% | 5.7561 | 0.47382 | 4.8085 | 6.7038 | 4.3346 | 7.1776 | 8.232% | 16.46% | 24.69% | 5.4683 | 6.0439
Ion Selective Electrode
F, wt.% | 4.3167 | 0.51827 | 3.2802 | 5.3533 | 2.7619 | 5.8715 | 12.01% | 24.01% | 36.02% | 4.1009 | 4.5325
4-Acid Digestion
Ag, ppm | 0.34479 | 0.023789 | 0.29721 | 0.39237 | 0.27342 | 0.41616 | 6.9% | 13.8% | 20.7% | 0.32755 | 0.36203
Al, wt.% | 6.7315 | 0.35038 | 6.0307 | 7.4322 | 5.6803 | 7.7826 | 5.205% | 10.41% | 15.62% | 6.3949 | 7.0681
As, ppm | 428.1965 | 13.229 | 401.7386 | 454.6545 | 388.5096 | 467.8834 | 3.089% | 6.179% | 9.268% | 406.7867 | 449.6064
Ba, ppm | 877.9652 | 26.857 | 824.2512 | 931.6793 | 797.3941 | 958.5364 | 3.059% | 6.118% | 9.177% | 834.067 | 921.8635
Be, ppm | 2.6025 | 0.13848 | 2.3255 | 2.8795 | 2.1871 | 3.018 | 5.321% | 10.64% | 15.96% | 2.4724 | 2.7326
Bi, ppm | 116.0927 | 5.4184 | 105.2559 | 126.9295 | 99.8375 | 132.3479 | 4.667% | 9.335% | 14% | 110.2881 | 121.8974
Ca, wt.% | 5.5864 | 0.16248 | 5.2615 | 5.9114 | 5.099 | 6.0738 | 2.908% | 5.817% | 8.725% | 5.3071 | 5.8657
Ce, ppm | 87.902 | 4.2501 | 79.4018 | 96.4023 | 75.1517 | 100.6524 | 4.835% | 9.67% | 14.51% | 83.5069 | 92.2971
Co, ppm | 1.8211 | 0.12072 | 1.5797 | 2.0626 | 1.459 | 2.1833 | 6.629% | 13.26% | 19.89% | 1.7301 | 1.9122
Cr, ppm | 88.1297 | 12.8269 | 62.4758 | 113.7835 | 49.6489 | 126.6105 | 14.55% | 29.11% | 43.66% | 83.7232 | 92.5362
Cs, ppm | 9.7143 | 0.49802 | 8.7182 | 10.7103 | 8.2202 | 11.2083 | 5.127% | 10.25% | 15.38% | 9.2286 | 10.2
Cu, ppm | 52.314 | 2.3899 | 47.5343 | 57.0937 | 45.1444 | 59.4836 | 4.568% | 9.137% | 13.7% | 49.6983 | 54.9297
Dy, ppm | 3.4002 | 0.33144 | 2.7373 | 4.0631 | 2.4058 | 4.3945 | 9.748% | 19.5% | 29.24% | 3.2302 | 3.5702
Er, ppm | 1.6277 | 0.21955 | 1.1886 | 2.0668 | 0.96907 | 2.2864 | 13.49% | 26.98% | 40.46% | 1.5463 | 1.7091
Eu, ppm | 1.2006 | 0.048175 | 1.1043 | 1.297 | 1.0561 | 1.3451 | 4.013% | 8.025% | 12.04% | 1.1406 | 1.2606
Fe, wt.% | 2.4203 | 0.08239 | 2.2556 | 2.5851 | 2.1732 | 2.6675 | 3.404% | 6.808% | 10.21% | 2.2993 | 2.5414
Ga, ppm | 17.8532 | 1.0006 | 15.852 | 19.8543 | 14.8514 | 20.8549 | 5.604% | 11.21% | 16.81% | 16.9605 | 18.7458
Gd, ppm | 4.8822 | 0.34808 | 4.186 | 5.5783 | 3.8379 | 5.9264 | 7.13% | 14.26% | 21.39% | 4.6381 | 5.1263
Hf, ppm | 3.6981 | 0.29506 | 3.108 | 4.2882 | 2.8129 | 4.5832 | 7.979% | 15.96% | 23.94% | 3.5132 | 3.883
Ho, ppm | 0.59266 | 0.085773 | 0.42111 | 0.76421 | 0.33534 | 0.84998 | 14.47% | 28.95% | 43.42% | 0.56303 | 0.62229
In, ppm | 0.081568 | 0.0080763 | 0.065416 | 0.097721 | 0.057339 | 0.1058 | 9.901% | 19.8% | 29.7% | 0.07749 | 0.085647
K, wt.% | 2.1957 | 0.056083 | 2.0835 | 2.3078 | 2.0274 | 2.3639 | 2.554% | 5.108% | 7.663% | 2.0859 | 2.3055
La, ppm | 43.7113 | 1.9578 | 39.7958 | 47.6269 | 37.838 | 49.5847 | 4.479% | 8.958% | 13.44% | 41.5258 | 45.8969
Li, ppm | 45.2898 | 1.4628 | 42.3642 | 48.2153 | 40.9014 | 49.6781 | 3.23% | 6.46% | 9.689% | 43.0253 | 47.5542
Lu, ppm | 0.23256 | 0.037979 | 0.1566 | 0.30851 | 0.11862 | 0.34649 | 16.33% | 32.66% | 48.99% | 0.22093 | 0.24418
Mg, wt.% | 0.26515 | 0.034427 | 0.1963 | 0.33401 | 0.16187 | 0.36843 | 12.98% | 25.97% | 38.95% | 0.25189 | 0.27841
Mn, wt.% | 0.0038318 | 0.00014319 | 0.0035455 | 0.0041182 | 0.0034023 | 0.0042614 | 3.737% | 7.473% | 11.21% | 0.0036403 | 0.0040234
Mo, ppm | 3.7641 | 0.29683 | 3.1704 | 4.3577 | 2.8736 | 4.6546 | 7.886% | 15.77% | 23.66% | 3.5759 | 3.9523
Na, wt.% | 0.13628 | 0.0086068 | 0.11907 | 0.15349 | 0.11046 | 0.1621 | 6.315% | 12.63% | 18.95% | 0.12947 | 0.14309
Nb, ppm | 11.6579 | 1.0419 | 9.574 | 13.7418 | 8.5321 | 14.7837 | 8.938% | 17.88% | 26.81% | 11.075 | 12.2408
Nd, ppm | 35.7422 | 1.6502 | 32.4419 | 39.0426 | 30.7918 | 40.6927 | 4.617% | 9.234% | 13.85% | 33.9551 | 37.5294
Ni, ppm | 12.8342 | 0.64694 | 11.5403 | 14.1281 | 10.8934 | 14.775 | 5.041% | 10.08% | 15.12% | 12.1925 | 13.4759
P, wt.% | 0.028171 | 0.0010784 | 0.026014 | 0.030328 | 0.024936 | 0.031406 | 3.828% | 7.656% | 11.48% | 0.026762 | 0.029579
Pb, ppm | 16.6372 | 1.2749 | 14.0875 | 19.187 | 12.8126 | 20.4619 | 7.663% | 15.33% | 22.99% | 15.8054 | 17.4691
Pr, ppm | 9.8199 | 0.44457 | 8.9307 | 10.709 | 8.4862 | 11.1536 | 4.527% | 9.054% | 13.58% | 9.3289 | 10.3109
Rb, ppm | 121.5177 | 8.0066 | 105.5045 | 137.5309 | 97.4979 | 145.5376 | 6.589% | 13.18% | 19.77% | 115.4418 | 127.5936
S, wt.% | 0.012163 | 0.0030131 | 0.0061369 | 0.018189 | 0.0031238 | 0.021202 | 24.77% | 49.54% | 74.32% | 0.011555 | 0.012771
Sb, ppm | 5.578 | 0.41763 | 4.7427 | 6.4132 | 4.3251 | 6.8309 | 7.487% | 14.97% | 22.46% | 5.2991 | 5.8569
Sc, ppm | 13.3377 | 0.76122 | 11.8152 | 14.8601 | 11.054 | 15.6213 | 5.707% | 11.41% | 17.12% | 12.6708 | 14.0046
Se, ppm | 0.91698 | 0.15467 | 0.60763 | 1.2263 | 0.45296 | 1.381 | 16.87% | 33.74% | 50.6% | 0.87113 | 0.96283
Sm, ppm | 6.2457 | 0.23193 | 5.7818 | 6.7095 | 5.5499 | 6.9415 | 3.713% | 7.427% | 11.14% | 5.9334 | 6.558
Sn, ppm | 14.8346 | 0.68576 | 13.4631 | 16.2061 | 12.7773 | 16.8919 | 4.623% | 9.245% | 13.87% | 14.0929 | 15.5764
Sr, ppm | 62.9178 | 2.4181 | 58.0816 | 67.7541 | 55.6634 | 70.1722 | 3.843% | 7.687% | 11.53% | 59.7719 | 66.0637
Ta, ppm | 0.88834 | 0.091077 | 0.70619 | 1.0705 | 0.61511 | 1.1616 | 10.25% | 20.51% | 30.76% | 0.84393 | 0.93276
Tb, ppm | 0.66301 | 0.085404 | 0.4922 | 0.83381 | 0.40679 | 0.91922 | 12.88% | 25.76% | 38.64% | 0.62986 | 0.69616
Te, ppm | 0.23328 | 0.032447 | 0.16839 | 0.29817 | 0.13594 | 0.33062 | 13.91% | 27.82% | 41.73% | 0.22162 | 0.24494
Th, ppm | 15.2727 | 0.68779 | 13.8971 | 16.6483 | 13.2093 | 17.3361 | 4.503% | 9.007% | 13.51% | 14.509 | 16.0363
Ti, wt.% | 0.34676 | 0.045243 | 0.25627 | 0.43724 | 0.21103 | 0.48249 | 13.05% | 26.1% | 39.14% | 0.32942 | 0.36409
Tl, ppm | 0.68671 | 0.050774 | 0.58517 | 0.78826 | 0.53439 | 0.83904 | 7.394% | 14.79% | 22.18% | 0.65238 | 0.72105
Tm, ppm | 0.21886 | 0.035403 | 0.14806 | 0.28967 | 0.11265 | 0.32507 | 16.18% | 32.35% | 48.53% | 0.20792 | 0.2298
U, ppm | 2.4915 | 0.15334 | 2.1848 | 2.7982 | 2.0315 | 2.9515 | 6.154% | 12.31% | 18.46% | 2.3669 | 2.6161
V, ppm | 91.0556 | 4.0649 | 82.9258 | 99.1855 | 78.8608 | 103.2504 | 4.464% | 8.928% | 13.39% | 86.5028 | 95.6084
W, ppm | 75.4289 | 2.8895 | 69.6499 | 81.2079 | 66.7604 | 84.0974 | 3.831% | 7.662% | 11.49% | 71.6575 | 79.2004
Y, ppm | 25.044 | 1.7775 | 21.4889 | 28.599 | 19.7114 | 30.3765 | 7.098% | 14.2% | 21.29% | 23.7918 | 26.2962
Yb, ppm | 1.618 | 0.18756 | 1.2429 | 1.9931 | 1.0553 | 2.1807 | 11.59% | 23.18% | 34.78% | 1.5371 | 1.6989
Zn, ppm | 28.9962 | 2.1561 | 24.6841 | 33.3084 | 22.528 | 35.4645 | 7.436% | 14.87% | 22.31% | 27.5464 | 30.4461
Zr, ppm | 129.5533 | 6.7033 | 116.1466 | 142.96 | 109.4433 | 149.6633 | 5.174% | 10.35% | 15.52% | 123.0756 | 136.031
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 780
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
ISE | ion specific electrode
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 780
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek, Perth, WA, Australia
Lucid Laboratories Private Limited, Hyderabad, Telangana, India
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan
UIS Analytical Services, Centurion, South Africa

## Oxidising Fusion XRF
| Analytical results for Al2O3 in OREAS 780 (Certified Value 13.61 wt.%) | [20 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 13.44 | 13.82 | 13.5 | 13.75 | 13.5388 | 13.4 | 8.85 | 13.66 | 13.17 | 13.51 | 14.06 | 13.5545 | 13.5296 | 13.4 | 13.6802 | 13.95 | 13.9 | 13.6 | 12.675 | 1
| 1 | 2 | 13.56 | 13.8 | 13.45 | 13.75 | 13.531 | 13.5 | 8.73 | 13.57 | 13.06 | 13.51 | 14.05 | 13.5146 | 13.5231 | 13.3 | 13.6837 | 13.75 | 13.75 | 13.55 | 12.767 | #N/A
| 1 | 3 | 13.37 | 13.83 | 13.75 | 13.8 | 13.5431 | 13.6 | 8.82 | 13.64 | 13.15 | 13.48 | 13.98 | 13.565 | 13.5426 | 13.4 | 13.6643 | 13.9 | 13.8 | 13.6 | 12.817 | 16
| 1 | 4 | 13.47 | 13.82 | 13.6 | 13.75 | 13.4762 | 13.5 | 8.99 | 13.5 | 13.05 | 13.59 | 14.01 | 13.5448 | 13.499 | 13.4 | 13.6373 | 14 | 13.8 | 13.6 | 12.741 | 13.6091
| 1 | 5 | 13.43 | 13.81 | 13.45 | 13.8 | 13.6503 | 13.4 | 8.72 | 13.56 | 13.13 | 13.53 | 14.14 | 13.4924 | 13.5497 | 13.5 | 13.6532 | 14.1 | 13.9 | 13.55 | 12.684 | 7
| 1 | 6 | 13.57 | 13.84 | 13.5 | 13.8 | 13.5393 | 13.4 | 9 | 13.59 | 13.17 | 13.48 | 14.07 | 13.5801 | 13.505 | 13.4 | 13.7188 | 13.75 | 13.7 | 13.65 | 12.882
| Mean |  | 13.4733 | 13.82 | 13.5417 | 13.775 | 13.5465 | 13.4667 | 8.8517 | 13.5867 | 13.1217 | 13.5167 | 14.0517 | 13.5419 | 13.5248 | 13.4 | 13.6729 | 13.9083 | 13.8083 | 13.5917 | 12.761
| Median |  | 13.455 | 13.82 | 13.5 | 13.775 | 13.5391 | 13.45 | 8.835 | 13.58 | 13.14 | 13.51 | 14.055 | 13.5496 | 13.5263 | 13.4 | 13.6722 | 13.925 | 13.8 | 13.6 | 12.754
| Std Dev. |  | 0.078145 | 0.014142 | 0.11583 | 0.027386 | 0.056736 | 0.08165 | 0.12189 | 0.05785 | 0.053821 | 0.040825 | 0.054924 | 0.032733 | 0.020116 | 0.063246 | 0.028297 | 0.13934 | 0.080104 | 0.037639 | 0.07942
| Rel.Std.Dev. |  | 0.58% | 0.1023% | 0.8554% | 0.1988% | 0.4188% | 0.6063% | 1.377% | 0.4258% | 0.4102% | 0.302% | 0.3909% | 0.2417% | 0.1487% | 0.472% | 0.207% | 1.002% | 0.5801% | 0.2769% | 0.6224%
| PDM3 |  | -0.9974% | 1.55% | -0.4953% | 1.219% | -0.4601% | -1.046% | -34.96% | -0.1646% | -3.581% | -0.679% | 3.252% | -0.4935% | -0.6191% | -1.536% | 0.469% | 2.199% | 1.464% | -0.1279% | -6.232%
| Z-Score (Absolute) |  | 0.46 | 1.8 | 0.02 | 1.51 | 0.01 | 0.51 | NA | 0.28 | 2.76 | 0.18 | 3.32 | 0.01 | 0.13 | 0.94 | 0.84 | 2.38 | 1.73 | 0.31 | 5.12
Results from laboratory 1.08 were removed due to their systematic bias.
| Analytical results for As in OREAS 780 (Certified Value 444 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 15 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | 500 | 400 | 300 | 523 | 424.1 | 450 | 300 | 600 | 500 | [10 empty] | 1
| 1 | 2 | 500 | 400 | 300 | 545 | 420.4 | 450 | 300 | 500 | 500 | [10 empty] | #N/A
| 1 | 3 | 600 | 400 | 300 | 538 | 420.4 | 380 | 300 | 600 | 500 | [10 empty] | 16
| 1 | 4 | 600 | 400 | 300 | 530 | 424.1 | 450 | 300 | 600 | 400 | [10 empty] | 444.1005
| 1 | 5 | 600 | 400 | 300 | 538 | 416.6 | 450 | 300 | 600 | 400 | [10 empty] | 8
| 1 | 6 | 500 | 400 | 300 | 530 | 397.6 | 450 | 300 | 600 | 500
| Mean |  | 550 | 400 | 300 | 534 | 417.2 | 438.3333 | 300 | 583.3333 | 466.6667
| Median |  | 550 | 400 | 300 | 534 | 420.4 | 450 | 300 | 600 | 500
| Std Dev. |  | 54.7723 | 0 | 0 | 7.823 | 10.0022 | 28.5774 | 0 | 40.8248 | 51.6398
| Rel.Std.Dev. |  | 9.959% | 0% | 0% | 1.465% | 2.397% | 6.52% | 0% | 6.999% | 11.07%
| PDM3 |  | 23.85% | -9.93% | -32.45% | 20.24% | -6.057% | -1.299% | -32.45% | 31.35% | 5.081%
| Z-Score (Absolute) |  | 0.79 | 0.31 | 1.04 | 0.67 | 0.18 | 0 | 1.04 | 1.04 | 0.18
| Analytical results for BaO in OREAS 780 (Certified Value 997 ppm) | [20 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 900 | 1000 | 900 | 1000 | 600 | 960 | 1038.65 | 940 | 1000 | 996.6714 | 1000 | 1000 | 1000 | 1040 | [5 empty] | 1
| 1 | 2 | 900 | 900 | 900 | 1000 | 600 | 980 | 1050 | 945 | 1000 | 965.3666 | 1100 | 1000 | 1100 | 1050 | [5 empty] | #N/A
| 1 | 3 | 1000 | 1100 | 1000 | 1000 | 900 | 960 | 1029.1 | 940 | 900 | 982.4975 | 1000 | 1100 | 1000 | 1040 | [5 empty] | 16
| 1 | 4 | 900 | 1000 | 900 | 1000 | 600 | 980 | 1017.75 | 950 | 1000 | 971.8421 | 1100 | 1000 | 1000 | 1060 | [5 empty] | 997.4881
| 1 | 5 | 1000 | 900 | 900 | 1000 | 800 | 1000 | 1044.1 | 955 | 1000 | 983.1427 | 1100 | 1100 | 1000 | 1060 | [5 empty] | 9
| 1 | 6 | 1000 | 1100 | 1000 | 1000 | < 100 | 960 | 1033.2 | 975 | 1000 | 936.749 | 1100 | 1000 | 1000 | 1060
| Mean |  | 950 | 1000 | 933.3333 | 1000 | 700 | 973.3333 | 1035.4667 | 950.8333 | 983.3333 | 972.7116 | 1066.6667 | 1033.3333 | 1016.6667 | 1051.6667
| Median |  | 950 | 1000 | 900 | 1000 | 600 | 970 | 1035.925 | 947.5 | 1000 | 977.1698 | 1100 | 1000 | 1000 | 1055
| Std Dev. |  | 54.7723 | 89.4427 | 51.6398 | 0 | 141.4214 | 16.3299 | 11.4494 | 13.1972 | 40.8248 | 20.6236 | 51.6398 | 51.6398 | 40.8248 | 9.8319
| Rel.Std.Dev. |  | 5.766% | 8.944% | 5.533% | 0% | 20.2% | 1.678% | 1.106% | 1.388% | 4.152% | 2.12% | 4.841% | 4.997% | 4.016% | 0.9349%
| PDM3 |  | -4.761% | 0.2518% | -6.432% | 0.2518% | -29.82% | -2.422% | 3.807% | -4.677% | -1.419% | -2.484% | 6.935% | 3.594% | 1.923% | 5.432%
| Z-Score (Absolute) |  | 0.68 | 0.14 | 0.95 | 0.14 | 6.54 | 0.3 | 0.72 | 0.67 | 0.14 | 0.31 | 1.23 | 0.68 | 0.41 | 0.98
| Analytical results for Bi in OREAS 780 (Indicative Value 114 ppm) | [20 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 100 | 100 | < 100 | 90 | 180 | 100 | 200 | 100 | [11 empty] | 1
| 1 | 2 | 100 | 100 | < 100 | 180 | 180 | 200 | 100 | 100 | [11 empty] | 1
| 1 | 3 | 100 | 100 | < 100 | 180 | 180 | 100 | 100 | 100 | [11 empty] | 16
| 1 | 4 | 100 | 100 | < 100 | 90 | 180 | 100 | 100 | 100 | [11 empty] | 114.0917
| 1 | 5 | 100 | 100 | < 100 | 90 | 180 | 100 | 200 | 100 | [11 empty] | 7
| 1 | 6 | 100 | 100 | < 100 | 180 | 180 | 100 | 100 | 100
| Mean |  | 100 | 100 |  | 135 | 180 | 116.6667 | 133.3333 | 100
| Median |  | 100 | 100 |  | 135 | 180 | 100 | 100 | 100
| Std Dev. |  | 0 | 0 |  | 49.295 | 0 | 40.8248 | 51.6398 | 0
| Rel.Std.Dev. |  | 0% | 0% |  | 36.51% | 0% | 34.99% | 38.73% | 0%
| PDM3 |  | -12.35% | -12.35% |  | 18.33% | 57.77% | 2.257% | 16.87% | -12.35%
| Z-Score (Absolute) |  | 0.34 | 0.34 | 2.36 | 1.06 | 2.88 | 0.34 | 1.01 | 0.34
| Analytical results for Ca in OREAS 780 (Certified Value 5.67 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 5.81 | 5.718 | 5.632 | 5.646 | 5.8185 | 5.75 | 5.689 | 5.52 | 5.718 | 5.682 | 5.668 | 5.7025 | 5.3988 | 5.517 | 5.6227 | 5.596 | 5.532 | 5.668 | 6.2157 | 1
| 1 | 2 | 5.818 | 5.725 | 5.539 | 5.682 | 5.7497 | 5.71 | 5.739 | 5.55 | 5.753 | 5.646 | 5.739 | 5.6875 | 5.3762 | 5.575 | 5.6805 | 5.675 | 5.56 | 5.71 | 6.1592 | 22
| 1 | 3 | 5.818 | 5.71 | 5.675 | 5.668 | 5.8028 | 5.68 | 5.796 | 5.53 | 5.66 | 5.61 | 5.753 | 5.7 | 5.3621 | 5.603 | 5.619 | 5.589 | 5.532 | 5.639 | 6.185 | 16
| 1 | 4 | 5.796 | 5.718 | 5.589 | 5.617 | 5.7077 | 5.75 | 6.154 | 5.58 | 5.639 | 5.567 | 5.718 | 5.7065 | 5.3769 | 5.539 | 5.6444 | 5.703 | 5.66 | 5.66 | 6.1492 | 5.6683
| 1 | 5 | 5.882 | 5.725 | 5.575 | 5.625 | 5.7274 | 5.86 | 5.753 | 5.6 | 5.703 | 5.675 | 5.675 | 5.7322 | 5.3747 | 5.489 | 5.6327 | 5.689 | 5.589 | 5.682 | 6.3093 | 10
| 1 | 6 | 5.76 | 5.71 | 5.575 | 5.589 | 5.8283 | 5.67 | 5.61 | 5.53 | 5.775 | 5.56 | 5.682 | 5.7311 | 5.4111 | 5.532 | 5.5916 | 5.525 | 5.718 | 5.646 | 6.1943
| Mean |  | 5.814 | 5.7177 | 5.5975 | 5.6378 | 5.7724 | 5.7367 | 5.7902 | 5.5517 | 5.708 | 5.6233 | 5.7058 | 5.71 | 5.3833 | 5.5425 | 5.6318 | 5.6295 | 5.5985 | 5.6675 | 6.2021
| Median |  | 5.814 | 5.718 | 5.582 | 5.6355 | 5.7763 | 5.73 | 5.746 | 5.54 | 5.7105 | 5.628 | 5.7 | 5.7045 | 5.3766 | 5.5355 | 5.6277 | 5.6355 | 5.5745 | 5.664 | 6.1897
| Std Dev. |  | 0.039779 | 0.0067132 | 0.048381 | 0.034383 | 0.050781 | 0.069186 | 0.18926 | 0.031885 | 0.052353 | 0.052891 | 0.035841 | 0.01795 | 0.01805 | 0.040869 | 0.029658 | 0.070279 | 0.075471 | 0.025875 | 0.057739
| Rel.Std.Dev. |  | 0.6842% | 0.1174% | 0.8643% | 0.6099% | 0.8797% | 1.206% | 3.269% | 0.5743% | 0.9172% | 0.9406% | 0.6281% | 0.3144% | 0.3353% | 0.7374% | 0.5266% | 1.248% | 1.348% | 0.4565% | 0.931%
| PDM3 |  | 2.57% | 0.8704% | -1.25% | -0.538% | 1.836% | 1.206% | 2.149% | -2.058% | 0.6999% | -0.7938% | 0.6617% | 0.735% | -5.028% | -2.22% | -0.6445% | -0.685% | -1.232% | -0.01462% | 9.417%
| Z-Score (Absolute) |  | 1.43 | 0.49 | 0.69 | 0.29 | 1.02 | 0.67 | 1.2 | 1.13 | 0.4 | 0.43 | 0.37 | 0.41 | 2.77 | 1.22 | 0.35 | 0.37 | 0.67 | 0 | 5.22
| Analytical results for Cd in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 2
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for CeO2 in OREAS 780 (Indicative Value 134 ppm) | [20 empty] | Indicative
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 13 | 15 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 200 | 200 | 180 | 30 | < 210 | 100 | 200 | < 100 | [11 empty] | 1
| 1 | 2 | 100 | 100 | 140 | 45 | 310 | 100 | 200 | < 100 | [11 empty] | 3
| 1 | 3 | 200 | 200 | 150 | 15 | 310 | < 100 | 200 | 100 | [11 empty] | 16
| 1 | 4 | 200 | < 100 | 160 | 50 | < 210 | 100 | 300 | 100 | [11 empty] | 133.8255
| 1 | 5 | 100 | < 100 | 160 | 50 | 210 | 100 | 300 | < 100 | [11 empty] | 9
| 1 | 6 | 100 | 100 | 150 | 40 | < 210 | 100 | 300 | 100
| Mean |  | 150 | 150 | 156.6667 | 38.3333 | 276.6667 | 100 | 250 | 100
| Median |  | 150 | 150 | 155 | 42.5 | 310 | 100 | 250 | 100
| Std Dev. |  | 54.7723 | 57.735 | 13.6626 | 13.6626 | 57.735 | 0 | 54.7723 | 0
| Rel.Std.Dev. |  | 36.51% | 38.49% | 8.721% | 35.64% | 20.87% | 0% | 21.91% | 0%
| PDM3 |  | 12.09% | 12.09% | 17.07% | -71.36% | 106.7% | -25.28% | 86.81% | -25.28%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 0.31 | 1.28 | 0.79 | 0.56 | 1.57 | 0.79
| Analytical results for Cl in OREAS 780 (Indicative Value 44.2 ppm) | [20 empty] | Indicative
Cl | Round | Replicate | Lab | [18 empty] | 1
| No | No | 13 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 1
1
| 1 | 1 | 50 | [18 empty] | 1
| 1 | 2 | < 50 | [18 empty] | 4
| 1 | 3 | 60 | [18 empty] | 16
| 1 | 4 | 80 | [18 empty] | 44.1667
| 1 | 5 | < 50 | [18 empty] | 10
| 1 | 6 | < 50
| Mean |  | 63.3333
| Median |  | 60
| Std Dev. |  | 15.2753
| Rel.Std.Dev. |  | 24.12%
| PDM3 |  | 43.4%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 780 (Certified Value < 100 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100 | [10 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100 | [10 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100 | [10 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100 | [10 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100 | [10 empty] | 11
| 1 | 6 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cr in OREAS 780 (Certified Value 114 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 109 | 68 | < 100 | 100 | 149.78 | 140 | < 70 | 140 | 140 | 116 | 116.3 | 92.4 | 140 | 99.7192 | 100 | 100 | 100 | 109 |  | 1
| 1 | 2 | 116 | 137 | < 100 | 100 | 149.37 | < 70 | < 70 | 140 | 140 | 109 | 121.1 | 106.1 | 140 | 99.0476 | 200 | 100 | 100 | 116 |  | #N/A
| 1 | 3 | 96 | 137 | < 100 | 100 | 127.81 | 70 | < 70 | 70 | 140 | 109 | 130 | 106.1 | 140 | 106.7885 | 200 | 100 | 100 | 123 |  | 16
| 1 | 4 | 123 | 68 | < 100 | 100 | 145.19 | 140 | < 70 | 140 | 140 | 116 | 123.2 | 102.6 | 140 | 100.3231 | 100 | 100 | 100 | 109 |  | 114.1236
| 1 | 5 | 109 | 68 | < 100 | 100 | 125.56 | 70 | < 70 | 140 | 70 | 109 | 123.2 | 106.1 | 70 | 103.5409 | 100 | 100 | 100 | 116 |  | 12
| 1 | 6 | 116 | 68 | < 100 | 100 | 145.26 | 140 | < 70 | 70 | 210 | 109 | 123.2 | 106.1 | 140 | 102.9639 | 100 | 100 | 100 | 109
| Mean |  | 111.5 | 91 |  | 100 | 140.495 | 112 |  | 116.6667 | 140 | 111.3333 | 122.8333 | 103.2333 | 128.3333 | 102.0639 | 133.3333 | 100 | 100 | 113.6667
| Median |  | 112.5 | 68 |  | 100 | 145.225 | 140 |  | 140 | 140 | 109 | 123.2 | 106.1 | 140 | 101.6435 | 100 | 100 | 100 | 112.5
| Std Dev. |  | 9.225 | 35.6314 |  | 0 | 10.8967 | 38.3406 |  | 36.1478 | 44.2719 | 3.6148 | 4.4148 | 5.4888 | 28.5774 | 2.9305 | 51.6398 | 0 | 0 | 5.7155
| Rel.Std.Dev. |  | 8.274% | 39.16% |  | 0% | 7.756% | 34.23% |  | 30.98% | 31.62% | 3.247% | 3.594% | 5.317% | 22.27% | 2.871% | 38.73% | 0% | 0% | 5.028%
| PDM3 |  | -2.299% | -20.26% |  | -12.38% | 23.11% | -1.861% |  | 2.228% | 22.67% | -2.445% | 7.632% | -9.543% | 12.45% | -10.57% | 16.83% | -12.38% | -12.38% | -0.4004%
| Z-Score (Absolute) |  | 0 | 1.12 | 3.37 | NA | 1.57 | 0.81 | 4.23 | 0.12 | 1.37 | 0 | 0.6 | 0.47 | 0.75 | 0.53 | NA | NA | NA | 0.12
| Analytical results for Cs in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 08 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | < 100 | < 90 | [17 empty] | 1
| 1 | 2 | < 100 | < 90 | [17 empty] | 5
| 1 | 3 | < 100 | 90 | [17 empty] | 16
| 1 | 4 | < 100 | < 90 | [17 empty] | < 100
| 1 | 5 | < 100 | < 90 | [17 empty] | 11
| 1 | 6 | < 100 | < 90
| Mean |  |  | 90
| Median |  |  | 90
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 780 (Certified Value 57 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 50 | 70 | < 100 | 40 | < 40 | < 160 | 67.9 | < 80 | 41.7944 | 60 | 70 | 50 | 48 | [6 empty] | 1
| 1 | 2 | < 50 | 50 | < 100 | < 40 | < 40 | < 160 | 67.9 | 80 | 44.8844 | 70 | 110 | 50 | 64 | [6 empty] | #N/A
| 1 | 3 | 50 | 50 | < 100 | 40 | < 40 | < 160 | 67.9 | 80 | 41.8666 | 60 | 90 | 70 | 56 | [6 empty] | 16
| 1 | 4 | 50 | 60 | < 100 | 40 | < 40 | < 160 | 59.9 | < 80 | 37.6894 | 70 | 90 | 70 | 56 | [6 empty] | 57.2731
| 1 | 5 | 70 | 70 | < 100 | < 40 | 40 | < 160 | 67.9 | < 80 | 54.195 | 70 | 90 | 60 | 64 | [6 empty] | 13
| 1 | 6 | 60 | 60 | < 100 | 48 | < 40 | < 160 | 63.9 | < 80 | 41.8433 | 60 | 80 | < 50 | 80
| Mean |  | 56 | 60 |  | 42 | 40 |  | 65.9 | 80 | 43.7122 | 65 | 88.3333 | 60 | 61.3333
| Median |  | 50 | 60 |  | 40 | 40 |  | 67.9 | 80 | 41.855 | 65 | 90 | 60 | 60
| Std Dev. |  | 8.9443 | 8.9443 |  | 4 |  |  | 3.3466 | 0 | 5.6234 | 5.4772 | 13.2916 | 10 | 10.9301
| Rel.Std.Dev. |  | 15.97% | 14.91% |  | 9.524% |  |  | 5.078% | 0% | 12.86% | 8.427% | 15.05% | 16.67% | 17.82%
| PDM3 |  | -2.223% | 4.761% |  | -26.67% | -30.16% |  | 15.06% | 39.68% | -23.68% | 13.49% | 54.23% | 4.761% | 7.089%
| Z-Score (Absolute) |  | 0.21 | 0.38 | 0.27 | 1.26 | 1.99 | 1.66 | 0.76 | 0.06 | 0.67 | 0.7 | 2.2 | 0 | 0.46
| Analytical results for Dy in OREAS 780 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 6
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 12
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 780 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 7
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 13
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for F in OREAS 780 (Certified Value 5.15 wt.%) | [20 empty] | Certified
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 08 | 13 | 15 | 17 | 22 | 23 | [9 empty] | wt.%
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 2
3
| 1 | 1 | 5.15 | 3.6775 | 5.4 | 4.9753 | 5.3 | 4.8 | 5.218 | 5 | 5.2 | 5.2 | [9 empty] | 1
| 1 | 2 | 5.19 | 3.5129 | 5.5 | 4.8401 | 5.28 | 4.9 | 5.1208 | 5.2 | 5.3 | 5.1 | [9 empty] | #N/A
| 1 | 3 | 5.25 | 3.5438 | 5.4 | 4.9 | 5.3 | 5 | 5.1646 | 5.1 | 5.2 | 5.1 | [9 empty] | 16
| 1 | 4 | 5.25 | 3.5941 | 5.3 | 4.8565 | 5.83 | 4.9 | 5.1046 | 5.1 | 5.2 | 5.2 | [9 empty] | 5.1467
| 1 | 5 | 5.37 | 3.5852 | 5.3 | 4.9834 | 5.84 | 4.8 | 5.0236 | 5 | 5.3 | 5.2 | [9 empty] | 14
| 1 | 6 | 4.98 | 3.5945 | 5.3 | 4.9469 | 5.37 | 5 | 5.2245 | 5 | 5.3 | 5.2
| Mean |  | 5.1983 | 3.5847 | 5.3667 | 4.917 | 5.4867 | 4.9 | 5.1427 | 5.0667 | 5.25 | 5.1667
| Median |  | 5.22 | 3.5896 | 5.35 | 4.9235 | 5.335 | 4.9 | 5.1427 | 5.05 | 5.25 | 5.2
| Std Dev. |  | 0.13029 | 0.055894 | 0.08165 | 0.060922 | 0.27156 | 0.089443 | 0.076093 | 0.08165 | 0.054772 | 0.05164
| Rel.Std.Dev. |  | 2.506% | 1.559% | 1.521% | 1.239% | 4.95% | 1.825% | 1.48% | 1.612% | 1.043% | 0.9995%
| PDM3 |  | 1.003% | -30.35% | 4.273% | -4.463% | 6.605% | -4.794% | -0.07849% | -1.556% | 2.007% | 0.3874%
| Z-Score (Absolute) |  | 0.19 | 6.89 | 0.93 | 1.04 | 1.46 | 1.12 | 0.05 | 0.39 | 0.42 | 0.05
| Analytical results for Fe in OREAS 780 (Certified Value 2.5 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.553 | 2.6089 | 2.44 | 2.45 | 2.5244 | 2.49 | 2.875 | 2.497 | 2.511 | 2.497 | 2.504 | 2.5232 | 2.3865 | 2.497 | 2.4823 | 2.55 | 2.55 | 2.48 | 2.4264 | 1
| 1 | 2 | 2.567 | 2.6159 | 2.44 | 2.45 | 2.5175 | 2.46 | 2.728 | 2.476 | 2.511 | 2.469 | 2.504 | 2.5217 | 2.3808 | 2.49 | 2.4969 | 2.52 | 2.55 | 2.49 | 2.4257 | 28
| 1 | 3 | 2.553 | 2.6019 | 2.5 | 2.48 | 2.534 | 2.46 | 2.966 | 2.483 | 2.483 | 2.455 | 2.511 | 2.5344 | 2.3865 | 2.49 | 2.4855 | 2.53 | 2.55 | 2.51 | 2.4481 | 16
| 1 | 4 | 2.546 | 2.6089 | 2.47 | 2.5 | 2.5273 | 2.58 | 2.707 | 2.469 | 2.49 | 2.476 | 2.504 | 2.518 | 2.3876 | 2.504 | 2.4817 | 2.56 | 2.57 | 2.47 | 2.4194 | 2.4984
| 1 | 5 | 2.553 | 2.6159 | 2.44 | 2.47 | 2.5409 | 2.46 | 2.98 | 2.504 | 2.504 | 2.448 | 2.511 | 2.5211 | 2.3883 | 2.497 | 2.4923 | 2.56 | 2.57 | 2.48 | 2.4425 | 15
| 1 | 6 | 2.574 | 2.6019 | 2.45 | 2.46 | 2.526 | 2.46 | 2.721 | 2.469 | 2.504 | 2.455 | 2.497 | 2.5281 | 2.3833 | 2.525 | 2.488 | 2.52 | 2.55 | 2.48 | 2.4257
| Mean |  | 2.5577 | 2.6089 | 2.4567 | 2.4683 | 2.5283 | 2.485 | 2.8295 | 2.483 | 2.5005 | 2.4667 | 2.5052 | 2.5244 | 2.3855 | 2.5005 | 2.4878 | 2.54 | 2.5567 | 2.485 | 2.4313
| Median |  | 2.553 | 2.6089 | 2.445 | 2.465 | 2.5266 | 2.46 | 2.8015 | 2.4795 | 2.504 | 2.462 | 2.504 | 2.5225 | 2.3865 | 2.497 | 2.4868 | 2.54 | 2.55 | 2.48 | 2.4261
| Std Dev. |  | 0.010539 | 0.006261 | 0.024221 | 0.019408 | 0.0081101 | 0.048062 | 0.12683 | 0.014683 | 0.011502 | 0.018074 | 0.0052694 | 0.0058874 | 0.0028776 | 0.013096 | 0.0059455 | 0.018974 | 0.010328 | 0.013784 | 0.011278
| Rel.Std.Dev. |  | 0.412% | 0.24% | 0.9859% | 0.7863% | 0.3208% | 1.934% | 4.483% | 0.5914% | 0.46% | 0.7327% | 0.2103% | 0.2332% | 0.1206% | 0.5237% | 0.239% | 0.747% | 0.404% | 0.5547% | 0.4639%
| PDM3 |  | 2.371% | 4.422% | -1.671% | -1.204% | 1.198% | -0.5371% | 13.25% | -0.6172% | 0.08328% | -1.271% | 0.2701% | 1.04% | -4.521% | 0.08328% | -0.4258% | 1.664% | 2.331% | -0.5371% | -2.686%
| Z-Score (Absolute) |  | 1.2 | 2.27 | 0.92 | 0.67 | 0.58 | 0.33 | 6.89 | 0.37 | 0 | 0.71 | 0.1 | 0.5 | 2.41 | 0 | 0.27 | 0.83 | 1.18 | 0.33 | 1.45
| Analytical results for Ga in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 08 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | < 100 | < 150 | [17 empty] | 1
| 1 | 2 | < 100 | < 150 | [17 empty] | 8
| 1 | 3 | < 100 | < 150 | [17 empty] | 16
| 1 | 4 | < 100 | < 150 | [17 empty] | < 100
| 1 | 5 | < 100 | < 150 | [17 empty] | 14
| 1 | 6 | < 100 | < 150
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Gd in OREAS 780 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 9
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 15
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 10
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 16
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for H2O- in OREAS 780 (Indicative Value 0.323 wt.%) | [20 empty] | Indicative
H2O- | Round | Replicate | Lab | [18 empty] | 1
| No | No | 20 | [18 empty] | wt.%
|  |  | BF*XRF | [18 empty] | 3
3
| 1 | 1 | 0.3397 | [18 empty] | 1
| 1 | 2 | 0.3459 | [18 empty] | 11
| 1 | 3 | 0.3219 | [18 empty] | 16
| 1 | 4 | 0.347 | [18 empty] | 0.32302
| 1 | 5 | 0.3467 | [18 empty] | 17
| 1 | 6 | 0.2369
| Mean |  | 0.32302
| Median |  | 0.3428
| Std Dev. |  | 0.043256
| Rel.Std.Dev. |  | 13.39%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 780 (Certified Value < 100 ppm) | [20 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 09 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 200 | 10 | 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 200 | 10 | < 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 200 | 10 | < 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 200 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 200 | 10 | 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | 16
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 200 | 10 | 200 | < 100 | < 100 | < 100 | < 100
| Mean | [6 empty] | 9.1667 | 133.3333
| Median | [6 empty] | 10 | 100
| Std Dev. | [6 empty] | 2.0412 | 57.735
| Rel.Std.Dev. | [6 empty] | 22.27% | 43.3%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Hg in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 12
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
In | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 1
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 780 (Certified Value 2.61 wt.%) | [20 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.619 | 2.72 | 2.61 | 2.69 | 2.473 | 2.64 | 2.22 | 2.51 | 2.57 | 2.64 | 2.67 | 2.648 | 2.603 | 2.61 | 2.3884 | 2.58 | 2.7 | 2.65 | 2.585 | 1
| 1 | 2 | 2.642 | 2.73 | 2.66 | 2.7 | 2.4922 | 2.65 | 2.24 | 2.51 | 2.56 | 2.65 | 2.67 | 2.6424 | 2.6189 | 2.59 | 2.4298 | 2.56 | 2.69 | 2.64 | 2.594 | #N/A
| 1 | 3 | 2.616 | 2.71 | 2.63 | 2.68 | 2.496 | 2.65 | 2.23 | 2.53 | 2.48 | 2.64 | 2.68 | 2.649 | 2.6168 | 2.59 | 2.4685 | 2.57 | 2.69 | 2.65 | 2.559 | 16
| 1 | 4 | 2.614 | 2.72 | 2.66 | 2.69 | 2.5011 | 2.64 | 2.23 | 2.51 | 2.53 | 2.66 | 2.67 | 2.646 | 2.6071 | 2.61 | 2.4966 | 2.59 | 2.68 | 2.63 | 2.59 | 2.6107
| 1 | 5 | 2.62 | 2.73 | 2.63 | 2.68 | 2.4859 | 2.62 | 2.21 | 2.53 | 2.54 | 2.65 | 2.67 | 2.6365 | 2.6086 | 2.61 | 2.5154 | 2.59 | 2.72 | 2.65 | 2.618 | 17
| 1 | 6 | 2.629 | 2.71 | 2.63 | 2.68 | 2.4841 | 2.64 | 2.24 | 2.54 | 2.55 | 2.63 | 2.66 | 2.648 | 2.601 | 2.6 | 2.4531 | 2.56 | 2.66 | 2.64 | 2.564
| Mean |  | 2.6233 | 2.72 | 2.6367 | 2.6867 | 2.4887 | 2.64 | 2.2283 | 2.5217 | 2.5383 | 2.645 | 2.67 | 2.645 | 2.6092 | 2.6017 | 2.4586 | 2.575 | 2.69 | 2.6433 | 2.585
| Median |  | 2.6195 | 2.72 | 2.63 | 2.685 | 2.489 | 2.64 | 2.23 | 2.52 | 2.545 | 2.645 | 2.67 | 2.647 | 2.6079 | 2.605 | 2.4608 | 2.575 | 2.69 | 2.645 | 2.5875
| Std Dev. |  | 0.010501 | 0.0089443 | 0.019664 | 0.008165 | 0.0099444 | 0.010954 | 0.01169 | 0.013292 | 0.031885 | 0.010488 | 0.0063246 | 0.0047751 | 0.0072227 | 0.0098319 | 0.045935 | 0.013784 | 0.02 | 0.008165 | 0.021503
| Rel.Std.Dev. |  | 0.4003% | 0.3288% | 0.7458% | 0.3039% | 0.3996% | 0.4149% | 0.5246% | 0.5271% | 1.256% | 0.3965% | 0.2369% | 0.1805% | 0.2768% | 0.3779% | 1.868% | 0.5353% | 0.7435% | 0.3089% | 0.8319%
| PDM3 |  | 0.4846% | 4.187% | 0.9953% | 2.911% | -4.672% | 1.123% | -14.65% | -3.41% | -2.771% | 1.315% | 2.272% | 1.314% | -0.05503% | -0.3453% | -5.824% | -1.367% | 3.038% | 1.251% | -0.9837%
| Z-Score (Absolute) |  | 0 | 1.4 | 0.19 | 0.92 | 1.95 | 0.24 | 5.71 | 1.47 | 1.23 | 0.31 | 0.67 | 0.31 | 0.2 | 0.31 | 2.38 | 0.7 | 0.96 | 0.29 | 0.55
| Analytical results for La2O3 in OREAS 780 (Indicative Value 89 ppm) | [20 empty] | Indicative
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 13 | 15 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 130 | 55 | < 200 | 100 | 100 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | 100 | 110 | 65 | < 200 | 100 | 100 | 100 | [11 empty] | 2
| 1 | 3 | < 100 | < 100 | 70 | 70 | < 200 | 100 | 100 | 100 | [11 empty] | 16
| 1 | 4 | < 100 | < 100 | 120 | 80 | 200 | 100 | 100 | 100 | [11 empty] | 89.4444
| 1 | 5 | < 100 | < 100 | 70 | 65 | < 200 | 100 | 100 | 100 | [11 empty] | 8
| 1 | 6 | < 100 | < 100 | 60 | 75 | < 200 | < 100 | 100 | < 100
| Mean |  |  | 100 | 93.3333 | 68.3333 | 200 | 100 | 100 | 100
| Median |  |  | 100 | 90 | 67.5 | 200 | 100 | 100 | 100
| Std Dev. |  |  |  | 30.1109 | 8.756 |  | 0 | 0 | 0
| Rel.Std.Dev. |  |  |  | 32.26% | 12.81% |  | 0% | 0% | 0%
| PDM3 |  |  | 11.8% | 4.348% | -23.6% | 123.6% | 11.8% | 11.8% | 11.8%
| Z-Score (Absolute) |  | 1.6 | 1.24 | 0.25 | 0.82 | 1.24 | 0.18 | 0.53 | 0.18
| Analytical results for MgO in OREAS 780 (Certified Value 0.546 wt.%) | [20 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.51 | 0.58 | 0.53 | 0.57 | 0.59869 | 0.52 | 0.28 | 0.47 | 0.53 | 0.58 | 0.57 | 0.54233 | 0.538 | 0.5 | 0.55118 | 0.56 | 0.6 | 0.52 | 0.477 | 1
| 1 | 2 | 0.51 | 0.57 | 0.53 | 0.57 | 0.60054 | 0.57 | 0.28 | 0.49 | 0.52 | 0.57 | 0.55 | 0.542 | 0.5325 | 0.5 | 0.54839 | 0.53 | 0.59 | 0.53 | 0.497 | #N/A
| 1 | 3 | 0.52 | 0.59 | 0.55 | 0.57 | 0.5871 | 0.56 | 0.29 | 0.5 | 0.52 | 0.59 | 0.55 | 0.54608 | 0.538 | 0.5 | 0.54672 | 0.54 | 0.59 | 0.51 | 0.481 | 16
| 1 | 4 | 0.51 | 0.58 | 0.53 | 0.55 | 0.59004 | 0.59 | 0.31 | 0.48 | 0.52 | 0.58 | 0.56 | 0.53912 | 0.5325 | 0.5 | 0.5535 | 0.57 | 0.59 | 0.54 | 0.487 | 0.54593
| 1 | 5 | 0.5 | 0.57 | 0.52 | 0.57 | 0.59332 | 0.59 | 0.3 | 0.48 | 0.53 | 0.57 | 0.55 | 0.53696 | 0.537 | 0.5 | 0.54712 | 0.56 | 0.59 | 0.53 | 0.485 | 18
| 1 | 6 | 0.51 | 0.59 | 0.54 | 0.56 | 0.5911 | 0.59 | 0.3 | 0.5 | 0.54 | 0.58 | 0.56 | 0.54232 | 0.525 | 0.5 | 0.55161 | 0.56 | 0.59 | 0.53 | 0.477
| Mean |  | 0.51 | 0.58 | 0.53333 | 0.565 | 0.59347 | 0.57 | 0.29333 | 0.48667 | 0.52667 | 0.57833 | 0.55667 | 0.54147 | 0.53383 | 0.5 | 0.54975 | 0.55333 | 0.59167 | 0.52667 | 0.484
| Median |  | 0.51 | 0.58 | 0.53 | 0.57 | 0.59221 | 0.58 | 0.295 | 0.485 | 0.525 | 0.58 | 0.555 | 0.54216 | 0.53475 | 0.5 | 0.54978 | 0.56 | 0.59 | 0.53 | 0.483
| Std Dev. |  | 0.0063246 | 0.0089443 | 0.010328 | 0.0083666 | 0.0052003 | 0.027568 | 0.012111 | 0.012111 | 0.008165 | 0.0075277 | 0.008165 | 0.0031262 | 0.0050266 | 0 | 0.0027388 | 0.015055 | 0.0040825 | 0.010328 | 0.0075631
| Rel.Std.Dev. |  | 1.24% | 1.542% | 1.936% | 1.481% | 0.8763% | 4.837% | 4.129% | 2.488% | 1.55% | 1.302% | 1.467% | 0.5774% | 0.9416% | 0% | 0.4982% | 2.721% | 0.69% | 1.961% | 1.563%
| PDM3 |  | -6.582% | 6.24% | -2.308% | 3.493% | 8.707% | 4.409% | -46.27% | -10.86% | -3.529% | 5.935% | 1.966% | -0.8181% | -2.216% | -8.414% | 0.6997% | 1.356% | 8.377% | -3.529% | -11.34%
| Z-Score (Absolute) |  | 1.1 | 1.06 | 0.38 | 0.6 | 1.47 | 0.75 | 7.77 | 1.82 | 0.58 | 1.01 | 0.34 | 0.13 | 0.36 | NA | 0.13 | 0.24 | 1.42 | 0.58 | 1.9
Results from laboratory 17 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Mn in OREAS 780 (Indicative Value 0.007 wt.%) | [20 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 3
3
| 1 | 1 | < 0.0039 | 0.008 | < 0.01 | < 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0046 | < 0.008 | < 0.008 | 0.00349 | 0.0045379 | < 0.007 | 0.002403 | < 0.01 | 0.01 | 0.01 | 0.005 |  | 1
| 1 | 2 | < 0.0039 | 0.015 | < 0.01 | < 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0077 | < 0.008 | < 0.008 | 0.00379 | 0.0041598 | < 0.007 | 0.0026207 | < 0.01 | 0.01 | 0.01 | 0.005 |  | 3
| 1 | 3 | < 0.0039 | 0.008 | < 0.01 | < 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0039 | < 0.008 | < 0.008 | 0.00426 | 0.0041598 | < 0.007 | 0.0037863 | < 0.01 | 0.01 | 0.01 | 0.0072 |  | 16
| 1 | 4 | < 0.0039 | 0.008 | < 0.01 | < 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0046 | < 0.008 | < 0.008 | 0.00349 | 0.0034034 | < 0.007 | 0.002766 | < 0.01 | 0.01 | 0.01 | 0.0058 |  | 0.006534
| 1 | 5 | < 0.0039 | 0.015 | < 0.01 | < 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0046 | < 0.008 | < 0.008 | 0.00387 | 0.0045379 | < 0.007 | 0.0042947 | < 0.01 | 0.01 | 0.01 | 0.0065 |  | 9
| 1 | 6 | < 0.0039 | 0.008 | < 0.01 | < 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0062 | < 0.008 | < 0.008 | 0.00426 | 0.0045379 | < 0.007 | 0.0032049 | < 0.01 | 0.01 | 0.01 | 0.0058
| Mean |  |  | 0.010333 | [5 empty] | 0.0052667 |  |  | 0.00386 | 0.0042228 |  | 0.0031793 |  | 0.01 | 0.01 | 0.0058833
| Median |  |  | 0.008 | [5 empty] | 0.0046 |  |  | 0.00383 | 0.0043488 |  | 0.0029855 |  | 0.01 | 0.01 | 0.0058
| Std Dev. |  |  | 0.0036148 | [5 empty] | 0.0014137 |  |  | 0.00034606 | 0.00044209 |  | 0.00073467 |  | 0 | 0 | 0.00085888
| Rel.Std.Dev. |  |  | 34.98% | [5 empty] | 26.84% |  |  | 8.965% | 10.47% |  | 23.11% |  | 0% | 0% | 14.6%
| PDM3 |  |  | 58.15% | [5 empty] | -19.4% |  |  | -40.92% | -35.37% |  | -51.34% |  | 53.05% | 53.05% | -9.958%
| Z-Score (Absolute) |  | 1.56 | 4.65 | 0.71 | 0.71 | 0.13 | 0.13 | 0.13 | 0.92 | 0.13 | 0.13 | 0.14 | 0.13 | 0.33 | 0.64 | 0.71 | 4.41 | 4.41 | 1.36
| Analytical results for Mo in OREAS 780 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | 53 | < 70 | < 50 | 110 | < 50 | [11 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | 33 | < 70 | < 50 | 70 | 60 | [11 empty] | 4
| 1 | 3 | < 50 | < 50 | < 100 | < 33 | < 70 | < 50 | 120 | 80 | [11 empty] | 16
| 1 | 4 | < 50 | 50 | < 100 | 33 | < 70 | < 50 | 110 | < 50 | [11 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 100 | 60 | < 70 | < 50 | 140 | < 50 | [11 empty] | 10
| 1 | 6 | < 50 | < 50 | < 100 | 53 | < 70 | < 50 | 120 | < 50
| Mean |  |  | 50 |  | 46.4 |  |  | 111.6667 | 70
| Median |  |  | 50 |  | 53 |  |  | 115 | 70
| Std Dev. |  |  |  |  | 12.5618 |  |  | 23.1661 | 14.1421
| Rel.Std.Dev. |  |  |  |  | 27.07% |  |  | 20.75% | 20.2%
PDM3
| Z-Score (Absolute) |  | 0.82 | 0.53 | 0.94 | 0.35 | 0.23 | 0.82 | 5.28 | 0.23
| Analytical results for Na2O in OREAS 780 (Indicative Value 0.194 wt.%) | [20 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 24 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.25 | 0.13 | 0.12 | 0.11 | 0.27 | 0.2 | 0.21 | 0.26 | 0.1545 | 0.15554 | 0.3 | 0.26722 | 0.18 | [6 empty] | 1
| 1 | 2 | 0.25 | 0.12 | 0.15 | 0.1 | 0.26 | 0.2 | 0.2 | 0.25 | 0.1571 | 0.1515 | 0.3 | 0.2522 | 0.183 | [6 empty] | 5
| 1 | 3 | 0.28 | 0.14 | 0.14 | 0.11 | 0.27 | 0.21 | 0.2 | 0.26 | 0.157 | 0.15807 | 0.3 | 0.26278 | 0.185 | [6 empty] | 16
| 1 | 4 | 0.29 | 0.13 | 0.19 | 0.12 | 0.26 | 0.2 | 0.2 | 0.27 | 0.1595 | 0.1616 | 0.3 | 0.22914 | 0.179 | [6 empty] | 0.19438
| 1 | 5 | 0.28 | 0.12 | 0.14 | 0.11 | 0.27 | 0.2 | 0.21 | 0.25 | 0.1525 | 0.15302 | 0.3 | 0.22125 | 0.18 | [6 empty] | 11
| 1 | 6 | 0.26 | 0.14 | 0.15 | 0.11 | 0.29 | 0.21 | 0.2 | 0.25 | 0.1565 | 0.15302 | 0.3 | 0.26109 | 0.185
| Mean |  | 0.26833 | 0.13 | 0.14833 | 0.11 | 0.27 | 0.20333 | 0.20333 | 0.25667 | 0.15618 | 0.15546 | 0.3 | 0.24895 | 0.182
| Median |  | 0.27 | 0.13 | 0.145 | 0.11 | 0.27 | 0.2 | 0.2 | 0.255 | 0.15675 | 0.15428 | 0.3 | 0.25664 | 0.1815
| Std Dev. |  | 0.017224 | 0.0089443 | 0.023166 | 0.0063246 | 0.010954 | 0.005164 | 0.005164 | 0.008165 | 0.0024087 | 0.0037981 | 0 | 0.019195 | 0.0026833
| Rel.Std.Dev. |  | 6.419% | 6.88% | 15.62% | 5.75% | 4.057% | 2.54% | 2.54% | 3.181% | 1.542% | 2.443% | 0% | 7.711% | 1.474%
| PDM3 |  | 38.04% | -33.12% | -23.69% | -43.41% | 38.9% | 4.605% | 4.605% | 32.04% | -19.65% | -20.03% | 54.34% | 28.07% | -6.37%
| Z-Score (Absolute) |  | 1.01 | 0.84 | 0.59 | 1.11 | 1.04 | 0.14 | 0.14 | 0.86 | 0.49 | 0.5 | NA | 0.75 | 0.14
Results from laboratory 17 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Nb in OREAS 780 (Certified Value < 50 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 13 | 17 | 21 | 22 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | #N/A
| 1 | 3 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | 19
| 1 | 6 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Nd in OREAS 780 (Indicative Value 61 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | 51 | < 170 | [17 empty] | 1
| 1 | 2 | 51 | < 170 | [17 empty] | 6
| 1 | 3 | 77 | < 170 | [17 empty] | 16
| 1 | 4 | 69 | < 170 | [17 empty] | 61.4426
| 1 | 5 | 69 | < 170 | [17 empty] | 12
| 1 | 6 | 51 | < 170
| Mean |  | 61.3333
| Median |  | 60
| Std Dev. |  | 11.6905
| Rel.Std.Dev. |  | 19.06%
| PDM3 |  | -0.1779%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 780 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 15.7 | < 80 | 70 | < 50 | < 50 | < 47 | [7 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 19.6 | < 80 | 170 | < 50 | < 50 | < 47 | [7 empty] | 7
| 1 | 3 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 23.6 | < 80 | 110 | < 50 | < 50 | < 47 | [7 empty] | 16
| 1 | 4 | 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 19.6 | < 80 | 50 | < 50 | < 50 | < 47 | [7 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 19.6 | < 80 | < 50 | < 50 | < 50 | < 47 | [7 empty] | 13
| 1 | 6 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 31.4 | < 80 | 50 | < 50 | < 50 | < 47
| Mean |  | 50 | [5 empty] | 21.5833 |  | 90
| Median |  | 50 | [5 empty] | 19.6 |  | 70
| Std Dev. | [7 empty] | 5.4194 |  | 50.9902
| Rel.Std.Dev. | [7 empty] | 25.11% |  | 56.66%
PDM3
| Z-Score (Absolute) |  | 0.59 | 0 | 3.54 | 1.87 | 0.76 | 7.59 | 0.48 | 2.02 | 7.67 | 0 | 0 | 0.2
| Analytical results for P2O5 in OREAS 780 (Certified Value 0.063 wt.%) | [20 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.05 | 0.07 | 0.05 | 0.06 | 0.06683 | 0.07 | 0.04 | 0.05 | 0.05 | 0.066 | 0.07 | 0.06489 | 0.06324 | 0.06 | 0.064031 | 0.06 | 0.06 | 0.06 | 0.058 | 1
| 1 | 2 | 0.051 | 0.06 | 0.05 | 0.06 | 0.06395 | 0.07 | 0.04 | 0.04 | 0.06 | 0.065 | 0.074 | 0.06538 | 0.06375 | 0.06 | 0.062321 | 0.06 | 0.06 | 0.06 | 0.063 | #N/A
| 1 | 3 | 0.053 | 0.08 | 0.05 | 0.06 | 0.06214 | 0.07 | 0.04 | 0.05 | 0.06 | 0.066 | 0.074 | 0.06489 | 0.06426 | 0.05 | 0.062711 | 0.06 | 0.06 | 0.06 | 0.067 | 16
| 1 | 4 | 0.053 | 0.07 | 0.05 | 0.07 | 0.06727 | 0.07 | 0.04 | 0.05 | 0.05 | 0.067 | 0.071 | 0.065355 | 0.06426 | 0.06 | 0.063442 | 0.07 | 0.07 | 0.06 | 0.062 | 0.063304
| 1 | 5 | 0.05 | 0.06 | 0.05 | 0.06 | 0.06087 | 0.07 | 0.04 | 0.05 | 0.05 | 0.065 | 0.073 | 0.062955 | 0.06273 | 0.06 | 0.063726 | 0.07 | 0.07 | 0.06 | 0.062 | 20
| 1 | 6 | 0.052 | 0.08 | 0.05 | 0.07 | 0.06438 | 0.08 | 0.04 | 0.05 | 0.06 | 0.067 | 0.071 | 0.06342 | 0.06426 | 0.05 | 0.063108 | 0.06 | 0.07 | 0.06 | 0.063
| Mean |  | 0.0515 | 0.07 | 0.05 | 0.063333 | 0.06424 | 0.071667 | 0.04 | 0.048333 | 0.055 | 0.066 | 0.072167 | 0.064482 | 0.06375 | 0.056667 | 0.063223 | 0.063333 | 0.065 | 0.06 | 0.0625
| Median |  | 0.0515 | 0.07 | 0.05 | 0.06 | 0.064165 | 0.07 | 0.04 | 0.05 | 0.055 | 0.066 | 0.072 | 0.06489 | 0.064005 | 0.06 | 0.063275 | 0.06 | 0.065 | 0.06 | 0.0625
| Std Dev. |  | 0.0013784 | 0.0089443 | 7.6012e-18 | 0.005164 | 0.0025203 | 0.0040825 | 0 | 0.0040825 | 0.0054772 | 0.00089443 | 0.0017224 | 0.0010355 | 0.0006451 | 0.005164 | 0.00063928 | 0.005164 | 0.0054772 | 0 | 0.002881
| Rel.Std.Dev. |  | 2.677% | 12.78% | 1.52e-14% | 8.154% | 3.923% | 5.696% | 0% | 8.447% | 9.959% | 1.355% | 2.387% | 1.606% | 1.012% | 9.113% | 1.011% | 8.154% | 8.427% | 0% | 4.61%
| PDM3 |  | -18.65% | 10.58% | -21.02% | 0.04658% | 1.479% | 13.21% | -36.81% | -23.65% | -13.12% | 4.259% | 14% | 1.861% | 0.7048% | -10.48% | -0.1273% | 0.04658% | 2.679% | -5.219% | -1.27%
| Z-Score (Absolute) |  | 2.39 | 1.35 | 2.7 | 0 | 0.18 | 1.69 | 4.72 | 3.03 | 1.69 | 0.54 | 1.79 | 0.23 | 0.08 | 1.35 | 0.02 | 0 | 0.34 | 0.67 | 0.17
| Analytical results for Pb in OREAS 780 (Certified Value < 50 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | < 46 | 923.7 | 27.8 | < 90 | < 50 | 90 | < 50 | < 74 | [8 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 46 | 914.4 | 37.1 | < 90 | < 50 | 60 | < 50 | < 74 | [8 empty] | #N/A
| 1 | 3 | < 50 | < 50 | < 100 | < 46 | 914.4 | 18.6 | < 90 | < 50 | 90 | < 50 | < 74 | [8 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 46 | 923.7 | 18.6 | < 90 | < 50 | 100 | < 50 | < 74 | [8 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 100 | < 46 | 914.4 | 27.8 | < 90 | < 50 | 120 | < 50 | < 74 | [8 empty] | 21
| 1 | 6 | < 50 | < 50 | < 100 | < 46 | 923.7 | 37.1 | < 90 | < 50 | 100 | < 50 | < 74
| Mean | [5 empty] | 919.05 | 27.8333 |  |  | 93.3333
| Median | [5 empty] | 919.05 | 27.8 |  |  | 95
| Std Dev. | [5 empty] | 5.0938 | 8.2735 |  |  | 19.6638
| Rel.Std.Dev. | [5 empty] | 0.5542% | 29.73% |  |  | 21.07%
PDM3
| Z-Score (Absolute) |  | 0.41 | 0.41 | 3.22 | 0.67 | 129.47 | 0 | 2.7 | 0.41 | 9.51 | 0.41 | 1.35
| Analytical results for Pr in OREAS 780 (Indicative Value 42.8 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 1
1
| 1 | 1 | 41 | < 170 | [17 empty] | 1
| 1 | 2 | < 41 | < 170 | [17 empty] | 8
| 1 | 3 | 58 | < 170 | [17 empty] | 16
| 1 | 4 | < 41 | < 170 | [17 empty] | 42.7633
| 1 | 5 | 50 | < 170 | [17 empty] | 14
| 1 | 6 | 66 | < 170
| Mean |  | 53.75
| Median |  | 54
| Std Dev. |  | 10.7199
| Rel.Std.Dev. |  | 19.94%
| PDM3 |  | 25.69%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 780 (Indicative Value 138 ppm) | [20 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 80 | 90 | < 100 | 183 | < 180 | 180 | 150 | 140 | [11 empty] | 1
| 1 | 2 | 80 | 90 | < 100 | 155 | < 180 | 190 | 140 | 120 | [11 empty] | 9
| 1 | 3 | 80 | 110 | < 100 | 146 | < 180 | 180 | 160 | 150 | [11 empty] | 16
| 1 | 4 | 80 | 110 | < 100 | 201 | < 180 | 180 | 150 | 130 | [11 empty] | 138.1468
| 1 | 5 | 90 | 140 | < 100 | 210 | < 180 | 180 | 160 | 120 | [11 empty] | 15
| 1 | 6 | 70 | 140 | < 100 | 137 | < 180 | 170 | 140 | 140
| Mean |  | 80 | 113.3333 |  | 172 |  | 180 | 150 | 133.3333
| Median |  | 80 | 110 |  | 169 |  | 180 | 150 | 135
| Std Dev. |  | 6.3246 | 22.5093 |  | 30.3183 |  | 6.3246 | 8.9443 | 12.1106
| Rel.Std.Dev. |  | 7.906% | 19.86% |  | 17.63% |  | 3.514% | 5.963% | 9.083%
| PDM3 |  | -42.09% | -17.96% |  | 24.51% |  | 30.3% | 8.58% | -3.484%
| Z-Score (Absolute) |  | 0.78 | 0.18 | 1.31 | 0.88 | 0.57 | 1.02 | 0.48 | 0.18
| Analytical results for S in OREAS 780 (Certified Value 97 ppm) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [5 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 40 | < 100 | 100 | 80 | < 40 | 50 | 112 | 101.955 | 124.58 | < 800 | 100 | 100 | < 100 | 212 | [5 empty] | 1
| 1 | 2 | < 40 | < 100 | 100 | 80 | < 40 | 70 | 112 | 99.248 | 138.88 | < 800 | < 100 | 100 | < 100 | 208 | [5 empty] | #N/A
| 1 | 3 | 80 | < 100 | 100 | 80 | < 40 | 50 | 120 | 103.957 | 132.75 | < 800 | 200 | 100 | 100 | 148 | [5 empty] | 16
| 1 | 4 | 80 | < 100 | 100 | 80 | < 40 | 60 | 108 | 106.119 | 153.17 | < 800 | 200 | 100 | < 100 | 168 | [5 empty] | 96.8498
| 1 | 5 | 80 | < 100 | 100 | 80 | < 40 | 60 | 108 | 104.117 | 128.66 | < 800 | < 100 | 100 | < 100 | 144 | [5 empty] | 22
| 1 | 6 | < 40 | < 100 | 100 | 80 | < 40 | 50 | 112 | 104.117 | 134.79 | < 800 | < 100 | 100 | < 100 | 156
| Mean |  | 80 |  | 100 | 80 |  | 56.6667 | 112 | 103.2522 | 135.4717 |  | 166.6667 | 100 | 100 | 172.6667
| Median |  | 80 |  | 100 | 80 |  | 55 | 112 | 104.037 | 133.77 |  | 200 | 100 | 100 | 162
| Std Dev. |  | 0 |  | 0 | 0 |  | 8.165 | 4.3818 | 2.3634 | 9.9774 |  | 57.735 | 0 |  | 30.0843
| Rel.Std.Dev. |  | 0% |  | 0% | 0% |  | 14.41% | 3.912% | 2.289% | 7.365% |  | 34.64% | 0% |  | 17.42%
| PDM3 |  | -17.4% |  | 3.253% | -17.4% |  | -41.49% | 15.64% | 6.611% | 39.88% |  | 72.09% | 3.253% | 3.253% | 78.28%
| Z-Score (Absolute) |  | 0.87 | 0.87 | 0 | 0.35 | 1.4 | 0.76 | 0.21 | 0.06 | 0.62 | 5.25 | 0.15 | 0 | 0.73 | 1.27
| Analytical results for Sb in OREAS 780 (Indicative Value 129 ppm) | [20 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 08 | 17 | 21 | 22 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | 80 | 350 | 60 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | 80 | 350 | < 50 | 60 | [12 empty] | 10
| 1 | 3 | < 50 | < 50 | < 100 | 80 | 330 | < 50 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 80 | 370 | < 50 | < 50 | [12 empty] | 128.6658
| 1 | 5 | < 50 | < 50 | < 100 | 170 | 390 | 70 | < 50 | [12 empty] | 16
| 1 | 6 | < 50 | < 50 | < 100 | 80 | 350 | 50 | < 50
| Mean |  |  |  |  | 98 | 356.6667 | 60 | 60
| Median |  |  |  |  | 80 | 350 | 60 | 60
| Std Dev. |  |  |  |  | 40.2492 | 20.6559 | 10
| Rel.Std.Dev. |  |  |  |  | 41.07% | 5.791% | 16.67%
| PDM3 |  |  |  |  | -23.83% | 177.2% | -53.37% | -53.37%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.29 | 1.85 | 12.11 | 0 | 0.45
| Analytical results for Sc in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Sc | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 11
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 12
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 780 (Certified Value 63.83 wt.%) | [20 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 62.58 | 63.62 | 63.9 | 64.9 | 64.5345 | 64 | 62.22 | 65.02 | 63.22 | 63.92 | 65.36 | 64.009 | 63.29 | 63.6 | 63.7489 | 63.6 | 64.3 | 64.1 | 61.582 | 1
| 1 | 2 | 63.15 | 63.63 | 63.7 | 65 | 64.4514 | 63.9 | 62.21 | 64.42 | 63.11 | 63.98 | 65.49 | 63.8827 | 63.265 | 63.3 | 63.8904 | 62.9 | 63.9 | 64.4 | 61.624 | #N/A
| 1 | 3 | 62.49 | 63.61 | 65.1 | 64.5 | 64.9205 | 63.7 | 62.15 | 64.3 | 62.76 | 63.84 | 65.44 | 64.0459 | 63.2689 | 63.4 | 63.7003 | 63.1 | 64.2 | 64.2 | 61.475 | 16
| 1 | 4 | 62.64 | 63.62 | 64.2 | 64.7 | 64.5716 | 63.9 | 62.14 | 64.28 | 62.59 | 64.09 | 65.3 | 63.994 | 63.2678 | 63.5 | 63.837 | 64 | 64 | 63.9 | 61.612 | 63.8312
| 1 | 5 | 62.78 | 63.63 | 63.7 | 64.8 | 64.759 | 63.4 | 62.18 | 64.8 | 62.96 | 64.04 | 65.46 | 63.8214 | 63.371 | 63.5 | 63.6542 | 64.3 | 64.3 | 64.1 | 61.39 | 23
| 1 | 6 | 63.16 | 63.61 | 63.9 | 64.7 | 64.6027 | 63.6 | 62.2 | 64.68 | 63.1 | 63.94 | 65.28 | 64.1166 | 63.2822 | 63.6 | 63.9412 | 62.3 | 63.6 | 64.3 | 61.464
| Mean |  | 62.8 | 63.62 | 64.0833 | 64.7667 | 64.6399 | 63.75 | 62.1833 | 64.5833 | 62.9567 | 63.9683 | 65.3883 | 63.9783 | 63.2908 | 63.4833 | 63.7953 | 63.3667 | 64.05 | 64.1667 | 61.5245
| Median |  | 62.71 | 63.62 | 63.9 | 64.75 | 64.5871 | 63.8 | 62.19 | 64.55 | 63.03 | 63.96 | 65.4 | 64.0015 | 63.2756 | 63.5 | 63.7929 | 63.35 | 64.1 | 64.15 | 61.5285
| Std Dev. |  | 0.29072 | 0.0089443 | 0.53072 | 0.17512 | 0.17066 | 0.22583 | 0.03266 | 0.29864 | 0.23939 | 0.089088 | 0.087731 | 0.1083 | 0.040456 | 0.1169 | 0.11238 | 0.74207 | 0.27386 | 0.17512 | 0.094935
| Rel.Std.Dev. |  | 0.4629% | 0.01406% | 0.8282% | 0.2704% | 0.264% | 0.3542% | 0.05252% | 0.4624% | 0.3802% | 0.1393% | 0.1342% | 0.1693% | 0.06392% | 0.1841% | 0.1762% | 1.171% | 0.4276% | 0.2729% | 0.1543%
| PDM3 |  | -1.616% | -0.3309% | 0.395% | 1.466% | 1.267% | -0.1272% | -2.582% | 1.178% | -1.37% | 0.2148% | 2.439% | 0.2304% | -0.8466% | -0.545% | -0.05623% | -0.7278% | 0.3428% | 0.5255% | -3.614%
| Z-Score (Absolute) |  | 1.57 | 0.28 | 0.45 | 1.53 | 1.33 | 0.07 | 2.54 | 1.24 | 1.32 | 0.27 | 2.51 | 0.29 | 0.79 | 0.49 | 0 | 0.67 | 0.4 | 0.58 | 3.57
| Analytical results for Sm in OREAS 780 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 28
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 7
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 780 (Indicative Value 49.3 ppm) | [20 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 14 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 1
1
| 1 | 1 | 60 | < 50 | < 100 | 55 | < 160 | < 80 | 60 | 60 | 80 | [10 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | 63 | < 160 | < 80 | < 50 | 60 | 50 | [10 empty] | 2
| 1 | 3 | < 50 | < 50 | < 100 | 55 | < 160 | < 80 | < 50 | 80 | 70 | [10 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | 55 | < 160 | < 80 | 50 | < 50 | < 50 | [10 empty] | 49.3239
| 1 | 5 | < 50 | < 50 | < 100 | 63 | < 160 | < 80 | < 50 | 80 | 50 | [10 empty] | 8
| 1 | 6 | 60 | < 50 | < 100 | < 39 | < 160 | < 80 | < 50 | 90 | 60
| Mean |  | 60 |  |  | 58.2 |  |  | 55 | 74 | 62
| Median |  | 60 |  |  | 55 |  |  | 55 | 80 | 60
| Std Dev. |  | 0 |  |  | 4.3818 |  |  | 7.0711 | 13.4164 | 13.0384
| Rel.Std.Dev. |  | 0% |  |  | 7.529% |  |  | 12.86% | 18.13% | 21.03%
| PDM3 |  | 21.64% |  |  | 18% |  |  | 11.51% | 50.03% | 25.7%
| Z-Score (Absolute) |  | 0.67 | 1.26 | 0 | 0.09 | 1.45 | 0.54 | 0.76 | 0.8 | 0.3
| Analytical results for Sr in OREAS 780 (Indicative Value 80 ppm) | [20 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [6 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 51 | < 100 | < 100 | < 80 | 80 | < 42 | 109.9 | < 80 | 63.3733 | 100 | 100 | 100 | < 80 | [6 empty] | 1
| 1 | 2 | 59 | < 100 | < 100 | < 80 | 80 | < 42 | 118.4 | < 80 | 65.0753 | 100 | 100 | 100 | < 80 | [6 empty] | 3
| 1 | 3 | 51 | < 100 | < 100 | < 80 | 80 | < 42 | 126.8 | < 80 | 63.3593 | 100 | 100 | 100 | < 80 | [6 empty] | 16
| 1 | 4 | 59 | < 100 | < 100 | < 80 | 80 | < 42 | 139.5 | < 80 | 63.8653 | 100 | 100 | < 100 | < 80 | [6 empty] | 80.4564
| 1 | 5 | 59 | < 100 | < 100 | < 80 | 80 | < 42 | 101.5 | < 80 | 65.7089 | 100 | 100 | < 100 | < 80 | [6 empty] | 9
| 1 | 6 | 51 | < 100 | < 100 | < 80 | 80 | < 42 | 109.9 | < 80 | 66.3625 | 100 | 100 | < 100 | < 80
| Mean |  | 55 |  |  |  | 80 |  | 117.6667 |  | 64.6241 | 100 | 100 | 100
| Median |  | 55 |  |  |  | 80 |  | 114.15 |  | 64.4703 | 100 | 100 | 100
| Std Dev. |  | 4.3818 |  |  |  | 0 |  | 13.7306 |  | 1.2761 | 0 | 0 | 0
| Rel.Std.Dev. |  | 7.967% |  |  |  | 0% |  | 11.67% |  | 1.975% | 0% | 0% | 0%
| PDM3 |  | -31.64% |  |  |  | -0.5673% |  | 46.25% |  | -19.68% | 24.29% | 24.29% | 24.29%
| Z-Score (Absolute) |  | 0.43 | 0 | 0 | 0.67 | 3.02 | 2.52 | 5.91 | 0.67 | 1.28 | NA | NA | NA | 0.67
Results from laboratories 21, 22 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for Ta in OREAS 780 (Indicative Value < 80 ppm) | [20 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 08 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | < 100 | < 80 | [17 empty] | 1
| 1 | 2 | < 100 | < 80 | [17 empty] | 4
| 1 | 3 | < 100 | < 80 | [17 empty] | 16
| 1 | 4 | < 100 | < 80 | [17 empty] | < 80
| 1 | 5 | < 100 | < 80 | [17 empty] | 10
| 1 | 6 | < 100 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Te in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 5
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 11
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 780 (Indicative Value 69 ppm) | [20 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 13 | 15 | 17 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 44 | 48.3 | 90 | [16 empty] | 1
| 1 | 2 | < 44 | 52.7 | 90 | [16 empty] | 6
| 1 | 3 | < 44 | 52.7 | < 90 | [16 empty] | 16
| 1 | 4 | < 44 | 52.7 | 90 | [16 empty] | 68.8403
| 1 | 5 | < 44 | 43.9 | < 90 | [16 empty] | 12
| 1 | 6 | < 44 | 48.3 | 180
| Mean |  |  | 49.7667 | 112.5
| Median |  |  | 50.5 | 90
| Std Dev. |  |  | 3.5926 | 45
| Rel.Std.Dev. |  |  | 7.219% | 40%
| PDM3 |  |  | -27.71% | 63.42%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.92
| Analytical results for TiO2 in OREAS 780 (Certified Value 0.688 wt.%) | [20 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.74 | 0.73 | 0.66 | 0.7 | 0.71561 | 0.73 | 0.51 | 0.834 | 0.69 | 0.7 | 0.68 | 0.64428 | 0.64821 | 0.69 | 0.69062 | 0.68 | 0.7 | 0.69 | 0.642 | 1
| 1 | 2 | 0.75 | 0.74 | 0.67 | 0.7 | 0.71379 | 0.71 | 0.5 | 0.834 | 0.67 | 0.68 | 0.68 | 0.63839 | 0.65382 | 0.68 | 0.69474 | 0.67 | 0.71 | 0.69 | 0.655 | #N/A
| 1 | 3 | 0.73 | 0.72 | 0.68 | 0.7 | 0.71764 | 0.72 | 0.49 | 0.851 | 0.66 | 0.68 | 0.68 | 0.63589 | 0.64974 | 0.7 | 0.6963 | 0.67 | 0.7 | 0.69 | 0.672 | 16
| 1 | 4 | 0.73 | 0.73 | 0.67 | 0.7 | 0.71209 | 0.7 | 0.48 | 0.834 | 0.68 | 0.69 | 0.68 | 0.64572 | 0.65229 | 0.68 | 0.68536 | 0.68 | 0.7 | 0.68 | 0.648 | 0.68751
| 1 | 5 | 0.74 | 0.74 | 0.67 | 0.7 | 0.71957 | 0.71 | 0.48 | 0.867 | 0.68 | 0.67 | 0.69 | 0.62911 | 0.65331 | 0.69 | 0.69262 | 0.68 | 0.71 | 0.68 | 0.647 | 24
| 1 | 6 | 0.74 | 0.72 | 0.66 | 0.7 | 0.71011 | 0.71 | 0.48 | 0.851 | 0.68 | 0.69 | 0.68 | 0.6434 | 0.65178 | 0.7 | 0.6956 | 0.68 | 0.7 | 0.69 | 0.647
| Mean |  | 0.73833 | 0.73 | 0.66833 | 0.7 | 0.7148 | 0.71333 | 0.49 | 0.84517 | 0.67667 | 0.685 | 0.68167 | 0.63947 | 0.65153 | 0.69 | 0.69254 | 0.67667 | 0.70333 | 0.68667 | 0.65183
| Median |  | 0.74 | 0.73 | 0.67 | 0.7 | 0.7147 | 0.71 | 0.485 | 0.8425 | 0.68 | 0.685 | 0.68 | 0.6409 | 0.65204 | 0.69 | 0.69368 | 0.68 | 0.7 | 0.69 | 0.6475
| Std Dev. |  | 0.0075277 | 0.0089443 | 0.0075277 | 1.2162e-16 | 0.0035167 | 0.010328 | 0.012649 | 0.013556 | 0.010328 | 0.010488 | 0.0040825 | 0.006308 | 0.0021577 | 0.0089443 | 0.0040886 | 0.005164 | 0.005164 | 0.005164 | 0.010722
| Rel.Std.Dev. |  | 1.02% | 1.225% | 1.126% | 1.737e-14% | 0.492% | 1.448% | 2.581% | 1.604% | 1.526% | 1.531% | 0.5989% | 0.9864% | 0.3312% | 1.296% | 0.5904% | 0.7631% | 0.7342% | 0.752% | 1.645%
| PDM3 |  | 7.392% | 6.18% | -2.789% | 1.817% | 3.97% | 3.756% | -28.73% | 22.93% | -1.577% | -0.3652% | -0.85% | -6.988% | -5.234% | 0.3621% | 0.7318% | -1.577% | 2.301% | -0.1227% | -5.189%
| Z-Score (Absolute) |  | 1.57 | 1.32 | 0.52 | NA | 0.87 | 0.82 | 5.87 | 4.78 | 0.27 | 0.02 | 0.12 | 1.39 | 1.03 | 0.12 | 0.2 | 0.27 | 0.52 | 0.02 | 1.02
Results from laboratory 4 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Tl in OREAS 780 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Tl | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 7
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 780 (Indicative Value < 17 ppm) | [20 empty] | Indicative
U | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 1
1
| 1 | 1 | < 17 | < 80 | [17 empty] | 1
| 1 | 2 | < 17 | < 80 | [17 empty] | 8
| 1 | 3 | < 17 | < 80 | [17 empty] | 16
| 1 | 4 | 25 | 80 | [17 empty] | < 17
| 1 | 5 | < 17 | < 80 | [17 empty] | 14
| 1 | 6 | < 17 | 80
| Mean |  | 25 | 80
| Median |  | 25 | 80
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V in OREAS 780 (Indicative Value 101 ppm) | [20 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 62 | 100 | 100 | < 100 | < 60 | 95 | 101 | < 110 | 92.4 | 110 | 82.0749 | 100 | 100 | 100 | 106 |  |  |  |  | 1
| 1 | 2 | 101 | 100 | 100 | < 100 | < 60 | 106 | 118 | < 110 | 92.4 | 110 | 96.4709 | 100 | 100 | 100 | 101 |  |  |  |  | 9
| 1 | 3 | 106 | 100 | 100 | < 100 | < 60 | 101 | 101 | < 110 | 86.8 | 110 | 95.2694 | 100 | 100 | 100 | 95 |  |  |  |  | 16
| 1 | 4 | 90 | 100 | 100 | < 100 | < 60 | 101 | 118 | < 110 | 84 | 110 | 72.766 | 100 | 100 | 100 | 95 |  |  |  |  | 100.6026
| 1 | 5 | 101 | 100 | 100 | < 100 | < 60 | 101 | 106 | < 110 | 86.8 | 110 | 92.1673 | 100 | 100 | 100 | 112 |  |  |  |  | 15
| 1 | 6 | 101 | 100 | 100 | < 100 | < 60 | 106 | 118 | < 110 | 98 | 110 | 85.6429 | 100 | 100 | 100 | 106
| Mean |  | 93.5 | 100 | 100 |  |  | 101.6667 | 110.3333 |  | 90.0667 | 110 | 87.3986 | 100 | 100 | 100 | 102.5
| Median |  | 101 | 100 | 100 |  |  | 101 | 112 |  | 89.6 | 110 | 88.9051 | 100 | 100 | 100 | 103.5
| Std Dev. |  | 16.3064 | 0 | 0 |  |  | 4.0825 | 8.5946 |  | 5.1376 | 0 | 9.0833 | 0 | 0 | 0 | 6.775
| Rel.Std.Dev. |  | 17.44% | 0% | 0% |  |  | 4.016% | 7.79% |  | 5.704% | 0% | 10.39% | 0% | 0% | 0% | 6.61%
| PDM3 |  | -7.06% | -0.599% | -0.599% |  |  | 1.058% | 9.672% |  | -10.47% | 9.341% | -13.12% | -0.599% | -0.599% | -0.599% | 1.886%
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 5.08 | 7.31 | 0.18 | 1.03 | 4.47 | 1.01 | 1.22 | 1.28 | 0 | 0 | 0 | 0.27
| Analytical results for W in OREAS 780 (Certified Value 81 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 13 | 15 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 70 | 60 | 79.29 | < 100 | 135 | 75.3 | 80 | 70 | 110 | 100 | [9 empty] | 1
| 1 | 2 | 80 | 80 | 55.51 | < 100 | 127 | 87.2 | 80 | 80 | 100 | 100 | [9 empty] | 26
| 1 | 3 | 90 | 80 | 71.37 | < 100 | 127 | 79.3 | < 80 | 80 | 100 | 90 | [9 empty] | 16
| 1 | 4 | 90 | 80 | 87.22 | < 100 | 127 | 79.3 | < 80 | 80 | 110 | 80 | [9 empty] | 81.3326
| 1 | 5 | 70 | 70 | 55.51 | < 100 | 127 | 83.3 | < 80 | 100 | 110 | 80 | [9 empty] | 25
| 1 | 6 | 70 | 70 | 55.51 | < 100 | 127 | 79.3 | < 80 | 70 | 90 | 80
| Mean |  | 78.3333 | 73.3333 | 67.4017 |  | 128.3333 | 80.6167 | 80 | 80 | 103.3333 | 88.3333
| Median |  | 75 | 75 | 63.44 |  | 127 | 79.3 | 80 | 80 | 105 | 85
| Std Dev. |  | 9.8319 | 8.165 | 13.9577 |  | 3.266 | 4.099 | 0 | 10.9545 | 8.165 | 9.8319
| Rel.Std.Dev. |  | 12.55% | 11.13% | 20.71% |  | 2.545% | 5.085% | 0% | 13.69% | 7.902% | 11.13%
| PDM3 |  | -3.688% | -9.835% | -17.13% |  | 57.79% | -0.8803% | -1.638% | -1.638% | 27.05% | 8.608%
| Z-Score (Absolute) |  | 0.05 | 0.38 | 0.76 | 1.88 | 3.16 | 0.09 | 1.69 | 0.05 | 1.56 | 0.59
| Analytical results for Y2O3 in OREAS 780 (Indicative Value 95 ppm) | [20 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 13 | 17 | 21 | 22 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | 50 | 70 | < 200 | 90 | 110 | 100 | [12 empty] | 1
| 1 | 2 | < 50 | 60 | 70 | < 200 | 100 | 80 | 110 | [12 empty] | 10
| 1 | 3 | < 50 | < 50 | 60 | < 200 | 90 | 90 | 130 | [12 empty] | 16
| 1 | 4 | < 50 | < 50 | 60 | < 200 | 100 | 140 | 110 | [12 empty] | 94.5833
| 1 | 5 | < 50 | < 50 | 70 | < 200 | 100 | 110 | 110 | [12 empty] | 16
| 1 | 6 | < 50 | < 50 | 80 | < 200 | 90 | 90 | 110
| Mean |  |  | 55 | 68.3333 |  | 95 | 103.3333 | 111.6667
| Median |  |  | 55 | 70 |  | 95 | 100 | 110
| Std Dev. |  |  | 7.0711 | 7.5277 |  | 5.4772 | 21.6025 | 9.8319
| Rel.Std.Dev. |  |  | 12.86% | 11.02% |  | 5.766% | 20.91% | 8.805%
| PDM3 |  |  | -41.85% | -27.75% |  | 0.4405% | 9.251% | 18.06%
| Z-Score (Absolute) |  | 2.83 | 2.43 | 1.08 | 0.2 | 0 | 0.34 | 0.67
| Analytical results for Zn in OREAS 780 (Indicative Value 34.1 ppm) | [20 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 1
1
| 1 | 1 | < 50 | < 50 | < 100 | < 40 | < 160 | 40.2 | < 80 | 24.4749 | < 50 | 50 | 50 | 32 | [7 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 40 | < 160 | 40.2 | < 80 | 26.7426 | < 50 | < 50 | 60 | 24 | [7 empty] | 11
| 1 | 3 | < 50 | < 50 | < 100 | < 40 | < 160 | 40.2 | < 80 | 26.7202 | < 50 | < 50 | 60 | 32 | [7 empty] | 16
| 1 | 4 | < 50 | 50 | < 100 | < 40 | < 160 | 40.2 | < 80 | 25.5367 | < 50 | < 50 | 50 | 40 | [7 empty] | 34.0651
| 1 | 5 | < 50 | 50 | < 100 | < 40 | < 160 | 40.2 | < 80 | 24.6538 | < 50 | 50 | < 50 | 48 | [7 empty] | 17
| 1 | 6 | < 50 | < 50 | < 100 | < 40 | < 160 | 36.2 | < 80 | 27.8862 | < 50 | < 50 | 60 | 48
| Mean |  |  | 50 |  |  |  | 39.5333 |  | 26.0024 |  | 50 | 56 | 37.3333
| Median |  |  | 50 |  |  |  | 40.2 |  | 26.1284 |  | 50 | 60 | 36
| Std Dev. |  |  | 0 |  |  |  | 1.633 |  | 1.3402 |  | 0 | 5.4772 | 9.6885
| Rel.Std.Dev. |  |  | 0% |  |  |  | 4.131% |  | 5.154% |  | 0% | 9.781% | 25.95%
| PDM3 |  |  | 46.78% |  |  |  | 16.05% |  | -23.67% |  | 46.78% | 64.39% | 9.594%
| Z-Score (Absolute) |  | 0.71 | 0.14 | 0.99 | 1.04 | 3.06 | 0.28 | 0.32 | 0.64 | 0.71 | 0.14 | 1.05 | 0.14
| Analytical results for Zr in OREAS 780 (Certified Value 200 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 08 | 09 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 200 | 100 | 100 | 70 | 220 | 192.5 | 185.1 | 150 | 204.5809 | 200 | 100 | 100 | 215 | [6 empty] | 1
| 1 | 2 | 200 | 100 | 100 | 70 | 220 | 194.7 | 199.9 | < 150 | 204.7198 | 200 | 100 | 200 | 237 | [6 empty] | #N/A
| 1 | 3 | 300 | 100 | 100 | 70 | 220 | 196.2 | 192.5 | 220 | 204.6885 | 200 | 100 | 200 | 244 | [6 empty] | 16
| 1 | 4 | 300 | 100 | 100 | 70 | 220 | 207.3 | 192.5 | 150 | 201.8311 | 200 | 100 | 100 | 229 | [6 empty] | 200.4677
| 1 | 5 | 200 | 100 | 100 | 70 | 150 | 188.8 | 166.6 | 220 | 206.319 | 200 | 100 | 200 | 207 | [6 empty] | 26
| 1 | 6 | 200 | 100 | 100 | 70 | 220 | 192.5 | 207.3 | 150 | 207.5426 | 200 | 100 | 200 | 237
| Mean |  | 233.3333 | 100 | 100 | 70 | 208.3333 | 195.3333 | 190.65 | 178 | 204.947 | 200 | 100 | 166.6667 | 228.1667
| Median |  | 200 | 100 | 100 | 70 | 220 | 193.6 | 192.5 | 150 | 204.7042 | 200 | 100 | 200 | 233
| Std Dev. |  | 51.6398 | 0 | 0 | 0 | 28.5774 | 6.3733 | 13.9917 | 38.3406 | 1.9277 | 0 | 0 | 51.6398 | 14.3446
| Rel.Std.Dev. |  | 22.13% | 0% | 0% | 0% | 13.72% | 3.263% | 7.339% | 21.54% | 0.9406% | 0% | 0% | 30.98% | 6.287%
| PDM3 |  | 16.39% | -50.12% | -50.12% | -65.08% | 3.924% | -2.561% | -4.897% | -11.21% | 2.234% | -0.2333% | -50.12% | -16.86% | 13.82%
| Z-Score (Absolute) |  | 0.95 | 2.02 | 2.02 | 2.6 | 0.43 | 0.1 | 0 | 0.67 | 0.32 | 0.21 | 2.02 | 0.53 | 0.84
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C722: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C758: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C776: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C813: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C831: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C867: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C904: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C922: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C940: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C958: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C976: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C994: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 780 (Indicative Value 0.323 wt.%) | [20 empty] | Indicative
H2O- | Round | Replicate | Lab | [18 empty] | 1
| No | No | 20 | [18 empty] | wt.%
|  |  | BF*XRF | [18 empty] | 3
3
| 1 | 1 | 0.3397 | [18 empty] | 1
| 1 | 2 | 0.3459 | [18 empty] | 11
| 1 | 3 | 0.3219 | [18 empty] | 16
| 1 | 4 | 0.347 | [18 empty] | 0.32302
| 1 | 5 | 0.3467 | [18 empty] | 20
| 1 | 6 | 0.2369
| Mean |  | 0.32302
| Median |  | 0.3428
| Std Dev. |  | 0.043256
| Rel.Std.Dev. |  | 13.39%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 780 (Certified Value 5.76 wt.%) | [20 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | OxBF*XRF | LOI*Furnace | LOI*TGA | OxBF*XRF | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | BF*XRF | LOI*Furnace | LOI*TGA | LOI*TGA | 2
3
| 1 | 1 | 4.34 | 6.26 | 5.75 | 5.47 | 6.1496 | 5.95 | 5.86 | 6.01 | 6.2 | 3.2125 | 4.31 | 4.4196 | 6.0509 | 4.73 | 5.7988 | 5.9744 | 5.75 | 3.79 | 4.74 | 1
| 1 | 2 | 4.2 | 6.25 | 5.85 | 5.32 | 6.029 | 5.83 | 5.95 | 6.04 | 6.09 | 3.1698 | 4.26 | 4.4074 | 6.0774 | 4.8 | 5.8876 | 6.0013 | 5.59 | 3.82 | 4.79 | #N/A
| 1 | 3 | 4.56 | 6.27 | 5.79 | 5.56 | 5.9563 | 5.93 | 5.92 | 5.99 | 6.21 | 3.2235 | 4.24 | 4.367 | 6.0804 | 4.96 | 5.9558 | 5.9811 | 5.73 | 3.85 | 4.54 | 16
| 1 | 4 | 4.92 | 6.26 | 5.74 | 5.44 | 5.6672 | 5.95 | 5.87 | 6.07 | 6.14 | 3.1917 | 4.32 | 4.3594 | 6.1402 | 4.75 | 5.88 | 6.0114 | 5.68 | 3.83 | 4.68 | 5.7561
| 1 | 5 | 4.61 | 6.25 | 5.86 | 5.51 | 6.0426 | 5.87 | 6.02 | 6.02 | 6.12 | 3.1286 | 4.27 | 4.4059 | 6.0517 | 4.85 | 6.0064 | 6.0093 | 5.76 | 3.77 | 4.57 | 28
| 1 | 6 | 4.28 | 6.27 | 5.7 | 5.34 | 6.1066 | 5.83 | 5.95 | 5.99 | 6.25 | 3.2013 | 4.33 | 4.4059 | 6.0374 | 4.65 | 5.9788 | 5.9489 | 5.56 | 3.82 | 4.61
| Mean |  | 4.485 | 6.26 | 5.7817 | 5.44 | 5.9919 | 5.8933 | 5.9283 | 6.02 | 6.1683 | 3.1879 | 4.2883 | 4.3942 | 6.073 | 4.79 | 5.9179 | 5.9877 | 5.6783 | 3.8133 | 4.655
| Median |  | 4.45 | 6.26 | 5.77 | 5.455 | 6.0358 | 5.9 | 5.935 | 6.015 | 6.17 | 3.1965 | 4.29 | 4.4059 | 6.0645 | 4.775 | 5.9217 | 5.9912 | 5.705 | 3.82 | 4.645
| Std Dev. |  | 0.26636 | 0.0089443 | 0.063692 | 0.094446 | 0.17244 | 0.057155 | 0.059133 | 0.030984 | 0.061128 | 0.034386 | 0.03656 | 0.024698 | 0.036858 | 0.10714 | 0.07682 | 0.024245 | 0.085186 | 0.028752 | 0.098539
| Rel.Std.Dev. |  | 5.939% | 0.1429% | 1.102% | 1.736% | 2.878% | 0.9698% | 0.9975% | 0.5147% | 0.991% | 1.079% | 0.8526% | 0.5621% | 0.6069% | 2.237% | 1.298% | 0.4049% | 1.5% | 0.754% | 2.117%
| PDM3 |  | -22.08% | 8.754% | 0.444% | -5.492% | 4.096% | 2.384% | 2.992% | 4.585% | 7.161% | -44.62% | -25.5% | -23.66% | 5.505% | -16.78% | 2.811% | 4.024% | -1.351% | -33.75% | -19.13%
| Z-Score (Absolute) |  | 2.56 | 0.94 | 0 | 0.67 | 0.41 | 0.22 | 0.29 | 0.47 | 0.76 | 5.12 | 2.95 | 2.74 | 0.57 | 1.96 | 0.27 | 0.41 | 0.2 | 3.88 | 2.22
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ISE
| Analytical results for F in OREAS 780 (Certified Value 4.32 wt.%) | [16 empty] | Certified
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 18 | 20 | 21 | 24 | wt.%
|  |  | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | 2
2
| 1 | 1 | 4.07 | 3.402 | 4.84 | 6.2547 | 3.6828 | 4.3 | 3.98 | 3.94 | 4.483 | 5.1137 | 4.6916 | 4.207 | 5.0176 | 4.5 | 5.158 | 1
| 1 | 2 | 4.01 | 3.242 | 4.83 | 6.082 | 3.8055 | 4.2 | 4.3 | 3.71 | 4.3768 | 5.0153 | 5.0548 | 4.37 | 5.0639 | 4.48 | 4.0044 | #N/A
| 1 | 3 | 3.96 | 3.272 | 4.74 | 6.1496 | 3.571 | 4.4 | 4.15 | 3.63 | 4.2602 | 5.0962 | 5.1156 | 4.35 | 4.8678 | 4.64 | 4.213 | 16
| 1 | 4 | 3.93 | 3.3209 | 4.52 | 6.0431 | 3.7581 | 4.1005 | 3.9 | 3.78 | 4.5382 | 5.0453 | 4.7485 | 4.29 | 4.7697 | 4.42 | 3.8391 | 4.3167
| 1 | 5 | 4.17 | 3.3123 | 4.76 | 6.0496 | 3.6828 | 4.4 | 4.17 | 3.83 | 4.2973 | 5.0923 | 4.8706 | 4.21 | 4.9798 | 4.68 | 4.8012 | 30
| 1 | 6 | 4.14 | 3.3213 | 4.77 | 6.1401 | 3.6029 | 4.2561 | 4.23 | 3.8 | 4.0385 | 5.1576 | 4.8184 | 4.56 | 4.9828 | 4.4 | 4.1948
| Mean |  | 4.0467 | 3.3117 | 4.7433 | 6.1199 | 3.6838 | 4.2761 | 4.1217 | 3.7817 | 4.3323 | 5.0867 | 4.8833 | 4.3312 | 4.9469 | 4.52 | 4.3684
| Median |  | 4.04 | 3.3166 | 4.765 | 6.111 | 3.6828 | 4.2781 | 4.16 | 3.79 | 4.3371 | 5.0943 | 4.8445 | 4.32 | 4.9813 | 4.49 | 4.2039
| Std Dev. |  | 0.096885 | 0.054342 | 0.11639 | 0.079686 | 0.088997 | 0.11687 | 0.15224 | 0.10572 | 0.1788 | 0.050377 | 0.16892 | 0.13116 | 0.10836 | 0.11524 | 0.5055
| Rel.Std.Dev. |  | 2.394% | 1.641% | 2.454% | 1.302% | 2.416% | 2.733% | 3.694% | 2.796% | 4.127% | 0.9904% | 3.459% | 3.028% | 2.191% | 2.55% | 11.57%
| PDM3 |  | -6.256% | -23.28% | 9.883% | 41.77% | -14.66% | -0.9404% | -4.518% | -12.39% | 0.362% | 17.84% | 13.12% | 0.335% | 14.6% | 4.71% | 1.198%
| Z-Score (Absolute) |  | 0.47 | 1.67 | 0.67 | 2.93 | 1.06 | 0.09 | 0.35 | 0.9 | 0 | 1.24 | 0.9 | 0 | 1.01 | 0.31 | 0.06
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 780 (Certified Value 0.345 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 3
3
| 1 | 1 | < 0.3 | 0.38 | 0.31 | < 0.5 | 0.34 | 0.7 | 0.4 | 0.34 | < 1 | < 0.5 | 0.4 | 0.4 | 0.36 | 0.32 | 0.29 | 1.5082 |  |  |  | 1
| 1 | 2 | < 0.3 | 0.35 | 0.32 | < 0.5 | 0.34 | 0.7 | 0.4 | 0.34 | < 1 | < 0.5 | 0.35 | 0.4 | 0.33 | 0.33 | 0.36 | 8.6383 |  |  |  | #N/A
| 1 | 3 | < 0.3 | 0.38 | 0.33 | < 0.5 | 0.36 | 0.6 | 0.4 | 0.33 | < 1 | < 0.5 | 0.4 | 0.4 | 0.34 | 0.33 | 0.33 | 4.5908 |  |  |  | 16
| 1 | 4 | 0.5 | 0.34 | 0.36 | < 0.5 | 0.35 | 0.6 | 0.3 | 0.34 | < 1 | < 0.5 | 0.36 | 0.3 | 0.36 | 0.31 | 0.32 | 4.3615 |  |  |  | 0.34479
| 1 | 5 | 0.3 | 0.4 | 0.32 | < 0.5 | 0.33 | 0.6 | 0.4 | 0.35 | < 1 | < 0.5 | 0.38 | 0.4 | 0.33 | 0.32 | 0.33 | 1.7466 |  |  |  | 32
| 1 | 6 | < 0.3 | 0.36 | 0.33 | < 0.5 | 0.35 | 0.7 | 0.4 | 0.36 | < 1 | < 0.5 | 0.35 | 0.3 | 0.34 | 0.37 | 0.33 | 3.7521
| Mean |  | 0.4 | 0.36833 | 0.32833 |  | 0.345 | 0.65 | 0.38333 | 0.34333 |  |  | 0.37333 | 0.36667 | 0.34333 | 0.33 | 0.32667 | 4.0996
| Median |  | 0.4 | 0.37 | 0.325 |  | 0.345 | 0.65 | 0.4 | 0.34 |  |  | 0.37 | 0.4 | 0.34 | 0.325 | 0.33 | 4.0568
| Std Dev. |  | 0.14142 | 0.022286 | 0.017224 |  | 0.010488 | 0.054772 | 0.040825 | 0.010328 |  |  | 0.023381 | 0.05164 | 0.013663 | 0.020976 | 0.022509 | 2.58
| Rel.Std.Dev. |  | 35.36% | 6.051% | 5.246% |  | 3.04% | 8.427% | 10.65% | 3.008% |  |  | 6.263% | 14.08% | 3.979% | 6.356% | 6.891% | 62.93%
| PDM3 |  | 16.01% | 6.828% | -4.773% |  | 0.06042% | 88.52% | 11.18% | -0.423% |  |  | 8.278% | 6.344% | -0.423% | -4.29% | -5.257% | 1089%
| Z-Score (Absolute) |  | NA | 1.01 | 0.61 | 3.78 | 0.07 | NA | NA | 0 | 6.34 | 3.78 | 1.21 | NA | 0 | 0.54 | 0.67 | NA
Results from laboratories 1, 10, 12 and 19 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.24 were removed due to their systematic bias.
| Analytical results for Al in OREAS 780 (Certified Value 6.73 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 6.182 | 7.064 | 6.71 | 6.56 | 6.54 | 6.89 | 6.73 | 6.57 | 7.0006 | 6.9213 | 5.75 | 6.43 | 7.2581 | 6.41 | 6.9 | 6.84 | 6.92 | 6.9518 |  | 1
| 1 | 2 | 6.1277 | 7.05 | 6.03 | 6.57 | 6.51 | 6.9 | 6.68 | 6.78 | 6.9992 | 7.0891 | 6.4 | 6.04 | 7.3241 | 6.5 | 6.88 | 6.79 | 7.3 | 6.6679 |  | 19
| 1 | 3 | 6.0673 | 7.0949 | 6.61 | 6.63 | 6.57 | 6.99 | 6.98 | 6.88 | 7.0975 | 7.0186 | 6.11 | 6.28 | 7.2807 | 6.38 | 6.72 | 7.02 | 6.85 | 6.5508 |  | 16
| 1 | 4 | 6.0319 | 7.0774 | 6.35 | 6.88 | 6.53 | 6.97 | 7.22 | 6.64 | 6.9548 | 6.9247 | 5.84 | 5.18 | 7.3822 | 6.49 | 6.88 | 7.07 | 6.92 | 6.5899 |  | 6.7315
| 1 | 5 | 6.0792 | 7.0897 | 6.25 | 6.65 | 6.51 | 6.84 | 6.79 | 6.7 | 7.0281 | 6.8349 | 6.3 | 5.62 | 7.3902 | 6.29 | 6.87 | 6.81 | 7 | 6.6635 |  | 33
| 1 | 6 | 6.1197 | 7.0721 | 6.43 | 6.76 | 6.53 | 7.05 | 6.96 | 6.64 | 7.07 | 6.867 | 6.16 | 5.31 | 7.2806 | 6.28 | 6.95 | 6.97 | 6.79 | 6.5933
| Mean |  | 6.1013 | 7.0747 | 6.3967 | 6.675 | 6.5317 | 6.94 | 6.8933 | 6.7017 | 7.025 | 6.9426 | 6.0933 | 5.81 | 7.3193 | 6.3917 | 6.8667 | 6.9167 | 6.9633 | 6.6695
| Median |  | 6.0995 | 7.0747 | 6.39 | 6.64 | 6.53 | 6.935 | 6.875 | 6.67 | 7.0143 | 6.923 | 6.135 | 5.83 | 7.3024 | 6.395 | 6.88 | 6.905 | 6.92 | 6.6284
| Std Dev. |  | 0.052932 | 0.016536 | 0.24614 | 0.12341 | 0.022286 | 0.076942 | 0.20077 | 0.11215 | 0.051923 | 0.095181 | 0.25438 | 0.51792 | 0.056106 | 0.094534 | 0.077374 | 0.1186 | 0.17963 | 0.14558
| Rel.Std.Dev. |  | 0.8675% | 0.2337% | 3.848% | 1.849% | 0.3412% | 1.109% | 2.912% | 1.673% | 0.7391% | 1.371% | 4.175% | 8.914% | 0.7665% | 1.479% | 1.127% | 1.715% | 2.58% | 2.183%
| PDM3 |  | -9.362% | 5.098% | -4.974% | -0.839% | -2.968% | 3.098% | 2.404% | -0.4429% | 4.361% | 3.137% | -9.48% | -13.69% | 8.733% | -5.048% | 2.008% | 2.751% | 3.444% | -0.9202%
| Z-Score (Absolute) |  | 2.19 | 0.93 | 1.24 | 0.35 | 0.81 | 0.5 | 0.35 | 0.26 | 0.77 | 0.51 | 2.22 | 3.13 | 1.72 | 1.26 | 0.26 | 0.43 | 0.58 | 0.37
| Analytical results for As in OREAS 780 (Certified Value 428 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
0
| 1 | 1 | 429.2 | 472.58 | 446 | 419 | 403.3 | 431 | 469 | 416 | 430.9 | 429.4152 | 432 | 464 | 416.8 | 444 | 423 | 423 | 324.1224 |  |  | 1
| 1 | 2 | 429.3 | 472.28 | 438 | 415 | 407.6 | 431 | 471 | 416 | 437.7 | 437.1358 | 434 | 449 | 406.2 | 438 | 422 | 449 | 311.2765 |  |  | #N/A
| 1 | 3 | 426.5 | 472.15 | 457 | 436 | 407.4 | 429 | 461 | 408 | 441.5 | 428.9895 | 420 | 455 | 416.9 | 424 | 430 | 423 | 317.6296 |  |  | 16
| 1 | 4 | 424.3 | 473.54 | 443 | 446 | 399.2 | 421 | 478 | 417 | 447.2 | 421.1733 | 433 | 430 | 410.4 | 431 | 429 | 426 | 311.1774 |  |  | 428.1965
| 1 | 5 | 423.2 | 471.88 | 446 | 435 | 404.39 | 428 | 476 | 416 | 440.1 | 421.1563 | 424 | 449 | 423 | 431 | 418 | 424 | 314.2514 |  |  | 34
| 1 | 6 | 424 | 473.84 | 442 | 440 | 404.6 | 413 | 476 | 418 | 436.5 | 437.148 | 425 | 441 | 415.3 | 442 | 425 | 423 | 311.9594
| Mean |  | 426.0833 | 472.7117 | 445.3333 | 431.8333 | 404.415 | 425.5 | 471.8333 | 415.1667 | 438.9833 | 429.1697 | 428 | 448 | 414.7667 | 435 | 424.5 | 428 | 315.0695
| Median |  | 425.4 | 472.43 | 444.5 | 435.5 | 404.495 | 428.5 | 473.5 | 416 | 438.9 | 429.2023 | 428.5 | 449 | 416.05 | 434.5 | 424 | 423.5 | 313.1054
| Std Dev. |  | 2.6858 | 0.79625 | 6.4395 | 12.1888 | 3.0831 | 7.1484 | 6.3061 | 3.6009 | 5.4452 | 7.1465 | 5.7619 | 11.6619 | 5.8174 | 7.642 | 4.5056 | 10.3537 | 5.0673
| Rel.Std.Dev. |  | 0.6304% | 0.1684% | 1.446% | 2.823% | 0.7624% | 1.68% | 1.337% | 0.8673% | 1.24% | 1.665% | 1.346% | 2.603% | 1.403% | 1.757% | 1.061% | 2.419% | 1.608%
| PDM3 |  | -0.4935% | 10.4% | 4.002% | 0.8493% | -5.554% | -0.6297% | 10.19% | -3.043% | 2.519% | 0.2273% | -0.0459% | 4.625% | -3.136% | 1.589% | -0.8633% | -0.0459% | -26.42%
| Z-Score (Absolute) |  | 0.12 | 2.75 | 1.06 | 0.24 | 1.45 | 0.15 | 2.69 | 0.79 | 0.67 | 0.07 | 0 | 1.23 | 0.81 | 0.43 | 0.21 | 0 | 6.93
| Analytical results for B in OREAS 780 (Indicative Value 23.1 ppm) | [20 empty] | Indicative
B | Round | Replicate | Lab | [18 empty] | 1
| No | No | 09 | [18 empty] | ppm
|  |  | 4A*OES | [18 empty] | 1
1
| 1 | 1 | 24.22 | [18 empty] | 1
| 1 | 2 | 22.68 | [18 empty] | 16
| 1 | 3 | 20.98 | [18 empty] | 16
| 1 | 4 | 20.94 | [18 empty] | 23.1017
| 1 | 5 | 22.97 | [18 empty] | 22
| 1 | 6 | 26.82
| Mean |  | 23.1017
| Median |  | 22.825
| Std Dev. |  | 2.2112
| Rel.Std.Dev. |  | 9.571%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 780 (Certified Value 878 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 851 | 950.68 | 900 | 860 | 827.9 | 840 | 890 | 878 | 896.7 | 869.5049 | 859 | 900 | 915.1843 | 867 | 846.153 | 880 | 900 | 870 | 919.7102 | 1
| 1 | 2 | 839 | 950.9265 | 900 | 860 | 818.5 | 887 | 913 | 866 | 914.8 | 881.7612 | 883 | 885 | 879.0803 | 862 | 852.1782 | 880 | 900 | 930 | 890.7198 | 20
| 1 | 3 | 840 | 950.929 | 900 | 880 | 811.6 | 894 | 900 | 854 | 912.5 | 861.602 | 865 | 888 | 888.7455 | 852 | 855.6581 | 860 | 910 | 870 | 885.9635 | 16
| 1 | 4 | 845 | 949.745 | 900 | 900 | 822.1 | 878 | 914 | 879 | 944.8 | 860.6465 | 863 | 865 | 913.4327 | 859 | 852.0034 | 870 | 900 | 880 | 881.2471 | 877.9652
| 1 | 5 | 850 | 949.815 | 910 | 890 | 812.51 | 863 | 905 | 878 | 941.2 | 867.4159 | 861 | 867 | 915.8565 | 852 | 865.5804 | 870 | 870 | 880 | 897.4321 | 35
| 1 | 6 | 846 | 950.1145 | 910 | 900 | 829.6 | 906 | 929 | 880 | 912.2 | 886.8564 | 858 | 857 | 913.553 | 866 | 769.7465 | 890 | 900 | 880 | 891.2369
| Mean |  | 845.1667 | 950.3683 | 903.3333 | 881.6667 | 820.3683 | 878 | 908.5 | 872.5 | 920.3667 | 871.2978 | 864.8333 | 877 | 904.3087 | 859.6667 | 840.2199 | 875 | 896.6667 | 885 | 894.3849
| Median |  | 845.5 | 950.3972 | 900 | 885 | 820.3 | 882.5 | 909 | 878 | 913.65 | 868.4604 | 862 | 876 | 913.4928 | 860.5 | 852.0908 | 875 | 900 | 880 | 890.9783
| Std Dev. |  | 4.9565 | 0.54446 | 5.164 | 18.3485 | 7.5773 | 23.622 | 13.3978 | 10.4259 | 18.712 | 10.7452 | 9.261 | 16.4803 | 16.1184 | 6.5929 | 35.1142 | 10.4881 | 13.6626 | 22.5832 | 13.5464
| Rel.Std.Dev. |  | 0.5864% | 0.05729% | 0.5717% | 2.081% | 0.9236% | 2.69% | 1.475% | 1.195% | 2.033% | 1.233% | 1.071% | 1.879% | 1.782% | 0.7669% | 4.179% | 1.199% | 1.524% | 2.552% | 1.515%
| PDM3 |  | -3.736% | 8.247% | 2.889% | 0.4216% | -6.56% | 0.003958% | 3.478% | -0.6225% | 4.83% | -0.7594% | -1.496% | -0.1099% | 3.001% | -2.084% | -4.299% | -0.3377% | 2.13% | 0.8013% | 1.87%
| Z-Score (Absolute) |  | 1.21 | 2.66 | 0.93 | 0.13 | 2.12 | 0 | 1.12 | 0.2 | 1.56 | 0.25 | 0.48 | 0.04 | 0.97 | 0.67 | 1.39 | 0.11 | 0.69 | 0.26 | 0.6
| Analytical results for Be in OREAS 780 (Certified Value 2.6 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.06 | 2.67 | 2.49 | 2.4 | < 5 | 2.6 | 2.5 | 2.61 | 2.5257 | 2.6 | 2.7 | 2.9946 | 2 | 2.6305 | 2.6 | 2.57 | 2.52 | 2.401 |  | 1
| 1 | 2 | 2.05 | 2.64 | 2.52 | 2.28 | < 5 | 2.7 | 2.5 | 2.61 | 2.658 | 2.7 | 2.53 | 2.4103 | 2 | 2.5922 | 2.58 | 2.51 | 2.67 | 2.3955 |  | #N/A
| 1 | 3 | 2.03 | 2.74 | 2.56 | 2.27 | < 5 | 2.6 | 2.5 | 2.59 | 2.5155 | 2.7 | 2.62 | 2.8009 | 2 | 2.6319 | 2.51 | 2.59 | 2.49 | 2.4552 |  | 16
| 1 | 4 | 2.09 | 2.62 | 2.64 | 2.4 | < 5 | 2.4 | 2.5 | 2.57 | 2.5596 | 2.7 | 2.55 | 2.7124 | 2 | 2.6361 | 2.54 | 2.55 | 2.53 | 2.5064 |  | 2.6025
| 1 | 5 | 2.09 | 2.68 | 2.56 | 2.44 | < 5 | 2.7 | 2.4 | 2.52 | 2.6097 | 2.8 | 2.61 | 2.7502 | 3 | 2.7174 | 2.57 | 2.48 | 2.57 | 2.4303 |  | 36
| 1 | 6 | 2.09 | 2.68 | 2.61 | 2.27 | < 5 | 2.6 | 2.5 | 2.6 | 2.6192 | 2.9 | 2.51 | 2.8598 | 3 | 2.6358 | 2.61 | 2.55 | 2.52 | 2.4428
| Mean |  | 2.0683 | 2.6717 | 2.5633 | 2.3433 |  | 2.6 | 2.4833 | 2.5833 | 2.5813 | 2.7333 | 2.5867 | 2.7547 | 2.3333 | 2.6406 | 2.5683 | 2.5417 | 2.55 | 2.4385
| Median |  | 2.075 | 2.675 | 2.56 | 2.34 |  | 2.6 | 2.5 | 2.595 | 2.5846 | 2.7 | 2.58 | 2.7756 | 2 | 2.6339 | 2.575 | 2.55 | 2.525 | 2.4365
| Std Dev. |  | 0.025626 | 0.041191 | 0.055377 | 0.078145 |  | 0.10954 | 0.040825 | 0.034448 | 0.056606 | 0.10328 | 0.070616 | 0.19554 | 0.5164 | 0.041128 | 0.037639 | 0.040208 | 0.064187 | 0.040572
| Rel.Std.Dev. |  | 1.239% | 1.542% | 2.16% | 3.335% |  | 4.213% | 1.644% | 1.333% | 2.193% | 3.779% | 2.73% | 7.098% | 22.13% | 1.558% | 1.465% | 1.582% | 2.517% | 1.664%
| PDM3 |  | -20.53% | 2.657% | -1.505% | -9.959% |  | -0.09644% | -4.579% | -0.7368% | -0.8161% | 5.027% | -0.6088% | 5.848% | -10.34% | 1.465% | -1.313% | -2.338% | -2.018% | -6.301%
| Z-Score (Absolute) |  | 4.77 | 1.01 | 0.02 | 2.13 | 0.63 | 0.33 | 0.79 | 0.17 | 0.15 | 1.61 | 0.2 | 1.81 | 2.23 | 0.72 | 0.02 | 0.23 | 0.15 | 1.22
| Analytical results for Bi in OREAS 780 (Certified Value 116 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 118.66 | 90.6761 | 129.5 | 117.5 | 95.59 | 116 | 114 | 113 | 116.53 | 114.9541 | 123 | 115 | 106.5869 | 112.1 | 113.6158 | 114.5 | 108 | 117 | 85.5424 | 1
| 1 | 2 | 123.59 | 90.9973 | 124.5 | 118.5 | 97.17 | 123 | 112 | 109 | 118.38 | 117.7138 | 119 | 112 | 108.9769 | 111.5 | 114.3116 | 110 | 112 | 127.5 | 85.6491 | 1
| 1 | 3 | 127.14 | 90.7794 | 129.5 | 118.5 | 95.06 | 117 | 112 | 110 | 117.91 | 112.592 | 116 | 115 | 108.4169 | 115.3 | 113.8278 | 110 | 113.5 | 116 | 86.5151 | 16
| 1 | 4 | 124.15 | 90.6795 | 123.5 | 122 | 97.57 | 119 | 117 | 111 | 121.78 | 113.5089 | 115 | 117 | 109.4269 | 113.1 | 116.5723 | 111 | 111 | 120 | 87.918 | 116.0927
| 1 | 5 | 130.77 | 90.7306 | 125 | 118.5 | 95.53 | 120 | 113 | 108 | 117.65 | 110.5581 | 114 | 115 | 111.0469 | 111 | 115.3801 | 110 | 109 | 118 | 88.5325 | 37
| 1 | 6 | 126.64 | 89.905 | 126 | 122 | 98.24 | 113 | 113 | 110 | 118.3 | 119.0616 | 118 | 115 | 112.7469 | 116.4 | 113.7042 | 113.5 | 110 | 121 | 87.1808
| Mean |  | 125.1583 | 90.628 | 126.3333 | 119.5 | 96.5267 | 118 | 113.5 | 110.1667 | 118.425 | 114.7314 | 117.5 | 114.8333 | 109.5336 | 113.2333 | 114.5686 | 111.5 | 110.5833 | 119.9167 | 86.8897
| Median |  | 125.395 | 90.7051 | 125.5 | 118.5 | 96.38 | 118 | 113 | 110 | 118.105 | 114.2315 | 117 | 115 | 109.2019 | 112.6 | 114.0697 | 110.5 | 110.5 | 119 | 86.8479
| Std Dev. |  | 4.0821 | 0.37349 | 2.582 | 1.9748 | 1.3007 | 3.4641 | 1.8708 | 1.7224 | 1.7737 | 3.1995 | 3.2711 | 1.6021 | 2.1385 | 2.1722 | 1.1784 | 2 | 2.0104 | 4.1523 | 1.2112
| Rel.Std.Dev. |  | 3.262% | 0.4121% | 2.044% | 1.653% | 1.348% | 2.936% | 1.648% | 1.563% | 1.498% | 2.789% | 2.784% | 1.395% | 1.952% | 1.918% | 1.029% | 1.794% | 1.818% | 3.463% | 1.394%
| PDM3 |  | 7.809% | -21.93% | 8.821% | 2.935% | -16.85% | 1.643% | -2.233% | -5.105% | 2.009% | -1.173% | 1.212% | -1.085% | -5.65% | -2.463% | -1.313% | -3.956% | -4.746% | 3.294% | -25.15%
| Z-Score (Absolute) |  | 1.79 | 4.05 | 1.99 | 0.83 | 3.05 | 0.58 | 0.18 | 0.74 | 0.65 | 0.03 | 0.5 | 0.04 | 0.85 | 0.23 | 0 | 0.52 | 0.67 | 0.9 | 4.68
| Analytical results for Ca in OREAS 780 (Certified Value 5.59 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 5.5922 | 5.7924 | 5.51 | 5.37 | 5.15 | 5.34 | 5.51 | 5.53 | 5.6267 | 5.5621 | 5.65 | 5.42 | 5.8549 | 5.73 | 5.58 | 5.74 | 5.51 | 6.2774 |  | 1
| 1 | 2 | 5.5032 | 5.8313 | 5.54 | 5.37 | 5.14 | 5.53 | 5.52 | 5.48 | 5.5529 | 5.5688 | 5.49 | 5.362 | 5.9239 | 5.69 | 5.46 | 5.76 | 6.1 | 6.3606 |  | 22
| 1 | 3 | 5.2025 | 5.8618 | 5.51 | 5.39 | 5.14 | 5.52 | 5.48 | 5.54 | 5.6028 | 5.4988 | 5.59 | 5.389 | 5.8935 | 5.74 | 5.48 | 5.69 | 5.59 | 6.0613 |  | 16
| 1 | 4 | 5.0313 | 5.8329 | 5.5 | 5.49 | 5.12 | 5.69 | 5.66 | 5.6 | 5.6023 | 5.5786 | 5.73 | 5.389 | 5.9729 | 5.68 | 5.49 | 5.64 | 5.75 | 6.1 |  | 5.5864
| 1 | 5 | 5.3386 | 5.7992 | 5.55 | 5.49 | 5.14 | 5.5 | 5.59 | 5.56 | 5.6513 | 5.4122 | 5.67 | 5.223 | 5.9539 | 5.66 | 5.56 | 5.56 | 5.59 | 6.0305 |  | 38
| 1 | 6 | 5.4422 | 5.862 | 5.58 | 5.62 | 5.12 | 5.3 | 5.59 | 5.54 | 5.6281 | 5.6753 | 5.79 | 5.173 | 5.87 | 5.65 | 5.6 | 5.65 | 5.69 | 6.1672
| Mean |  | 5.3517 | 5.8299 | 5.5317 | 5.455 | 5.135 | 5.48 | 5.5583 | 5.5417 | 5.6107 | 5.5493 | 5.6533 | 5.326 | 5.9115 | 5.6917 | 5.5283 | 5.6733 | 5.705 | 6.1662
| Median |  | 5.3904 | 5.8321 | 5.525 | 5.44 | 5.14 | 5.51 | 5.555 | 5.54 | 5.6147 | 5.5654 | 5.66 | 5.3755 | 5.9087 | 5.685 | 5.525 | 5.67 | 5.64 | 6.1336
| Std Dev. |  | 0.20692 | 0.029691 | 0.030605 | 0.098336 | 0.012247 | 0.14184 | 0.066758 | 0.0392 | 0.033707 | 0.087872 | 0.10539 | 0.10207 | 0.046843 | 0.03656 | 0.058793 | 0.073121 | 0.21107 | 0.12958
| Rel.Std.Dev. |  | 3.866% | 0.5093% | 0.5533% | 1.803% | 0.2385% | 2.588% | 1.201% | 0.7074% | 0.6008% | 1.583% | 1.864% | 1.916% | 0.7924% | 0.6424% | 1.063% | 1.289% | 3.7% | 2.102%
| PDM3 |  | -4.202% | 4.359% | -0.9798% | -2.352% | -8.08% | -1.905% | -0.5025% | -0.8008% | 0.4346% | -0.6642% | 1.198% | -4.661% | 5.819% | 1.884% | -1.04% | 1.556% | 2.123% | 10.38%
| Z-Score (Absolute) |  | 1.24 | 1.7 | 0.14 | 0.61 | 2.58 | 0.45 | 0.03 | 0.07 | 0.35 | 0.03 | 0.61 | 1.4 | 2.2 | 0.85 | 0.16 | 0.74 | 0.93 | 3.77
| Analytical results for Cd in OREAS 780 (Indicative Value 0.042 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.04 | < 0.02 | < 0.02 | < 0.2 | 0.08 | < 0.3 | < 0.05 | 0.04 | < 0.5 | < 0.5 | 0.1 | 0.008209 | < 0.1 | 0.03775 | 0.02 | < 0.02 | < 0.02 | 1.6212 |  | 1
| 1 | 2 | 0.05 | < 0.02 | < 0.02 | < 0.2 | 0.08 | < 0.3 | < 0.05 | 0.04 | < 0.5 | < 0.5 | 0.09 | 0.024435 | < 0.1 | 0.03287 | < 0.02 | < 0.02 | < 0.02 | 1.4477 |  | 2
| 1 | 3 | 0.05 | < 0.02 | < 0.02 | < 0.2 | 0.09 | < 0.3 | < 0.05 | 0.03 | < 0.5 | < 0.5 | 0.09 | 0.040976 | < 0.1 | 0.04374 | < 0.02 | < 0.02 | < 0.02 | 1.2933 |  | 16
| 1 | 4 | 0.05 | < 0.02 | 0.02 | < 0.2 | 0.09 | < 0.3 | < 0.05 | 0.03 | < 0.5 | < 0.5 | 0.07 | 0.24462 | < 0.1 | 0.03388 | 0.02 | < 0.02 | < 0.02 | 1.1556 |  | 0.041745
| 1 | 5 | 0.05 | < 0.02 | < 0.02 | < 0.2 | 0.09 | < 0.3 | < 0.05 | < 0.02 | < 0.5 | < 0.5 | 0.09 | 0.24477 | < 0.1 | 0.03903 | < 0.02 | < 0.02 | < 0.02 | 1.5969 |  | 23
| 1 | 6 | 0.06 | 0.02 | < 0.02 | < 0.2 | 0.08 | < 0.3 | < 0.05 | 0.03 | < 0.5 | < 0.5 | 0.1 | 0.030108 | < 0.1 | 0.03294 | < 0.02 | 0.02 | < 0.02 | 1.2473
| Mean |  | 0.05 | 0.02 | 0.02 |  | 0.085 |  |  | 0.034 |  |  | 0.09 | 0.098853 |  | 0.036702 | 0.02 | 0.02 |  | 1.3937
| Median |  | 0.05 | 0.02 | 0.02 |  | 0.085 |  |  | 0.03 |  |  | 0.09 | 0.035542 |  | 0.035815 | 0.02 | 0.02 |  | 1.3705
| Std Dev. |  | 0.0063246 |  |  |  | 0.0054772 |  |  | 0.0054772 |  |  | 0.010954 | 0.11346 |  | 0.0043093 | 0 |  |  | 0.19192
| Rel.Std.Dev. |  | 12.65% |  |  |  | 6.444% |  |  | 16.11% |  |  | 12.17% | 114.8% |  | 11.74% | 0% |  |  | 13.77%
| PDM3 |  | 19.78% | -52.09% | -52.09% |  | 103.6% |  |  | -18.55% |  |  | 115.6% | 136.8% |  | -12.08% | -52.09% | -52.09% |  | 3239%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.67 | 0.88 | 0.62 | 1.76 | 0.44 | 0.35 | 3.52 | 3.52 | 0.7 | 0.86 | 0 | 0.23 | 0.64 | 0.67 | 0.7 | 23.64
| Analytical results for Ce in OREAS 780 (Certified Value 88 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 69.3 | 93.7292 | 81.6 | 84.8 | 86 | 89.2 | 89.35 | 82.8 | 93.52 | 95.1463 | 92 | 83.6489 | 89.7 | 86 | 79.2 | 78.7155 |  |  |  | 1
| 1 | 2 | 71.3 | 93.7741 | 70.2 | 84.4 | 87 | 89.3 | 89.46 | 90.1 | 91.47 | 90.1282 | 87 | 83.3885 | 89.2 | 89.8 | 86.7 | 77.9882 |  |  |  | 23
| 1 | 3 | 68 | 93.5136 | 77.3 | 86.9 | 88 | 87.9 | 90.35 | 84.9 | 93.29 | 91.3911 | 91 | 82.5552 | 90.7 | 88.4 | 81.1 | 81.9021 |  |  |  | 16
| 1 | 4 | 67.3 | 93.6959 | 71.1 | 90 | 89 | 85.4 | 91.25 | 84.7 | 82.69 | 94.5061 | 90 | 82.4479 | 89.4 | 89.5 | 83.6 | 91.9166 |  |  |  | 87.902
| 1 | 5 | 72.6 | 93.7842 | 75.4 | 85.9 | 86 | 90.2 | 90.58 | 80.9 | 89.17 | 95.1529 | 91 | 83.6334 | 90.6 | 87.7 | 82.7 | 82.8139 |  |  |  | 39
| 1 | 6 | 73.5 | 93.4105 | 75.9 | 84.9 | 89 | 91.2 | 91.32 | 84.7 | 87.27 | 94.4503 | 86 | 73.6315 | 92.1 | 88.2 | 77.7 | 86.7232
| Mean |  | 70.3333 | 93.6512 | 75.25 | 86.15 | 87.5 | 88.8667 | 90.385 | 84.6833 | 89.5683 | 93.4625 | 89.5 | 81.5509 | 90.2833 | 88.2667 | 81.8333 | 83.3433
| Median |  | 70.3 | 93.7126 | 75.65 | 85.4 | 87.5 | 89.25 | 90.465 | 84.7 | 90.32 | 94.4782 | 90.5 | 82.9718 | 90.15 | 88.3 | 81.9 | 82.358
| Std Dev. |  | 2.5224 | 0.15346 | 4.188 | 2.0926 | 1.3784 | 2.0235 | 0.84718 | 3.074 | 4.1433 | 2.1525 | 2.429 | 3.9152 | 1.0834 | 1.3677 | 3.2284 | 5.2417
| Rel.Std.Dev. |  | 3.586% | 0.1639% | 5.565% | 2.429% | 1.575% | 2.277% | 0.9373% | 3.63% | 4.626% | 2.303% | 2.714% | 4.801% | 1.2% | 1.55% | 3.945% | 6.289%
| PDM3 |  | -19.99% | 6.54% | -14.39% | -1.993% | -0.4574% | 1.097% | 2.825% | -3.662% | 1.896% | 6.326% | 1.818% | -7.225% | 2.709% | 0.4148% | -6.904% | -5.186%
| Z-Score (Absolute) |  | 4.15 | 1.36 | 2.99 | 0.41 | 0.09 | 0.23 | 0.59 | 0.76 | 0.4 | 1.32 | 0.38 | 1.5 | 0.57 | 0.09 | 1.43 | 1.07
| Analytical results for Co in OREAS 780 (Certified Value 1.82 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2 | 1.6226 | 1.9 | 1.9 | 6.01 | 1.54 | 1 | 2 | 1.8 | 1.7635 | 2 | 2 | 4.7522 | 1.6 | 1.8191 | 1.8 | 1.8 | 1.8 | 4.0664 | 1
| 1 | 2 | 2.1 | 1.6304 | 1.9 | 1.9 | 6.02 | 1.78 | 1 | 2 | 1.8 | 1.8296 | 2 | 2.1 | 18.5146 | 1.6 | 1.864 | 1.7 | 1.9 | 2 | 3.1548 | #N/A
| 1 | 3 | 1.8 | 1.6921 | 1.8 | 1.9 | 5.91 | 1.64 | 2 | 2 | 1.9 | 1.6693 | 2 | 2 | 3.345 | 1.7 | 1.8556 | 1.8 | 1.9 | 1.9 | 3.314 | 16
| 1 | 4 | 1.7 | 1.7287 | 1.8 | 1.9 | 6.03 | 1.78 | 2 | 2 | 1.8 | 1.6841 | 2 | 2 | 3.2414 | 1.7 | 1.8125 | 1.8 | 1.9 | 1.9 | 3.0429 | 1.8211
| 1 | 5 | 3.6 | 1.6461 | 1.9 | 1.9 | 6.07 | 1.79 | 2 | 2 | 1.9 | 1.7309 | 2 | 2 | 4.0975 | 1.6 | 1.8473 | 1.7 | 1.9 | 1.9 | 3.2695 | 40
| 1 | 6 | 1.7 | 1.6366 | 2 | 1.9 | 5.94 | 1.82 | 1 | 2 | 1.8 | 1.737 | 2 | 2 | 3.5443 | 1.6 | 1.8457 | 1.8 | 1.9 | 1.9 | 3.0285
| Mean |  | 2.15 | 1.6594 | 1.8833 | 1.9 | 5.9967 | 1.725 | 1.5 | 2 | 1.8333 | 1.7357 | 2 | 2.0167 | 6.2492 | 1.6333 | 1.8407 | 1.7667 | 1.8833 | 1.9 | 3.3127
| Median |  | 1.9 | 1.6413 | 1.9 | 1.9 | 6.015 | 1.78 | 1.5 | 2 | 1.8 | 1.734 | 2 | 2 | 3.8209 | 1.6 | 1.8465 | 1.8 | 1.9 | 1.9 | 3.2121
| Std Dev. |  | 0.7287 | 0.041887 | 0.075277 | 2.4324e-16 | 0.059889 | 0.11023 | 0.54772 | 0 | 0.05164 | 0.057758 | 0 | 0.040825 | 6.035 | 0.05164 | 0.020476 | 0.05164 | 0.040825 | 0.063246 | 0.38689
| Rel.Std.Dev. |  | 33.89% | 2.524% | 3.997% | 1.28e-14% | 0.9987% | 6.39% | 36.51% | 0% | 2.817% | 3.328% | 0% | 2.024% | 96.57% | 3.162% | 1.112% | 2.923% | 2.168% | 3.329% | 11.68%
| PDM3 |  | 18.06% | -8.879% | 3.416% | 4.332% | 229.3% | -5.278% | -17.63% | 9.823% | 0.6708% | -4.689% | 9.823% | 10.74% | 243.1% | -10.31% | 1.076% | -2.99% | 3.416% | 4.332% | 81.9%
| Z-Score (Absolute) |  | 1.28 | 1.07 | 0 | 0.08 | 19.75 | 0.76 | NA | NA | 0.24 | 0.71 | NA | 0.64 | 20.96 | 1.2 | 0.2 | 0.56 | 0 | 0.08 | 6.86
Results from laboratories 10, 12 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 780 (Certified Value 88 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 84.5 | 65.437 | 92 | 84 | 100.9 | 79 | 66 | 73 | 96 | 89.651 | 56 | 117.4185 | 89 | 106.7333 | 88 | 89 | 94 | 94.1802 |  | 1
| 1 | 2 | 84.6 | 66.0315 | 92 | 83 | 104.72 | 76 | 73 | 69 | 102 | 91.816 | 48 | 77.9825 | 89 | 109.4031 | 90 | 81 | 97 | 102.7439 |  | #N/A
| 1 | 3 | 84.2 | 65.5385 | 97 | 84 | 101.9 | 78 | 67 | 78 | 102 | 91.5556 | 57 | 75.6543 | 89 | 111.4204 | 88 | 83 | 89 | 93.0732 |  | 16
| 1 | 4 | 85.6 | 64.92 | 91 | 87 | 106.6 | 78 | 66 | 71 | 100 | 90.1662 | 60 | 91.8794 | 90 | 109.4353 | 94 | 82 | 91 | 151.729 |  | 88.1297
| 1 | 5 | 86.6 | 67.441 | 95 | 84 | 100.4 | 83 | 66 | 65 | 101 | 93.3079 | 60 | 130.4576 | 89 | 111.7911 | 89 | 86 | 94 | 103.4833 |  | 41
| 1 | 6 | 85.9 | 65.8945 | 96 | 84 | 103.3 | 81 | 68 | 73 | 99 | 89.5369 | 55 | 79.3741 | 89 | 109.2757 | 89 | 82 | 96 | 118.0885
| Mean |  | 85.2333 | 65.8771 | 93.8333 | 84.3333 | 102.97 | 79.1667 | 67.6667 | 71.5 | 100 | 91.0056 | 56 | 95.4611 | 89.1667 | 109.6765 | 89.6667 | 83.8333 | 93.5 | 110.5497
| Median |  | 85.1 | 65.7165 | 93.5 | 84 | 102.6 | 78.5 | 66.5 | 72 | 100.5 | 90.8609 | 56.5 | 85.6268 | 89 | 109.4192 | 89 | 82.5 | 94 | 103.1136
| Std Dev. |  | 0.94375 | 0.85957 | 2.4833 | 1.3663 | 2.3826 | 2.4833 | 2.7325 | 4.3704 | 2.2804 | 1.4803 | 4.4272 | 23.1358 | 0.40825 | 1.8149 | 2.2509 | 3.0605 | 3.0166 | 22.0763
| Rel.Std.Dev. |  | 1.107% | 1.305% | 2.646% | 1.62% | 2.314% | 3.137% | 4.038% | 6.112% | 2.28% | 1.627% | 7.906% | 24.24% | 0.4578% | 1.655% | 2.51% | 3.651% | 3.226% | 19.97%
| PDM3 |  | -3.286% | -25.25% | 6.472% | -4.308% | 16.84% | -10.17% | -23.22% | -18.87% | 13.47% | 3.263% | -36.46% | 8.319% | 1.177% | 24.45% | 1.744% | -4.875% | 6.094% | 25.44%
| Z-Score (Absolute) |  | 0.35 | 1.95 | 0.37 | 0.42 | 1.12 | 0.85 | 1.8 | 1.48 | 0.88 | 0.13 | 2.77 | 0.5 | 0.02 | 1.68 | 0.02 | 0.46 | 0.34 | 1.75
| Analytical results for Cs in OREAS 780 (Certified Value 9.71 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 9.9 | 8.8705 | 9.59 | 9.58 | 9.26 | 9.77 | 9.8 | 9.58 | 9.4 | 8.89 | 9.8598 | 9.8 | 9.869 | 10.35 | 9.68 | 10.15 | 8.744 |  |  | 1
| 1 | 2 | 10.3 | 8.7018 | 9.09 | 9.31 | 9.69 | 9.72 | 9.6 | 9.82 | 9 | 8.37 | 9.3135 | 10.2 | 9.8387 | 10.2 | 10.1 | 11.1 | 8.7088 |  |  | 5
| 1 | 3 | 10.6 | 8.8071 | 9.41 | 9.59 | 9.4 | 9.72 | 9.6 | 9.78 | 9.5 | 8.4 | 9.4544 | 9.8 | 9.6952 | 10.35 | 10.15 | 10.15 | 9.3062 |  |  | 16
| 1 | 4 | 10.1 | 8.5921 | 9.08 | 10.1 | 8.98 | 10 | 9.8 | 9.8 | 9.4 | 7.36 | 9.7137 | 10.2 | 9.7115 | 10.15 | 10.35 | 10.6 | 10.2485 |  |  | 9.7143
| 1 | 5 | 10.5 | 8.5353 | 9.29 | 9.27 | 8.97 | 10.1 | 9.6 | 9.78 | 9.4 | 7.63 | 9.8211 | 9.7 | 9.929 | 10.25 | 9.96 | 10.7 | 9.5615 |  |  | 42
| 1 | 6 | 10.5 | 8.8362 | 9.48 | 9.44 | 9.32 | 10.1 | 9.8 | 9.8 | 9.5 | 7.46 | 9.7564 | 9.7 | 9.7163 | 10.45 | 10 | 9.82 | 9.9494
| Mean |  | 10.3167 | 8.7238 | 9.3233 | 9.5483 | 9.27 | 9.9017 | 9.7 | 9.76 | 9.3667 | 8.0183 | 9.6532 | 9.9 | 9.7933 | 10.2917 | 10.04 | 10.42 | 9.4197
| Median |  | 10.4 | 8.7544 | 9.35 | 9.51 | 9.29 | 9.885 | 9.7 | 9.79 | 9.4 | 8 | 9.735 | 9.8 | 9.7775 | 10.3 | 10.05 | 10.375 | 9.4338
| Std Dev. |  | 0.27142 | 0.13741 | 0.20877 | 0.30103 | 0.27203 | 0.1853 | 0.10954 | 0.089443 | 0.18619 | 0.6205 | 0.21911 | 0.23664 | 0.098433 | 0.11143 | 0.22334 | 0.46454 | 0.62649
| Rel.Std.Dev. |  | 2.631% | 1.575% | 2.239% | 3.153% | 2.935% | 1.871% | 1.129% | 0.9164% | 1.988% | 7.738% | 2.27% | 2.39% | 1.005% | 1.083% | 2.224% | 4.458% | 6.651%
| PDM3 |  | 6.201% | -10.2% | -4.024% | -1.708% | -4.573% | 1.929% | -0.1469% | 0.4707% | -3.578% | -17.46% | -0.6292% | 1.912% | 0.8135% | 5.944% | 3.353% | 7.265% | -3.032%
| Z-Score (Absolute) |  | 1.25 | 1.97 | 0.76 | 0.31 | 0.87 | 0.41 | 0 | 0.12 | 0.67 | 3.4 | 0.09 | 0.4 | 0.19 | 1.2 | 0.69 | 1.46 | 0.57
| Analytical results for Cu in OREAS 780 (Certified Value 52 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 48.5 | 59.38 | 54.6 | 53.3 | 74.11 | 52 | 50 | 54 | 51 | 52.864 | 50 | 49 | 47.9917 | 51.8 | 50.5397 | 54.9 | 50.4 | 49 | 57.3711 | 1
| 1 | 2 | 48.7 | 56.98 | 52.5 | 51.2 | 77.79 | 52 | 48 | 53 | 51 | 53.7762 | 53 | 50 | 47.5089 | 51.3 | 51.1902 | 55.6 | 52.9 | 55.1 | 57.7463 | #N/A
| 1 | 3 | 48.9 | 57.31 | 55.1 | 50.5 | 76.17 | 55 | 51 | 52 | 51 | 52.4582 | 52 | 49 | 47.153 | 52.6 | 51.3687 | 56.4 | 51.4 | 50.7 | 56.0913 | 16
| 1 | 4 | 43.9 | 57.27 | 54.2 | 56.1 | 76.81 | 53 | 53 | 55 | 51 | 52.7377 | 51 | 49 | 51.8283 | 52.6 | 50.2657 | 54.6 | 52.8 | 53 | 56.9387 | 52.314
| 1 | 5 | 46.2 | 58.42 | 54.4 | 53.5 | 75.56 | 51 | 49 | 55 | 52 | 52.0247 | 50 | 48 | 49.5941 | 52.7 | 51.4231 | 55 | 50.9 | 50.6 | 56.5663 | 43
| 1 | 6 | 43.4 | 57.57 | 54.7 | 55.5 | 74.47 | 53 | 51 | 52 | 51 | 54.5316 | 51 | 49 | 50.3349 | 51.9 | 51.5879 | 56.1 | 52.1 | 54.5 | 56.7512
| Mean |  | 46.6 | 57.8217 | 54.25 | 53.35 | 75.8183 | 52.6667 | 50.3333 | 53.5 | 51.1667 | 53.0654 | 51.1667 | 49 | 49.0685 | 52.15 | 51.0625 | 55.4333 | 51.75 | 52.15 | 56.9108
| Median |  | 47.35 | 57.44 | 54.5 | 53.4 | 75.865 | 52.5 | 50.5 | 53.5 | 51 | 52.8009 | 51 | 49 | 48.7929 | 52.25 | 51.2794 | 55.3 | 51.75 | 51.85 | 56.8449
| Std Dev. |  | 2.49 | 0.90822 | 0.9094 | 2.2341 | 1.3998 | 1.3663 | 1.7512 | 1.3784 | 0.40825 | 0.92246 | 1.169 | 0.63246 | 1.8307 | 0.56833 | 0.53374 | 0.7174 | 1.0213 | 2.4238 | 0.58761
| Rel.Std.Dev. |  | 5.343% | 1.571% | 1.676% | 4.188% | 1.846% | 2.594% | 3.479% | 2.576% | 0.7979% | 1.738% | 2.285% | 1.291% | 3.731% | 1.09% | 1.045% | 1.294% | 1.973% | 4.648% | 1.033%
| PDM3 |  | -10.92% | 10.53% | 3.701% | 1.98% | 44.93% | 0.6741% | -3.786% | 2.267% | -2.193% | 1.436% | -2.193% | -6.335% | -6.204% | -0.3135% | -2.392% | 5.963% | -1.078% | -0.3135% | 8.787%
| Z-Score (Absolute) |  | 2.77 | 2.83 | 1.05 | 0.6 | 11.82 | 0.26 | 0.91 | 0.67 | 0.49 | 0.46 | 0.49 | 1.57 | 1.54 | 0 | 0.54 | 1.64 | 0.2 | 0 | 2.38
| Analytical results for Dy in OREAS 780 (Certified Value 3.4 ppm) | [20 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 2.73 | 3.7279 | 3.02 | 3.6 | 3.56 | 3.55 | 3.43 | 4.017 | 5.0964 | 3.151 | [9 empty] | 1
| 1 | 2 | 2.93 | 3.5433 | 2.95 | 3.5 | 3.54 | 3.55 | 3.35 | 3.7593 | 5.28 | 3.099 | [9 empty] | 6
| 1 | 3 | 2.87 | 3.6983 | 3.02 | 3.4 | 3.73 | 3.5 | 3.41 | 3.7373 | 5.0156 | 3.1187 | [9 empty] | 16
| 1 | 4 | 2.73 | 3.6959 | 2.94 | 3.5 | 3.72 | 3.6 | 3.31 | 3.8428 | 5.0366 | 3.569 | [9 empty] | 3.4002
| 1 | 5 | 3.01 | 3.6301 | 2.87 | 3.7 | 3.63 | 3.25 | 3.34 | 3.9236 | 5.1049 | 3.3212 | [9 empty] | 44
| 1 | 6 | 3 | 3.7492 | 2.94 | 3.4 | 3.76 | 3.7 | 3.29 | 3.6983 | 4.482 | 2.9969
| Mean |  | 2.8783 | 3.6741 | 2.9567 | 3.5167 | 3.6567 | 3.525 | 3.355 | 3.8297 | 5.0026 | 3.2093
| Median |  | 2.9 | 3.6971 | 2.945 | 3.5 | 3.675 | 3.55 | 3.345 | 3.8011 | 5.0665 | 3.1349
| Std Dev. |  | 0.12561 | 0.075701 | 0.056804 | 0.1169 | 0.093524 | 0.15083 | 0.055045 | 0.12234 | 0.2715 | 0.20533
| Rel.Std.Dev. |  | 4.364% | 2.06% | 1.921% | 3.324% | 2.558% | 4.279% | 1.641% | 3.195% | 5.427% | 6.398%
| PDM3 |  | -15.35% | 8.057% | -13.04% | 3.426% | 7.544% | 3.671% | -1.328% | 12.63% | 47.13% | -5.613%
| Z-Score (Absolute) |  | 1.83 | 0.44 | 1.6 | 0.01 | 0.39 | 0.01 | 0.47 | 0.88 | 4.21 | 0.88
| Analytical results for Er in OREAS 780 (Certified Value 1.63 ppm) | [20 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 1.35 | 1.7971 | 1.26 | 1.6 | 1.69 | 1.65 | 1.67 | 2.1859 | 2.9644 | 1.4057 | [9 empty] | 1
| 1 | 2 | 1.45 | 1.7903 | 1.28 | 1.5 | 1.68 | 1.65 | 1.62 | 1.8717 | 2.9705 | 1.4394 | [9 empty] | 26
| 1 | 3 | 1.43 | 1.8088 | 1.31 | 2 | 1.87 | 1.6 | 1.71 | 1.8497 | 2.9523 | 1.4259 | [9 empty] | 16
| 1 | 4 | 1.38 | 1.8953 | 1.26 | 1.6 | 1.71 | 1.6 | 1.68 | 1.9584 | 2.9778 | 1.6719 | [9 empty] | 1.6277
| 1 | 5 | 1.48 | 1.8201 | 1.2 | 1.6 | 1.71 | 1.6 | 1.69 | 2.0161 | 2.9051 | 1.4952 | [9 empty] | 45
| 1 | 6 | 1.49 | 1.8697 | 1.26 | 1.8 | 1.79 | 1.6 | 1.67 | 1.7964 | 2.5723 | 1.3594
| Mean |  | 1.43 | 1.8302 | 1.2617 | 1.6833 | 1.7417 | 1.6167 | 1.6733 | 1.9464 | 2.8904 | 1.4662
| Median |  | 1.44 | 1.8145 | 1.26 | 1.6 | 1.71 | 1.6 | 1.675 | 1.915 | 2.9584 | 1.4326
| Std Dev. |  | 0.055498 | 0.042539 | 0.036009 | 0.18348 | 0.073869 | 0.02582 | 0.030111 | 0.1413 | 0.15797 | 0.11006
| Rel.Std.Dev. |  | 3.881% | 2.324% | 2.854% | 10.9% | 4.241% | 1.597% | 1.799% | 7.26% | 5.465% | 7.506%
| PDM3 |  | -12.15% | 12.44% | -22.49% | 3.416% | 7% | -0.6793% | 2.802% | 19.58% | 77.57% | -9.92%
| Z-Score (Absolute) |  | 0.92 | 0.56 | 1.54 | 0.02 | 0.23 | 0.23 | 0.02 | 0.99 | 4.49 | 0.79
| Analytical results for Eu in OREAS 780 (Certified Value 1.2 ppm) | [20 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 1.12 | 2.5057 | 1.12 | 1.2 | 1.22 | 1.25 | 1.29 | 1.238 | 1.2219 | 1.5551 | [9 empty] | 1
| 1 | 2 | 1.19 | 2.4157 | 1.1 | 1.21 | 1.22 | 1.25 | 1.24 | 1.1396 | 1.2143 | 1.4557 | [9 empty] | 7
| 1 | 3 | 1.17 | 2.407 | 1.13 | 1.14 | 1.24 | 1.2 | 1.27 | 1.1842 | 1.195 | 1.5678 | [9 empty] | 16
| 1 | 4 | 1.11 | 2.3894 | 1.15 | 1.18 | 1.25 | 1.2 | 1.22 | 1.2586 | 1.1832 | 1.6522 | [9 empty] | 1.2006
| 1 | 5 | 1.21 | 2.4052 | 1.11 | 1.15 | 1.27 | 1.2 | 1.24 | 1.2424 | 1.2015 | 1.6141 | [9 empty] | 46
| 1 | 6 | 1.22 | 2.4264 | 1.15 | 1.25 | 1.23 | 1.25 | 1.22 | 1.1774 | 1.0071 | 1.479
| Mean |  | 1.17 | 2.4249 | 1.1267 | 1.1883 | 1.2383 | 1.225 | 1.2467 | 1.2067 | 1.1705 | 1.554
| Median |  | 1.18 | 2.4113 | 1.125 | 1.19 | 1.235 | 1.225 | 1.24 | 1.2111 | 1.1982 | 1.5615
| Std Dev. |  | 0.046043 | 0.041441 | 0.020656 | 0.040702 | 0.019408 | 0.027386 | 0.028048 | 0.046492 | 0.081234 | 0.075809
| Rel.Std.Dev. |  | 3.935% | 1.709% | 1.833% | 3.425% | 1.567% | 2.236% | 2.25% | 3.853% | 6.94% | 4.878%
| PDM3 |  | -2.55% | 102% | -6.159% | -1.023% | 3.142% | 2.031% | 3.836% | 0.5088% | -2.51% | 29.43%
| Z-Score (Absolute) |  | 0.81 | 21.4 | 1.58 | 0.49 | 0.4 | 0.16 | 0.55 | 0.16 | 0.8 | 5.99
| Analytical results for Fe in OREAS 780 (Certified Value 2.42 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.2232 | 2.561 | 2.42 | 2.33 | 2.51 | 2.34 | 2.44 | 2.35 | 2.43 | 2.3794 | 2.45 | 2.36 | 2.6534 | 2.28 | 2.39 | 2.46 | 2.45 | 2.6 |  | 1
| 1 | 2 | 2.2088 | 2.5468 | 2.42 | 2.33 | 2.52 | 2.33 | 2.48 | 2.41 | 2.43 | 2.4357 | 2.48 | 2.34 | 2.6558 | 2.28 | 2.37 | 2.43 | 2.58 | 2.6 |  | 28
| 1 | 3 | 2.1936 | 2.5507 | 2.45 | 2.39 | 2.51 | 2.31 | 2.44 | 2.55 | 2.45 | 2.3806 | 2.4 | 2.36 | 2.6386 | 2.31 | 2.36 | 2.48 | 2.43 | 2.5 |  | 16
| 1 | 4 | 2.2294 | 2.5668 | 2.39 | 2.45 | 2.5 | 2.34 | 2.53 | 2.38 | 2.43 | 2.3768 | 2.42 | 2.24 | 2.6531 | 2.31 | 2.36 | 2.48 | 2.46 | 2.6 |  | 2.4203
| 1 | 5 | 2.3559 | 2.5707 | 2.41 | 2.38 | 2.49 | 2.32 | 2.5 | 2.43 | 2.45 | 2.3246 | 2.35 | 2.28 | 2.6421 | 2.28 | 2.35 | 2.4 | 2.48 | 2.5 |  | 47
| 1 | 6 | 2.2525 | 2.5393 | 2.45 | 2.43 | 2.43 | 2.31 | 2.48 | 2.36 | 2.46 | 2.4385 | 2.43 | 2.24 | 2.6285 | 2.33 | 2.39 | 2.44 | 2.46 | 2.5
| Mean |  | 2.2439 | 2.5559 | 2.4233 | 2.385 | 2.4933 | 2.325 | 2.4783 | 2.4133 | 2.4417 | 2.3892 | 2.4217 | 2.3033 | 2.6452 | 2.2983 | 2.37 | 2.4483 | 2.4767 | 2.55
| Median |  | 2.2263 | 2.5559 | 2.42 | 2.385 | 2.505 | 2.325 | 2.48 | 2.395 | 2.44 | 2.38 | 2.425 | 2.31 | 2.6476 | 2.295 | 2.365 | 2.45 | 2.46 | 2.55
| Std Dev. |  | 0.058342 | 0.012232 | 0.023381 | 0.049699 | 0.03266 | 0.013784 | 0.034881 | 0.073394 | 0.013292 | 0.042629 | 0.04446 | 0.057155 | 0.01072 | 0.02137 | 0.016733 | 0.031252 | 0.053166 | 0.054772
| Rel.Std.Dev. |  | 2.6% | 0.4786% | 0.9648% | 2.084% | 1.31% | 0.5929% | 1.407% | 3.041% | 0.5444% | 1.784% | 1.836% | 2.481% | 0.4053% | 0.9298% | 0.706% | 1.276% | 2.147% | 2.148%
| PDM3 |  | -7.29% | 5.6% | 0.1236% | -1.46% | 3.016% | -3.939% | 2.396% | -0.2896% | 0.8811% | -1.285% | 0.05475% | -4.834% | 9.292% | -5.041% | -2.08% | 1.157% | 2.327% | 5.357%
| Z-Score (Absolute) |  | 2.19 | 1.64 | 0.01 | 0.46 | 0.87 | 1.2 | 0.68 | 0.11 | 0.23 | 0.41 | 0.01 | 1.46 | 2.73 | 1.52 | 0.64 | 0.32 | 0.66 | 1.56
| Analytical results for Ga in OREAS 780 (Certified Value 17.9 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 16.29 | 17.6808 | 19.25 | 18.1 | 17.11 | 17.7 | 18.4 | 17.53 | 17.6 | 18.87 | 20.1772 | 19 | 18.0242 | 17.15 | 16.05 | 16.55 | 15.3994 |  |  | 1
| 1 | 2 | 16.28 | 17.3339 | 18.8 | 18.05 | 17.38 | 18 | 18.1 | 17.78 | 16.8 | 18.59 | 19.1394 | 19 | 17.7584 | 16.7 | 16.8 | 18.35 | 12.9903 |  |  | 8
| 1 | 3 | 16 | 17.5774 | 19.4 | 18.4 | 17.45 | 17.8 | 18.2 | 18.04 | 17.2 | 18.83 | 19.2767 | 18 | 18.139 | 16.75 | 16.6 | 17.3 | 14.8165 |  |  | 16
| 1 | 4 | 16.5 | 17.423 | 18.65 | 20.4 | 17.18 | 17.6 | 18.6 | 18.48 | 17.2 | 18.26 | 20.0987 | 18 | 17.9891 | 17.95 | 16.9 | 17.6 | 14.1311 |  |  | 17.8532
| 1 | 5 | 16.88 | 17.4417 | 18.95 | 19.15 | 17.24 | 18 | 18.4 | 18.05 | 17.2 | 18.45 | 20.2842 | 19 | 18.3076 | 16.35 | 16.45 | 17.5 | 15.4396 |  |  | 48
| 1 | 6 | 16.45 | 17.7977 | 19.25 | 19.35 | 16.59 | 17.8 | 18.2 | 18.28 | 17.4 | 17.94 | 19.8707 | 19 | 18.0321 | 16.9 | 16.5 | 16.95 | 14.1646
| Mean |  | 16.4 | 17.5424 | 19.05 | 18.9083 | 17.1583 | 17.8167 | 18.3167 | 18.0267 | 17.2333 | 18.49 | 19.8078 | 18.6667 | 18.0417 | 16.9667 | 16.55 | 17.375 | 14.4902
| Median |  | 16.37 | 17.5095 | 19.1 | 18.775 | 17.21 | 17.8 | 18.3 | 18.045 | 17.2 | 18.52 | 19.9847 | 19 | 18.0281 | 16.825 | 16.55 | 17.4 | 14.4905
| Std Dev. |  | 0.29305 | 0.17517 | 0.29496 | 0.90852 | 0.30551 | 0.16021 | 0.18348 | 0.33998 | 0.26583 | 0.35412 | 0.48597 | 0.5164 | 0.18087 | 0.54833 | 0.3 | 0.6138 | 0.92942
| Rel.Std.Dev. |  | 1.787% | 0.9986% | 1.548% | 4.805% | 1.781% | 0.8992% | 1.002% | 1.886% | 1.543% | 1.915% | 2.453% | 2.766% | 1.003% | 3.232% | 1.813% | 3.533% | 6.414%
| PDM3 |  | -8.139% | -1.741% | 6.704% | 5.91% | -3.892% | -0.2044% | 2.596% | 0.9719% | -3.472% | 3.567% | 10.95% | 4.557% | 1.056% | -4.965% | -7.299% | -2.678% | -18.84%
| Z-Score (Absolute) |  | 1.28 | 0.14 | 1.37 | 1.23 | 0.52 | 0.14 | 0.64 | 0.35 | 0.45 | 0.81 | 2.13 | NA | 0.36 | 0.71 | 1.13 | 0.3 | 3.19
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 780 (Certified Value 4.88 ppm) | [20 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 4.4 | 10.0742 | 4.41 | 5 | 4.83 | 4.6 | 5.3 | 5.5053 | 5.3581 | 4.5788 | [9 empty] | 1
| 1 | 2 | 4.81 | 10.0655 | 4.69 | 4.9 | 4.91 | 4.2 | 5.2 | 4.9772 | 5.542 | 4.2409 | [9 empty] | 9
| 1 | 3 | 4.73 | 10.0642 | 4.93 | 4.9 | 4.87 | 4.6 | 5.4 | 5.1957 | 5.3561 | 4.592 | [9 empty] | 16
| 1 | 4 | 4.34 | 9.8811 | 4.57 | 4.7 | 4.87 | 4.6 | 5 | 5.3902 | 5.2325 | 4.4667 | [9 empty] | 4.8822
| 1 | 5 | 4.89 | 10.3733 | 4.56 | 4.9 | 4.95 | 4.4 | 5.1 | 5.5228 | 5.3313 | 4.6436 | [9 empty] | 49
| 1 | 6 | 4.85 | 9.8872 | 4.8 | 5.2 | 4.99 | 4.8 | 5.1 | 5.2998 | 4.7046 | 4.3999
| Mean |  | 4.67 | 10.0576 | 4.66 | 4.9333 | 4.9033 | 4.5333 | 5.1833 | 5.3152 | 5.2541 | 4.487
| Median |  | 4.77 | 10.0648 | 4.63 | 4.9 | 4.89 | 4.6 | 5.15 | 5.345 | 5.3437 | 4.5228
| Std Dev. |  | 0.23908 | 0.17902 | 0.18655 | 0.1633 | 0.058878 | 0.20656 | 0.1472 | 0.20669 | 0.28719 | 0.15003
| Rel.Std.Dev. |  | 5.12% | 1.78% | 4.003% | 3.31% | 1.201% | 4.556% | 2.84% | 3.889% | 5.466% | 3.344%
| PDM3 |  | -4.346% | 106% | -4.551% | 1.048% | 0.4334% | -7.145% | 6.169% | 8.869% | 7.618% | -8.095%
| Z-Score (Absolute) |  | 0.56 | 11.54 | 0.58 | 0.03 | 0.03 | 0.86 | 0.59 | 0.89 | 0.75 | 0.97
| Analytical results for Ge in OREAS 780 (Indicative Value 0.22 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 0.09 | 0.17 | 0.2 | 0.25 | 0.3 | 1.9 | 1.8 | < 0.2 | 0.39273 | 3.8 | 0.13 | 0.16 | 0.18 | [6 empty] | 1
| 1 | 2 | 0.1 | 0.13 | 0.21 | 0.25 | 0.3 | 1.9 | 2 | < 0.2 | 0.41998 | 4 | 0.15 | 0.19 | 0.26 | [6 empty] | 10
| 1 | 3 | 0.1 | 0.13 | 0.21 | 0.24 | 0.3 | 1.9 | 2 | < 0.2 | 0.38746 | 4 | 0.16 | 0.18 | 0.26 | [6 empty] | 16
| 1 | 4 | 0.11 | 0.11 | 0.19 | 0.23 | 0.4 | 1.9 | 1.9 | < 0.2 | 0.41522 | 3.9 | 0.14 | 0.2 | 0.27 | [6 empty] | 0.2195
| 1 | 5 | 0.17 | 0.14 | 0.19 | 0.24 | 0.3 | 1.9 | 1.9 | < 0.2 | 0.36725 | 4 | 0.17 | 0.18 | 0.22 | [6 empty] | 22
| 1 | 6 | 0.13 | 0.16 | 0.2 | 0.23 | 0.4 | 1.9 | 2 | < 0.2 | 0.36017 | 4 | 0.16 | 0.18 | 0.14
| Mean |  | 0.11667 | 0.14 | 0.2 | 0.24 | 0.33333 | 1.9 | 1.9333 |  | 0.39047 | 3.95 | 0.15167 | 0.18167 | 0.22167
| Median |  | 0.105 | 0.135 | 0.2 | 0.24 | 0.3 | 1.9 | 1.95 |  | 0.3901 | 4 | 0.155 | 0.18 | 0.24
| Std Dev. |  | 0.029439 | 0.021909 | 0.0089443 | 0.0089443 | 0.05164 | 2.4324e-16 | 0.08165 |  | 0.024311 | 0.083666 | 0.01472 | 0.013292 | 0.052313
| Rel.Std.Dev. |  | 25.23% | 15.65% | 4.472% | 3.727% | 15.49% | 1.28e-14% | 4.223% |  | 6.226% | 2.118% | 9.705% | 7.316% | 23.6%
| PDM3 |  | -46.85% | -36.22% | -8.882% | 9.341% | 51.86% | 765.6% | 780.8% |  | 77.89% | 1700% | -30.9% | -17.23% | 0.9887%
| Z-Score (Absolute) |  | 0.67 | 0.52 | 0.14 | 0.12 | 0.72 | 10.78 | 10.99 | 0.78 | 1.08 | 23.94 | 0.45 | 0.26 | 0
| Analytical results for Hf in OREAS 780 (Certified Value 3.7 ppm) | [20 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 3.04 | 4.1 | 3.7 | 3.31 | 4 | 3.8 | 3.67 | 3.8 | 4.02 | 4.0866 | 3.2 | 5.7118 | 3.8 | 3.8 | 3.7 |  |  |  |  | 1
| 1 | 2 | 3.26 | 4 | 3.7 | 3.27 | 4 | 3.9 | 3.8 | 3.8 | 3.93 | 3.7253 | 3.2 | 5.8829 | 3.8 | 3.7 | 3.9 |  |  |  |  | #N/A
| 1 | 3 | 3.24 | 4.2 | 3.7 | 3.18 | 3.9 | 3.9 | 4.27 | 3.6 | 4 | 3.9098 | 3.3 | 5.8464 | 3.8 | 3.7 | 3.4 |  |  |  |  | 16
| 1 | 4 | 3.23 | 4 | 3.7 | 3.24 | 4.1 | 3.7 | 3.78 | 3.6 | 3.88 | 3.9838 | 3.1 | 5.7364 | 3.7 | 3.7 | 3.5 |  |  |  |  | 3.6981
| 1 | 5 | 3.31 | 4.1 | 3.7 | 3.36 | 3.9 | 3.8 | 3.79 | 3.6 | 3.95 | 4.1348 | 3.1 | 5.8212 | 3.7 | 3.7 | 3.5 |  |  |  |  | 50
| 1 | 6 | 3.32 | 4.1 | 3.7 | 3.27 | 4 | 3.7 | 3.85 | 3.6 | 3.93 | 3.8972 | 3.2 | 5.7765 | 3.8 | 3.7 | 3.6
| Mean |  | 3.2333 | 4.0833 | 3.7 | 3.2717 | 3.9833 | 3.8 | 3.86 | 3.6667 | 3.9517 | 3.9563 | 3.1833 | 5.7959 | 3.7667 | 3.7167 | 3.6
| Median |  | 3.25 | 4.1 | 3.7 | 3.27 | 4 | 3.8 | 3.795 | 3.6 | 3.94 | 3.9468 | 3.2 | 5.7988 | 3.8 | 3.7 | 3.55
| Std Dev. |  | 0.10152 | 0.075277 | 4.8648e-16 | 0.061128 | 0.075277 | 0.089443 | 0.20938 | 0.10328 | 0.051153 | 0.14738 | 0.075277 | 0.065923 | 0.05164 | 0.040825 | 0.17889
| Rel.Std.Dev. |  | 3.14% | 1.844% | 1.315e-14% | 1.868% | 1.89% | 2.354% | 5.424% | 2.817% | 1.294% | 3.725% | 2.365% | 1.137% | 1.371% | 1.098% | 4.969%
| PDM3 |  | -12.57% | 10.42% | 0.05231% | -11.53% | 7.714% | 2.756% | 4.379% | -0.8491% | 6.858% | 6.982% | -13.92% | 56.73% | 1.855% | 0.503% | -2.652%
| Z-Score (Absolute) |  | 1.94 | 1.15 | 0.24 | 1.8 | 0.79 | 0.12 | 0.34 | 0.36 | 0.67 | 0.69 | 2.13 | 7.4 | 0 | 0.18 | 0.61
| Analytical results for Hg in OREAS 780 (Indicative Value 0.055 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 14 | [17 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [17 empty] | 3
3
| 1 | 1 | 0.05 | < 1 | [17 empty] | 1
| 1 | 2 | 0.04 | < 1 | [17 empty] | 12
| 1 | 3 | 0.08 | < 1 | [17 empty] | 16
| 1 | 4 | 0.05 | < 1 | [17 empty] | 0.055
| 1 | 5 | 0.06 | < 1 | [17 empty] | 23
| 1 | 6 | 0.05 | < 1
| Mean |  | 0.055
| Median |  | 0.05
| Std Dev. |  | 0.013784
| Rel.Std.Dev. |  | 25.06%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 780 (Certified Value 0.59 ppm) | [20 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 0.43 | 0.68002 | 0.48 | 0.6 | 0.64 | 0.6 | 0.62 | 0.73785 | 1.0072 | 0.563 | [9 empty] | 1
| 1 | 2 | 0.46 | 0.64921 | 0.49 | 0.6 | 0.65 | 0.62 | 0.61 | 0.6605 | 1.0132 | 0.5574 | [9 empty] | 20
| 1 | 3 | 0.44 | 0.63064 | 0.49 | 0.6 | 0.66 | 0.62 | 0.63 | 0.66831 | 1.0241 | 0.5487 | [9 empty] | 16
| 1 | 4 | 0.43 | 0.66836 | 0.51 | 0.6 | 0.73 | 0.58 | 0.6 | 0.7246 | 1.0018 | 0.6378 | [9 empty] | 0.59266
| 1 | 5 | 0.47 | 0.65471 | 0.48 | 0.6 | 0.66 | 0.6 | 0.63 | 0.72966 | 1.0273 | 0.5708 | [9 empty] | 51
| 1 | 6 | 0.46 | 0.68586 | 0.49 | 0.6 | 0.62 | 0.58 | 0.6 | 0.67292 | 0.88509 | 0.5273
| Mean |  | 0.44833 | 0.66147 | 0.49 | 0.6 | 0.66 | 0.6 | 0.615 | 0.69897 | 0.99313 | 0.5675
| Median |  | 0.45 | 0.66153 | 0.49 | 0.6 | 0.655 | 0.6 | 0.615 | 0.69876 | 1.0102 | 0.5602
| Std Dev. |  | 0.017224 | 0.020653 | 0.010954 | 0 | 0.037417 | 0.017889 | 0.013784 | 0.035239 | 0.053813 | 0.037536
| Rel.Std.Dev. |  | 3.842% | 3.122% | 2.236% | 0% | 5.669% | 2.981% | 2.241% | 5.042% | 5.419% | 6.614%
| PDM3 |  | -24.35% | 11.61% | -17.32% | 1.239% | 11.36% | 1.239% | 3.77% | 17.94% | 67.57% | -4.245%
| Z-Score (Absolute) |  | 2.37 | 0.66 | 1.77 | NA | 0.64 | 0.21 | 0 | 1.19 | 5.37 | 0.67
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 780 (Certified Value 0.082 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.115 | 0.079 | 0.078 | 0.11 | < 0.1 | 0.07 | 0.08 | 0.1 | 0.086 | 0.12175 | 0.12 | 0.07756 | 0.083 | 0.074 | 0.08 |  |  |  |  | 1
| 1 | 2 | 0.124 | 0.079 | 0.082 | 0.11 | 0.1 | 0.07 | 0.09 | 0.1 | 0.085 | 0.12817 | 0.12 | 0.07432 | 0.072 | 0.075 | 0.084 |  |  |  |  | 1
| 1 | 3 | 0.123 | 0.087 | 0.08 | 0.11 | < 0.1 | 0.08 | 0.08 | 0.1 | 0.09 | 0.11445 | 0.13 | 0.07701 | 0.079 | 0.075 | 0.084 |  |  |  |  | 16
| 1 | 4 | 0.129 | 0.076 | 0.085 | 0.1 | < 0.1 | 0.07 | 0.08 | 0.1 | 0.082 | 0.14429 | 0.12 | 0.07306 | 0.084 | 0.083 | 0.079 |  |  |  |  | 0.081568
| 1 | 5 | 0.133 | 0.092 | 0.082 | 0.11 | < 0.1 | 0.07 | 0.08 | 0.1 | 0.086 | 0.12613 | 0.12 | 0.07643 | 0.074 | 0.072 | 0.081 |  |  |  |  | 52
| 1 | 6 | 0.124 | 0.081 | 0.083 | 0.1 | < 0.1 | 0.07 | 0.08 | 0.1 | 0.087 | 0.13654 | 0.11 | 0.07671 | 0.075 | 0.085 | 0.08
| Mean |  | 0.12467 | 0.082333 | 0.081667 | 0.10667 | 0.1 | 0.071667 | 0.081667 | 0.1 | 0.086 | 0.12855 | 0.12 | 0.075848 | 0.077833 | 0.077333 | 0.081333
| Median |  | 0.124 | 0.08 | 0.082 | 0.11 | 0.1 | 0.07 | 0.08 | 0.1 | 0.086 | 0.12715 | 0.12 | 0.07657 | 0.077 | 0.075 | 0.0805
| Std Dev. |  | 0.0060882 | 0.0059889 | 0.0024221 | 0.005164 |  | 0.0040825 | 0.0040825 | 1.5202e-17 | 0.0026077 | 0.010608 | 0.0063246 | 0.001759 | 0.0049565 | 0.0053166 | 0.0021602
| Rel.Std.Dev. |  | 4.884% | 7.274% | 2.966% | 4.841% |  | 5.696% | 4.999% | 1.52e-14% | 3.032% | 8.252% | 5.27% | 2.319% | 6.368% | 6.875% | 2.656%
| PDM3 |  | 52.84% | 0.9381% | 0.1208% | 30.77% | 22.6% | -12.14% | 0.1208% | 22.6% | 5.433% | 57.6% | 47.12% | -7.012% | -4.579% | -5.192% | -0.2879%
| Z-Score (Absolute) |  | 5.32 | 0.04 | 0.04 | 3.07 | NA | 1.29 | 0.04 | 2.24 | 0.5 | 5.8 | 4.74 | 0.77 | 0.52 | 0.58 | 0.08
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 780 (Certified Value 2.2 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.1362 | 2.2665 | 2.24 | 2.11 | 2.4 | 2.11 | 2.21 | 2.19 | 2.2393 | 2.1986 | 2.22 | 2.18 | 2.2954 | 2.18 | 2.3 | 2.16 | 2.17 | 2.2743 |  | 1
| 1 | 2 | 2.1379 | 2.2594 | 2.17 | 2.08 | 2.37 | 2.15 | 2.22 | 2.23 | 2.2399 | 2.2407 | 2.27 | 2.14 | 2.2767 | 2.2 | 2.17 | 2.17 | 2.28 | 2.2011 |  | #N/A
| 1 | 3 | 2.1146 | 2.2476 | 2.25 | 2.13 | 2.33 | 2.1 | 2.21 | 2.18 | 2.2452 | 2.1814 | 2.28 | 2.16 | 2.2871 | 2.16 | 2.16 | 2.24 | 2.13 | 2.1738 |  | 16
| 1 | 4 | 2.1207 | 2.2491 | 2.19 | 2.18 | 2.36 | 2.11 | 2.22 | 2.27 | 2.2243 | 2.1561 | 2.23 | 2.1 | 2.3055 | 2.15 | 2.16 | 2.24 | 2.18 | 2.1861 |  | 2.1957
| 1 | 5 | 2.1183 | 2.2654 | 2.19 | 2.1 | 2.31 | 2.11 | 2.23 | 2.19 | 2.2279 | 2.134 | 2.26 | 2.14 | 2.2915 | 2.18 | 2.17 | 2.17 | 2.19 | 2.2036 |  | 53
| 1 | 6 | 2.1148 | 2.2502 | 2.21 | 2.15 | 2.35 | 2.1 | 2.23 | 2.22 | 2.2331 | 2.2346 | 2.23 | 2.09 | 2.2864 | 2.19 | 2.21 | 2.23 | 2.17 | 2.202
| Mean |  | 2.1237 | 2.2564 | 2.2083 | 2.125 | 2.3533 | 2.1133 | 2.22 | 2.2133 | 2.2349 | 2.1909 | 2.2483 | 2.135 | 2.2905 | 2.1767 | 2.195 | 2.2017 | 2.1867 | 2.2068
| Median |  | 2.1195 | 2.2548 | 2.2 | 2.12 | 2.355 | 2.11 | 2.22 | 2.205 | 2.2362 | 2.19 | 2.245 | 2.14 | 2.2893 | 2.18 | 2.17 | 2.2 | 2.175 | 2.2016
| Std Dev. |  | 0.010565 | 0.0084991 | 0.031252 | 0.036194 | 0.031411 | 0.018619 | 0.0089443 | 0.033862 | 0.0079377 | 0.042398 | 0.024833 | 0.034496 | 0.0096829 | 0.018619 | 0.054681 | 0.038687 | 0.050067 | 0.035041
| Rel.Std.Dev. |  | 0.4975% | 0.3767% | 1.415% | 1.703% | 1.335% | 0.881% | 0.4029% | 1.53% | 0.3552% | 1.935% | 1.104% | 1.616% | 0.4228% | 0.8554% | 2.491% | 1.757% | 2.29% | 1.588%
| PDM3 |  | -3.276% | 2.764% | 0.5762% | -3.219% | 7.18% | -3.75% | 1.108% | 0.8039% | 1.788% | -0.2173% | 2.398% | -2.764% | 4.316% | -0.866% | -0.03103% | 0.2726% | -0.4106% | 0.5072%
| Z-Score (Absolute) |  | 1.86 | 1.21 | 0.09 | 1.83 | 3.45 | 2.1 | 0.36 | 0.21 | 0.71 | 0.31 | 1.02 | 1.6 | 1.99 | 0.64 | 0.21 | 0.06 | 0.41 | 0.06
| Analytical results for La in OREAS 780 (Certified Value 43.7 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 36.1 | 45.4183 | 37 | 40.7 | 36.92 | 44.1 | 43.9 | 45.42 | 42.5773 | 43.5 | 44.89 | 45.4503 | 46 | 41.7679 | 45.2 | 43.2 | 41 | 38.9116 |  | 1
| 1 | 2 | 37 | 45.3422 | 30.8 | 39.3 | 36.3 | 45.5 | 43.9 | 45.07 | 44.4094 | 45.3 | 43.58 | 42.9938 | 45 | 41.736 | 45.2 | 46.1 | 44.6 | 38.9138 |  | 21
| 1 | 3 | 35.2 | 45.3288 | 35.4 | 41.7 | 36.74 | 43.4 | 43.6 | 45.34 | 43.0758 | 43.6 | 44.45 | 43.8203 | 46 | 41.0248 | 46 | 45.3 | 41.6 | 42.3461 |  | 16
| 1 | 4 | 34.8 | 45.485 | 32.9 | 43.4 | 36.36 | 44 | 41.2 | 45.43 | 42.9391 | 43.6 | 39.14 | 45.3425 | 45 | 41.1445 | 45.2 | 45.9 | 43.5 | 46.5962 |  | 43.7113
| 1 | 5 | 37.1 | 45.4662 | 32.2 | 41.2 | 36.42 | 41.7 | 44 | 45.82 | 41.9382 | 43.4 | 42.56 | 45.8323 | 45 | 41.62 | 45.9 | 45 | 43.1 | 42.2566 |  | 54
| 1 | 6 | 37.6 | 45.6219 | 34.6 | 41.1 | 36.81 | 43.4 | 44.5 | 46.54 | 42.8428 | 43.9 | 41.57 | 44.9645 | 45 | 36.6564 | 46.9 | 44.6 | 38.5 | 44.8849
| Mean |  | 36.3 | 45.4437 | 33.8167 | 41.2333 | 36.5917 | 43.6833 | 43.5167 | 45.6033 | 42.9638 | 43.8833 | 42.6983 | 44.734 | 45.3333 | 40.6583 | 45.7333 | 45.0167 | 42.05 | 42.3182
| Median |  | 36.55 | 45.4423 | 33.75 | 41.15 | 36.58 | 43.7 | 43.9 | 45.425 | 42.8909 | 43.6 | 43.07 | 45.1535 | 45 | 41.3823 | 45.55 | 45.15 | 42.35 | 42.3013
| Std Dev. |  | 1.1243 | 0.10782 | 2.2719 | 1.3382 | 0.26294 | 1.2384 | 1.172 | 0.51802 | 0.81472 | 0.71391 | 2.1262 | 1.096 | 0.5164 | 1.9851 | 0.6802 | 1.0496 | 2.1732 | 3.1012
| Rel.Std.Dev. |  | 3.097% | 0.2373% | 6.718% | 3.245% | 0.7186% | 2.835% | 2.693% | 1.136% | 1.896% | 1.627% | 4.98% | 2.45% | 1.139% | 4.882% | 1.487% | 2.332% | 5.168% | 7.328%
| PDM3 |  | -16.96% | 3.963% | -22.64% | -5.669% | -16.29% | -0.06405% | -0.4453% | 4.328% | -1.71% | 0.3935% | -2.317% | 2.339% | 3.711% | -6.985% | 4.626% | 2.986% | -3.801% | -3.187%
| Z-Score (Absolute) |  | 2.47 | 0.79 | 3.36 | 0.72 | 2.37 | 0.16 | 0.1 | 0.84 | 0.1 | 0.23 | 0.19 | 0.53 | 0.75 | 0.92 | 0.89 | 0.63 | 0.42 | 0.33
| Analytical results for Li in OREAS 780 (Certified Value 45.3 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 44 | 33.8717 | 47.5 | 44.1 | 45.43 | 45 | 44 | 43.6 | 43.7 | 43.828 | 45 | 46.8 | 41.05 | 46.5 | 42.7957 | 46.3 | 48 | 46.4 | 61.5272 | 1
| 1 | 2 | 44 | 33.8051 | 47.1 | 44.1 | 46.07 | 46 | 43 | 44 | 44.8 | 44.89 | 45.5 | 45.2 | 39.98 | 46.6 | 42.3386 | 46 | 47 | 48.9 | 64.8744 | 22
| 1 | 3 | 44 | 33.9531 | 48.5 | 45 | 45.91 | 47 | 46 | 43.8 | 44.3 | 44.3525 | 43.5 | 45.5 | 39.39 | 46.8 | 42.5829 | 44.9 | 48.3 | 45.5 | 63.7243 | 16
| 1 | 4 | 44 | 33.9231 | 46.7 | 46.4 | 46.05 | 46 | 47 | 44.1 | 44 | 43.6972 | 46.5 | 45.1 | 40.79 | 46.6 | 42.216 | 45.3 | 47.6 | 46.3 | 65.7185 | 45.2898
| 1 | 5 | 46 | 34.0865 | 48.1 | 44.9 | 45.73 | 44 | 44 | 44 | 44.8 | 43.0681 | 45 | 45.7 | 40.02 | 46.6 | 43.3459 | 46.1 | 46.3 | 46.8 | 63.4631 | 55
| 1 | 6 | 44 | 33.7196 | 47.5 | 46 | 45.84 | 44 | 45 | 44 | 44.4 | 44.9145 | 46 | 43.8 | 39.84 | 46.3 | 42.1778 | 46.6 | 47.1 | 46.4 | 65.1765
| Mean |  | 44.3333 | 33.8932 | 47.5667 | 45.0833 | 45.8383 | 45.3333 | 44.8333 | 43.9167 | 44.3333 | 44.125 | 45.25 | 45.35 | 40.1783 | 46.5667 | 42.5761 | 45.8667 | 47.3833 | 46.7167 | 64.0807
| Median |  | 44 | 33.8974 | 47.5 | 44.95 | 45.875 | 45.5 | 44.5 | 44 | 44.35 | 44.0902 | 45.25 | 45.35 | 40 | 46.6 | 42.4607 | 46.05 | 47.35 | 46.4 | 64.2993
| Std Dev. |  | 0.8165 | 0.12663 | 0.6532 | 0.95376 | 0.23752 | 1.2111 | 1.472 | 0.18348 | 0.43665 | 0.72782 | 1.0368 | 0.97314 | 0.62198 | 0.1633 | 0.44398 | 0.64083 | 0.73052 | 1.1514 | 1.5193
| Rel.Std.Dev. |  | 1.842% | 0.3736% | 1.373% | 2.116% | 0.5182% | 2.671% | 3.283% | 0.4178% | 0.9849% | 1.649% | 2.291% | 2.146% | 1.548% | 0.3507% | 1.043% | 1.397% | 1.542% | 2.465% | 2.371%
| PDM3 |  | -2.112% | -25.16% | 5.027% | -0.4558% | 1.211% | 0.09621% | -1.008% | -3.032% | -2.112% | -2.572% | -0.08779% | 0.133% | -11.29% | 2.819% | -5.992% | 1.274% | 4.623% | 3.151% | 41.49%
| Z-Score (Absolute) |  | 0.55 | 6.81 | 1.39 | 0.1 | 0.35 | 0.05 | 0.25 | 0.8 | 0.55 | 0.67 | 0 | 0.06 | 3.04 | 0.79 | 1.6 | 0.37 | 1.28 | 0.88 | 11.29
| Analytical results for Lu in OREAS 780 (Certified Value 0.23 ppm) | [20 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [10 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.21 | 0.21 | 0.3 | 0.23 | 0.24 | 0.25 | 0.28274 | 0.39644 | 0.161 | [10 empty] | 1
| 1 | 2 | 0.22 | 0.21 | 0.2 | 0.25 | 0.24 | 0.25 | 0.25671 | 0.39687 | 0.1614 | [10 empty] | 23
| 1 | 3 | 0.22 | 0.23 | 0.3 | 0.27 | 0.24 | 0.26 | 0.26381 | 0.38245 | 0.1583 | [10 empty] | 16
| 1 | 4 | 0.22 | 0.21 | 0.3 | 0.26 | 0.24 | 0.25 | 0.28853 | 0.3893 | 0.1623 | [10 empty] | 0.23256
| 1 | 5 | 0.23 | 0.21 | 0.3 | 0.31 | 0.24 | 0.25 | 0.28754 | 0.38791 | 0.1673 | [10 empty] | 56
| 1 | 6 | 0.23 | 0.21 | 0.2 | 0.28 | 0.24 | 0.25 | 0.26375 | 0.34583 | 0.154
| Mean |  | 0.22167 | 0.21333 | 0.26667 | 0.26667 | 0.24 | 0.25167 | 0.27385 | 0.38313 | 0.16072
| Median |  | 0.22 | 0.21 | 0.3 | 0.265 | 0.24 | 0.25 | 0.27328 | 0.3886 | 0.1612
| Std Dev. |  | 0.0075277 | 0.008165 | 0.05164 | 0.027325 | 0 | 0.0040825 | 0.013988 | 0.019072 | 0.0044124
| Rel.Std.Dev. |  | 3.396% | 3.827% | 19.36% | 10.25% | 0% | 1.622% | 5.108% | 4.978% | 2.745%
| PDM3 |  | -4.683% | -8.266% | 14.67% | 14.67% | 3.201% | 8.217% | 17.75% | 64.75% | -30.89%
| Z-Score (Absolute) |  | 0.62 | 0.84 | NA | 0.54 | 0.15 | 0.15 | 0.72 | 3.55 | 2.2
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 780 (Certified Value 0.265 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.2023 | 0.28255 | 0.25 | 0.26 | 0.3 | 0.22 | 0.26 | 0.292 | 0.2912 | 0.26843 | 0.24 | 0.261 | 0.3495 | 0.22 | 0.26 | 0.29 | 0.28 | 0.229 |  | 1
| 1 | 2 | 0.2039 | 0.28801 | 0.24 | 0.27 | 0.29 | 0.22 | 0.27 | 0.29 | 0.2956 | 0.27339 | 0.23 | 0.259 | 0.34453 | 0.23 | 0.26 | 0.28 | 0.3 | 0.2238 |  | 24
| 1 | 3 | 0.1985 | 0.28157 | 0.25 | 0.27 | 0.3 | 0.22 | 0.27 | 0.295 | 0.3084 | 0.27249 | 0.24 | 0.271 | 0.33898 | 0.22 | 0.27 | 0.29 | 0.28 | 0.2259 |  | 16
| 1 | 4 | 0.2028 | 0.2813 | 0.24 | 0.28 | 0.3 | 0.22 | 0.27 | 0.289 | 0.2926 | 0.27068 | 0.24 | 0.251 | 0.35068 | 0.22 | 0.26 | 0.29 | 0.28 | 0.2499 |  | 0.26515
| 1 | 5 | 0.2038 | 0.28975 | 0.24 | 0.27 | 0.29 | 0.22 | 0.27 | 0.292 | 0.3023 | 0.26594 | 0.23 | 0.259 | 0.35125 | 0.21 | 0.26 | 0.28 | 0.28 | 0.2296 |  | 57
| 1 | 6 | 0.2044 | 0.28368 | 0.25 | 0.28 | 0.29 | 0.22 | 0.27 | 0.29 | 0.3005 | 0.27287 | 0.24 | 0.253 | 0.33867 | 0.22 | 0.27 | 0.28 | 0.28 | 0.2357
| Mean |  | 0.20262 | 0.28448 | 0.245 | 0.27167 | 0.295 | 0.22 | 0.26833 | 0.29133 | 0.29843 | 0.27063 | 0.23667 | 0.259 | 0.3456 | 0.22 | 0.26333 | 0.285 | 0.28333 | 0.23232
| Median |  | 0.2033 | 0.28311 | 0.245 | 0.27 | 0.295 | 0.22 | 0.27 | 0.291 | 0.29805 | 0.27158 | 0.24 | 0.259 | 0.34701 | 0.22 | 0.26 | 0.285 | 0.28 | 0.2293
| Std Dev. |  | 0.0021591 | 0.0035538 | 0.0054772 | 0.0075277 | 0.0054772 | 0 | 0.0040825 | 0.0021602 | 0.0065258 | 0.0029296 | 0.005164 | 0.0070427 | 0.0057625 | 0.0063246 | 0.005164 | 0.0054772 | 0.008165 | 0.0095153
| Rel.Std.Dev. |  | 1.066% | 1.249% | 2.236% | 2.771% | 1.857% | 0% | 1.521% | 0.7415% | 2.187% | 1.083% | 2.182% | 2.719% | 1.667% | 2.875% | 1.961% | 1.922% | 2.882% | 4.096%
| PDM3 |  | -23.58% | 7.288% | -7.6% | 2.457% | 11.26% | -17.03% | 1.2% | 9.874% | 12.55% | 2.066% | -10.74% | -2.32% | 30.34% | -17.03% | -0.6861% | 7.485% | 6.857% | -12.38%
| Z-Score (Absolute) |  | 1.95 | 0.44 | 0.71 | 0.06 | 0.74 | 1.44 | 0.03 | 0.64 | 0.84 | 0.03 | 0.96 | 0.31 | 2.22 | 1.44 | 0.18 | 0.45 | 0.4 | 1.08
| Analytical results for Mn in OREAS 780 (Certified Value 0.004 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0024 | 0.0037935 | 0.004 | 0.0039 | 0.003799 | 0.003181 | 0.0039 | 0.0038 | 0.0039 | 0.0040374 | 0.0038 | 0.00383 | 0.004415 | 0.0028 | 0.0039 | 0.0039 | 0.0039 | < 0.1 |  | 1
| 1 | 2 | 0.0019 | 0.0037088 | 0.0039 | 0.0037 | 0.003729 | 0.003217 | 0.0038 | 0.0037 | 0.0039 | 0.0040871 | 0.0036 | 0.00386 | 0.004218 | 0.0029 | 0.0037 | 0.0038 | 0.0045 | < 0.1 |  | 3
| 1 | 3 | 0.0024 | 0.0038352 | 0.0041 | 0.0039 | 0.00385 | 0.00352 | 0.0037 | 0.0038 | 0.004 | 0.0040712 | 0.0038 | 0.0042 | 0.004202 | 0.0031 | 0.0038 | 0.0039 | 0.0039 | < 0.1 |  | 16
| 1 | 4 | 0.0022 | 0.0038385 | 0.004 | 0.0039 | 0.003726 | 0.003489 | 0.0039 | 0.0037 | 0.0039 | 0.0039169 | 0.0038 | 0.00412 | 0.004295 | 0.0031 | 0.0036 | 0.004 | 0.0039 | < 0.1 |  | 0.0038318
| 1 | 5 | 0.0024 | 0.0038099 | 0.004 | 0.0038 | 0.003845 | 0.003558 | 0.0037 | 0.0039 | 0.004 | 0.0039243 | 0.0036 | 0.00405 | 0.004193 | 0.0028 | 0.0036 | 0.0038 | 0.0039 | < 0.1 |  | 58
| 1 | 6 | 0.002 | 0.0037684 | 0.005 | 0.004 | 0.003765 | 0.003532 | 0.0037 | 0.0037 | 0.0039 | 0.0038314 | 0.0038 | 0.00383 | 0.004178 | 0.0029 | 0.0037 | 0.0039 | 0.0039 | < 0.1
| Mean |  | 0.0022167 | 0.0037924 | 0.0041667 | 0.0038667 | 0.0037857 | 0.0034162 | 0.0037833 | 0.0037667 | 0.0039333 | 0.003978 | 0.0037333 | 0.0039817 | 0.0042502 | 0.0029333 | 0.0037167 | 0.0038833 | 0.004
| Median |  | 0.0023 | 0.0038017 | 0.004 | 0.0039 | 0.003782 | 0.0035045 | 0.00375 | 0.00375 | 0.0039 | 0.0039809 | 0.0038 | 0.003955 | 0.00421 | 0.0029 | 0.0037 | 0.0039 | 0.0039
| Std Dev. |  | 0.00022286 | 4.8644e-05 | 0.00041312 | 0.00010328 | 5.4837e-05 | 0.00017005 | 9.8319e-05 | 8.165e-05 | 5.164e-05 | 0.0001022 | 0.00010328 | 0.00016266 | 9.0566e-05 | 0.00013663 | 0.0001169 | 7.5277e-05 | 0.00024495
| Rel.Std.Dev. |  | 10.05% | 1.283% | 9.915% | 2.671% | 1.449% | 4.978% | 2.599% | 2.168% | 1.313% | 2.569% | 2.766% | 4.085% | 2.131% | 4.658% | 3.145% | 1.938% | 6.124%
| PDM3 |  | -42.15% | -1.03% | 8.738% | 0.9087% | -1.205% | -10.85% | -1.266% | -1.701% | 2.649% | 3.815% | -2.571% | 3.91% | 10.92% | -23.45% | -3.006% | 1.344% | 4.388%
| Z-Score (Absolute) |  | 7.54 | 0 | 1.79 | 0.36 | 0.03 | 1.8 | 0.04 | 0.12 | 0.67 | 0.89 | 0.28 | 0.91 | 2.19 | 4.11 | 0.36 | 0.44 | 0.99 | NA
Results from laboratory 1.24 were removed due to their systematic bias.
| Analytical results for Mo in OREAS 780 (Certified Value 3.76 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 3.7 | 3.2333 | 3.89 | 3.53 | 2.32 | 4.29 | 4 | 3.7 | 3.8 | < 50 | 3.8 | 4.1 | 4.0312 | 3.7 | 3.6268 | 3.59 | 3.47 | 3.48 | 28361 | 1
| 1 | 2 | 3.8 | 3.3454 | 4.01 | 3.88 | 2.36 | 4.11 | 4 | 3.5 | 3.8 | < 50 | 3.9 | 4.2 | 3.9859 | 3.6 | 3.6061 | 3.57 | 3.44 | 3.61 | 30729 | 4
| 1 | 3 | 4 | 3.1185 | 4 | 3.57 | 2.34 | 3.68 | 4 | 3.6 | 3.8 | < 50 | 3.6 | 4.3 | 3.6854 | 3.8 | 3.7899 | 3.99 | 3.49 | 3.48 | 31795 | 16
| 1 | 4 | 4.4 | 2.9499 | 3.76 | 3.68 | 2.4 | 3.71 | 4 | 3.5 | 3.9 | < 50 | 4.1 | 4.2 | 4.467 | 3.7 | 3.5678 | 3.74 | 3.65 | 3.55 | 31669 | 3.7641
| 1 | 5 | 4.6 | 2.9555 | 4.12 | 3.49 | 2.41 | 3.83 | 4 | 3.5 | 3.7 | < 50 | 3.8 | 4 | 7.0391 | 3.6 | 3.8054 | 5.89 | 3.52 | 3.63 | 29947 | 59
| 1 | 6 | 4.2 | 3.5984 | 3.95 | 3.69 | 2.46 | 3.77 | 4 | 3.6 | 3.9 | < 50 | 3.7 | 4.2 | 4.0558 | 3.7 | 3.7308 | 3.66 | 4.05 | 3.58 | 31853
| Mean |  | 4.1167 | 3.2002 | 3.955 | 3.64 | 2.3817 | 3.8983 | 4 | 3.5667 | 3.8167 |  | 3.8167 | 4.1667 | 4.5441 | 3.6833 | 3.6878 | 4.0733 | 3.6033 | 3.555 | 30725.6667
| Median |  | 4.1 | 3.1759 | 3.975 | 3.625 | 2.38 | 3.8 | 4 | 3.55 | 3.8 |  | 3.8 | 4.2 | 4.0435 | 3.7 | 3.6788 | 3.7 | 3.505 | 3.565 | 31199
| Std Dev. |  | 0.34881 | 0.24898 | 0.12211 | 0.14227 | 0.051543 | 0.24596 | 0 | 0.08165 | 0.075277 |  | 0.17224 | 0.10328 | 1.2475 | 0.075277 | 0.1009 | 0.9029 | 0.23062 | 0.064109 | 1378.5654
| Rel.Std.Dev. |  | 8.473% | 7.78% | 3.087% | 3.908% | 2.164% | 6.309% | 0% | 2.289% | 1.972% |  | 4.513% | 2.479% | 27.45% | 2.044% | 2.736% | 22.17% | 6.4% | 1.803% | 4.487%
| PDM3 |  | 9.367% | -14.98% | 5.072% | -3.297% | -36.73% | 3.566% | 6.267% | -5.245% | 1.397% |  | 1.397% | 10.7% | 20.72% | -2.145% | -2.027% | 8.216% | -4.271% | -5.555% | 816184%
| Z-Score (Absolute) |  | 0.81 | 1.66 | 0.37 | 0.48 | 3.87 | 0.22 | NA | 0.67 | 0 | 57.14 | 0 | 0.94 | 1.96 | 0.36 | 0.35 | 0.69 | 0.58 | 0.71 | NA
Results from laboratory 1.24 were removed due to their systematic bias. Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 780 (Certified Value 0.136 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.093 | 0.10611 | 0.14 | 0.14 | 0.24 | 0.12 | 0.12 | 0.132 | 0.1367 | 0.13481 | 0.14 | 0.131 | 0.34592 | 0.15 | 0.13 | 0.15 | 0.14 | 0.168 |  | 1
| 1 | 2 | 0.0931 | 0.10722 | 0.14 | 0.14 | 0.24 | 0.12 | 0.13 | 0.131 | 0.1389 | 0.13901 | 0.13 | 0.13 | 0.33567 | 0.15 | 0.13 | 0.14 | 0.15 | 0.1702 |  | #N/A
| 1 | 3 | 0.0918 | 0.10446 | 0.14 | 0.14 | 0.24 | 0.12 | 0.13 | 0.133 | 0.143 | 0.13591 | 0.13 | 0.137 | 0.33477 | 0.16 | 0.13 | 0.15 | 0.14 | 0.1793 |  | 16
| 1 | 4 | 0.0933 | 0.10874 | 0.14 | 0.14 | 0.24 | 0.12 | 0.13 | 0.134 | 0.1376 | 0.13514 | 0.13 | 0.134 | 0.34611 | 0.14 | 0.13 | 0.15 | 0.14 | 0.2041 |  | 0.13628
| 1 | 5 | 0.0949 | 0.10511 | 0.14 | 0.14 | 0.24 | 0.12 | 0.13 | 0.138 | 0.1405 | 0.13464 | 0.13 | 0.135 | 0.33895 | 0.15 | 0.13 | 0.14 | 0.14 | 0.1709 |  | 60
| 1 | 6 | 0.0935 | 0.10443 | 0.14 | 0.14 | 0.24 | 0.12 | 0.13 | 0.132 | 0.1398 | 0.1369 | 0.13 | 0.13 | 0.33845 | 0.16 | 0.13 | 0.15 | 0.14 | 0.1776
| Mean |  | 0.093267 | 0.10601 | 0.14 | 0.14 | 0.24 | 0.12 | 0.12833 | 0.13333 | 0.13942 | 0.13607 | 0.13167 | 0.13283 | 0.33998 | 0.15167 | 0.13 | 0.14667 | 0.14167 | 0.17835
| Median |  | 0.0932 | 0.10561 | 0.14 | 0.14 | 0.24 | 0.12 | 0.13 | 0.1325 | 0.13935 | 0.13552 | 0.13 | 0.1325 | 0.3387 | 0.15 | 0.13 | 0.15 | 0.14 | 0.17425
| Std Dev. |  | 0.00099733 | 0.0017091 | 0 | 0 | 0 | 0 | 0.0040825 | 0.0025033 | 0.0022409 | 0.0016644 | 0.0040825 | 0.0029269 | 0.0049393 | 0.0075277 | 0 | 0.005164 | 0.0040825 | 0.013368
| Rel.Std.Dev. |  | 1.069% | 1.612% | 0% | 0% | 0% | 0% | 3.181% | 1.877% | 1.607% | 1.223% | 3.101% | 2.203% | 1.453% | 4.963% | 0% | 3.521% | 2.882% | 7.496%
| PDM3 |  | -31.56% | -22.21% | 2.729% | 2.729% | 76.11% | -11.95% | -5.832% | -2.163% | 2.301% | -0.1564% | -3.386% | -2.53% | 149.5% | 11.29% | -4.609% | 7.621% | 3.952% | 30.87%
| Z-Score (Absolute) |  | 3.6 | 2.57 | 0.18 | 0.18 | 8.27 | 1.44 | 0.76 | 0.36 | 0.14 | 0.14 | 0.49 | 0.4 | 16.36 | 1.13 | 0.63 | 0.72 | 0.32 | 3.29
| Analytical results for Nb in OREAS 780 (Certified Value 11.7 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 4.78 | 12.3 | 11.6 | 14.75 | 10.6 | 7.4 | 12.05 | 10.2 | 11.8 | 13.394 | 10.2 | 13.9537 | 10.9 | 11 | 11.4 | 7.9198 |  |  |  | 1
| 1 | 2 | 5.42 | 12.2 | 11.8 | 14.22 | 11.4 | 7.7 | 10.55 | 9.6 | 11.9 | 12.4987 | 10.2 | 13.842 | 11.1 | 11.6 | 13 | 6.8513 |  |  |  | #N/A
| 1 | 3 | 4.83 | 12.3 | 11.6 | 15.03 | 11.8 | 7.4 | 11.34 | 10.2 | 12 | 12.3544 | 10.4 | 13.7723 | 11.1 | 11.4 | 11.3 | 7.5493 |  |  |  | 16
| 1 | 4 | 5.81 | 11.8 | 11.8 | 14.75 | 11.8 | 7.3 | 10.78 | 10.2 | 12.3 | 12.7825 | 10.4 | 13.7065 | 11.1 | 11.7 | 11.8 | 10.0845 |  |  |  | 11.6579
| 1 | 5 | 5.85 | 12 | 11.9 | 14.43 | 11.7 | 7.3 | 10.95 | 10.1 | 12.2 | 13.3251 | 10.6 | 14.1259 | 11.3 | 11.3 | 11.8 | 7.2025 |  |  |  | 61
| 1 | 6 | 5.79 | 12.2 | 11.6 | 15.52 | 11.3 | 7.4 | 10.77 | 10.2 | 12.2 | 12.3382 | 10.3 | 14.0347 | 11.1 | 11.6 | 11.6 | 8.0659
| Mean |  | 5.4133 | 12.1333 | 11.7167 | 14.7833 | 11.4333 | 7.4167 | 11.0733 | 10.0833 | 12.0667 | 12.7821 | 10.35 | 13.9059 | 11.1 | 11.4333 | 11.8167 | 7.9455
| Median |  | 5.605 | 12.2 | 11.7 | 14.75 | 11.55 | 7.4 | 10.865 | 10.2 | 12.1 | 12.6406 | 10.35 | 13.8979 | 11.1 | 11.5 | 11.7 | 7.7346
| Std Dev. |  | 0.49625 | 0.19664 | 0.13292 | 0.45781 | 0.45898 | 0.1472 | 0.54628 | 0.24014 | 0.19664 | 0.47529 | 0.15166 | 0.16058 | 0.12649 | 0.2582 | 0.61455 | 1.1399
| Rel.Std.Dev. |  | 9.167% | 1.621% | 1.134% | 3.097% | 4.014% | 1.985% | 4.933% | 2.382% | 1.63% | 3.718% | 1.465% | 1.155% | 1.14% | 2.258% | 5.201% | 14.35%
| PDM3 |  | -53.57% | 4.078% | 0.5042% | 26.81% | -1.926% | -36.38% | -5.014% | -13.51% | 3.506% | 9.644% | -11.22% | 19.28% | -4.786% | -1.926% | 1.362% | -31.84%
| Z-Score (Absolute) |  | 4.55 | 0.53 | 0.21 | 2.53 | 0 | 3.04 | 0.27 | 1.02 | 0.48 | 1.02 | 0.82 | 1.87 | 0.25 | 0 | 0.29 | 2.64
| Analytical results for Nd in OREAS 780 (Certified Value 35.7 ppm) | [20 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 31.29 | 37.7881 | 34.3 | 35.8 | 36.36 | 36.8 | 35.5 | 36.0256 | 34.833 | 38.0944 | [9 empty] | 1
| 1 | 2 | 33.84 | 37.3923 | 35.5 | 35.9 | 36.05 | 37.7 | 34.8 | 34.2481 | 35.1879 | 35.8499 | [9 empty] | 6
| 1 | 3 | 32.89 | 37.29 | 35.6 | 36.6 | 36.24 | 36.5 | 35.6 | 34.8934 | 35.159 | 37.8096 | [9 empty] | 16
| 1 | 4 | 31.19 | 37.3293 | 35.8 | 35.1 | 35.87 | 36.9 | 32.6 | 36.0348 | 34.6021 | 40.0315 | [9 empty] | 35.7422
| 1 | 5 | 34.34 | 37.4297 | 33.6 | 36 | 36.97 | 36.7 | 33.9 | 36.3299 | 34.9285 | 38.9428 | [9 empty] | 62
| 1 | 6 | 34.48 | 37.0692 | 35.2 | 36.7 | 37.15 | 36.9 | 33.9 | 35.6845 | 30.8771 | 36.0691
| Mean |  | 33.005 | 37.3831 | 35 | 36.0167 | 36.44 | 36.9167 | 34.3833 | 35.536 | 34.2646 | 37.7996
| Median |  | 33.365 | 37.3608 | 35.35 | 35.95 | 36.3 | 36.85 | 34.35 | 35.855 | 34.8808 | 37.952
| Std Dev. |  | 1.4767 | 0.23511 | 0.86487 | 0.58452 | 0.51155 | 0.41191 | 1.1444 | 0.80156 | 1.6736 | 1.623
| Rel.Std.Dev. |  | 4.474% | 0.6289% | 2.471% | 1.623% | 1.404% | 1.116% | 3.328% | 2.256% | 4.884% | 4.294%
| PDM3 |  | -7.658% | 4.591% | -2.077% | 0.7678% | 1.952% | 3.286% | -3.802% | -0.5769% | -4.134% | 5.756%
| Z-Score (Absolute) |  | 1.48 | 0.86 | 0.41 | 0.13 | 0.35 | 0.61 | 0.74 | 0.13 | 0.8 | 1.08
| Analytical results for Ni in OREAS 780 (Certified Value 12.8 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 12.5 | 11.99 | 13.2 | 12.2 | 11.65 | 12 | 13 | 13 | 13 | 12.7019 | 14 | 13 | 13.4993 | 13.1 | 13.5076 | 12.5 | 12.8 | 12.2 | 9.773 | 1
| 1 | 2 | 12.1 | 11.8 | 13.1 | 12.2 | 11.62 | 13 | 12 | 13 | 13 | 13.4098 | 14 | 12 | 9.5789 | 12.8 | 13.5675 | 12.3 | 13.1 | 13.7 | 11.4412 | 7
| 1 | 3 | 11.8 | 12.01 | 13.6 | 12.3 | 11.66 | 13 | 13 | 13 | 14 | 13.3895 | 14 | 12 | 9.9265 | 13.2 | 13.3528 | 12.5 | 13 | 12.8 | 10.6818 | 16
| 1 | 4 | 12.2 | 12.23 | 13.3 | 12.7 | 11.9 | 13 | 13 | 13 | 13 | 12.6524 | 14 | 14 | 12.351 | 13.2 | 13.5261 | 13 | 13.4 | 13 | 10.3162 | 12.8342
| 1 | 5 | 21 | 12.23 | 13.4 | 12.4 | 11.69 | 13 | 12 | 13 | 14 | 12.5921 | 12 | 12 | 13.4807 | 13.5 | 13.3465 | 12 | 12.8 | 12.7 | 10.1375 | 63
| 1 | 6 | 12.3 | 11.91 | 14.2 | 12.8 | 11.65 | 13 | 13 | 13 | 14 | 13.2763 | 14 | 13 | 13.1977 | 13.2 | 13.0401 | 12.3 | 13.3 | 12.6 | 10.0685
| Mean |  | 13.65 | 12.0283 | 13.4667 | 12.4333 | 11.695 | 12.8333 | 12.6667 | 13 | 13.5 | 13.0037 | 13.6667 | 12.6667 | 12.0057 | 13.1667 | 13.3901 | 12.4333 | 13.0667 | 12.8333 | 10.403
| Median |  | 12.25 | 12 | 13.35 | 12.35 | 11.655 | 13 | 13 | 13 | 13.5 | 12.9891 | 14 | 12.5 | 12.7744 | 13.2 | 13.4302 | 12.4 | 13.05 | 12.75 | 10.2268
| Std Dev. |  | 3.6082 | 0.17279 | 0.39833 | 0.2582 | 0.10291 | 0.40825 | 0.5164 | 0 | 0.54772 | 0.39293 | 0.8165 | 0.8165 | 1.7977 | 0.22509 | 0.19471 | 0.33267 | 0.25033 | 0.50067 | 0.59032
| Rel.Std.Dev. |  | 26.43% | 1.437% | 2.958% | 2.077% | 0.8799% | 3.181% | 4.077% | 0% | 4.057% | 3.022% | 5.974% | 6.446% | 14.97% | 1.71% | 1.454% | 2.676% | 1.916% | 3.901% | 5.675%
| PDM3 |  | 6.356% | -6.279% | 4.928% | -3.124% | -8.876% | -0.006903% | -1.306% | 1.292% | 5.188% | 1.32% | 6.486% | -1.306% | -6.456% | 2.59% | 4.331% | -3.124% | 1.811% | -0.006903% | -18.94%
| Z-Score (Absolute) |  | 1.38 | 1.36 | 1.07 | 0.67 | 1.92 | 0 | 0.28 | 0.28 | 1.12 | 0.29 | 1.4 | 0.28 | 1.4 | 0.56 | 0.94 | 0.67 | 0.39 | 0 | 4.1
| Analytical results for P in OREAS 780 (Certified Value 0.028 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0268 | 0.03122 | 0.028 | 0.027 | 0.029 | 0.028 | 0.026 | 0.028 | 0.0278 | 0.028236 | 0.033 | 0.029 | 0.027981 | 0.026 | 0.029 | 0.028 | 0.028 | < 0.1 |  | 1
| 1 | 2 | 0.0269 | 0.030913 | 0.028 | 0.027 | 0.029 | 0.03 | 0.027 | 0.029 | 0.0277 | 0.028152 | 0.029 | 0.029 | 0.028123 | 0.027 | 0.029 | 0.028 | 0.029 | < 0.1 |  | #N/A
| 1 | 3 | 0.0269 | 0.030968 | 0.029 | 0.028 | 0.029 | 0.0289 | 0.027 | 0.029 | 0.0278 | 0.028418 | 0.03 | 0.028 | 0.028305 | 0.026 | 0.029 | 0.028 | 0.029 | < 0.1 |  | 16
| 1 | 4 | 0.0278 | 0.031211 | 0.028 | 0.028 | 0.03 | 0.0301 | 0.028 | 0.029 | 0.0276 | 0.028257 | 0.031 | 0.026 | 0.028382 | 0.027 | 0.028 | 0.029 | 0.029 | < 0.1 |  | 0.028171
| 1 | 5 | 0.0275 | 0.03099 | 0.028 | 0.027 | 0.028 | 0.0293 | 0.026 | 0.029 | 0.0279 | 0.027391 | 0.03 | 0.027 | 0.028136 | 0.026 | 0.029 | 0.028 | 0.03 | < 0.1 |  | 64
| 1 | 6 | 0.0273 | 0.031259 | 0.028 | 0.029 | 0.029 | 0.0288 | 0.026 | 0.029 | 0.0278 | 0.028973 | 0.03 | 0.027 | 0.028135 | 0.026 | 0.029 | 0.028 | 0.028 | < 0.1
| Mean |  | 0.0272 | 0.031093 | 0.028167 | 0.027667 | 0.029 | 0.029183 | 0.026667 | 0.028833 | 0.027767 | 0.028238 | 0.0305 | 0.027667 | 0.028177 | 0.026333 | 0.028833 | 0.028167 | 0.028833
| Median |  | 0.0271 | 0.0311 | 0.028 | 0.0275 | 0.029 | 0.0291 | 0.0265 | 0.029 | 0.0278 | 0.028246 | 0.03 | 0.0275 | 0.028136 | 0.026 | 0.029 | 0.028 | 0.029
| Std Dev. |  | 0.0004 | 0.00015242 | 0.00040825 | 0.0008165 | 0.00063246 | 0.00079352 | 0.0008165 | 0.00040825 | 0.00010328 | 0.00050929 | 0.0013784 | 0.0012111 | 0.00014366 | 0.0005164 | 0.00040825 | 0.00040825 | 0.00075277
| Rel.Std.Dev. |  | 1.471% | 0.4902% | 1.449% | 2.951% | 2.181% | 2.719% | 3.062% | 1.416% | 0.372% | 1.804% | 4.519% | 4.377% | 0.5098% | 1.961% | 1.416% | 1.449% | 2.611%
| PDM3 |  | -3.446% | 10.38% | -0.0144% | -1.789% | 2.944% | 3.595% | -5.339% | 2.352% | -1.434% | 0.2383% | 8.268% | -1.789% | 0.02255% | -6.522% | 2.352% | -0.0144% | 2.352%
| Z-Score (Absolute) |  | 1.09 | 3.11 | 0.04 | 0.58 | 0.85 | 1.05 | 1.66 | 0.67 | 0.47 | 0.03 | 2.47 | 0.58 | 0.03 | 2.02 | 0.67 | 0.04 | 0.67 | 23.48
| Analytical results for Pb in OREAS 780 (Certified Value 16.6 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 17 | 18.1 | 17 | 43.2 | 14 | 16 | 18 | 17.3 | 16.656 | 16 | 16.8 | 18.2065 | 14.2 | 17.0197 | 16.2 | 16.4 | 16.2 | 15.5315 |  | 1
| 1 | 2 | 17 | 17.7 | 16.6 | 43.8 | 14 | 15 | 18 | 17.6 | 16.9481 | 17 | 16.8 | 19.6276 | 14.7 | 17.0618 | 16.3 | 16.8 | 17.6 | 15.6386 |  | #N/A
| 1 | 3 | 19 | 18.1 | 16.8 | 43 | 15 | 16 | 18 | 17.7 | 16.841 | 16 | 17.1 | 17.4683 | 13.9 | 16.705 | 15.8 | 17.1 | 15.9 | 15.6604 |  | 16
| 1 | 4 | 18 | 17.8 | 17.2 | 43.4 | 15 | 16 | 18 | 17.4 | 16.6845 | 15 | 17.2 | 19.2314 | 14.1 | 17.0176 | 16 | 17.3 | 16.4 | 14.9201 |  | 16.6372
| 1 | 5 | 20 | 17.5 | 17.2 | 44.6 | 15 | 15 | 18 | 17.2 | 16.1134 | 15 | 17.2 | 18.8066 | 14.9 | 16.7903 | 15.6 | 16.8 | 16.4 | 15.64 |  | 65
| 1 | 6 | 17 | 18 | 17.4 | 43.5 | 14 | 15 | 18 | 17.8 | 17.0327 | 17 | 17.1 | 17.9239 | 13.9 | 16.8331 | 16.5 | 17 | 16.7 | 15.8722
| Mean |  | 18 | 17.8667 | 17.0333 | 43.5833 | 14.5 | 15.5 | 18 | 17.5 | 16.7126 | 16 | 17.0333 | 18.5441 | 14.2833 | 16.9046 | 16.0667 | 16.9 | 16.5333 | 15.5438
| Median |  | 17.5 | 17.9 | 17.1 | 43.45 | 14.5 | 15.5 | 18 | 17.5 | 16.7628 | 16 | 17.1 | 18.5066 | 14.15 | 16.9254 | 16.1 | 16.9 | 16.4 | 15.6393
| Std Dev. |  | 1.2649 | 0.24221 | 0.29439 | 0.56716 | 0.54772 | 0.54772 | 0 | 0.23664 | 0.32785 | 0.89443 | 0.18619 | 0.82113 | 0.42151 | 0.14749 | 0.33267 | 0.30984 | 0.58538 | 0.32522
| Rel.Std.Dev. |  | 7.027% | 1.356% | 1.728% | 1.301% | 3.777% | 3.534% | 0% | 1.352% | 1.962% | 5.59% | 1.093% | 4.428% | 2.951% | 0.8725% | 2.071% | 1.833% | 3.541% | 2.092%
| PDM3 |  | 8.191% | 7.39% | 2.381% | 162% | -12.85% | -6.836% | 8.191% | 5.186% | 0.4531% | -3.83% | 2.381% | 11.46% | -14.15% | 1.607% | -3.43% | 1.579% | -0.6246% | -6.572%
| Z-Score (Absolute) |  | 0.79 | 0.7 | 0.09 | 19.28 | 1.74 | 1.01 | 0.79 | 0.43 | 0.14 | 0.65 | 0.09 | 1.19 | 1.89 | 0 | 0.6 | 0 | 0.27 | 0.98
| Analytical results for Pr in OREAS 780 (Certified Value 9.82 ppm) | [20 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 8.68 | 10.5039 | 9.28 | 10.1 | 9.74 | 10.2 | 10.13 | 10.2452 | 9.5146 | 8.9319 | [9 empty] | 1
| 1 | 2 | 9.37 | 10.4209 | 9.4 | 10.2 | 9.81 | 10.5 | 10.01 | 9.7195 | 9.5368 | 8.9564 | [9 empty] | 8
| 1 | 3 | 9.15 | 10.4565 | 9.75 | 9.9 | 10.03 | 10.1 | 10.02 | 9.8883 | 9.5212 | 9.4084 | [9 empty] | 16
| 1 | 4 | 8.64 | 10.4459 | 9.91 | 9.7 | 9.91 | 10.1 | 9.19 | 10.0924 | 9.4505 | 10.4372 | [9 empty] | 9.8199
| 1 | 5 | 9.49 | 10.3193 | 9.44 | 10.2 | 9.92 | 10.2 | 9.7 | 10.2377 | 9.5959 | 9.6255 | [9 empty] | 66
| 1 | 6 | 9.56 | 10.3282 | 10.03 | 10.2 | 10.04 | 10.1 | 9.59 | 10.1214 | 8.4366 | 9.9804
| Mean |  | 9.1483 | 10.4124 | 9.635 | 10.05 | 9.9083 | 10.2 | 9.7733 | 10.0507 | 9.3426 | 9.5566
| Median |  | 9.26 | 10.4334 | 9.595 | 10.15 | 9.915 | 10.15 | 9.855 | 10.1069 | 9.5179 | 9.517
| Std Dev. |  | 0.40326 | 0.07386 | 0.30468 | 0.20736 | 0.11856 | 0.15492 | 0.35308 | 0.20765 | 0.4463 | 0.5886
| Rel.Std.Dev. |  | 4.408% | 0.7093% | 3.162% | 2.063% | 1.197% | 1.519% | 3.613% | 2.066% | 4.777% | 6.159%
| PDM3 |  | -6.838% | 6.034% | -1.883% | 2.344% | 0.9009% | 3.871% | -0.4738% | 2.351% | -4.86% | -2.681%
| Z-Score (Absolute) |  | 1.89 | 1.56 | 0.56 | 0.57 | 0.18 | 0.98 | 0.18 | 0.57 | 1.36 | 0.78
| Analytical results for Rb in OREAS 780 (Certified Value 122 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
0
| 1 | 1 | 133.5 | 120.453 | 113.5 | 118.5 | 117 | 133 | 116 | 125.55 | 125 | 124 | 123.43 | 135.5 | 125.5595 | 125.5 | 116 | 119.5 | 110.9694 |  |  | 1
| 1 | 2 | 142.7 | 120.0845 | 98.7 | 114 | 118 | 134 | 115 | 128.37 | 127 | 118 | 118.1097 | 135.1 | 122.7234 | 124 | 120.5 | 131.5 | 105.2094 |  |  | 9
| 1 | 3 | 139.6 | 119.9595 | 109 | 115 | 118 | 130 | 115 | 128.74 | 122 | 118 | 119.8235 | 134.8 | 130.7742 | 128 | 119.5 | 123 | 115.4825 |  |  | 16
| 1 | 4 | 147.2 | 119.6094 | 103 | 129 | 120 | 125 | 118 | 129.89 | 125 | 100 | 122.6698 | 135.6 | 127.3131 | 125 | 120.5 | 125.5 | 127.1118 |  |  | 121.5177
| 1 | 5 | 149.6 | 119.247 | 102 | 112 | 117 | 134 | 118 | 127.92 | 122 | 111 | 123.927 | 131 | 130.748 | 127.5 | 119 | 127.5 | 120.7901 |  |  | 67
| 1 | 6 | 143.5 | 119.8468 | 106.5 | 117 | 121 | 133 | 119 | 127.04 | 124 | 112 | 123.4542 | 134.7 | 122.9435 | 131.5 | 118.5 | 113 | 118.5518
| Mean |  | 142.6833 | 119.8667 | 105.45 | 117.5833 | 118.5 | 131.5 | 116.8333 | 127.9183 | 124.1667 | 113.8333 | 121.9024 | 134.45 | 126.6769 | 126.9167 | 119 | 123.3333 | 116.3525
| Median |  | 143.1 | 119.9031 | 104.75 | 116 | 118 | 133 | 117 | 128.145 | 124.5 | 115 | 123.0499 | 134.95 | 126.4363 | 126.5 | 119.25 | 124.25 | 117.0172
| Std Dev. |  | 5.7052 | 0.41241 | 5.3287 | 6.0367 | 1.6432 | 3.5071 | 1.7224 | 1.4931 | 1.9408 | 8.2563 | 2.3722 | 1.7283 | 3.594 | 2.7096 | 1.6733 | 6.4859 | 7.6707
| Rel.Std.Dev. |  | 3.999% | 0.3441% | 5.053% | 5.134% | 1.387% | 2.667% | 1.474% | 1.167% | 1.563% | 7.253% | 1.946% | 1.285% | 2.837% | 2.135% | 1.406% | 5.259% | 6.593%
| PDM3 |  | 17.42% | -1.359% | -13.22% | -3.238% | -2.483% | 8.215% | -3.855% | 5.267% | 2.18% | -6.324% | 0.3165% | 10.64% | 4.246% | 4.443% | -2.072% | 1.494% | -4.251%
| Z-Score (Absolute) |  | 2.79 | 0.27 | 2.21 | 0.58 | 0.46 | 1.29 | 0.68 | 0.81 | 0.3 | 1.09 | 0 | 1.69 | 0.64 | 0.67 | 0.39 | 0.19 | 0.75
| Analytical results for Re in OREAS 780 (Indicative Value 0.002 ppm) | [20 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | < 0.003 | < 0.002 | 0.002 | 0.003 | 0.003 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | < 0.005 | 0.002 | < 0.002 | 0.002 | [6 empty] | 1
| 1 | 2 | 0.003 | < 0.002 | 0.002 | 0.003 | 0.004 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | 0.006 | < 0.002 | < 0.002 | 0.002 | [6 empty] | 16
| 1 | 3 | 0.004 | < 0.002 | 0.002 | 0.003 | 0.002 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | < 0.005 | 0.002 | < 0.002 | 0.002 | [6 empty] | 16
| 1 | 4 | 0.003 | < 0.002 | 0.002 | 0.003 | 0.003 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | < 0.005 | 0.002 | < 0.002 | < 0.002 | [6 empty] | 0.0024062
| 1 | 5 | 0.003 | < 0.002 | 0.002 | 0.003 | 0.003 | < 0.05 | 0.002 | < 0.1 | < 0.1 | 0.005 | 0.003 | < 0.002 | < 0.002 | [6 empty] | 22
| 1 | 6 | 0.003 | < 0.002 | 0.002 | 0.003 | 0.003 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | 0.005 | 0.002 | < 0.002 | 0.002
| Mean |  | 0.0032 |  | 0.002 | 0.003 | 0.003 |  | 0.002 |  |  | 0.0053333 | 0.0022 |  | 0.002
| Median |  | 0.003 |  | 0.002 | 0.003 | 0.003 |  | 0.002 |  |  | 0.005 | 0.002 |  | 0.002
| Std Dev. |  | 0.00044721 |  | 0 | 4.7507e-19 | 0.00063246 |  |  |  |  | 0.00057735 | 0.00044721 |  | 0
| Rel.Std.Dev. |  | 13.98% |  | 0% | 1.584e-14% | 21.08% |  |  |  |  | 10.83% | 20.33% |  | 0%
| PDM3 |  | 32.99% |  | -16.88% | 24.68% | 24.68% |  | -16.88% |  |  | 121.6% | -8.571% |  | -16.88%
| Z-Score (Absolute) |  | 0 | 1.03 | 0.49 | 0.04 | 0.04 | 11.91 | 0.94 | 25.4 | 25.4 | 0.54 | 0.49 | 1.03 | 0.67
| Analytical results for S in OREAS 780 (Certified Value 0.012 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0122 | 0.011507 | 0.02 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.0121 | < 0.02 | 0.02 | 0.011 | < 0.05 | 0.01 | 0.01 | 0.01 | < 0.1 |  |  | 1
| 1 | 2 | 0.0115 | 0.011033 | 0.01 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.0122 | < 0.02 | 0.015 | 0.01 | < 0.05 | 0.01 | 0.01 | 0.02 | < 0.1 |  |  | #N/A
| 1 | 3 | 0.0126 | 0.011161 | 0.01 | 0.03 | 0.01 | 0.01 | 0.02 | 0.013 | 0.0122 | < 0.02 | 0.015 | 0.012 | < 0.05 | 0.02 | 0.01 | 0.01 | < 0.1 |  |  | 16
| 1 | 4 | 0.0113 | 0.011317 | 0.02 | 0.02 | 0.01 | 0.02 | 0.02 | 0.013 | 0.0123 | < 0.02 | 0.015 | 0.008 | < 0.05 | 0.01 | 0.01 | 0.01 | < 0.1 |  |  | 0.012163
| 1 | 5 | 0.012 | 0.011705 | 0.01 | 0.02 | 0.01 | 0.02 | 0.02 | 0.013 | 0.0125 | < 0.02 | 0.015 | 0.011 | < 0.05 | 0.01 | 0.01 | 0.01 | < 0.1 |  |  | 68
| 1 | 6 | 0.011 | 0.011117 | 0.01 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.012 | < 0.02 | 0.015 | 0.011 | < 0.05 | 0.01 | 0.02 | 0.01 | < 0.1
| Mean |  | 0.011767 | 0.011307 | 0.013333 | 0.021667 | 0.01 | 0.013333 | 0.02 | 0.013 | 0.012217 |  | 0.015833 | 0.0105 |  | 0.011667 | 0.011667 | 0.011667
| Median |  | 0.01175 | 0.011239 | 0.01 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.0122 |  | 0.015 | 0.011 |  | 0.01 | 0.01 | 0.01
| Std Dev. |  | 0.00060222 | 0.00025729 | 0.005164 | 0.0040825 | 0 | 0.005164 | 0 | 0 | 0.00017224 |  | 0.0020412 | 0.0013784 |  | 0.0040825 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 5.118% | 2.276% | 38.73% | 18.84% | 0% | 38.73% | 0% | 0% | 1.41% |  | 12.89% | 13.13% |  | 34.99% | 34.99% | 34.99%
| PDM3 |  | -3.259% | -7.041% | 9.622% | 78.14% | -17.78% | 9.622% | 64.43% | 6.881% | 0.4408% |  | 30.18% | -13.67% |  | -4.081% | -4.081% | -4.081%
| Z-Score (Absolute) |  | 0.27 | 0.55 | 0.67 | 5.71 | 1.34 | 0.67 | 4.7 | 0.47 | 0 | 1.34 | 2.18 | 1.04 | 7.72 | 0.33 | 0.33 | 0.33 | 22.82
| Analytical results for Sb in OREAS 780 (Certified Value 5.58 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.71 | 4.9229 | 6.11 | 5.35 | < 5 | 4.88 | 5.7 | 5.4 | 5.94 | 5.0285 | 5.9 | 6.63 | 5.5571 | 6.1 | 5.5547 | 5.84 | 5.45 | 5.19 | 5.3618 | 1
| 1 | 2 | 5.21 | 5.1572 | 6.02 | 5.41 | < 5 | 4.81 | 6 | 5.6 | 6.02 | 5.2806 | 5.8 | 6.35 | 5.3362 | 5.7 | 5.6265 | 5.78 | 5.85 | 5.73 | 4.1844 | 10
| 1 | 3 | 5.01 | 4.8898 | 6.36 | 5.42 | < 5 | 4.74 | 5.9 | 5.4 | 5.97 | 5.1274 | 5.8 | 6.16 | 5.3961 | 5.7 | 5.5913 | 5.63 | 5.8 | 5.47 | 7.203 | 16
| 1 | 4 | 5.38 | 5.0121 | 6.11 | 5.61 | < 5 | 4.42 | 5.8 | 5.4 | 6.05 | 5.2757 | 5.7 | 6.35 | 5.6625 | 5.8 | 5.6472 | 5.79 | 5.83 | 5.47 | 4.9956 | 5.578
| 1 | 5 | 5.61 | 5.2318 | 6.38 | 5.41 | < 5 | 4.69 | 5.7 | 5.7 | 6.14 | 5.0305 | 5.7 | 6.23 | 5.5998 | 5.9 | 5.7457 | 5.71 | 5.74 | 5.4 | 5.6935 | 69
| 1 | 6 | 5.37 | 5.1802 | 6.4 | 5.57 | < 5 | 4.78 | 5.7 | 5.4 | 5.87 | 5.2717 | 5.8 | 6.17 | 5.651 | 5.9 | 5.679 | 5.76 | 5.72 | 5.26 | 4.6457
| Mean |  | 5.215 | 5.0657 | 6.23 | 5.4617 |  | 4.72 | 5.8 | 5.4833 | 5.9983 | 5.1691 | 5.7833 | 6.315 | 5.5338 | 5.85 | 5.6407 | 5.7517 | 5.7317 | 5.42 | 5.3473
| Median |  | 5.29 | 5.0846 | 6.235 | 5.415 |  | 4.76 | 5.75 | 5.4 | 5.995 | 5.1996 | 5.8 | 6.29 | 5.5784 | 5.85 | 5.6369 | 5.77 | 5.77 | 5.435 | 5.1787
| Std Dev. |  | 0.3176 | 0.14371 | 0.16805 | 0.10323 |  | 0.16037 | 0.12649 | 0.13292 | 0.093684 | 0.12249 | 0.075277 | 0.17547 | 0.13654 | 0.15166 | 0.067171 | 0.073052 | 0.14689 | 0.18974 | 1.0519
| Rel.Std.Dev. |  | 6.09% | 2.837% | 2.697% | 1.89% |  | 3.398% | 2.181% | 2.424% | 1.562% | 2.37% | 1.302% | 2.779% | 2.467% | 2.592% | 1.191% | 1.27% | 2.563% | 3.501% | 19.67%
| PDM3 |  | -6.507% | -9.184% | 11.69% | -2.085% |  | -15.38% | 3.981% | -1.697% | 7.536% | -7.33% | 3.682% | 13.21% | -0.7926% | 4.877% | 1.125% | 3.114% | 2.755% | -2.832% | -4.135%
| Z-Score (Absolute) |  | 0.81 | 1.19 | 1.76 | 0.18 | 7.68 | 2.06 | 0.67 | 0.13 | 1.18 | 0.92 | 0.63 | 1.98 | 0 | 0.8 | 0.27 | 0.55 | 0.5 | 0.29 | 0.47
| Analytical results for Sc in OREAS 780 (Certified Value 13.3 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 12.13 | 13.67 | 14.1 | 13.4 | 11.86 | 12 | 13 | 13.8 | 13.7193 | 13 | 13.9 | 14.4677 | 12.9 | 13.0103 | 14.1 | 12.6 | 13.3 | 11.9731 |  | 1
| 1 | 2 | 12.06 | 13.44 | 12.6 | 13.4 | 11.99 | 12 | 12 | 13.9 | 13.8485 | 13 | 13.5 | 13.4428 | 12.9 | 13.1063 | 13.8 | 13 | 14.8 | 13.5238 |  | 11
| 1 | 3 | 12 | 13.46 | 14 | 13.3 | 11.95 | 12 | 12 | 14 | 13.7132 | 13 | 13.8 | 13.6003 | 12.5 | 13.3107 | 14.4 | 13 | 13.6 | 14.5888 |  | 16
| 1 | 4 | 12.18 | 13.55 | 12.9 | 14.3 | 11.62 | 12 | 13 | 13.8 | 13.3888 | 13 | 12.6 | 14.0549 | 12.7 | 13.1068 | 14.5 | 13.2 | 14.2 | 16.141 |  | 13.3377
| 1 | 5 | 12.29 | 13.65 | 13.2 | 13.2 | 10.67 | 12 | 13 | 13.9 | 13.6825 | 13 | 13.2 | 14.1529 | 11.3 | 12.8237 | 14 | 12.8 | 14 | 14.6925 |  | 70
| 1 | 6 | 12.22 | 13.54 | 13.8 | 13.5 | 12.42 | 12 | 12 | 13.9 | 13.4592 | 13 | 13.1 | 14.143 | 12.4 | 11.6411 | 14.8 | 12.9 | 13.4 | 14.9298
| Mean |  | 12.1467 | 13.5517 | 13.4333 | 13.5167 | 11.7517 | 12 | 12.5 | 13.8833 | 13.6353 | 13 | 13.35 | 13.9769 | 12.45 | 12.8332 | 14.2667 | 12.9167 | 13.8833 | 14.3082
| Median |  | 12.155 | 13.545 | 13.5 | 13.4 | 11.905 | 12 | 12.5 | 13.9 | 13.6978 | 13 | 13.35 | 14.099 | 12.6 | 13.0583 | 14.25 | 12.95 | 13.8 | 14.6407
| Std Dev. |  | 0.10614 | 0.094534 | 0.62183 | 0.39707 | 0.59023 | 0 | 0.54772 | 0.075277 | 0.17469 | 0 | 0.48477 | 0.38279 | 0.59917 | 0.60496 | 0.36697 | 0.20412 | 0.56716 | 1.417
| Rel.Std.Dev. |  | 0.8739% | 0.6976% | 4.629% | 2.938% | 5.023% | 0% | 4.382% | 0.5422% | 1.281% | 0% | 3.631% | 2.739% | 4.813% | 4.714% | 2.572% | 1.58% | 4.085% | 9.904%
| PDM3 |  | -8.93% | 1.604% | 0.7173% | 1.342% | -11.89% | -10.03% | -6.28% | 4.091% | 2.231% | -2.532% | 0.09246% | 4.793% | -6.655% | -3.783% | 6.965% | -3.156% | 4.091% | 7.276%
| Z-Score (Absolute) |  | 2.01 | 0.05 | 0.12 | 0 | 2.59 | NA | NA | 0.54 | 0.17 | NA | 0.24 | 0.67 | 1.56 | 1 | 1.1 | 0.88 | 0.54 | 1.16
Results from laboratories 10, 12 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 780 (Certified Value 0.92 ppm) | [20 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
2
| 1 | 1 | 0.96 | 0.91488 | 2 | 1 | 2.16 | 0.9 | 1 | 0.6 | < 5 | < 5 | < 1 | 1.9006 | 1 | 1 | 1 | < 0.1 |  |  |  | 1
| 1 | 2 | 1 | 0.92533 | 2 | 1 | 2.1 | 0.9 | 1 | 0.6 | < 5 | < 5 | < 1 | 2.0987 | 1 | 1 | 1 | < 0.1 |  |  |  | 12
| 1 | 3 | 0.94 | 0.92461 | 2 | < 1 | 2.17 | 0.7 | 1 | 0.6 | < 5 | < 5 | < 1 | 1.8062 | 1 | 1 | 1 | < 0.1 |  |  |  | 16
| 1 | 4 | 1.09 | 0.92973 | 2 | 1 | 2.26 | 0.5 | 1 | 0.5 | < 5 | < 5 | < 1 | 1.5284 | 1 | 1 | 1 | < 0.1 |  |  |  | 0.91698
| 1 | 5 | 1.12 | 0.92546 | 2 | 1 | 2.27 | 0.8 | 1 | 0.6 | < 5 | < 5 | < 1 | 2.1632 | 1 | 1 | 1 | < 0.1 |  |  |  | 71
| 1 | 6 | 1.02 | 0.88696 | 3 | 2 | 2.18 | 0.6 | 1 | < 0.5 | < 5 | < 5 | < 1 | 1.6957 | 1 | 1 | 1 | < 0.1
| Mean |  | 1.0217 | 0.91783 | 2.1667 | 1.2 | 2.19 | 0.73333 | 1 | 0.58 |  |  |  | 1.8655 | 1 | 1 | 1
| Median |  | 1.01 | 0.92497 | 2 | 1 | 2.175 | 0.75 | 1 | 0.6 |  |  |  | 1.8534 | 1 | 1 | 1
| Std Dev. |  | 0.07111 | 0.0159 | 0.40825 | 0.44721 | 0.064498 | 0.1633 | 0 | 0.044721 |  |  |  | 0.24094 | 0 | 0 | 0
| Rel.Std.Dev. |  | 6.96% | 1.732% | 18.84% | 37.27% | 2.945% | 22.27% | 0% | 7.711% |  |  |  | 12.92% | 0% | 0% | 0%
| PDM3 |  | 11.42% | 0.09244% | 136.3% | 30.86% | 138.8% | -20.03% | 9.054% | -36.75% |  |  |  | 103.4% | 9.054% | 9.054% | 9.054%
| Z-Score (Absolute) |  | 0.04 | 0.15 | 2.12 | 0.15 | 2.16 | 0.48 | 0 | 0.86 | 2.73 | 2.73 | 0.91 | 1.57 | 0 | 0 | 0 | 1.73
| Analytical results for Sm in OREAS 780 (Certified Value 6.25 ppm) | [20 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 5.46 | 6.6882 | 6.04 | 6.1 | 6.4 | 5.9 | 6.3 | 6.6756 | 6.2908 | 6.1345 | [9 empty] | 1
| 1 | 2 | 5.83 | 6.4761 | 6.06 | 6.3 | 6.3 | 6.3 | 6.2 | 6.2956 | 6.3254 | 5.8372 | [9 empty] | 28
| 1 | 3 | 5.73 | 6.7038 | 6.4 | 6.3 | 6.39 | 6.1 | 6.2 | 6.4312 | 6.1279 | 6.0619 | [9 empty] | 16
| 1 | 4 | 5.47 | 6.6099 | 6.31 | 6.1 | 6.39 | 6 | 5.8 | 6.5318 | 6.0969 | 6.6012 | [9 empty] | 6.2457
| 1 | 5 | 6 | 6.7211 | 5.81 | 6.4 | 6.59 | 6.15 | 6.1 | 6.64 | 6.2385 | 6.4234 | [9 empty] | 72
| 1 | 6 | 6 | 6.6779 | 6.28 | 6.7 | 6.54 | 6.05 | 5.9 | 6.5041 | 5.4936 | 5.9502
| Mean |  | 5.7483 | 6.6462 | 6.15 | 6.3167 | 6.435 | 6.0833 | 6.0833 | 6.5131 | 6.0955 | 6.1681
| Median |  | 5.78 | 6.683 | 6.17 | 6.3 | 6.395 | 6.075 | 6.15 | 6.518 | 6.1832 | 6.0982
| Std Dev. |  | 0.2426 | 0.091593 | 0.21945 | 0.22286 | 0.10821 | 0.13663 | 0.19408 | 0.13925 | 0.30815 | 0.29056
| Rel.Std.Dev. |  | 4.22% | 1.378% | 3.568% | 3.528% | 1.682% | 2.246% | 3.19% | 2.138% | 5.055% | 4.711%
| PDM3 |  | -7.963% | 6.413% | -1.532% | 1.137% | 3.031% | -2.599% | -2.599% | 4.281% | -2.404% | -1.243%
| Z-Score (Absolute) |  | 2.37 | 2.82 | 0.05 | 0.91 | 1.6 | 0.44 | 0.44 | 2.05 | 0.37 | 0.05
| Analytical results for Sn in OREAS 780 (Certified Value 14.8 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 12.45 | 17.5 | 15.6 | < 5 | 12.8 | 15 | 14.4 | 15.8 | 14.7582 | 14.1 | 14.7 | 12.3552 | 14.5 | 16.0636 | 14.6 | 14.1 | 13.9 | 10.8106 |  | 1
| 1 | 2 | 13.95 | 17.2 | 15.5 | < 5 | 12.62 | 16 | 14.5 | 15.4 | 14.9708 | 14.8 | 14.5 | 14.5695 | 14.4 | 16.3232 | 14.7 | 14.4 | 15.6 | 10.7891 |  | 2
| 1 | 3 | 13.54 | 17.4 | 15.7 | < 5 | 12.44 | 15 | 14.6 | 15.8 | 15.0182 | 13.7 | 14.6 | 11.9944 | 14.4 | 16.1774 | 14.6 | 14.5 | 14.4 | 10.8747 |  | 16
| 1 | 4 | 14.34 | 16.8 | 15.8 | < 5 | 12.01 | 16 | 14.6 | 15.3 | 15.0087 | 14.1 | 11.8 | 15.9738 | 14.5 | 16.3268 | 14.4 | 14.8 | 14.9 | 11.4996 |  | 14.8346
| 1 | 5 | 14.58 | 17.2 | 15.4 | < 5 | 12.25 | 15 | 14.8 | 15.8 | 14.2245 | 13.8 | 14.5 | 14.4087 | 14.5 | 16.8585 | 14.4 | 14.6 | 14.8 | 11.0672 |  | 73
| 1 | 6 | 14.02 | 17.5 | 16 | < 5 | 11.89 | 16 | 14.6 | 15.4 | 15.8095 | 13.9 | 14.3 | 14.411 | 14.8 | 16.301 | 14.9 | 14.7 | 14.4 | 11.3437
| Mean |  | 13.8133 | 17.2667 | 15.6667 |  | 12.335 | 15.5 | 14.5833 | 15.5833 | 14.965 | 14.0667 | 14.0667 | 13.9521 | 14.5167 | 16.3417 | 14.6 | 14.5167 | 14.6667 | 11.0641
| Median |  | 13.985 | 17.3 | 15.65 |  | 12.345 | 15.5 | 14.6 | 15.6 | 14.9897 | 14 | 14.5 | 14.4099 | 14.5 | 16.3121 | 14.6 | 14.55 | 14.6 | 10.971
| Std Dev. |  | 0.7563 | 0.26583 | 0.21602 |  | 0.35195 | 0.54772 | 0.13292 | 0.24014 | 0.51168 | 0.39328 | 1.1183 | 1.5014 | 0.1472 | 0.27335 | 0.18974 | 0.24833 | 0.5785 | 0.29786
| Rel.Std.Dev. |  | 5.475% | 1.54% | 1.379% |  | 2.853% | 3.534% | 0.9114% | 1.541% | 3.419% | 2.796% | 7.95% | 10.76% | 1.014% | 1.673% | 1.3% | 1.711% | 3.944% | 2.692%
| PDM3 |  | -6.885% | 16.39% | 5.609% |  | -16.85% | 4.485% | -1.694% | 5.047% | 0.8787% | -5.177% | -5.177% | -5.949% | -2.143% | 10.16% | -1.582% | -2.143% | -1.132% | -25.42%
| Z-Score (Absolute) |  | 0.84 | 3.28 | 1.37 | 14.35 | 2.61 | NA | 0.08 | 1.27 | 0.54 | 0.54 | 0.54 | 0.67 | 0 | 2.18 | 0.1 | 0 | 0.18 | 4.12
Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 780 (Certified Value 63 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 64 | 61.143 | 68.9 | 62.7 | 63.7 | 59 | 62 | 64.5 | 64 | 63.5505 | 58 | 63.2 | 61.1081 | 63 | 62.7725 | 65.2 | 63.9 | 60.4 | 60.791 | 1
| 1 | 2 | 64 | 62.2735 | 65.6 | 63.8 | 63.99 | 59 | 62.4 | 62.3 | 65.06 | 64.9134 | 62.5 | 60.9 | 61.4601 | 65 | 60.0269 | 64.1 | 65.9 | 67.1 | 59.4036 | 3
| 1 | 3 | 63 | 61.364 | 68.3 | 64.4 | 63.56 | 61 | 62.9 | 64.5 | 66.04 | 63.9084 | 60.5 | 62.6 | 60.7891 | 63 | 61.7764 | 64.7 | 65.1 | 61.8 | 61.2289 | 16
| 1 | 4 | 64 | 61.782 | 65.9 | 66.1 | 62.14 | 59 | 64.1 | 65 | 66.03 | 63.0065 | 61.5 | 57.2 | 62.9451 | 65 | 60.5444 | 64.1 | 66.3 | 64 | 59.4678 | 62.9178
| 1 | 5 | 65 | 62.2045 | 68.4 | 64.4 | 61.07 | 56 | 64.1 | 62.5 | 65.57 | 62.6231 | 59.5 | 59.9 | 61.9771 | 62 | 62.1432 | 63.3 | 64.9 | 63.3 | 60.6081 | 74
| 1 | 6 | 64 | 62.2545 | 68.6 | 65.4 | 62.17 | 58 | 63.6 | 63.9 | 64.6 | 65.7123 | 61 | 59.4 | 59.5021 | 62 | 58.8176 | 64.8 | 65 | 62.9 | 61.3054
| Mean |  | 64 | 61.8369 | 67.6167 | 64.4667 | 62.7717 | 58.6667 | 63.1833 | 63.7833 | 65.2167 | 63.9524 | 60.5 | 60.5333 | 61.2969 | 63.3333 | 61.0135 | 64.3667 | 65.1833 | 63.25 | 60.4675
| Median |  | 64 | 61.9932 | 68.35 | 64.4 | 62.865 | 59 | 63.25 | 64.2 | 65.315 | 63.7294 | 60.75 | 60.4 | 61.2841 | 63 | 61.1604 | 64.4 | 65.05 | 63.1 | 60.6996
| Std Dev. |  | 0.63246 | 0.49159 | 1.4634 | 1.1928 | 1.151 | 1.633 | 0.88863 | 1.1286 | 0.81845 | 1.1702 | 1.5811 | 2.2051 | 1.1594 | 1.3663 | 1.481 | 0.67429 | 0.84004 | 2.2687 | 0.84113
| Rel.Std.Dev. |  | 0.9882% | 0.795% | 2.164% | 1.85% | 1.834% | 2.784% | 1.406% | 1.769% | 1.255% | 1.83% | 2.613% | 3.643% | 1.891% | 2.157% | 2.427% | 1.048% | 1.289% | 3.587% | 1.391%
| PDM3 |  | 1.72% | -1.718% | 7.468% | 2.462% | -0.2323% | -6.757% | 0.422% | 1.376% | 3.654% | 1.644% | -3.843% | -3.79% | -2.576% | 0.6604% | -3.027% | 2.303% | 3.601% | 0.5279% | -3.895%
| Z-Score (Absolute) |  | 0.36 | 0.67 | 2.08 | 0.58 | 0.23 | 2.19 | 0.03 | 0.25 | 0.94 | 0.34 | 1.31 | 1.3 | 0.93 | 0.04 | 1.07 | 0.53 | 0.92 | 0 | 1.33
| Analytical results for Ta in OREAS 780 (Certified Value 0.89 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 0.37 | 0.93 | 0.83 | 0.8 | 0.8 | 0.6 | 0.81 | 0.75 | 0.91 | 1.012 | 0.9 | 1.0705 | 0.78 | 0.81 | 0.9 |  |  |  |  | 1
| 1 | 2 | 0.42 | 0.92 | 0.84 | 0.82 | 0.9 | 0.58 | 0.82 | 0.75 | 0.93 | 1.052 | 0.8 | 1.0773 | 0.82 | 0.86 | 0.97 |  |  |  |  | 4
| 1 | 3 | 0.38 | 0.9 | 0.85 | 0.81 | 0.9 | 0.61 | 0.92 | 0.7 | 0.95 | 0.99979 | 0.9 | 1.0596 | 0.78 | 0.87 | 0.88 |  |  |  |  | 16
| 1 | 4 | 0.56 | 0.92 | 0.87 | 0.86 | 0.9 | 0.61 | 0.82 | 0.75 | 0.9 | 1.0064 | 0.9 | 1.0658 | 0.79 | 0.86 | 0.92 |  |  |  |  | 0.88834
| 1 | 5 | 0.46 | 0.92 | 0.84 | 0.84 | 0.9 | 0.58 | 0.86 | 0.75 | 0.95 | 1.0111 | 0.9 | 1.0872 | 0.8 | 0.85 | 0.88 |  |  |  |  | 75
| 1 | 6 | 0.46 | 0.88 | 0.85 | 0.88 | 0.9 | 0.62 | 0.87 | 0.75 | 0.96 | 0.95361 | 0.9 | 1.0852 | 0.82 | 0.88 | 0.92
| Mean |  | 0.44167 | 0.91167 | 0.84667 | 0.835 | 0.88333 | 0.6 | 0.85 | 0.74167 | 0.93333 | 1.0058 | 0.88333 | 1.0743 | 0.79833 | 0.855 | 0.91167
| Median |  | 0.44 | 0.92 | 0.845 | 0.83 | 0.9 | 0.605 | 0.84 | 0.75 | 0.94 | 1.0087 | 0.9 | 1.0739 | 0.795 | 0.86 | 0.91
| Std Dev. |  | 0.069402 | 0.018348 | 0.013663 | 0.030822 | 0.040825 | 0.016733 | 0.041952 | 0.020412 | 0.024221 | 0.031512 | 0.040825 | 0.010915 | 0.018348 | 0.02429 | 0.033714
| Rel.Std.Dev. |  | 15.71% | 2.013% | 1.614% | 3.691% | 4.622% | 2.789% | 4.936% | 2.752% | 2.595% | 3.133% | 4.622% | 1.016% | 2.298% | 2.841% | 3.698%
| PDM3 |  | -50.28% | 2.626% | -4.691% | -6.005% | -0.5639% | -32.46% | -4.316% | -16.51% | 5.065% | 13.22% | -0.5639% | 20.93% | -10.13% | -3.753% | 2.626%
| Z-Score (Absolute) |  | 4.47 | 0.67 | 0.04 | 0.16 | NA | 2.73 | 0 | 1.18 | 0.91 | 1.7 | NA | 2.45 | 0.56 | 0.05 | 0.67
Results from laboratories 10 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 780 (Certified Value 0.66 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | < 0.2 | 0.98695 | 0.59 | 0.6 | 0.67 | 0.66 | 0.67 | 0.73187 | 0.83536 | 0.5212 | [9 empty] | 1
| 1 | 2 | < 0.2 | 0.94732 | 0.61 | 0.7 | 0.68 | 0.66 | 0.73 | 0.68115 | 0.81517 | 0.4823 | [9 empty] | 20
| 1 | 3 | < 0.2 | 0.96703 | 0.61 | 0.7 | 0.69 | 0.66 | 0.68 | 0.67218 | 0.7967 | 0.5158 | [9 empty] | 16
| 1 | 4 | < 0.2 | 0.97992 | 0.64 | 0.7 | 0.67 | 0.66 | 0.63 | 0.72071 | 0.79885 | 0.5092 | [9 empty] | 0.66301
| 1 | 5 | < 0.2 | 0.95184 | 0.59 | 0.7 | 0.68 | 0.64 | 0.65 | 0.70709 | 0.82173 | 0.5413 | [9 empty] | 76
| 1 | 6 | < 0.2 | 0.96934 | 0.63 | 0.7 | 0.68 | 0.66 | 0.64 | 0.70137 | 0.69699 | 0.5007
| Mean |  |  | 0.96707 | 0.61167 | 0.68333 | 0.67833 | 0.65667 | 0.66667 | 0.70239 | 0.79413 | 0.51175
| Median |  |  | 0.96819 | 0.61 | 0.7 | 0.68 | 0.66 | 0.66 | 0.70423 | 0.80701 | 0.5125
| Std Dev. |  |  | 0.015412 | 0.020412 | 0.040825 | 0.0075277 | 0.008165 | 0.036148 | 0.022767 | 0.049738 | 0.019874
| Rel.Std.Dev. |  |  | 1.594% | 3.337% | 5.974% | 1.11% | 1.243% | 5.422% | 3.241% | 6.263% | 3.883%
| PDM3 |  |  | 45.86% | -7.743% | 3.066% | 2.312% | -0.9561% | 0.5522% | 5.941% | 19.78% | -22.81%
| Z-Score (Absolute) |  | 6.95 | 3.68 | 0.67 | NA | 0.14 | 0.12 | 0 | 0.44 | 1.56 | 1.9
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 780 (Certified Value 0.23 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.18 | 0.28 | 0.22 | 0.48 | 0.2 | 0.2 | 0.2 | < 5 | 0.2 | 0.2 | 0.36804 | 0.3 | 0.23788 | 0.27 | 0.23 | 0.22 |  |  |  | 1
| 1 | 2 | 0.14 | 0.18 | 0.23 | 0.49 | 0.3 | 0.2 | 0.2 | < 5 | 0.2 | 0.23 | 0.29104 | 0.3 | 0.25166 | 0.24 | 0.2 | 0.24 |  |  |  | 5
| 1 | 3 | 0.32 | 0.19 | 0.24 | 0.51 | 0.3 | 0.2 | 0.2 | < 5 | 0.2 | 0.23 | 0.19486 | 0.4 | 0.24991 | 0.28 | 0.17 | 0.25 |  |  |  | 16
| 1 | 4 | 0.19 | 0.25 | 0.22 | 0.5 | 0.3 | 0.2 | 0.2 | < 5 | 0.3 | 0.24 | 0.29336 | 0.2 | 0.23799 | 0.27 | 0.21 | 0.22 |  |  |  | 0.23328
| 1 | 5 | 0.27 | 0.26 | 0.22 | 0.48 | 0.2 | 0.2 | 0.2 | < 5 | 0.3 | 0.22 | 0.27152 | 0.4 | 0.2499 | 0.26 | 0.22 | 0.25 |  |  |  | 77
| 1 | 6 | 0.23 | 0.22 | 0.23 | 0.5 | 0.2 | 0.2 | 0.3 | < 5 | 0.3 | 0.22 | 0.32812 | 0.3 | 0.27013 | 0.28 | 0.24 | 0.24
| Mean |  | 0.22167 | 0.23 | 0.22667 | 0.49333 | 0.25 | 0.2 | 0.21667 |  | 0.25 | 0.22333 | 0.29116 | 0.31667 | 0.24958 | 0.26667 | 0.21167 | 0.23667
| Median |  | 0.21 | 0.235 | 0.225 | 0.495 | 0.25 | 0.2 | 0.2 |  | 0.25 | 0.225 | 0.2922 | 0.3 | 0.24991 | 0.27 | 0.215 | 0.24
| Std Dev. |  | 0.065549 | 0.04 | 0.008165 | 0.012111 | 0.054772 | 3.0405e-17 | 0.040825 |  | 0.054772 | 0.013663 | 0.058208 | 0.075277 | 0.011816 | 0.015055 | 0.024833 | 0.013663
| Rel.Std.Dev. |  | 29.57% | 17.39% | 3.602% | 2.455% | 21.91% | 1.52e-14% | 18.84% |  | 21.91% | 6.118% | 19.99% | 23.77% | 4.734% | 5.646% | 11.73% | 5.773%
| PDM3 |  | -4.979% | -1.406% | -2.835% | 111.5% | 7.167% | -14.27% | -7.122% |  | 7.167% | -4.264% | 24.81% | 35.74% | 6.986% | 14.31% | -9.265% | 1.451%
| Z-Score (Absolute) |  | 0.67 | 0.3 | 0.45 | 11.54 | NA | NA | NA | 101.75 | NA | 0.6 | 2.45 | NA | 0.58 | 1.35 | 1.12 | 0
Results from laboratories 10, 12, 13, 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 780 (Certified Value 15.3 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 15.33 | 13.95 | 14.5 | 12.69 | 15.8 | 14.6 | 15.14 | 14.1 | 16.31 | 15.4388 | 15.4 | 14.9636 | 15.95 | 14.9 | 15 | 14.9765 |  |  |  | 1
| 1 | 2 | 16.14 | 12.3 | 14.15 | 12.61 | 15.9 | 14.7 | 15.29 | 13.1 | 15.97 | 14.5339 | 15.5 | 14.825 | 15.9 | 14.9 | 16.05 | 15.8217 |  |  |  | 6
| 1 | 3 | 15.98 | 13.65 | 14.6 | 13.18 | 15.6 | 14.9 | 15.35 | 14.5 | 16.37 | 14.8772 | 15.3 | 14.8138 | 16.35 | 16.15 | 14.85 | 15.2757 |  |  |  | 16
| 1 | 4 | 15.37 | 13 | 14.8 | 12.48 | 15.6 | 14.8 | 15.23 | 14.3 | 16.9 | 15.2282 | 15.4 | 14.9337 | 16 | 16.05 | 15.6 | 14.2126 |  |  |  | 15.2727
| 1 | 5 | 16.53 | 12.45 | 14.5 | 12.36 | 15.9 | 14.5 | 15.14 | 13.9 | 16.66 | 15.3702 | 14.6 | 14.9734 | 15.8 | 15.15 | 15.3 | 15.0256 |  |  |  | 78
| 1 | 6 | 16.79 | 13.3 | 14.5 | 12.52 | 16.2 | 14.7 | 15.24 | 14.3 | 15.93 | 14.9888 | 15.4 | 13.4019 | 16.55 | 15.15 | 14.4 | 14.8548
| Mean |  | 16.0233 | 13.1083 | 14.5083 | 12.64 | 15.8333 | 14.7 | 15.2317 | 14.0333 | 16.3567 | 15.0728 | 15.2667 | 14.6519 | 16.0917 | 15.3833 | 15.2 | 15.0278
| Median |  | 16.06 | 13.15 | 14.5 | 12.565 | 15.85 | 14.7 | 15.235 | 14.2 | 16.34 | 15.1085 | 15.4 | 14.8794 | 15.975 | 15.15 | 15.15 | 15.001
| Std Dev. |  | 0.59456 | 0.65377 | 0.21075 | 0.28754 | 0.22509 | 0.14142 | 0.082805 | 0.50067 | 0.37977 | 0.34126 | 0.33267 | 0.61623 | 0.29226 | 0.56716 | 0.58224 | 0.52704
| Rel.Std.Dev. |  | 3.711% | 4.987% | 1.453% | 2.275% | 1.422% | 0.9621% | 0.5436% | 3.568% | 2.322% | 2.264% | 2.179% | 4.206% | 1.816% | 3.687% | 3.831% | 3.507%
| PDM3 |  | 4.915% | -14.17% | -5.005% | -17.24% | 3.671% | -3.75% | -0.2686% | -8.115% | 7.098% | -1.309% | -0.03939% | -4.065% | 5.362% | 0.7245% | -0.4759% | -1.603%
| Z-Score (Absolute) |  | 1.08 | 2.46 | 0.76 | 3.03 | 0.84 | 0.53 | 0.12 | 1.34 | 1.48 | 0.08 | 0.16 | 0.59 | 1.16 | 0.3 | 0.08 | 0.13
| Analytical results for Ti in OREAS 780 (Certified Value 0.347 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.1641 | 0.43316 | 0.368 | 0.334 | 0.36 | 0.25 | 0.323 | 0.288 | 0.3226 | 0.34139 | 0.338 | 0.375 | 0.41537 | 0.292 | 0.39463 | 0.333 | 0.356 | 0.365 | 0.28 | 1
| 1 | 2 | 0.1713 | 0.43909 | 0.357 | 0.335 | 0.35 | 0.26 | 0.341 | 0.288 | 0.3332 | 0.34087 | 0.335 | 0.377 | 0.41002 | 0.318 | 0.39948 | 0.342 | 0.357 | 0.391 | 0.2622 | 22
| 1 | 3 | 0.1571 | 0.42723 | 0.367 | 0.339 | 0.36 | 0.26 | 0.345 | 0.302 | 0.3602 | 0.34131 | 0.323 | 0.381 | 0.40976 | 0.308 | 0.41421 | 0.334 | 0.359 | 0.36 | 0.2742 | 16
| 1 | 4 | 0.1797 | 0.43316 | 0.354 | 0.349 | 0.36 | 0.26 | 0.359 | 0.289 | 0.3278 | 0.34832 | 0.345 | 0.384 | 0.41827 | 0.315 | 0.40028 | 0.338 | 0.363 | 0.373 | 0.2862 | 0.34676
| 1 | 5 | 0.175 | 0.43909 | 0.36 | 0.329 | 0.35 | 0.24 | 0.331 | 0.311 | 0.3407 | 0.33665 | 0.336 | 0.381 | 0.41609 | 0.3 | 0.41363 | 0.342 | 0.352 | 0.371 | 0.2726 | 79
| 1 | 6 | 0.1799 | 0.42723 | 0.361 | 0.34 | 0.36 | 0.24 | 0.34 | 0.302 | 0.34 | 0.34686 | 0.347 | 0.378 | 0.40507 | 0.302 | 0.40865 | 0.348 | 0.365 | 0.369 | 0.2338
| Mean |  | 0.17118 | 0.43316 | 0.36117 | 0.33767 | 0.35667 | 0.25167 | 0.33983 | 0.29667 | 0.33742 | 0.34257 | 0.33733 | 0.37933 | 0.41243 | 0.30583 | 0.40515 | 0.3395 | 0.35867 | 0.3715 | 0.26817
| Median |  | 0.17315 | 0.43316 | 0.3605 | 0.337 | 0.36 | 0.255 | 0.3405 | 0.2955 | 0.3366 | 0.34135 | 0.337 | 0.3795 | 0.4127 | 0.305 | 0.40447 | 0.34 | 0.358 | 0.37 | 0.2734
| Std Dev. |  | 0.0090731 | 0.0053073 | 0.0054924 | 0.006802 | 0.005164 | 0.0098319 | 0.012303 | 0.0097091 | 0.013163 | 0.0042988 | 0.0085479 | 0.003266 | 0.0049649 | 0.0097656 | 0.0081559 | 0.005648 | 0.004761 | 0.010616 | 0.018644
| Rel.Std.Dev. |  | 5.3% | 1.225% | 1.521% | 2.014% | 1.448% | 3.907% | 3.62% | 3.273% | 3.901% | 1.255% | 2.534% | 0.861% | 1.204% | 3.193% | 2.013% | 1.664% | 1.327% | 2.858% | 6.952%
| PDM3 |  | -50.63% | 24.92% | 4.156% | -2.621% | 2.858% | -27.42% | -1.996% | -14.44% | -2.693% | -1.208% | -2.717% | 9.395% | 18.94% | -11.8% | 16.84% | -2.092% | 3.435% | 7.136% | -22.66%
| Z-Score (Absolute) |  | 3.59 | 1.99 | 0.45 | 0.05 | 0.36 | 1.88 | 0 | 0.92 | 0.05 | 0.06 | 0.05 | 0.84 | 1.55 | 0.72 | 1.39 | 0.01 | 0.4 | 0.67 | 1.53
| Analytical results for Tl in OREAS 780 (Certified Value 0.69 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.626 | 0.76 | 0.71 | 0.63 | 0.7 | 0.71 | 0.75 | 0.7 | 0.84 | 0.64678 | 0.8 | 0.7288 | 0.69 | 0.61 | 0.68 | 0.5787 |  |  |  | 1
| 1 | 2 | 0.674 | 0.74 | 0.7 | 0.64 | 0.71 | 0.71 | 0.73 | 0.7 | 0.82 | 0.62024 | 0.7 | 0.73019 | 0.66 | 0.62 | 0.72 | 0.5621 |  |  |  | 7
| 1 | 3 | 0.695 | 0.79 | 0.71 | 0.61 | 0.71 | 0.7 | 0.75 | 0.7 | 0.86 | 0.61482 | 0.8 | 0.71242 | 0.67 | 0.64 | 0.66 | 0.694 |  |  |  | 16
| 1 | 4 | 0.653 | 0.75 | 0.72 | 0.61 | 0.74 | 0.72 | 0.75 | 0.7 | 0.81 | 0.63648 | 0.7 | 0.72874 | 0.66 | 0.65 | 0.68 | 0.6148 |  |  |  | 0.68671
| 1 | 5 | 0.689 | 0.75 | 0.73 | 0.6 | 0.72 | 0.71 | 0.76 | 0.7 | 0.84 | 0.64507 | 0.7 | 0.7176 | 0.67 | 0.63 | 0.67 | 0.7545 |  |  |  | 80
| 1 | 6 | 0.68 | 0.75 | 0.73 | 0.59 | 0.72 | 0.72 | 0.74 | 0.7 | 0.83 | 0.63564 | 0.7 | 0.70902 | 0.7 | 0.61 | 0.69 | 0.7768
| Mean |  | 0.6695 | 0.75667 | 0.71667 | 0.61333 | 0.71667 | 0.71167 | 0.74667 | 0.7 | 0.83333 | 0.63317 | 0.73333 | 0.72113 | 0.675 | 0.62667 | 0.68333 | 0.66348
| Median |  | 0.677 | 0.75 | 0.715 | 0.61 | 0.715 | 0.71 | 0.75 | 0.7 | 0.835 | 0.63606 | 0.7 | 0.72317 | 0.67 | 0.625 | 0.68 | 0.6544
| Std Dev. |  | 0.02579 | 0.017512 | 0.012111 | 0.018619 | 0.013663 | 0.0075277 | 0.010328 | 1.2162e-16 | 0.017512 | 0.013025 | 0.05164 | 0.0093145 | 0.016432 | 0.01633 | 0.020656 | 0.091535
| Rel.Std.Dev. |  | 3.852% | 2.314% | 1.69% | 3.036% | 1.906% | 1.058% | 1.383% | 1.737e-14% | 2.101% | 2.057% | 7.042% | 1.292% | 2.434% | 2.606% | 3.023% | 13.8%
| PDM3 |  | -2.507% | 10.19% | 4.362% | -10.69% | 4.362% | 3.634% | 8.73% | 1.935% | 21.35% | -7.797% | 6.789% | 5.011% | -1.706% | -8.744% | -0.4923% | -3.383%
| Z-Score (Absolute) |  | 0.61 | 1.29 | 0.42 | 1.83 | 0.42 | 0.31 | 1.07 | NA | 2.95 | 1.4 | NA | 0.51 | 0.49 | 1.54 | 0.31 | 0.74
Results from laboratories 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 780 (Certified Value 0.22 ppm) | [20 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [10 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.16 | 0.18 | 0.2 | 0.24 | 0.22 | 0.24 | 0.28236 | 0.41717 | 0.1988 | [10 empty] | 1
| 1 | 2 | 0.17 | 0.18 | 0.2 | 0.25 | 0.21 | 0.25 | 0.25528 | 0.42257 | 0.1945 | [10 empty] | 24
| 1 | 3 | 0.17 | 0.2 | 0.2 | 0.25 | 0.21 | 0.25 | 0.25623 | 0.41752 | 0.1987 | [10 empty] | 16
| 1 | 4 | 0.17 | 0.18 | 0.2 | 0.24 | 0.2 | 0.25 | 0.28082 | 0.4207 | 0.2281 | [10 empty] | 0.21886
| 1 | 5 | 0.18 | 0.17 | 0.2 | 0.25 | 0.21 | 0.25 | 0.28477 | 0.41406 | 0.2025 | [10 empty] | 81
| 1 | 6 | 0.18 | 0.18 | 0.2 | 0.25 | 0.22 | 0.24 | 0.34827 | 0.35867 | 0.1882
| Mean |  | 0.17167 | 0.18167 | 0.2 | 0.24667 | 0.21167 | 0.24667 | 0.28462 | 0.40845 | 0.2018
| Median |  | 0.17 | 0.18 | 0.2 | 0.25 | 0.21 | 0.25 | 0.28159 | 0.41734 | 0.19875
| Std Dev. |  | 0.0075277 | 0.0098319 | 3.0405e-17 | 0.005164 | 0.0075277 | 0.005164 | 0.033876 | 0.024565 | 0.013777
| Rel.Std.Dev. |  | 4.385% | 5.412% | 1.52e-14% | 2.094% | 3.556% | 2.094% | 11.9% | 6.014% | 6.827%
| PDM3 |  | -21.56% | -16.99% | -8.618% | 12.7% | -3.287% | 12.7% | 30.05% | 86.62% | -7.795%
| Z-Score (Absolute) |  | 1.04 | 0.86 | NA | 0.32 | 0.32 | 0.32 | 1 | 3.23 | 0.49
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 780 (Certified Value 2.49 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.21 | 2.0026 | 2.4 | 2.5 | 2.35 | 2.8 | 2.45 | 2.59 | < 50 | 2.6 | 2.66 | 2.5336 | 2.6 | 2.8012 | 2.4 | 2.5 | 2.5 | 2.319 |  | 1
| 1 | 2 | 2.32 | 1.9038 | 2.2 | 2.4 | 2.45 | 2.7 | 2.45 | 2.56 | < 50 | 2.3 | 2.6 | 2.6337 | 2.6 | 2.7356 | 2.4 | 2.5 | 2.6 | 2.3787 |  | 8
| 1 | 3 | 2.35 | 1.909 | 2.2 | 2.5 | 2.42 | 2.6 | 2.48 | 2.93 | < 50 | 2.5 | 2.59 | 2.3571 | 2.5 | 3.016 | 2.4 | 2.6 | 2.3 | 2.2807 |  | 16
| 1 | 4 | 2.3 | 1.9292 | 2.2 | 2.6 | 2.37 | 2.7 | 2.45 | 3.41 | < 50 | 2.4 | 2.36 | 2.5001 | 2.6 | 2.7426 | 2.4 | 2.6 | 2.4 | 2.915 |  | 2.4915
| 1 | 5 | 2.41 | 1.9388 | 2.3 | 2.5 | 2.24 | 2.7 | 2.41 | 2.63 | < 50 | 2.4 | 2.56 | 2.5815 | 2.6 | 2.8067 | 2.4 | 2.6 | 2.4 | 2.5085 |  | 82
| 1 | 6 | 2.42 | 1.9449 | 2.2 | 2.5 | 2.25 | 2.7 | 2.45 | 2.64 | < 50 | 2.5 | 2.54 | 2.5306 | 2.6 | 2.6651 | 2.5 | 2.5 | 2.3 | 2.3763
| Mean |  | 2.335 | 1.938 | 2.25 | 2.5 | 2.3467 | 2.7 | 2.4483 | 2.7933 |  | 2.45 | 2.5517 | 2.5227 | 2.5833 | 2.7945 | 2.4167 | 2.55 | 2.4167 | 2.463
| Median |  | 2.335 | 1.934 | 2.2 | 2.5 | 2.36 | 2.7 | 2.45 | 2.635 |  | 2.45 | 2.575 | 2.5321 | 2.6 | 2.7719 | 2.4 | 2.55 | 2.4 | 2.3775
| Std Dev. |  | 0.077653 | 0.035495 | 0.083666 | 0.063246 | 0.08641 | 0.063246 | 0.022286 | 0.33013 |  | 0.10488 | 0.10245 | 0.09371 | 0.040825 | 0.12014 | 0.040825 | 0.054772 | 0.1169 | 0.23451
| Rel.Std.Dev. |  | 3.326% | 1.831% | 3.718% | 2.53% | 3.682% | 2.342% | 0.9103% | 11.82% |  | 4.281% | 4.015% | 3.715% | 1.58% | 4.299% | 1.689% | 2.148% | 4.837% | 9.521%
| PDM3 |  | -6.281% | -22.21% | -9.693% | 0.3413% | -5.813% | 8.369% | -1.732% | 12.11% |  | -1.666% | 2.415% | 1.254% | 3.686% | 12.16% | -3.003% | 2.348% | -3.003% | -1.142%
| Z-Score (Absolute) |  | 1.15 | 4.26 | 1.82 | 0.14 | 1.06 | 1.71 | 0.26 | 2.45 | 176.6 | 0.25 | 0.55 | 0.32 | 0.8 | 2.45 | 0.51 | 0.54 | 0.51 | 0.14
| Analytical results for V in OREAS 780 (Certified Value 91 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 81 | 94 | 89 | 85.51 | 87 | 90 | 87 | 94 | 89.6407 | 92 | 98 | 94.4739 | 90 | 97.8294 | 93 | 92 | 89 | 84.7283 |  | 1
| 1 | 2 | 83 | 93 | 89 | 82.03 | 90 | 90 | 87 | 92 | 89.5322 | 94 | 96 | 89.2741 | 89 | 98.7909 | 93 | 90 | 96 | 82.609 |  | 9
| 1 | 3 | 82 | 96 | 92 | 83.08 | 90 | 91 | 86 | 92 | 88.355 | 94 | 99 | 89.4414 | 89 | 100.4603 | 91 | 92 | 90 | 84.1506 |  | 16
| 1 | 4 | 84 | 93 | 93 | 84.44 | 90 | 92 | 88 | 88 | 88.7178 | 96 | 88 | 92.6068 | 89 | 98.3029 | 91 | 91 | 92 | 85.5061 |  | 91.0556
| 1 | 5 | 84 | 95 | 91 | 88.63 | 88 | 91 | 88 | 92 | 88.2079 | 94 | 96 | 94.9392 | 89 | 101.2803 | 91 | 91 | 92 | 83.3145 |  | 83
| 1 | 6 | 84 | 94 | 93 | 84.7 | 89 | 92 | 86 | 96 | 90.5606 | 94 | 95 | 92.5353 | 89 | 96.3016 | 93 | 92 | 91 | 79.923
| Mean |  | 83 | 94.1667 | 91.1667 | 84.7317 | 89 | 91 | 87 | 92.3333 | 89.169 | 94 | 95.3333 | 92.2118 | 89.1667 | 98.8276 | 92 | 91.3333 | 91.6667 | 83.3719
| Median |  | 83.5 | 94 | 91.5 | 84.57 | 89.5 | 91 | 87 | 92 | 89.125 | 94 | 96 | 92.571 | 89 | 98.5469 | 92 | 91.5 | 91.5 | 83.7326
| Std Dev. |  | 1.2649 | 1.169 | 1.8348 | 2.2767 | 1.2649 | 0.89443 | 0.89443 | 2.6583 | 0.90342 | 1.2649 | 3.8816 | 2.4134 | 0.40825 | 1.8075 | 1.0954 | 0.8165 | 2.4221 | 1.9739
| Rel.Std.Dev. |  | 1.524% | 1.241% | 2.013% | 2.687% | 1.421% | 0.9829% | 1.028% | 2.879% | 1.013% | 1.346% | 4.072% | 2.617% | 0.4578% | 1.829% | 1.191% | 0.894% | 2.642% | 2.368%
| PDM3 |  | -8.847% | 3.417% | 0.122% | -6.945% | -2.258% | -0.06107% | -4.454% | 1.403% | -2.072% | 3.234% | 4.698% | 1.27% | -2.074% | 8.535% | 1.037% | 0.305% | 0.6711% | -8.438%
| Z-Score (Absolute) |  | 2.57 | 0.91 | 0.03 | 2.03 | 0.7 | 0.08 | 1.32 | 0.34 | 0.65 | 0.86 | 1.27 | 0.3 | 0.65 | 2.36 | 0.23 | 0.03 | 0.13 | 2.45
| Analytical results for W in OREAS 780 (Certified Value 75 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 70.44 | 74.7836 | 89 | 73.4 | 71.31 | 77 | 79 | 75.4 | 73.7 | 73.6906 | 72.5 | 81 | 79.4108 | 72.7 | 77.8518 | 74.4 | 75.8 | 73.2 |  | 1
| 1 | 2 | 68.42 | 74.5098 | 85 | 71.5 | 75.33 | 75 | 78.4 | 73.4 | 77 | 74.5701 | 74 | 77.1 | 78.9145 | 73.6 | 78.244 | 73.5 | 73.8 | 78.9 |  | 26
| 1 | 3 | 66.91 | 74.8302 | 87.1 | 73.8 | 75.49 | 76 | 78.5 | 74.6 | 77.4 | 73.6762 | 75.5 | 81 | 72.6221 | 72.9 | 76.787 | 73 | 77.1 | 71.1 |  | 16
| 1 | 4 | 71.43 | 75.1819 | 87.3 | 78.6 | 78.9 | 79 | 80.4 | 74.9 | 76.6 | 73.8916 | 78 | 83.7 | 76.1171 | 73.8 | 76.7168 | 72.9 | 75.3 | 74.9 |  | 75.4289
| 1 | 5 | 70.69 | 75.0757 | 87.2 | 76.2 | 71.38 | 77 | 79.8 | 76.6 | 77.3 | 71.6787 | 77.5 | 79.9 | 75.9033 | 72.4 | 80.2651 | 72.8 | 73.4 | 73.5 |  | 84
| 1 | 6 | 71.64 | 74.8714 | 81.4 | 76.1 | 73.75 | 77 | 78.8 | 74.9 | 79.9 | 75.4278 | 76.5 | 78.9 | 77.8186 | 72.9 | 77.2229 | 75 | 75.3 | 73
| Mean |  | 69.9217 | 74.8754 | 86.1667 | 74.9333 | 74.36 | 76.8333 | 79.15 | 74.9667 | 76.9833 | 73.8225 | 75.6667 | 80.2667 | 76.7977 | 73.05 | 77.8479 | 73.6 | 75.1167 | 74.1
| Median |  | 70.565 | 74.8508 | 87.15 | 74.95 | 74.54 | 77 | 78.9 | 74.9 | 77.15 | 73.7911 | 76 | 80.45 | 76.9678 | 72.9 | 77.5374 | 73.25 | 75.3 | 73.35
| Std Dev. |  | 1.8661 | 0.23588 | 2.6583 | 2.5232 | 2.877 | 1.3292 | 0.7893 | 1.0443 | 1.9854 | 1.247 | 2.1134 | 2.2313 | 2.492 | 0.53944 | 1.3263 | 0.90554 | 1.3527 | 2.648
| Rel.Std.Dev. |  | 2.669% | 0.315% | 3.085% | 3.367% | 3.869% | 1.73% | 0.9972% | 1.393% | 2.579% | 1.689% | 2.793% | 2.78% | 3.245% | 0.7385% | 1.704% | 1.23% | 1.801% | 3.574%
| PDM3 |  | -7.301% | -0.7338% | 14.24% | -0.657% | -1.417% | 1.862% | 4.933% | -0.6129% | 2.061% | -2.13% | 0.3152% | 6.414% | 1.815% | -3.154% | 3.207% | -2.425% | -0.414% | -1.762%
| Z-Score (Absolute) |  | 2.16 | 0.07 | 4.69 | 0.05 | 0.29 | 0.76 | 1.73 | 0.03 | 0.82 | 0.51 | 0.26 | 2.2 | 0.74 | 0.84 | 1.18 | 0.61 | 0.03 | 0.4
| Analytical results for Y in OREAS 780 (Certified Value 25 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 22.18 | 23.52 | 24.3 | 25.2 | 21 | 24.9 | 26.7 | 25.75 | 24.7342 | 24.5 | 26.2 | 29.1992 | 26 | 38.8688 | 27.5 | 25.5 | 25 | 22.3757 |  | 1
| 1 | 2 | 22.33 | 23.4908 | 22 | 25.2 | 24 | 25.5 | 25.8 | 28.04 | 25.1594 | 25.6 | 25.4 | 25.3625 | 25.9 | 38.4041 | 26.8 | 30 | 26.4 | 21.2322 |  | #N/A
| 1 | 3 | 21.1 | 23.633 | 23.5 | 24.9 | 22 | 25.8 | 26.3 | 27.13 | 24.9103 | 25.2 | 26.9 | 25.0097 | 27.3 | 38.3188 | 26.9 | 27.4 | 24.4 | 23.232 |  | 16
| 1 | 4 | 21.15 | 23.2975 | 22.1 | 26.1 | 24 | 25.2 | 26.7 | 26.9 | 24.9401 | 25.5 | 24.9 | 26.831 | 28.2 | 37.8111 | 26.1 | 26.7 | 25.1 | 26.143 |  | 25.044
| 1 | 5 | 22.25 | 23.3958 | 22.9 | 25.1 | 22 | 25.8 | 26.3 | 31.24 | 24.9352 | 25.5 | 24.9 | 27.3295 | 24.6 | 38.1786 | 27.1 | 26.9 | 25.2 | 23.3989 |  | 85
| 1 | 6 | 22.43 | 23.391 | 23.4 | 25.7 | 22 | 26.5 | 26.8 | 26.17 | 25.2721 | 25.7 | 24.5 | 25.8697 | 25.9 | 33.7707 | 27.8 | 26.6 | 23.7 | 23.4725
| Mean |  | 21.9067 | 23.4547 | 23.0333 | 25.3667 | 22.5 | 25.6167 | 26.4333 | 27.5383 | 24.9919 | 25.3333 | 25.4667 | 26.6003 | 26.3167 | 37.5587 | 27.0333 | 27.1833 | 24.9667 | 23.309
| Median |  | 22.215 | 23.4433 | 23.15 | 25.2 | 22 | 25.65 | 26.5 | 27.015 | 24.9377 | 25.5 | 25.15 | 26.3503 | 25.95 | 38.2487 | 27 | 26.8 | 25.05 | 23.3154
| Std Dev. |  | 0.61138 | 0.11797 | 0.88468 | 0.44572 | 1.2247 | 0.55648 | 0.37771 | 1.9798 | 0.19265 | 0.44121 | 0.9136 | 1.545 | 1.2576 | 1.8871 | 0.59217 | 1.5145 | 0.90037 | 1.6283
| Rel.Std.Dev. |  | 2.791% | 0.503% | 3.841% | 1.757% | 5.443% | 2.172% | 1.429% | 7.189% | 0.7708% | 1.742% | 3.587% | 5.808% | 4.779% | 5.024% | 2.191% | 5.571% | 3.606% | 6.986%
| PDM3 |  | -12.53% | -6.346% | -8.028% | 1.289% | -10.16% | 2.287% | 5.548% | 9.96% | -0.2079% | 1.155% | 1.688% | 6.214% | 5.082% | 49.97% | 7.944% | 8.542% | -0.3086% | -6.927%
| Z-Score (Absolute) |  | 1.69 | 0.94 | 1.15 | 0.02 | 1.4 | 0.1 | 0.49 | 1.02 | 0.2 | 0.04 | 0.02 | 0.57 | 0.43 | 5.85 | 0.78 | 0.85 | 0.22 | 1.02
| Analytical results for Yb in OREAS 780 (Certified Value 1.62 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 1.25 | 1.6833 | 1.3 | 1.6 | 1.67 | 1.55 | 1.6 | 1.991 | 2.729 | 1.6551 | [9 empty] | 1
| 1 | 2 | 1.35 | 1.6097 | 1.36 | 1.6 | 1.7 | 1.6 | 1.7 | 1.8076 | 2.7716 | 1.6217 | [9 empty] | 27
| 1 | 3 | 1.31 | 1.7974 | 1.43 | 1.7 | 1.77 | 1.55 | 1.7 | 1.7967 | 2.7065 | 1.6561 | [9 empty] | 16
| 1 | 4 | 1.37 | 1.9309 | 1.29 | 1.7 | 2.01 | 1.55 | 1.7 | 1.8405 | 2.7843 | 1.6538 | [9 empty] | 1.618
| 1 | 5 | 1.38 | 1.803 | 1.28 | 1.7 | 1.66 | 1.45 | 1.7 | 1.9106 | 2.7875 | 1.6814 | [9 empty] | 86
| 1 | 6 | 1.37 | 1.7818 | 1.33 | 1.6 | 1.76 | 1.55 | 1.6 | 1.7763 | 2.4433 | 1.6342
| Mean |  | 1.3383 | 1.7677 | 1.3317 | 1.65 | 1.7617 | 1.5417 | 1.6667 | 1.8538 | 2.7037 | 1.6504
| Median |  | 1.36 | 1.7896 | 1.315 | 1.65 | 1.73 | 1.55 | 1.7 | 1.824 | 2.7503 | 1.6544
| Std Dev. |  | 0.049967 | 0.11051 | 0.056362 | 0.054772 | 0.12983 | 0.04916 | 0.05164 | 0.082012 | 0.13158 | 0.020559
| Rel.Std.Dev. |  | 3.733% | 6.251% | 4.232% | 3.32% | 7.37% | 3.189% | 3.098% | 4.424% | 4.867% | 1.246%
| PDM3 |  | -17.28% | 9.253% | -17.7% | 1.979% | 8.88% | -4.717% | 3.009% | 14.57% | 67.1% | 2.002%
| Z-Score (Absolute) |  | 1.91 | 0.65 | 1.95 | 0.05 | 0.62 | 0.7 | 0.05 | 1.16 | 6.24 | 0.05
| Analytical results for Zn in OREAS 780 (Certified Value 29 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 25 |  | 31 | 29 | 37.13 | 27 | 28 | 27 | 29 | 30.6621 | 24 | 28 | 29.4262 | 27 | 32.9915 | 30 | 31 | 29 | 68.3555 | 1
| 1 | 2 | 24 |  | 31 | 30 | 36.49 | 28 | 28 | 29 | 29 | 31.6043 | 26 | 28 | 26.9462 | 27 | 30.8918 | 29 | 30 | 31 | 63.7377 | 11
| 1 | 3 | 24 |  | 32 | 30 | 33.68 | 29 | 30 | 30 | 29 | 30.6838 | 24 | 30 | 26.4262 | 28 | 32.7441 | 30 | 31 | 29 | 61.5614 | 16
| 1 | 4 | 26 |  | 31 | 31 | 35.59 | 28 | 30 | 28 | 29 | 30.5444 | 30 | 28 | 27.1962 | 30 | 32.709 | 29 | 31 | 29 | 62.513 | 28.9962
| 1 | 5 | 26 |  | 32 | 30 | 34.61 | 29 | 28 | 29 | 29 | 30.1979 | 24 | 30 | 26.8162 | 31 | 33.0927 | 29 | 30 | 30 | 63.792 | 87
| 1 | 6 | 25 |  | 32 | 31 | 34.29 | 27 | 28 | 29 | 28 | 31.2938 | 24 | 29 | 27.6962 | 29 | 31.7173 | 29 | 31 | 30 | 61.5422
| Mean |  | 25 |  | 31.5 | 30.1667 | 35.2983 | 28 | 28.6667 | 28.6667 | 28.8333 | 30.831 | 25.3333 | 28.8333 | 27.4178 | 28.6667 | 32.3577 | 29.3333 | 30.6667 | 29.6667 | 63.5836
| Median |  | 25 |  | 31.5 | 30 | 35.1 | 28 | 28 | 29 | 29 | 30.6729 | 24 | 28.5 | 27.0712 | 28.5 | 32.7266 | 29 | 31 | 29.5 | 63.1253
| Std Dev. |  | 0.89443 |  | 0.54772 | 0.75277 | 1.3394 | 0.89443 | 1.0328 | 1.0328 | 0.40825 | 0.51872 | 2.4221 | 0.98319 | 1.0703 | 1.633 | 0.86876 | 0.5164 | 0.5164 | 0.8165 | 2.5393
| Rel.Std.Dev. |  | 3.578% |  | 1.739% | 2.495% | 3.794% | 3.194% | 3.603% | 3.603% | 1.416% | 1.682% | 9.561% | 3.41% | 3.904% | 5.696% | 2.685% | 1.76% | 1.684% | 2.752% | 3.994%
| PDM3 |  | -13.78% |  | 8.635% | 4.036% | 21.73% | -3.436% | -1.137% | -1.137% | -0.5618% | 6.328% | -12.63% | -0.5618% | -5.444% | -1.137% | 11.59% | 1.163% | 5.761% | 2.312% | 119.3%
| Z-Score (Absolute) |  | 2.07 | NA | 1.22 | 0.55 | 3.14 | 0.55 | 0.21 | 0.21 | 0.13 | 0.88 | 1.9 | 0.13 | 0.84 | 0.21 | 1.66 | 0.13 | 0.8 | 0.3 | 17.45
| Analytical results for Zr in OREAS 780 (Certified Value 130 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 108.9 | 141.772 | 151 | 128.5 | 151 | 126 | 128 | 129 | 135.4 | 116.3953 | 128 | 131 | 127.0223 | 104.7 | 208.5007 | 127 | 124 | 125.5 | 153.2385 | 1
| 1 | 2 | 111.8 | 141.772 | 141.5 | 130 | 156.3 | 126 | 134 | 129 | 128.2 | 120.8923 | 127 | 138 | 123.4323 | 110.5 | 207.182 | 127.5 | 128.5 | 136 | 152.6701 | #N/A
| 1 | 3 | 107.1 | 143.62 | 147 | 130 | 157.5 | 126 | 128 | 129 | 129.3 | 118.1788 | 122 | 133 | 123.8123 | 107.1 | 210.4376 | 128.5 | 129 | 129.5 | 152.7362 | 16
| 1 | 4 | 108.1 | 142.444 | 140 | 135 | 161.2 | 125 | 130 | 129 | 133.3 | 118.5105 | 125 | 125 | 124.7823 | 109 | 208.9256 | 129 | 130.5 | 127 | 149.1052 | 129.5533
| 1 | 5 | 110.4 | 143.608 | 141 | 129.5 | 154 | 123 | 131 | 132 | 129.3 | 115.4704 | 127 | 126 | 125.9123 | 101.4 | 210.7644 | 126 | 129 | 127.5 | 152.6865 | 88
| 1 | 6 | 110.8 | 145.54 | 143.5 | 132 | 159.4 | 124 | 132 | 130 | 128.5 | 119.9295 | 124 | 124 | 120.2823 | 108.6 | 207.5536 | 128.5 | 125.5 | 125 | 138.6111
| Mean |  | 109.5167 | 143.126 | 144 | 130.8333 | 156.5667 | 125 | 130.5 | 129.6667 | 130.6667 | 118.2295 | 125.5 | 129.5 | 124.2073 | 106.8833 | 208.894 | 127.75 | 127.75 | 128.4167 | 149.8413
| Median |  | 109.65 | 143.026 | 142.5 | 130 | 156.9 | 125.5 | 130.5 | 129 | 129.3 | 118.3446 | 126 | 128.5 | 124.2973 | 107.85 | 208.7131 | 128 | 128.75 | 127.25 | 152.6783
| Std Dev. |  | 1.7815 | 1.4447 | 4.2308 | 2.3381 | 3.6882 | 1.2649 | 2.3452 | 1.2111 | 2.9615 | 2.0513 | 2.2583 | 5.4681 | 2.3377 | 3.3283 | 1.4673 | 1.1292 | 2.4648 | 4.0425 | 5.704
| Rel.Std.Dev. |  | 1.627% | 1.009% | 2.938% | 1.787% | 2.356% | 1.012% | 1.797% | 0.934% | 2.266% | 1.735% | 1.799% | 4.222% | 1.882% | 3.114% | 0.7024% | 0.8839% | 1.929% | 3.148% | 3.807%
| PDM3 |  | -15.47% | 10.48% | 11.15% | 0.988% | 20.85% | -3.515% | 0.7307% | 0.08751% | 0.8594% | -8.741% | -3.129% | -0.04114% | -4.126% | -17.5% | 61.24% | -1.392% | -1.392% | -0.8773% | 15.66%
| Z-Score (Absolute) |  | 2.99 | 2.04 | 2.17 | 0.2 | 4.06 | 0.67 | 0.15 | 0.02 | 0.17 | 1.69 | 0.6 | 0 | 0.79 | 3.39 | 11.9 | 0.26 | 0.26 | 0.16 | 3.05
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Al in OREAS 780 (Indicative Value 7.51 wt.%) |  |  |  |  | Indicative
Al | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 2
2
| 1 | 1 | 7.3071 |  |  | 1
| 1 | 2 | 7.5233 |  |  | 19
| 1 | 3 | 7.62 |  |  | 16
| 1 | 4 | 7.5005 |  |  | 7.5135
| 1 | 5 | 7.5271 |  |  | 25
| 1 | 6 | 7.6025
| Mean |  | 7.5135
| Median |  | 7.5252
| Std Dev. |  | 0.11164
| Rel.Std.Dev. |  | 1.486%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 780 (Indicative Value 991 ppm) |  |  |  |  | Indicative
Ba | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 0
0
| 1 | 1 | 987.0967 |  |  | 1
| 1 | 2 | 1004.2467 |  |  | 20
| 1 | 3 | 1006.0767 |  |  | 16
| 1 | 4 | 987.8667 |  |  | 991.07
| 1 | 5 | 981.2467 |  |  | 26
| 1 | 6 | 979.8867
| Mean |  | 991.07
| Median |  | 987.4817
| Std Dev. |  | 11.3705
| Rel.Std.Dev. |  | 1.147%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 780 (Indicative Value 96 ppm) |  |  |  |  | Indicative
Bi | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 0
0
| 1 | 1 | 93.7841 |  |  | 1
| 1 | 2 | 94.8922 |  |  | 1
| 1 | 3 | 97.418 |  |  | 16
| 1 | 4 | 96.6662 |  |  | 95.9809
| 1 | 5 | 97.1905 |  |  | 27
| 1 | 6 | 95.9341
| Mean |  | 95.9809
| Median |  | 96.3002
| Std Dev. |  | 1.415
| Rel.Std.Dev. |  | 1.474%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ca in OREAS 780 (Indicative Value 5.94 wt.%) |  |  |  |  | Indicative
Ca | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 2
2
| 1 | 1 | 5.9524 |  |  | 1
| 1 | 2 | 5.8351 |  |  | 22
| 1 | 3 | 5.8847 |  |  | 16
| 1 | 4 | 6.0861 |  |  | 5.9437
| 1 | 5 | 5.9234 |  |  | 28
| 1 | 6 | 5.9806
| Mean |  | 5.9437
| Median |  | 5.9379
| Std Dev. |  | 0.086489
| Rel.Std.Dev. |  | 1.455%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 780 (Indicative Value 90 ppm) |  |  |  |  | Indicative
Ce | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 0
0
| 1 | 1 | 88.6896 |  |  | 1
| 1 | 2 | 90.2729 |  |  | 23
| 1 | 3 | 91.756 |  |  | 16
| 1 | 4 | 88.5419 |  |  | 89.7254
| 1 | 5 | 89.5338 |  |  | 29
| 1 | 6 | 89.5583
| Mean |  | 89.7254
| Median |  | 89.5461
| Std Dev. |  | 1.1799
| Rel.Std.Dev. |  | 1.315%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 780 (Indicative Value 10.3 ppm) |  |  |  |  | Indicative
Cs | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 1
1
| 1 | 1 | 10.463 |  |  | 1
| 1 | 2 | 10.1425 |  |  | 5
| 1 | 3 | 10.1774 |  |  | 16
| 1 | 4 | 10.3916 |  |  | 10.2984
| 1 | 5 | 10.5192 |  |  | 30
| 1 | 6 | 10.0968
| Mean |  | 10.2984
| Median |  | 10.2845
| Std Dev. |  | 0.18115
| Rel.Std.Dev. |  | 1.759%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 780 (Indicative Value 4.65 ppm) |  |  |  |  | Indicative
Dy | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 4.7962 |  |  | 1
| 1 | 2 | 4.6937 |  |  | 6
| 1 | 3 | 4.5538 |  |  | 16
| 1 | 4 | 4.6459 |  |  | 4.6496
| 1 | 5 | 4.6299 |  |  | 31
| 1 | 6 | 4.5784
| Mean |  | 4.6496
| Median |  | 4.6379
| Std Dev. |  | 0.087271
| Rel.Std.Dev. |  | 1.877%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 780 (Indicative Value 2.75 ppm) |  |  |  |  | Indicative
Er | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 2.7817 |  |  | 1
| 1 | 2 | 2.7742 |  |  | 26
| 1 | 3 | 2.8255 |  |  | 16
| 1 | 4 | 2.6969 |  |  | 2.7516
| 1 | 5 | 2.7145 |  |  | 32
| 1 | 6 | 2.7171
| Mean |  | 2.7516
| Median |  | 2.7456
| Std Dev. |  | 0.049886
| Rel.Std.Dev. |  | 1.813%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 780 (Indicative Value 1.28 ppm) |  |  |  |  | Indicative
Eu | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 1.2854 |  |  | 1
| 1 | 2 | 1.3114 |  |  | 7
| 1 | 3 | 1.3733 |  |  | 16
| 1 | 4 | 1.239 |  |  | 1.2825
| 1 | 5 | 1.2527 |  |  | 33
| 1 | 6 | 1.2334
| Mean |  | 1.2825
| Median |  | 1.2691
| Std Dev. |  | 0.05342
| Rel.Std.Dev. |  | 4.165%
| PDM3 |  | 3.997e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 780 (Indicative Value 2.33 wt.%) |  |  |  |  | Indicative
Fe | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 2
2
| 1 | 1 | 2.3119 |  |  | 1
| 1 | 2 | 2.2688 |  |  | 28
| 1 | 3 | 2.3388 |  |  | 16
| 1 | 4 | 2.3108 |  |  | 2.3257
| 1 | 5 | 2.3855 |  |  | 34
| 1 | 6 | 2.3385
| Mean |  | 2.3257
| Median |  | 2.3252
| Std Dev. |  | 0.038888
| Rel.Std.Dev. |  | 1.672%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 780 (Indicative Value 5.7 ppm) |  |  |  |  | Indicative
Gd | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 5.6203 |  |  | 1
| 1 | 2 | 5.7675 |  |  | 9
| 1 | 3 | 5.9952 |  |  | 16
| 1 | 4 | 5.8772 |  |  | 5.6984
| 1 | 5 | 5.4717 |  |  | 25
| 1 | 6 | 5.4589
| Mean |  | 5.6984
| Median |  | 5.6939
| Std Dev. |  | 0.21895
| Rel.Std.Dev. |  | 3.842%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 780 (Indicative Value 0.96 ppm) |  |  |  |  | Indicative
Ho | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.98101 |  |  | 1
| 1 | 2 | 0.93909 |  |  | 20
| 1 | 3 | 1.0004 |  |  | 16
| 1 | 4 | 0.92396 |  |  | 0.95733
| 1 | 5 | 0.9541 |  |  | 26
| 1 | 6 | 0.9454
| Mean |  | 0.95733
| Median |  | 0.94975
| Std Dev. |  | 0.028346
| Rel.Std.Dev. |  | 2.961%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 780 (Indicative Value 42.4 ppm) |  |  |  |  | Indicative
La | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 1
1
| 1 | 1 | 42.465 |  |  | 1
| 1 | 2 | 43.8235 |  |  | 21
| 1 | 3 | 43.3772 |  |  | 16
| 1 | 4 | 41.4623 |  |  | 42.3842
| 1 | 5 | 41.2845 |  |  | 27
| 1 | 6 | 41.8929
| Mean |  | 42.3842
| Median |  | 42.1789
| Std Dev. |  | 1.0358
| Rel.Std.Dev. |  | 2.444%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 780 (Indicative Value 44.4 ppm) |  |  |  |  | Indicative
Li | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 1
1
| 1 | 1 | 44.76 |  |  | 1
| 1 | 2 | 43.26 |  |  | 22
| 1 | 3 | 44.41 |  |  | 16
| 1 | 4 | 44.8 |  |  | 44.3917
| 1 | 5 | 43.94 |  |  | 28
| 1 | 6 | 45.18
| Mean |  | 44.3917
| Median |  | 44.585
| Std Dev. |  | 0.6941
| Rel.Std.Dev. |  | 1.564%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 780 (Indicative Value 0.37 ppm) |  |  |  |  | Indicative
Lu | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.37957 |  |  | 1
| 1 | 2 | 0.37185 |  |  | 23
| 1 | 3 | 0.39597 |  |  | 16
| 1 | 4 | 0.36108 |  |  | 0.37201
| 1 | 5 | 0.36209 |  |  | 29
| 1 | 6 | 0.36149
| Mean |  | 0.37201
| Median |  | 0.36697
| Std Dev. |  | 0.013856
| Rel.Std.Dev. |  | 3.725%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mg in OREAS 780 (Indicative Value 0.36 wt.%) |  |  |  |  | Indicative
Mg | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 3
3
| 1 | 1 | 0.34991 |  |  | 1
| 1 | 2 | 0.3602 |  |  | 24
| 1 | 3 | 0.36789 |  |  | 16
| 1 | 4 | 0.36022 |  |  | 0.3596
| 1 | 5 | 0.35765 |  |  | 30
| 1 | 6 | 0.36173
| Mean |  | 0.3596
| Median |  | 0.36021
| Std Dev. |  | 0.0058604
| Rel.Std.Dev. |  | 1.63%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 780 (Indicative Value 34.2 ppm) |  |  |  |  | Indicative
Nd | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 1
1
| 1 | 1 | 33.4097 |  |  | 1
| 1 | 2 | 33.5281 |  |  | 6
| 1 | 3 | 35.0153 |  |  | 16
| 1 | 4 | 33.7574 |  |  | 34.1847
| 1 | 5 | 34.4933 |  |  | 31
| 1 | 6 | 34.9043
| Mean |  | 34.1847
| Median |  | 34.1254
| Std Dev. |  | 0.70953
| Rel.Std.Dev. |  | 2.076%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 780 (Indicative Value 9.28 ppm) |  |  |  |  | Indicative
Pr | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 9.4437 |  |  | 1
| 1 | 2 | 9.5208 |  |  | 8
| 1 | 3 | 9.2127 |  |  | 16
| 1 | 4 | 9.1562 |  |  | 9.2794
| 1 | 5 | 9.1516 |  |  | 32
| 1 | 6 | 9.1912
| Mean |  | 9.2794
| Median |  | 9.202
| Std Dev. |  | 0.16064
| Rel.Std.Dev. |  | 1.731%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Si in OREAS 780 (Indicative Value 30.79 wt.%) |  |  |  |  | Indicative
Si | Round | Replicate | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 22 | wt.%
|  |  | PF*OES | PF*OES | BF*OES | 2
2
| 1 | 1 | 28.2597 | 30.62 | 30.01 | 1
| 1 | 2 | 28.3286 | 30.81 | 30.43 | 27
| 1 | 3 | 27.3909 | 30.99 | 30.76 | 16
| 1 | 4 | 27.9656 | 31.41 | 30.62 | 30.7864
| 1 | 5 | 26.7888 | 30.48 | 31.32 | 33
| 1 | 6 | 27.4959 | 30.9 | 31.09
| Mean |  | 27.7049 | 30.8683 | 30.705
| Median |  | 27.7308 | 30.855 | 30.69
| Std Dev. |  | 0.59098 | 0.32406 | 0.46805
| Rel.Std.Dev. |  | 2.133% | 1.05% | 1.524%
| PDM3 |  | -10.01% | 0.266% | -0.2645%
| Z-Score (Absolute) |  | 12.36 | 0.67 | 0
| Analytical results for Sm in OREAS 780 (Indicative Value 6.21 ppm) |  |  |  |  | Indicative
Sm | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 6.1622 |  |  | 1
| 1 | 2 | 6.0364 |  |  | 28
| 1 | 3 | 6.9113 |  |  | 16
| 1 | 4 | 6.0516 |  |  | 6.2081
| 1 | 5 | 6.026 |  |  | 34
| 1 | 6 | 6.0611
| Mean |  | 6.2081
| Median |  | 6.0563
| Std Dev. |  | 0.34797
| Rel.Std.Dev. |  | 5.605%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 780 (Indicative Value 65 ppm) |  |  |  |  | Indicative
Sr | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 0
0
| 1 | 1 | 63.5467 |  |  | 1
| 1 | 2 | 64.8787 |  |  | 3
| 1 | 3 | 64.5307 |  |  | 16
| 1 | 4 | 65.8227 |  |  | 64.567
| 1 | 5 | 63.7717 |  |  | 25
| 1 | 6 | 64.8517
| Mean |  | 64.567
| Median |  | 64.6912
| Std Dev. |  | 0.82823
| Rel.Std.Dev. |  | 1.283%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 780 (Indicative Value 0.82 ppm) |  |  |  |  | Indicative
Tb | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.83022 |  |  | 1
| 1 | 2 | 0.80276 |  |  | 20
| 1 | 3 | 0.86258 |  |  | 16
| 1 | 4 | 0.79666 |  |  | 0.81625
| 1 | 5 | 0.8052 |  |  | 26
| 1 | 6 | 0.80005
| Mean |  | 0.81625
| Median |  | 0.80398
| Std Dev. |  | 0.025657
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 780 (Indicative Value 12.7 ppm) |  |  |  |  | Indicative
Th | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 1
1
| 1 | 1 | 12.6937 |  |  | 1
| 1 | 2 | 12.6494 |  |  | 6
| 1 | 3 | 13.4342 |  |  | 16
| 1 | 4 | 12.2729 |  |  | 12.6574
| 1 | 5 | 12.518 |  |  | 27
| 1 | 6 | 12.3762
| Mean |  | 12.6574
| Median |  | 12.5837
| Std Dev. |  | 0.41258
| Rel.Std.Dev. |  | 3.26%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 780 (Indicative Value 0.445 wt.%) |  |  |  |  | Indicative
Ti | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 3
3
| 1 | 1 | 0.44277 |  |  | 1
| 1 | 2 | 0.44362 |  |  | 22
| 1 | 3 | 0.44575 |  |  | 16
| 1 | 4 | 0.44572 |  |  | 0.44477
| 1 | 5 | 0.44316 |  |  | 28
| 1 | 6 | 0.44762
| Mean |  | 0.44477
| Median |  | 0.44467
| Std Dev. |  | 0.0018913
| Rel.Std.Dev. |  | 0.4252%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 780 (Indicative Value 0.49 ppm) |  |  |  |  | Indicative
Tl | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.4942 |  |  | 1
| 1 | 2 | 0.47965 |  |  | 7
| 1 | 3 | 0.52261 |  |  | 16
| 1 | 4 | 0.46333 |  |  | 0.48572
| 1 | 5 | 0.47675 |  |  | 29
| 1 | 6 | 0.47779
| Mean |  | 0.48572
| Median |  | 0.47872
| Std Dev. |  | 0.020564
| Rel.Std.Dev. |  | 4.234%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 780 (Indicative Value 0.39 ppm) |  |  |  |  | Indicative
Tm | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.38053 |  |  | 1
| 1 | 2 | 0.39574 |  |  | 24
| 1 | 3 | 0.4 |  |  | 16
| 1 | 4 | 0.39516 |  |  | 0.39164
| 1 | 5 | 0.39477 |  |  | 30
| 1 | 6 | 0.38361
| Mean |  | 0.39164
| Median |  | 0.39496
| Std Dev. |  | 0.0077035
| Rel.Std.Dev. |  | 1.967%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 780 (Indicative Value 96 ppm) |  |  |  |  | Indicative
V | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 0
0
| 1 | 1 | 93.671 |  |  | 1
| 1 | 2 | 95.408 |  |  | 9
| 1 | 3 | 97.316 |  |  | 16
| 1 | 4 | 97.559 |  |  | 96.3155
| 1 | 5 | 97.406 |  |  | 31
| 1 | 6 | 96.533
| Mean |  | 96.3155
| Median |  | 96.9245
| Std Dev. |  | 1.5232
| Rel.Std.Dev. |  | 1.581%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 780 (Indicative Value 68 ppm) |  |  |  |  | Indicative
W | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 0
0
| 1 | 1 | 68.7453 |  |  | 1
| 1 | 2 | 68.3903 |  |  | 26
| 1 | 3 | 68.5658 |  |  | 16
| 1 | 4 | 67.0929 |  |  | 68.2666
| 1 | 5 | 69.9108 |  |  | 32
| 1 | 6 | 66.8945
| Mean |  | 68.2666
| Median |  | 68.4781
| Std Dev. |  | 1.1223
| Rel.Std.Dev. |  | 1.644%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 780 (Indicative Value 2.47 ppm) |  |  |  |  | Indicative
Yb | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 2.5465 |  |  | 1
| 1 | 2 | 2.4589 |  |  | 27
| 1 | 3 | 2.5116 |  |  | 16
| 1 | 4 | 2.4787 |  |  | 2.4747
| 1 | 5 | 2.4707 |  |  | 33
| 1 | 6 | 2.3819
| Mean |  | 2.4747
| Median |  | 2.4747
| Std Dev. |  | 0.05553
| Rel.Std.Dev. |  | 2.244%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 780 (Indicative Value 43.3 ppm) |  |  |  |  | Indicative
Zn | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 1
1
| 1 | 1 | 38.1675 |  |  | 1
| 1 | 2 | 39.6165 |  |  | 11
| 1 | 3 | 37.653 |  |  | 16
| 1 | 4 | 73.5525 |  |  | 43.2722
| 1 | 5 | 40.026 |  |  | 34
| 1 | 6 | 30.618
| Mean |  | 43.2722
| Median |  | 38.892
| Std Dev. |  | 15.2221
| Rel.Std.Dev. |  | 35.18%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.