## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 782
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Oxidising Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 8.9097 | 8.8041 | 9.0153 | 8.8524 | 8.967
As, Arsenic (ppm) | 297.1907 | 187.2559 | 407.1254 | IND | IND
BaO, Barium oxide (ppm) | 1204.7531 | 1115.2034 | 1294.3028 | 1163.9692 | 1245.537
Ca, Calcium (wt.%) | 16.4541 | 16.293 | 16.6153 | 16.3828 | 16.5254
Co, Cobalt (ppm) | < 100 | IND | IND | IND | IND
Cr, Chromium (ppm) | 77.8954 | 57.3229 | 98.4679 | 55.9564 | 99.8343
Cu, Copper (ppm) | 84.7538 | 65.1347 | 104.3728 | IND | IND
F, Fluorine (wt.%) | 15.1962 | 14.8095 | 15.583 | 15.0494 | 15.343
Fe, Iron (wt.%) | 1.8293 | 1.7968 | 1.8618 | 1.8135 | 1.8451
HfO2, Hafnium dioxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 1.7124 | 1.6823 | 1.7426 | 1.6961 | 1.7288
MgO, Magnesium oxide (wt.%) | 0.34794 | 0.32853 | 0.36735 | 0.33778 | 0.35811
Nb, Niobium (ppm) | < 50 | IND | IND | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.041204 | 0.033019 | 0.049389 | 0.039335 | 0.043072
S, Sulphur (ppm) | 204.2475 | 167.8016 | 240.6934 | 185.828 | 222.667
SiO2, Silicon dioxide (wt.%) | 51.2225 | 50.7289 | 51.7161 | 51.0104 | 51.4345
TiO2, Titanium dioxide (wt.%) | 0.44998 | 0.4372 | 0.46276 | 0.44118 | 0.45879
W, Tungsten (ppm) | 187.223 | 165.0938 | 209.3522 | 172.4531 | 201.993
Zr, Zirconium (ppm) | 114.6922 | 90.5255 | 138.8588 | IND | IND
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 5.1378 | 4.4215 | 5.8541 | 4.989 | 5.2866
Ion Selective Electrode
F, Fluorine (wt.%) | 14.5697 | 13.0156 | 16.1237 | 14.1647 | 14.9747
4-Acid Digestion
Ag, Silver (ppm) | 0.3465 | 0.31764 | 0.37536 | 0.30986 | 0.38314
Al, Aluminium (wt.%) | 4.6344 | 4.4778 | 4.791 | 4.5647 | 4.7041
As, Arsenic (ppm) | 276.1314 | 266.4479 | 285.8149 | 270.7614 | 281.5014
Ba, Barium (ppm) | 1065.1963 | 1031.9562 | 1098.4364 | 1048.1331 | 1082.2594
Be, Beryllium (ppm) | 2.6781 | 2.5772 | 2.7789 | 2.5862 | 2.77
Bi, Bismuth (ppm) | 247.7663 | 237.0673 | 258.4652 | 241.6987 | 253.8339
Ca, Calcium (wt.%) | 16.4202 | 15.9449 | 16.8954 | 16.1411 | 16.6993
Ce, Cerium (ppm) | 64.1815 | 61.6614 | 66.7016 | 62.6494 | 65.7137
Co, Cobalt (ppm) | 2.0359 | 1.8149 | 2.2569 | 1.9284 | 2.1435
Cr, Chromium (ppm) | 59.9281 | 53.8736 | 65.9825 | 56.7833 | 63.0728
Cs, Caesium (ppm) | 6.5309 | 6.2747 | 6.787 | 6.3686 | 6.6932
Cu, Copper (ppm) | 80.0911 | 76.8199 | 83.3622 | 77.7769 | 82.4052
Dy, Dysprosium (ppm) | 3.5857 | 3.3714 | 3.8 | 3.4292 | 3.7423
Er, Erbium (ppm) | 1.8888 | 1.6967 | 2.0809 | 1.6957 | 2.0819
Eu, Europium (ppm) | 0.97856 | 0.8923 | 1.0648 | 0.91552 | 1.0416
Fe, Iron (wt.%) | 1.8248 | 1.7686 | 1.881 | 1.7949 | 1.8547
Ga, Gallium (ppm) | 11.834 | 11.2019 | 12.4661 | 11.4161 | 12.2518
Gd, Gadolinium (ppm) | 4.4464 | 4.1648 | 4.7279 | 4.2832 | 4.6095
Hf, Hafnium (ppm) | 2.5261 | 2.3348 | 2.7175 | 2.4128 | 2.6394
Ho, Holmium (ppm) | 0.67187 | 0.61698 | 0.72676 | 0.63517 | 0.70857
In, Indium (ppm) | 0.062192 | 0.052736 | 0.071647 | 0.058396 | 0.065987
K, Potassium (wt.%) | 1.4728 | 1.4327 | 1.5129 | 1.4524 | 1.4931
La, Lanthanum (ppm) | 32.3096 | 30.864 | 33.7553 | 31.6745 | 32.9447
Li, Lithium (ppm) | 39.0289 | 37.3069 | 40.7509 | 38.0816 | 39.9762
Lu, Lutetium (ppm) | 0.26335 | 0.2228 | 0.3039 | 0.24778 | 0.27893
Mg, Magnesium (wt.%) | 0.1948 | 0.18153 | 0.20807 | 0.18913 | 0.20047
Mn, Manganese (wt.%) | 0.0052702 | 0.0050415 | 0.005499 | 0.0051432 | 0.0053972
Mo, Molybdenum (ppm) | 3.6836 | 3.4325 | 3.9348 | 3.4569 | 3.9104
Na, Sodium (wt.%) | 0.099765 | 0.088902 | 0.11063 | 0.097685 | 0.10185
Nb, Niobium (ppm) | 7.8287 | 7.1449 | 8.5125 | 7.4771 | 8.1803
Nd, Neodymium (ppm) | 26.6697 | 25.3682 | 27.9711 | 26.0794 | 27.2599
Ni, Nickel (ppm) | 11.2944 | 10.5811 | 12.0078 | 10.6811 | 11.9078
P, Phosphorus (wt.%) | 0.020028 | 0.018968 | 0.021088 | 0.019286 | 0.020769
Pb, Lead (ppm) | 14.5388 | 13.7202 | 15.3574 | 13.8277 | 15.25
Pr, Praseodymium (ppm) | 7.0455 | 6.6328 | 7.4582 | 6.7864 | 7.3046
Rb, Rubidium (ppm) | 83.3108 | 79.7557 | 86.8658 | 81.7942 | 84.8273
S, Sulphur (wt.%) | 0.024754 | 0.022471 | 0.027036 | 0.023808 | 0.0257
Sb, Antimony (ppm) | 5.1611 | 4.8538 | 5.4684 | 4.9819 | 5.3403
Sc, Scandium (ppm) | 9.0161 | 8.5505 | 9.4817 | 8.8083 | 9.2239
Se, Selenium (ppm) | 0.99641 | 0.88955 | 1.1033 | IND | IND
Sm, Samarium (ppm) | 4.9983 | 4.6547 | 5.3418 | 4.7985 | 5.1981
Sn, Tin (ppm) | 9.858 | 9.3296 | 10.3864 | 9.477 | 10.239
Sr, Strontium (ppm) | 51.0203 | 48.9919 | 53.0487 | 49.976 | 52.0646
Ta, Tantalum (ppm) | 0.59342 | 0.52302 | 0.66382 | 0.55788 | 0.62896
Tb, Terbium (ppm) | 0.6477 | 0.57627 | 0.71914 | 0.60819 | 0.68722
Te, Tellurium (ppm) | 0.20099 | 0.16211 | 0.23986 | 0.17533 | 0.22664
Th, Thorium (ppm) | 9.9392 | 9.5397 | 10.3386 | 9.6014 | 10.277
Ti, Titanium (wt.%) | 0.24042 | 0.2287 | 0.25215 | 0.23312 | 0.24773
Tl, Thallium (ppm) | 0.47314 | 0.44075 | 0.50554 | 0.45704 | 0.48924
Tm, Thulium (ppm) | 0.26488 | 0.24669 | 0.28307 | 0.25065 | 0.27911
U, Uranium (ppm) | 1.9413 | 1.8329 | 2.0496 | 1.8854 | 1.9971
V, Vanadium (ppm) | 61.8007 | 59.6676 | 63.9338 | 60.2361 | 63.3653
W, Tungsten (ppm) | 178.4765 | 171.4033 | 185.5496 | 174.9562 | 181.9967
Y, Yttrium (ppm) | 45.4846 | 43.626 | 47.3431 | 44.4913 | 46.4779
Yb, Ytterbium (ppm) | 1.8636 | 1.7251 | 2.0022 | 1.7126 | 2.0147
Zn, Zinc (ppm) | 21.6232 | 20.3222 | 22.9243 | 20.6677 | 22.5788
Zr, Zirconium (ppm) | 85.2807 | 80.989 | 89.5723 | 83.3131 | 87.2482
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 782
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Oxidising Fusion XRF
Bi | ppm | 265.6452 | In | ppm | < 100 | Se | ppm | < 100
Cd | ppm | < 100 | La2O3 | ppm | 88.5714 | Sm | ppm | < 170
CeO2 | ppm | 118.7656 | Mn | wt.% | 0.0078275 | Sn | ppm | 47.8147
Cl | ppm | 65 | Mo | ppm | 53.1536 | Sr | ppm | 87.8169
Cs | ppm | 117.9023 | Na2O | wt.% | 0.13844 | Ta | ppm | < 80
Dy | ppm | 246.8706 | Nd | ppm | 29.2924 | Te | ppm | < 100
Eu | ppm | < 170 | Ni | ppm | < 50 | Th | ppm | 83.4871
Ga | ppm | < 100 | Pb | ppm | 52.0194 | Tl | ppm | < 100
Gd | ppm | < 170 | Pr | ppm | 55.1785 | U | ppm | 155.4726
Ge | ppm | < 100 | Rb | ppm | 103.2679 | V | ppm | 66.77
H2O- | wt.% | 0.26265 | Sb | ppm | 116.1111 | Y2O3 | ppm | 89.8889
Hg | ppm | < 100 | Sc | ppm | < 100 | Zn | ppm | 33.6057
Thermogravimetry
H2O- | wt.% | 0.26265
4-Acid Digestion
B | ppm | 9.0233 | Ge | ppm | 0.19356 | Re | ppm | 0.00335
Cd | ppm | 0.028113 | Hg | ppm | 0.12333
Borate / Peroxide Fusion ICP
Al | wt.% | 4.9824 | Gd | ppm | 4.6956 | Sr | ppm | 36.591
Ba | ppm | 1132.94 | Ho | ppm | 0.90697 | Tb | ppm | 0.71244
Bi | ppm | 200.2811 | La | ppm | 30.3868 | Th | ppm | 8.0807
Ca | wt.% | 17.0122 | Li | ppm | 39.1617 | Ti | wt.% | 0.2904
Ce | ppm | 63.0155 | Lu | ppm | 0.33342 | Tl | ppm | 0.30213
Cs | ppm | 6.5524 | Mg | wt.% | 0.24156 | Tm | ppm | 0.35602
Dy | ppm | 4.1123 | Nd | ppm | 25.1293 | V | ppm | 62.9495
Er | ppm | 2.5944 | Pr | ppm | 6.7728 | W | ppm | 166.0043
Eu | ppm | 1.0531 | Si | wt.% | 23.2654 | Yb | ppm | 2.3057
Fe | wt.% | 1.5643 | Sm | ppm | 4.726 | Zn | ppm | 46.0425
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 782
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Oxidising Fusion XRF
Al2O3, wt.% | 8.9097 | 0.14387 | 8.6219 | 9.1974 | 8.4781 | 9.3413 | 1.615% | 3.23% | 4.844% | 8.4642 | 9.3552
As, ppm | 297.1907 | 57.2345 | 182.7217 | 411.6596 | 125.4872 | 468.8941 | 19.26% | 38.52% | 57.78% | 282.3311 | 312.0502
BaO, ppm | 1204.7531 | 70.1262 | 1064.5008 | 1345.0054 | 994.3746 | 1415.1316 | 5.821% | 11.64% | 17.46% | 1144.5154 | 1264.9907
Ca, wt.% | 16.4541 | 0.1669 | 16.1203 | 16.7879 | 15.9534 | 16.9548 | 1.014% | 2.029% | 3.043% | 15.6314 | 17.2768
Co, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cr, ppm | 77.8954 | 13.5454 | 50.8046 | 104.9861 | 37.2593 | 118.5314 | 17.39% | 34.78% | 52.17% | 74.0006 | 81.7901
Cu, ppm | 84.7538 | 12.6469 | 59.4599 | 110.0477 | 46.813 | 122.6946 | 14.92% | 29.84% | 44.77% | 80.5161 | 88.9915
F, wt.% | 15.1962 | 0.36449 | 14.4672 | 15.9252 | 14.1028 | 16.2897 | 2.399% | 4.797% | 7.196% | 14.4364 | 15.956
Fe, wt.% | 1.8293 | 0.049048 | 1.7312 | 1.9274 | 1.6822 | 1.9764 | 2.681% | 5.362% | 8.044% | 1.7378 | 1.9208
HfO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 1.7124 | 0.042582 | 1.6272 | 1.7976 | 1.5847 | 1.8402 | 2.487% | 4.973% | 7.46% | 1.6268 | 1.798
MgO, wt.% | 0.34794 | 0.023409 | 0.30112 | 0.39476 | 0.27772 | 0.41817 | 6.728% | 13.46% | 20.18% | 0.33054 | 0.36534
Nb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
P2O5, wt.% | 0.041204 | 0.0062331 | 0.028737 | 0.05367 | 0.022504 | 0.059903 | 15.13% | 30.26% | 45.38% | 0.039143 | 0.043264
S, ppm | 204.2475 | 40.1897 | 123.8681 | 284.6269 | 83.6784 | 324.8166 | 19.68% | 39.35% | 59.03% | 194.0351 | 214.4599
SiO2, wt.% | 51.2225 | 0.60628 | 50.0099 | 52.435 | 49.4036 | 53.0413 | 1.184% | 2.367% | 3.551% | 48.6614 | 53.7836
TiO2, wt.% | 0.44998 | 0.019372 | 0.41124 | 0.48873 | 0.39187 | 0.5081 | 4.305% | 8.61% | 12.91% | 0.42748 | 0.47248
W, ppm | 187.223 | 23.3147 | 140.5935 | 233.8525 | 117.2788 | 257.1673 | 12.45% | 24.91% | 37.36% | 177.8619 | 196.5842
Zr, ppm | 114.6922 | 29.9465 | 54.7991 | 174.5852 | 24.8526 | 204.5317 | 26.11% | 52.22% | 78.33% | 108.9575 | 120.4268
Thermogravimetry
LOI1000, wt.% | 5.1378 | 1.2918 | 2.5543 | 7.7214 | 1.2625 | 9.0131 | 25.14% | 50.29% | 75.43% | 4.8809 | 5.3947
Ion Selective Electrode
F, wt.% | 14.5697 | 2.1323 | 10.3051 | 18.8342 | 8.1728 | 20.9665 | 14.64% | 29.27% | 43.91% | 13.8412 | 15.2982
4-Acid Digestion
Ag, ppm | 0.3465 | 0.016091 | 0.31432 | 0.37868 | 0.29823 | 0.39477 | 4.644% | 9.288% | 13.93% | 0.32917 | 0.36382
Al, wt.% | 4.6344 | 0.20181 | 4.2308 | 5.038 | 4.029 | 5.2398 | 4.355% | 8.709% | 13.06% | 4.4027 | 4.8661
As, ppm | 276.1314 | 10.6644 | 254.8025 | 297.4603 | 244.1381 | 308.1247 | 3.862% | 7.724% | 11.59% | 262.3248 | 289.938
Ba, ppm | 1065.1963 | 38.7316 | 987.733 | 1142.6595 | 949.0014 | 1181.3912 | 3.636% | 7.272% | 10.91% | 1011.9365 | 1118.4561
Be, ppm | 2.6781 | 0.12786 | 2.4224 | 2.9338 | 2.2945 | 3.0617 | 4.774% | 9.549% | 14.32% | 2.5442 | 2.812
Bi, ppm | 247.7663 | 14.0069 | 219.7525 | 275.78 | 205.7457 | 289.7869 | 5.653% | 11.31% | 16.96% | 235.378 | 260.1546
Ca, wt.% | 16.4202 | 0.64419 | 15.1318 | 17.7086 | 14.4876 | 18.3527 | 3.923% | 7.846% | 11.77% | 15.5992 | 17.2412
Ce, ppm | 64.1815 | 2.4722 | 59.2371 | 69.126 | 56.7649 | 71.5982 | 3.852% | 7.704% | 11.56% | 60.9725 | 67.3906
Co, ppm | 2.0359 | 0.25981 | 1.5163 | 2.5555 | 1.2565 | 2.8154 | 12.76% | 25.52% | 38.28% | 1.9341 | 2.1377
Cr, ppm | 59.9281 | 10.1348 | 39.6585 | 80.1976 | 29.5237 | 90.3324 | 16.91% | 33.82% | 50.73% | 56.9317 | 62.9245
Cs, ppm | 6.5309 | 0.2172 | 6.0965 | 6.9653 | 5.8793 | 7.1825 | 3.326% | 6.652% | 9.977% | 6.2043 | 6.8574
Cu, ppm | 80.0911 | 3.5203 | 73.0506 | 87.1316 | 69.5303 | 90.6518 | 4.395% | 8.791% | 13.19% | 76.0865 | 84.0956
Dy, ppm | 3.5857 | 0.15299 | 3.2797 | 3.8917 | 3.1268 | 4.0447 | 4.267% | 8.533% | 12.8% | 3.4064 | 3.765
Er, ppm | 1.8888 | 0.12043 | 1.6479 | 2.1296 | 1.5275 | 2.2501 | 6.376% | 12.75% | 19.13% | 1.7943 | 1.9832
Eu, ppm | 0.97856 | 0.052991 | 0.87258 | 1.0845 | 0.81959 | 1.1375 | 5.415% | 10.83% | 16.25% | 0.92964 | 1.0275
Fe, wt.% | 1.8248 | 0.049187 | 1.7264 | 1.9232 | 1.6772 | 1.9723 | 2.695% | 5.391% | 8.086% | 1.7335 | 1.916
Ga, ppm | 11.834 | 0.78818 | 10.2576 | 13.4104 | 9.4694 | 14.1985 | 6.66% | 13.32% | 19.98% | 11.2423 | 12.4257
Gd, ppm | 4.4464 | 0.18385 | 4.0787 | 4.8141 | 3.8948 | 4.9979 | 4.135% | 8.27% | 12.4% | 4.224 | 4.6687
Hf, ppm | 2.5261 | 0.16968 | 2.1868 | 2.8655 | 2.0171 | 3.0351 | 6.717% | 13.43% | 20.15% | 2.3998 | 2.6524
Ho, ppm | 0.67187 | 0.045398 | 0.58107 | 0.76266 | 0.53568 | 0.80806 | 6.757% | 13.51% | 20.27% | 0.63828 | 0.70546
In, ppm | 0.062192 | 0.0073337 | 0.047524 | 0.076859 | 0.04019 | 0.084193 | 11.79% | 23.58% | 35.38% | 0.059082 | 0.065301
K, wt.% | 1.4728 | 0.037932 | 1.3969 | 1.5487 | 1.359 | 1.5866 | 2.575% | 5.151% | 7.726% | 1.3992 | 1.5464
La, ppm | 32.3096 | 1.4571 | 29.3955 | 35.2237 | 27.9385 | 36.6808 | 4.51% | 9.019% | 13.53% | 30.6941 | 33.9251
Li, ppm | 39.0289 | 1.9252 | 35.1784 | 42.8793 | 33.2532 | 44.8046 | 4.933% | 9.866% | 14.8% | 37.0774 | 40.9803
Lu, ppm | 0.26335 | 0.026022 | 0.21131 | 0.31539 | 0.18529 | 0.34142 | 9.881% | 19.76% | 29.64% | 0.25018 | 0.27652
Mg, wt.% | 0.1948 | 0.016836 | 0.16113 | 0.22847 | 0.14429 | 0.24531 | 8.643% | 17.29% | 25.93% | 0.18506 | 0.20454
Mn, wt.% | 0.0052702 | 0.00020877 | 0.0048527 | 0.0056878 | 0.0046439 | 0.0058965 | 3.961% | 7.923% | 11.88% | 0.0050067 | 0.0055337
Mo, ppm | 3.6836 | 0.21324 | 3.2571 | 4.1101 | 3.0439 | 4.3234 | 5.789% | 11.58% | 17.37% | 3.4995 | 3.8678
Na, wt.% | 0.099765 | 0.013787 | 0.072191 | 0.12734 | 0.058403 | 0.14113 | 13.82% | 27.64% | 41.46% | 0.094777 | 0.10475
Nb, ppm | 7.8287 | 0.71955 | 6.3896 | 9.2678 | 5.67 | 9.9873 | 9.191% | 18.38% | 27.57% | 7.4373 | 8.2201
Nd, ppm | 26.6697 | 1.1484 | 24.3729 | 28.9665 | 23.2244 | 30.1149 | 4.306% | 8.612% | 12.92% | 25.3362 | 28.0031
Ni, ppm | 11.2944 | 0.7975 | 9.6995 | 12.8894 | 8.902 | 13.6869 | 7.061% | 14.12% | 21.18% | 10.7297 | 11.8592
P, wt.% | 0.020028 | 0.00078244 | 0.018463 | 0.021593 | 0.01768 | 0.022375 | 3.907% | 7.814% | 11.72% | 0.019026 | 0.021029
Pb, ppm | 14.5388 | 1.0579 | 12.4231 | 16.6546 | 11.3652 | 17.7125 | 7.276% | 14.55% | 21.83% | 13.8119 | 15.2658
Pr, ppm | 7.0455 | 0.29186 | 6.4618 | 7.6292 | 6.1699 | 7.9211 | 4.142% | 8.285% | 12.43% | 6.6932 | 7.3978
Rb, ppm | 83.3108 | 3.2261 | 76.8586 | 89.763 | 73.6324 | 92.9891 | 3.872% | 7.745% | 11.62% | 79.1452 | 87.4763
S, wt.% | 0.024754 | 0.004131 | 0.016492 | 0.033016 | 0.012361 | 0.037147 | 16.69% | 33.38% | 50.06% | 0.023516 | 0.025992
Sb, ppm | 5.1611 | 0.3837 | 4.3937 | 5.9285 | 4.01 | 6.3122 | 7.435% | 14.87% | 22.3% | 4.903 | 5.4191
Sc, ppm | 9.0161 | 0.35211 | 8.3119 | 9.7203 | 7.9598 | 10.0724 | 3.905% | 7.811% | 11.72% | 8.5653 | 9.4669
Se, ppm | 0.99641 | 0.11131 | 0.77378 | 1.219 | 0.66247 | 1.3304 | 11.17% | 22.34% | 33.51% | 0.94659 | 1.0462
Sm, ppm | 4.9983 | 0.19213 | 4.614 | 5.3825 | 4.4219 | 5.5747 | 3.844% | 7.688% | 11.53% | 4.7483 | 5.2482
Sn, ppm | 9.858 | 0.62525 | 8.6075 | 11.1085 | 7.9823 | 11.7338 | 6.343% | 12.69% | 19.03% | 9.3651 | 10.3509
Sr, ppm | 51.0203 | 2.415 | 46.1904 | 55.8502 | 43.7754 | 58.2652 | 4.733% | 9.467% | 14.2% | 48.4693 | 53.5713
Ta, ppm | 0.59342 | 0.064418 | 0.46458 | 0.72226 | 0.40017 | 0.78667 | 10.86% | 21.71% | 32.57% | 0.56375 | 0.62309
Tb, ppm | 0.6477 | 0.036776 | 0.57415 | 0.72126 | 0.53738 | 0.75803 | 5.678% | 11.36% | 17.03% | 0.61532 | 0.68009
Te, ppm | 0.20099 | 0.022089 | 0.15681 | 0.24516 | 0.13472 | 0.26725 | 10.99% | 21.98% | 32.97% | 0.19094 | 0.21104
Th, ppm | 9.9392 | 0.45854 | 9.0221 | 10.8562 | 8.5636 | 11.3148 | 4.613% | 9.227% | 13.84% | 9.4422 | 10.4361
Ti, wt.% | 0.24042 | 0.016139 | 0.20815 | 0.2727 | 0.19201 | 0.28884 | 6.713% | 13.43% | 20.14% | 0.2284 | 0.25245
Tl, ppm | 0.47314 | 0.037332 | 0.39848 | 0.54781 | 0.36115 | 0.58514 | 7.89% | 15.78% | 23.67% | 0.44949 | 0.4968
Tm, ppm | 0.26488 | 0.014569 | 0.23574 | 0.29402 | 0.22117 | 0.30859 | 5.5% | 11% | 16.5% | 0.25164 | 0.27812
U, ppm | 1.9413 | 0.0867 | 1.7679 | 2.1147 | 1.6812 | 2.2014 | 4.466% | 8.932% | 13.4% | 1.8442 | 2.0383
V, ppm | 61.8007 | 2.2026 | 57.3956 | 66.2058 | 55.193 | 68.4084 | 3.564% | 7.128% | 10.69% | 58.7107 | 64.8907
W, ppm | 178.4765 | 6.5701 | 165.3363 | 191.6166 | 158.7662 | 198.1867 | 3.681% | 7.362% | 11.04% | 169.5526 | 187.4003
Y, ppm | 45.4846 | 2.3265 | 40.8315 | 50.1376 | 38.505 | 52.4641 | 5.115% | 10.23% | 15.34% | 43.2103 | 47.7588
Yb, ppm | 1.8636 | 0.096882 | 1.6699 | 2.0574 | 1.573 | 2.1543 | 5.199% | 10.4% | 15.6% | 1.7705 | 1.9568
Zn, ppm | 21.6232 | 1.9674 | 17.6884 | 25.558 | 15.721 | 27.5254 | 9.099% | 18.2% | 27.3% | 20.5421 | 22.7044
Zr, ppm | 85.2807 | 5.2834 | 74.7138 | 95.8476 | 69.4303 | 101.131 | 6.195% | 12.39% | 18.59% | 81.0166 | 89.5447
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 782
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
ISE | ion specific electrode
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
Text Values:
I.N.F. | Insufficient Sample (Lab 07)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 782
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek, Perth, WA, Australia
Lucid Laboratories Private Limited, Hyderabad, Telangana, India
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan
UIS Analytical Services, Centurion, South Africa

## Oxidising Fusion XRF
| Analytical results for Al2O3 in OREAS 782 (Certified Value 8.91 wt.%) | [20 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 8.82 | 9.81 | 8.89 | 8.98 | 8.8887 | 8.76 | 1.96 | 9.17 | 8.9 | 8.83 | 9.05 | 8.8591 | 8.8269 | 8.8 | 8.9903 | 9.15 | 9.05 | 8.93 | 8.192 | 1
| 1 | 2 | 8.72 | 9.81 | 8.78 | 9.05 | 8.8638 | 8.81 | 1.97 | 9.12 | 8.77 | 8.77 | 9.06 | 8.8107 | 8.7888 | 8.9 | 8.9913 | 9.14 | 8.99 | 8.94 | 7.948 | #N/A
| 1 | 3 | 8.77 | 9.83 | 8.86 | 9.02 | 8.8757 | 8.84 | 2.05 | 9.11 | 8.59 | 8.85 | 9.06 | 8.8652 | 8.7757 | 8.8 | 9.0059 | 9.12 | 9.11 | 8.91 | 8.042 | 16
| 1 | 4 | 8.8 | 9.8 | 8.85 | 9.05 | 8.8452 | 8.78 | 2.08 | 9.18 | 8.58 | 8.81 | 9.05 | 8.8569 | 8.7953 | 8.8 | 9.0187 | 9.11 | 9.03 | 8.96 | 7.928 | 8.9097
| 1 | 5 | 8.83 | 9.82 | 8.97 | 9.06 | 8.9203 | 8.8 | 1.9 | 9.09 | 8.6 | 8.83 | 9.07 | 8.8556 | 8.7647 | 8.8 | 9.0272 | 9.08 | 9.13 | 8.97 | 8.194 | 7
| 1 | 6 | 8.76 | 9.79 | 8.89 | 8.95 | 8.8323 | 8.73 | 1.94 | 9.17 | 8.49 | 8.78 | 9.08 | 8.8609 | 8.7802 | 8.8 | 8.9911 | 8.99 | 9.04 | 8.91 | 7.932
| Mean |  | 8.7833 | 9.81 | 8.8733 | 9.0183 | 8.871 | 8.7867 | 1.9833 | 9.14 | 8.655 | 8.8117 | 9.0617 | 8.8514 | 8.7886 | 8.8167 | 9.0041 | 9.0983 | 9.0583 | 8.9367 | 8.0393
| Median |  | 8.785 | 9.81 | 8.875 | 9.035 | 8.8698 | 8.79 | 1.965 | 9.145 | 8.595 | 8.82 | 9.06 | 8.858 | 8.7845 | 8.8 | 8.9986 | 9.115 | 9.045 | 8.935 | 7.995
| Std Dev. |  | 0.041312 | 0.014142 | 0.062183 | 0.04446 | 0.03158 | 0.038816 | 0.068313 | 0.037947 | 0.15057 | 0.031252 | 0.01169 | 0.020218 | 0.021553 | 0.040825 | 0.015957 | 0.058452 | 0.052313 | 0.025033 | 0.12609
| Rel.Std.Dev. |  | 0.4703% | 0.1442% | 0.7008% | 0.493% | 0.356% | 0.4418% | 3.444% | 0.4152% | 1.74% | 0.3547% | 0.129% | 0.2284% | 0.2452% | 0.463% | 0.1772% | 0.6425% | 0.5775% | 0.2801% | 1.568%
| PDM3 |  | -1.418% | 10.1% | -0.4081% | 1.219% | -0.4342% | -1.381% | -77.74% | 2.585% | -2.859% | -1.1% | 1.706% | -0.6543% | -1.359% | -1.044% | 1.059% | 2.117% | 1.668% | 0.3028% | -9.769%
| Z-Score (Absolute) |  | 0.54 | 5.73 | 0.01 | 0.89 | 0.01 | 0.52 | NA | 1.64 | 1.33 | 0.37 | 1.16 | 0.13 | 0.51 | 0.34 | 0.81 | 1.38 | 1.14 | 0.39 | 5.09
Results from laboratory 1.08 were removed due to their systematic bias.
| Analytical results for As in OREAS 782 (Certified Value 297 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 15 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | 300 | 400 | < 100 | 333 | 268.9 | 230 | 200 | 500 | 400 | [10 empty] | 1
| 1 | 2 | 300 | 300 | < 100 | 341 | 272.7 | 230 | 100 | 400 | 300 | [10 empty] | #N/A
| 1 | 3 | 300 | 200 | < 100 | 333 | 265.1 | 230 | 200 | 500 | 300 | [10 empty] | 16
| 1 | 4 | 400 | 200 | < 100 | 333 | 268.9 | 230 | 100 | 400 | 300 | [10 empty] | 297.1907
| 1 | 5 | 300 | 300 | < 100 | 341 | 272.7 | 230 | 100 | 500 | 400 | [10 empty] | 8
| 1 | 6 | 300 | 300 | < 100 | 333 | 272.7 | 230 | 100 | 500 | 400
| Mean |  | 316.6667 | 283.3333 |  | 335.6667 | 270.1667 | 230 | 133.3333 | 466.6667 | 350
| Median |  | 300 | 300 |  | 333 | 270.8 | 230 | 100 | 500 | 350
| Std Dev. |  | 40.8248 | 75.2773 |  | 4.1312 | 3.1027 | 0 | 51.6398 | 51.6398 | 54.7723
| Rel.Std.Dev. |  | 12.89% | 26.57% |  | 1.231% | 1.148% | 0% | 38.73% | 11.07% | 15.65%
| PDM3 |  | 6.553% | -4.663% |  | 12.95% | -9.093% | -22.61% | -55.14% | 57.03% | 17.77%
| Z-Score (Absolute) |  | 0.4 | 0 | 2.8 | 0.63 | 0.16 | 0.67 | 1.8 | 2.2 | 0.8
| Analytical results for BaO in OREAS 782 (Certified Value 1205 ppm) | [20 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 1100 | 1200 | 1200 | 1200 | 300 | 1200 | 1274 | 1095 | 1200 | 1203.7114 | 1300 | 1300 | 1300 | 1210 | [5 empty] | 1
| 1 | 2 | 1100 | 1200 | 1200 | 1200 | 300 | 1160 | 1274.9 | 1140 | 1200 | 1204.4511 | 1200 | 1200 | 1300 | 1200 | [5 empty] | #N/A
| 1 | 3 | 1000 | 1300 | 1200 | 1200 | 200 | 1180 | 1258.1 | 1105 | 1200 | 1201.8327 | 1300 | 1300 | 1300 | 1240 | [5 empty] | 16
| 1 | 4 | 1100 | 1100 | 1200 | 1200 | 200 | 1150 | 1254 | 1145 | 1200 | 1202.9086 | 1300 | 1300 | 1200 | 1230 | [5 empty] | 1204.7531
| 1 | 5 | 1100 | 1300 | 1200 | 1200 | 200 | 1170 | 1269 | 1150 | 1300 | 1175.3946 | 1300 | 1300 | 1200 | 1240 | [5 empty] | 9
| 1 | 6 | 1000 | 1100 | 1200 | 1200 | 200 | 1130 | 1258.1 | 1105 | 1200 | 1174.3425 | 1200 | 1300 | 1300 | 1170
| Mean |  | 1066.6667 | 1200 | 1200 | 1200 | 233.3333 | 1165 | 1264.6833 | 1123.3333 | 1216.6667 | 1193.7735 | 1266.6667 | 1283.3333 | 1266.6667 | 1215
| Median |  | 1100 | 1200 | 1200 | 1200 | 200 | 1165 | 1263.55 | 1122.5 | 1200 | 1202.3706 | 1300 | 1300 | 1300 | 1220
| Std Dev. |  | 51.6398 | 89.4427 | 0 | 0 | 51.6398 | 24.2899 | 9.0623 | 24.2212 | 40.8248 | 14.6732 | 51.6398 | 40.8248 | 51.6398 | 27.3861
| Rel.Std.Dev. |  | 4.841% | 7.454% | 0% | 0% | 22.13% | 2.085% | 0.7166% | 2.156% | 3.355% | 1.229% | 4.077% | 3.181% | 4.077% | 2.254%
| PDM3 |  | -11.46% | -0.3945% | -0.3945% | -0.3945% | -80.63% | -3.3% | 4.974% | -6.758% | 0.9889% | -0.9114% | 5.139% | 6.523% | 5.139% | 0.8505%
| Z-Score (Absolute) |  | 1.8 | 0 | 0 | 0 | 13.08 | 0.47 | 0.88 | 1.04 | 0.23 | 0.08 | 0.9 | 1.13 | 0.9 | 0.2
| Analytical results for Bi in OREAS 782 (Indicative Value 266 ppm) | [20 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 200 | 200 | < 100 | 270 | 360 | 300 | 300 | 300 | [11 empty] | 1
| 1 | 2 | 200 | 200 | < 100 | 270 | 360 | 300 | 300 | 200 | [11 empty] | 1
| 1 | 3 | 200 | 200 | < 100 | 270 | 360 | 300 | 300 | 200 | [11 empty] | 16
| 1 | 4 | 200 | 200 | < 100 | 270 | 360 | 300 | 300 | 200 | [11 empty] | 265.6452
| 1 | 5 | 200 | 200 | < 100 | 270 | 360 | 300 | 300 | 200 | [11 empty] | 7
| 1 | 6 | 200 | 200 | < 100 | 270 | 450 | 300 | 300 | 200
| Mean |  | 200 | 200 |  | 270 | 375 | 300 | 300 | 216.6667
| Median |  | 200 | 200 |  | 270 | 360 | 300 | 300 | 200
| Std Dev. |  | 0 | 0 |  | 0 | 36.7423 | 0 | 0 | 40.8248
| Rel.Std.Dev. |  | 0% | 0% |  | 0% | 9.798% | 0% | 0% | 18.84%
| PDM3 |  | -24.71% | -24.71% |  | 1.639% | 41.17% | 12.93% | 12.93% | -18.44%
| Z-Score (Absolute) |  | 0.58 | 0.58 | 2.6 | 0.35 | 1.76 | 0.77 | 0.77 | 0.35
| Analytical results for Ca in OREAS 782 (Certified Value 16.45 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 17.024 | 17.524 | 16.366 | 16.509 | 16.6781 | 16.52 | 17.281 | 16.31 | 16.474 | 16.274 | 16.359 | 16.7764 | 15.8267 | 16.138 | 16.6419 | 16.581 | 16.366 | 16.438 | 16.31 | 1
| 1 | 2 | 17.031 | 17.524 | 16.224 | 16.581 | 16.6922 | 16.58 | 17.288 | 16.25 | 16.502 | 16.467 | 16.352 | 16.7589 | 15.8364 | 16.209 | 16.6496 | 16.366 | 16.224 | 16.366 | 16.4072 | 22
| 1 | 3 | 16.86 | 17.517 | 16.366 | 16.366 | 16.6918 | 16.65 | 17.925 | 16.43 | 16.388 | 16.359 | 16.374 | 16.7435 | 15.755 | 16.274 | 16.6706 | 16.509 | 16.366 | 16.438 | 16.3929 | 16
| 1 | 4 | 16.988 | 17.531 | 16.366 | 16.438 | 16.7227 | 16.73 | 17.246 | 16.47 | 16.274 | 16.202 | 16.252 | 16.8053 | 15.8156 | 16.231 | 16.7361 | 16.438 | 16.295 | 16.366 | 16.4523 | 16.4541
| 1 | 5 | 17.038 | 17.517 | 16.438 | 16.438 | 16.6441 | 16.58 | 17.353 | 16.43 | 16.395 | 16.231 | 16.324 | 16.7596 | 15.7951 | 16.202 | 16.6926 | 16.438 | 16.438 | 16.438 | 16.3951 | 10
| 1 | 6 | 17.045 | 17.531 | 16.295 | 16.438 | 16.732 | 16.52 | 17.16 | 16.54 | 16.395 | 16.274 | 16.295 | 16.7967 | 15.7524 | 16.252 | 16.6191 | 16.366 | 16.366 | 16.509 | 16.4594
| Mean |  | 16.9977 | 17.524 | 16.3425 | 16.4617 | 16.6935 | 16.5967 | 17.3755 | 16.405 | 16.4047 | 16.3012 | 16.326 | 16.7734 | 15.7969 | 16.2177 | 16.6683 | 16.4497 | 16.3425 | 16.4258 | 16.4028
| Median |  | 17.0275 | 17.524 | 16.366 | 16.438 | 16.692 | 16.58 | 17.2845 | 16.43 | 16.395 | 16.274 | 16.338 | 16.768 | 15.8054 | 16.22 | 16.6601 | 16.438 | 16.366 | 16.438 | 16.4012
| Std Dev. |  | 0.07031 | 0.006261 | 0.073587 | 0.07391 | 0.031664 | 0.081158 | 0.2765 | 0.10654 | 0.079774 | 0.097035 | 0.045848 | 0.023954 | 0.036123 | 0.047323 | 0.041623 | 0.083715 | 0.073587 | 0.05389 | 0.053744
| Rel.Std.Dev. |  | 0.4136% | 0.03573% | 0.4503% | 0.449% | 0.1897% | 0.489% | 1.591% | 0.6494% | 0.4863% | 0.5953% | 0.2808% | 0.1428% | 0.2287% | 0.2918% | 0.2497% | 0.5089% | 0.4503% | 0.3281% | 0.3277%
| PDM3 |  | 3.303% | 6.502% | -0.6784% | 0.04581% | 1.455% | 0.8663% | 5.6% | -0.2986% | -0.3006% | -0.9296% | -0.7787% | 1.94% | -3.994% | -1.437% | 1.302% | -0.02712% | -0.6784% | -0.172% | -0.3119%
| Z-Score (Absolute) |  | 3.08 | 5.92 | 0.45 | 0.19 | 1.44 | 0.92 | 5.12 | 0.11 | 0.12 | 0.67 | 0.54 | 1.87 | 3.39 | 1.12 | 1.31 | 0.13 | 0.45 | 0 | 0.13
| Analytical results for Cd in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 2
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for CeO2 in OREAS 782 (Indicative Value 119 ppm) | [20 empty] | Indicative
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 13 | 15 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 200 | 100 | 70 | 50 | 210 | 100 | 200 | 100 | [11 empty] | 1
| 1 | 2 | 100 | < 100 | 130 | 40 | 310 | < 100 | 200 | < 100 | [11 empty] | 3
| 1 | 3 | 100 | 100 | 120 | 70 | < 210 | 100 | 300 | 100 | [11 empty] | 16
| 1 | 4 | 100 | 100 | 100 | 60 | 210 | < 100 | 200 | < 100 | [11 empty] | 118.7656
| 1 | 5 | 200 | 100 | 130 | 40 | < 210 | < 100 | 200 | 100 | [11 empty] | 9
| 1 | 6 | 200 | < 100 | 130 | 60 | < 210 | < 100 | 200 | 100
| Mean |  | 150 | 100 | 113.3333 | 53.3333 | 243.3333 | 100 | 216.6667 | 100
| Median |  | 150 | 100 | 125 | 55 | 210 | 100 | 200 | 100
| Std Dev. |  | 54.7723 | 0 | 24.2212 | 12.1106 | 57.735 | 0 | 40.8248 | 0
| Rel.Std.Dev. |  | 36.51% | 0% | 21.37% | 22.71% | 23.73% | 0% | 18.84% | 0%
| PDM3 |  | 26.3% | -15.8% | -4.574% | -55.09% | 104.9% | -15.8% | 82.43% | -15.8%
| Z-Score (Absolute) |  | 0.91 | 0.26 | 0.26 | 0.79 | 1.34 | 0.56 | 2.08 | 0.26
| Analytical results for Cl in OREAS 782 (Indicative Value 65 ppm) | [20 empty] | Indicative
Cl | Round | Replicate | Lab | [18 empty] | 1
| No | No | 13 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | 80 | [18 empty] | 1
| 1 | 2 | 60 | [18 empty] | 4
| 1 | 3 | 70 | [18 empty] | 16
| 1 | 4 | 60 | [18 empty] | 65
| 1 | 5 | 70 | [18 empty] | 10
| 1 | 6 | 50
| Mean |  | 65
| Median |  | 65
| Std Dev. |  | 10.4881
| Rel.Std.Dev. |  | 16.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 782 (Certified Value < 100 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 37 | < 39 | 70 | 100 | < 100 | < 100 | [10 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 37 | < 39 | 70 | < 100 | < 100 | < 100 | [10 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 37 | < 39 | 70 | 100 | < 100 | < 100 | [10 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100 | [10 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 37 | < 39 | 70 | < 100 | < 100 | < 100 | [10 empty] | 11
| 1 | 6 | < 100 | < 100 | < 100 | < 37 | < 39 | < 70 | < 100 | < 100 | < 100
| Mean | [6 empty] | 70 | 100
| Median | [6 empty] | 70 | 100
| Std Dev. | [6 empty] | 0 | 0
| Rel.Std.Dev. | [6 empty] | 0% | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cr in OREAS 782 (Certified Value 78 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 89 | 68 | < 100 | 100 | 109.41 | < 70 | < 70 | 70 | 70 | 82 | 92.4 | 78.7 | 70 | 80.9369 | 100 | 100 | 100 | 55 |  | 1
| 1 | 2 | 103 | 68 | < 100 | 100 | 127.4 | < 70 | < 70 | 70 | < 70 | 75 | 88.9 | 85.5 | 140 | 85.187 | 100 | 100 | 100 | 89 |  | #N/A
| 1 | 3 | 109 | 68 | < 100 | 100 | 91.69 | < 70 | < 70 | 70 | 70 | 75 | 99.2 | 71.8 | 140 | 78.2246 | 300 | 100 | 100 | 68 |  | 16
| 1 | 4 | 109 | 68 | < 100 | 100 | 105.64 | < 70 | < 70 | 70 | 70 | 75 | 102.6 | 82.1 | 140 | 64.6555 | 100 | 100 | 100 | 62 |  | 77.8954
| 1 | 5 | 96 | 68 | < 100 | 100 | 125.21 | < 70 | < 70 | 70 | < 70 | 82 | 102.6 | 68.4 | 140 | 60.7704 | 100 | 100 | 100 | 103 |  | 12
| 1 | 6 | 82 | 68 | < 100 | 100 | 109.41 | < 70 | < 70 | 70 | < 70 | 68 | 88.9 | 78.7 | 140 | 66.4758 | 100 | 100 | 100 | 75
| Mean |  | 98 | 68 |  | 100 | 111.46 |  |  | 70 | 70 | 76.1667 | 95.7667 | 77.5333 | 128.3333 | 72.7084 | 133.3333 | 100 | 100 | 75.3333
| Median |  | 99.5 | 68 |  | 100 | 109.41 |  |  | 70 | 70 | 75 | 95.8 | 78.7 | 140 | 72.3502 | 100 | 100 | 100 | 71.5
| Std Dev. |  | 11.0272 | 0 |  | 0 | 13.2384 |  |  | 0 | 0 | 5.2694 | 6.4933 | 6.377 | 28.5774 | 10.0006 | 81.6497 | 0 | 0 | 17.8736
| Rel.Std.Dev. |  | 11.25% | 0% |  | 0% | 11.88% |  |  | 0% | 0% | 6.918% | 6.78% | 8.225% | 22.27% | 13.75% | 61.24% | 0% | 0% | 23.73%
| PDM3 |  | 25.81% | -12.7% |  | 28.38% | 43.09% |  |  | -10.14% | -10.14% | -2.219% | 22.94% | -0.4648% | 64.75% | -6.659% | 71.17% | 28.38% | 28.38% | -3.289%
| Z-Score (Absolute) |  | 0.73 | 0.17 | 0.73 | NA | 1.14 | 1.21 | 1.21 | 0.17 | 0.69 | 0.07 | 0.66 | 0.11 | 1.56 | 0.04 | NA | NA | NA | 0.04
Results from laboratories 4, 21, 22 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 782 (Indicative Value 118 ppm) | [20 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 08 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | < 100 | < 90 | [17 empty] | 1
| 1 | 2 | < 100 | 90 | [17 empty] | 5
| 1 | 3 | < 100 | 90 | [17 empty] | 16
| 1 | 4 | < 100 | 190 | [17 empty] | 117.9023
| 1 | 5 | < 100 | 90 | [17 empty] | 11
| 1 | 6 | < 100 | 190
| Mean |  |  | 130
| Median |  |  | 90
| Std Dev. |  |  | 54.7723
| Rel.Std.Dev. |  |  | 42.13%
| PDM3 |  |  | 10.26%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 782 (Certified Value 85 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 150 | 80 | < 100 | 72 | 48 | < 160 | 87.9 | 80 | 69.7528 | 110 | 290 | 90 | 80 | [6 empty] | 1
| 1 | 2 | 70 | 100 | < 100 | 80 | 48 | < 160 | 95.9 | 80 | 72.0973 | 100 | 140 | 120 | 80 | [6 empty] | #N/A
| 1 | 3 | 80 | 80 | < 100 | 72 | 48 | < 160 | 79.9 | 80 | 63.5214 | 100 | 120 | 100 | 80 | [6 empty] | 16
| 1 | 4 | 100 | 80 | < 100 | 56 | 48 | < 160 | 91.9 | 80 | 72.3446 | 100 | 150 | 90 | 88 | [6 empty] | 84.7538
| 1 | 5 | 80 | 100 | < 100 | 64 | 64 | < 160 | 87.9 | 80 | 73.6228 | 90 | 140 | 100 | 80 | [6 empty] | 13
| 1 | 6 | 100 | 90 | < 100 | 64 | 48 | < 160 | 99.9 | 80 | 73.4035 | 130 | 130 | 80 | 80
| Mean |  | 96.6667 | 88.3333 |  | 68 | 50.6667 |  | 90.5667 | 80 | 70.7904 | 105 | 161.6667 | 96.6667 | 81.3333
| Median |  | 90 | 85 |  | 68 | 48 |  | 89.9 | 80 | 72.2209 | 100 | 140 | 95 | 80
| Std Dev. |  | 28.7518 | 9.8319 |  | 8.3905 | 6.532 |  | 7.0048 | 0 | 3.8181 | 13.784 | 63.692 | 13.6626 | 3.266
| Rel.Std.Dev. |  | 29.74% | 11.13% |  | 12.34% | 12.89% |  | 7.734% | 0% | 5.394% | 13.13% | 39.4% | 14.13% | 4.016%
| PDM3 |  | 14.06% | 4.223% |  | -19.77% | -40.22% |  | 6.859% | -5.609% | -16.48% | 23.89% | 90.75% | 14.06% | -4.036%
| Z-Score (Absolute) |  | 0.78 | 0.36 | 1.58 | 0.67 | 1.55 | 0.07 | 0.47 | 0.07 | 0.53 | 1.2 | 4.07 | 0.78 | 0
| Analytical results for Dy in OREAS 782 (Indicative Value 247 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | 260 | [18 empty] | 1
| 1 | 2 | 260 | [18 empty] | 6
| 1 | 3 | 170 | [18 empty] | 16
| 1 | 4 | 260 | [18 empty] | 246.8706
| 1 | 5 | 260 | [18 empty] | 12
| 1 | 6 | 260
| Mean |  | 245
| Median |  | 260
| Std Dev. |  | 36.7423
| Rel.Std.Dev. |  | 15%
| PDM3 |  | -0.7577%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 782 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 7
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 13
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for F in OREAS 782 (Certified Value 15.2 wt.%) | [20 empty] | Certified
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 08 | 13 | 15 | 17 | 22 | 23 | [9 empty] | wt.%
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 2
3
| 1 | 1 | 15.45 | 19.5552 | 15.8 | 15.1644 | 15.8 | 14.7 | 14.992 | 15.1 | 15.2 | 15 | [9 empty] | 1
| 1 | 2 | 15.34 | 19.3494 | 15.6 | 15.1152 | 15.85 | 14.6 | 15.1179 | 15.2 | 15.1 | 15 | [9 empty] | #N/A
| 1 | 3 | 15.22 | 19.3494 | 15.5 | 15.0488 | 15.79 | 14.6 | 15.0292 | 15 | 15.4 | 14.9 | [9 empty] | 16
| 1 | 4 | 15.44 | 19.2404 | 15.7 | 14.9937 | 15.82 | 14.6 | 15.1475 | 14.7 | 15.1 | 15.2 | [9 empty] | 15.1962
| 1 | 5 | 15.37 | 19.188 | 15.4 | 15.236 | 15.84 | 14.4 | 15.075 | 15.2 | 15.2 | 14.8 | [9 empty] | 14
| 1 | 6 | 15.5 | 19.1324 | 15.4 | 15.0383 | 15.83 | 14.6 | 15.0883 | 15 | 15.5 | 14.8
| Mean |  | 15.3867 | 19.3025 | 15.5667 | 15.0994 | 15.8217 | 14.5833 | 15.075 | 15.0333 | 15.25 | 14.95
| Median |  | 15.405 | 19.2949 | 15.55 | 15.082 | 15.825 | 14.6 | 15.0816 | 15.05 | 15.2 | 14.95
| Std Dev. |  | 0.099933 | 0.15109 | 0.1633 | 0.090114 | 0.023166 | 0.098319 | 0.057027 | 0.18619 | 0.16432 | 0.15166
| Rel.Std.Dev. |  | 0.6495% | 0.7828% | 1.049% | 0.5968% | 0.1464% | 0.6742% | 0.3783% | 1.239% | 1.077% | 1.014%
| PDM3 |  | 1.253% | 27.02% | 2.438% | -0.6372% | 4.116% | -4.033% | -0.7979% | -1.072% | 0.3539% | -1.62%
| Z-Score (Absolute) |  | 0.65 | 12.75 | 1.21 | 0.23 | 2 | 1.83 | 0.31 | 0.44 | 0.23 | 0.69
| Analytical results for Fe in OREAS 782 (Certified Value 1.83 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.882 | 2.0494 | 1.82 | 1.8 | 1.8315 | 1.78 | 1.854 | 1.84 | 1.833 | 1.777 | 1.812 | 1.8957 | 1.7297 | 1.861 | 1.8762 | 1.88 | 1.85 | 1.82 | 1.7563 | 1
| 1 | 2 | 1.875 | 2.0494 | 1.78 | 1.8 | 1.8386 | 1.77 | 1.875 | 1.833 | 1.847 | 1.784 | 1.812 | 1.8904 | 1.7236 | 1.854 | 1.8867 | 1.86 | 1.9 | 1.82 | 1.7549 | 28
| 1 | 3 | 1.875 | 2.0424 | 1.81 | 1.8 | 1.8409 | 1.83 | 1.847 | 1.798 | 1.826 | 1.784 | 1.812 | 1.896 | 1.719 | 1.854 | 1.8761 | 1.87 | 1.95 | 1.82 | 1.7605 | 16
| 1 | 4 | 1.882 | 2.0564 | 1.8 | 1.8 | 1.8603 | 1.8 | 1.882 | 1.805 | 1.805 | 1.763 | 1.798 | 1.8963 | 1.7297 | 1.84 | 1.8849 | 1.88 | 1.86 | 1.84 | 1.7675 | 1.8293
| 1 | 5 | 1.875 | 2.0424 | 1.82 | 1.8 | 1.8341 | 1.86 | 1.868 | 1.84 | 1.819 | 1.77 | 1.805 | 1.8977 | 1.7161 | 1.847 | 1.8864 | 1.86 | 1.94 | 1.84 | 1.7675 | 15
| 1 | 6 | 1.868 | 2.0564 | 1.79 | 1.78 | 1.8373 | 1.78 | 1.833 | 1.854 | 1.826 | 1.798 | 1.812 | 1.8953 | 1.7236 | 1.861 | 1.8728 | 1.85 | 1.91 | 1.83 | 1.7514
| Mean |  | 1.8762 | 2.0494 | 1.8033 | 1.7967 | 1.8404 | 1.8033 | 1.8598 | 1.8283 | 1.826 | 1.7793 | 1.8085 | 1.8952 | 1.7236 | 1.8528 | 1.8805 | 1.8667 | 1.9017 | 1.8283 | 1.7597
| Median |  | 1.875 | 2.0494 | 1.805 | 1.8 | 1.8379 | 1.79 | 1.861 | 1.8365 | 1.826 | 1.7805 | 1.812 | 1.8958 | 1.7236 | 1.854 | 1.8805 | 1.865 | 1.905 | 1.825 | 1.7584
| Std Dev. |  | 0.0052694 | 0.006261 | 0.01633 | 0.008165 | 0.010295 | 0.035024 | 0.018476 | 0.021988 | 0.014 | 0.012258 | 0.0058566 | 0.0025165 | 0.0055201 | 0.0081833 | 0.0061401 | 0.012111 | 0.040702 | 0.0098319 | 0.0067202
| Rel.Std.Dev. |  | 0.2809% | 0.3055% | 0.9055% | 0.4545% | 0.5594% | 1.942% | 0.9934% | 1.203% | 0.7667% | 0.6889% | 0.3238% | 0.1328% | 0.3203% | 0.4417% | 0.3265% | 0.6488% | 2.14% | 0.5378% | 0.3819%
| PDM3 |  | 2.562% | 12.03% | -1.419% | -1.784% | 0.6095% | -1.419% | 1.669% | -0.05279% | -0.1803% | -2.731% | -1.137% | 3.605% | -5.777% | 1.287% | 2.8% | 2.043% | 3.956% | -0.05279% | -3.806%
| Z-Score (Absolute) |  | 1.01 | 4.71 | 0.53 | 0.67 | 0.26 | 0.53 | 0.66 | 0.01 | 0.06 | 1.05 | 0.43 | 1.42 | 2.23 | 0.51 | 1.11 | 0.82 | 1.56 | 0 | 1.46
| Analytical results for Ga in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 08 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | < 100 | < 150 | [17 empty] | 1
| 1 | 2 | < 100 | < 150 | [17 empty] | 8
| 1 | 3 | < 100 | < 150 | [17 empty] | 16
| 1 | 4 | < 100 | < 150 | [17 empty] | < 100
| 1 | 5 | < 100 | < 150 | [17 empty] | 14
| 1 | 6 | < 100 | < 150
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Gd in OREAS 782 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 9
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 15
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 10
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 16
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for H2O- in OREAS 782 (Indicative Value 0.263 wt.%) | [20 empty] | Indicative
H2O- | Round | Replicate | Lab | [18 empty] | 1
| No | No | 20 | [18 empty] | wt.%
|  |  | BF*XRF | [18 empty] | 3
3
| 1 | 1 | 0.2875 | [18 empty] | 1
| 1 | 2 | 0.2807 | [18 empty] | 11
| 1 | 3 | 0.2125 | [18 empty] | 16
| 1 | 4 | 0.2518 | [18 empty] | 0.26265
| 1 | 5 | 0.2706 | [18 empty] | 17
| 1 | 6 | 0.2728
| Mean |  | 0.26265
| Median |  | 0.2717
| Std Dev. |  | 0.027358
| Rel.Std.Dev. |  | 10.42%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 782 (Certified Value < 100 ppm) | [20 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 09 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 200 | 20 | 200 | < 100 | < 100 | < 100 | < 100 | [8 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 200 |  | 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 200 |  | 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 200 | 5 | 100 | < 100 | < 100 | < 100 | < 100 | [8 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 200 | 20 | 200 | < 100 | < 100 | < 100 | < 100 | [8 empty] | 16
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 200 |  | 100 | < 100 | < 100 | < 100 | < 100
| Mean | [6 empty] | 15 | 133.3333
| Median | [6 empty] | 20 | 100
| Std Dev. | [6 empty] | 8.6603 | 51.6398
| Rel.Std.Dev. | [6 empty] | 57.74% | 38.73%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Hg in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 12
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
In | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 1
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 782 (Certified Value 1.71 wt.%) | [20 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.717 | 1.87 | 1.76 | 1.77 | 1.6445 | 1.71 | 0.46 | 1.67 | 1.66 | 1.72 | 1.73 | 1.743 | 1.7003 | 1.7 | 1.6376 | 1.71 | 1.77 | 1.76 | 1.66 | 1
| 1 | 2 | 1.714 | 1.87 | 1.76 | 1.78 | 1.6426 | 1.73 | 0.47 | 1.64 | 1.66 | 1.73 | 1.72 | 1.7345 | 1.6942 | 1.7 | 1.5812 | 1.72 | 1.76 | 1.76 | 1.669 | #N/A
| 1 | 3 | 1.717 | 1.86 | 1.74 | 1.76 | 1.6349 | 1.79 | 0.47 | 1.63 | 1.63 | 1.73 | 1.73 | 1.7455 | 1.6912 | 1.71 | 1.6402 | 1.71 | 1.77 | 1.74 | 1.671 | 16
| 1 | 4 | 1.724 | 1.88 | 1.72 | 1.78 | 1.6258 | 1.76 | 0.49 | 1.67 | 1.63 | 1.72 | 1.72 | 1.741 | 1.6968 | 1.72 | 1.6077 | 1.7 | 1.76 | 1.74 | 1.67 | 1.7124
| 1 | 5 | 1.722 | 1.86 | 1.76 | 1.78 | 1.6565 | 1.7 | 0.45 | 1.67 | 1.64 | 1.74 | 1.73 | 1.7355 | 1.6901 | 1.72 | 1.6131 | 1.7 | 1.78 | 1.74 | 1.652 | 17
| 1 | 6 | 1.706 | 1.88 | 1.76 | 1.76 | 1.6454 | 1.73 | 0.46 | 1.66 | 1.64 | 1.71 | 1.72 | 1.7395 | 1.6952 | 1.71 | 1.6049 | 1.69 | 1.77 | 1.75 | 1.673
| Mean |  | 1.7167 | 1.87 | 1.75 | 1.7717 | 1.6416 | 1.7367 | 0.46667 | 1.6567 | 1.6433 | 1.725 | 1.725 | 1.7398 | 1.6946 | 1.71 | 1.6141 | 1.705 | 1.7683 | 1.7483 | 1.6658
| Median |  | 1.717 | 1.87 | 1.76 | 1.775 | 1.6436 | 1.73 | 0.465 | 1.665 | 1.64 | 1.725 | 1.725 | 1.7403 | 1.6947 | 1.71 | 1.6104 | 1.705 | 1.77 | 1.745 | 1.6695
| Std Dev. |  | 0.006377 | 0.0089443 | 0.016733 | 0.0098319 | 0.010398 | 0.033267 | 0.013663 | 0.017512 | 0.013663 | 0.010488 | 0.0054772 | 0.0042622 | 0.0037396 | 0.0089443 | 0.022117 | 0.010488 | 0.0075277 | 0.0098319 | 0.0081343
| Rel.Std.Dev. |  | 0.3715% | 0.4783% | 0.9562% | 0.555% | 0.6334% | 1.916% | 2.928% | 1.057% | 0.8314% | 0.608% | 0.3175% | 0.245% | 0.2207% | 0.5231% | 1.37% | 0.6151% | 0.4257% | 0.5624% | 0.4883%
| PDM3 |  | 0.2484% | 9.203% | 2.195% | 3.46% | -4.134% | 1.416% | -72.75% | -3.255% | -4.034% | 0.7351% | 0.7351% | 1.601% | -1.038% | -0.1409% | -5.739% | -0.4329% | 3.266% | 2.098% | -2.72%
| Z-Score (Absolute) |  | 0.1 | 3.55 | 0.69 | 1.21 | 1.89 | 0.38 | NA | 1.53 | 1.84 | 0.1 | 0.1 | 0.45 | 0.62 | 0.26 | 2.54 | 0.38 | 1.13 | 0.65 | 1.31
Results from laboratory 1.08 were removed due to their systematic bias.
| Analytical results for La2O3 in OREAS 782 (Indicative Value 89 ppm) | [20 empty] | Indicative
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 13 | 15 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 80 | 60 | < 200 | 100 | 100 | 100 | [11 empty] | 1
| 1 | 2 | < 100 | < 100 | 60 | 60 | < 200 | < 100 | 100 | 100 | [11 empty] | 2
| 1 | 3 | < 100 | 100 | 60 | 50 | < 200 | 100 | 100 | 100 | [11 empty] | 16
| 1 | 4 | < 100 | < 100 | 70 | 60 | < 200 | 100 | 100 | 100 | [11 empty] | 88.5714
| 1 | 5 | < 100 | 100 | 50 | 60 | < 200 | 100 | 200 | 100 | [11 empty] | 8
| 1 | 6 | < 100 | < 100 | 60 | 50 | 200 | < 100 | 200 | 100
| Mean |  |  | 100 | 63.3333 | 56.6667 | 200 | 100 | 133.3333 | 100
| Median |  |  | 100 | 60 | 60 | 200 | 100 | 100 | 100
| Std Dev. |  |  | 0 | 10.328 | 5.164 |  | 0 | 51.6398 | 0
| Rel.Std.Dev. |  |  | 0% | 16.31% | 9.113% |  | 0% | 38.73% | 0%
| PDM3 |  |  | 12.9% | -28.49% | -36.02% | 125.8% | 12.9% | 50.54% | 12.9%
| Z-Score (Absolute) |  | 0.78 | 0.26 | 0.36 | 0.57 | 1.3 | 0.26 | 1.82 | 0.78
| Analytical results for MgO in OREAS 782 (Certified Value 0.348 wt.%) | [20 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.33 | 0.45 | 0.34 | 0.35 | 0.41301 | 0.35 | 0.07 | 0.31 | 0.35 | 0.38 | 0.35 | 0.35959 | 0.3455 | 0.4 | 0.36111 | 0.34 | 0.36 | 0.35 | 0.328 | 1
| 1 | 2 | 0.34 | 0.45 | 0.34 | 0.35 | 0.4164 | 0.4 | 0.07 | 0.3 | 0.35 | 0.39 | 0.36 | 0.35959 | 0.355 | 0.4 | 0.36907 | 0.35 | 0.36 | 0.34 | 0.308 | #N/A
| 1 | 3 | 0.33 | 0.46 | 0.35 | 0.35 | 0.42354 | 0.33 | 0.07 | 0.31 | 0.34 | 0.4 | 0.36 | 0.35798 | 0.347 | 0.4 | 0.36395 | 0.33 | 0.38 | 0.34 | 0.32 | 16
| 1 | 4 | 0.33 | 0.44 | 0.35 | 0.35 | 0.41309 | 0.39 | 0.07 | 0.3 | 0.33 | 0.4 | 0.36 | 0.36388 | 0.347 | 0.4 | 0.36758 | 0.34 | 0.37 | 0.35 | 0.313 | 0.34794
| 1 | 5 | 0.34 | 0.46 | 0.36 | 0.34 | 0.42124 | 0.38 | 0.07 | 0.3 | 0.33 | 0.38 | 0.36 | 0.36065 | 0.346 | 0.4 | 0.3649 | 0.34 | 0.37 | 0.34 | 0.307 | 18
| 1 | 6 | 0.32 | 0.44 | 0.35 | 0.32 | 0.40756 | 0.38 | 0.07 | 0.3 | 0.32 | 0.39 | 0.36 | 0.36118 | 0.347 | 0.4 | 0.36579 | 0.32 | 0.36 | 0.34 | 0.316
| Mean |  | 0.33167 | 0.45 | 0.34833 | 0.34333 | 0.41581 | 0.37167 | 0.07 | 0.30333 | 0.33667 | 0.39 | 0.35833 | 0.36048 | 0.34792 | 0.4 | 0.3654 | 0.33667 | 0.36667 | 0.34333 | 0.31533
| Median |  | 0.33 | 0.45 | 0.35 | 0.35 | 0.41474 | 0.38 | 0.07 | 0.3 | 0.335 | 0.39 | 0.36 | 0.36012 | 0.347 | 0.4 | 0.36535 | 0.34 | 0.365 | 0.34 | 0.3145
| Std Dev. |  | 0.0075277 | 0.0089443 | 0.0075277 | 0.012111 | 0.0058803 | 0.026394 | 0 | 0.005164 | 0.012111 | 0.0089443 | 0.0040825 | 0.0019964 | 0.0035273 | 6.0809e-17 | 0.0027964 | 0.010328 | 0.008165 | 0.005164 | 0.0078909
| Rel.Std.Dev. |  | 2.27% | 1.988% | 2.161% | 3.527% | 1.414% | 7.102% | 0% | 1.702% | 3.597% | 2.293% | 1.139% | 0.5538% | 1.014% | 1.52e-14% | 0.7653% | 3.068% | 2.227% | 1.504% | 2.502%
| PDM3 |  | -4.678% | 29.33% | 0.1125% | -1.325% | 19.5% | 6.819% | -79.88% | -12.82% | -3.241% | 12.09% | 2.987% | 3.603% | -0.00726% | 14.96% | 5.018% | -3.241% | 5.382% | -1.325% | -9.372%
| Z-Score (Absolute) |  | 0.66 | 4.07 | 0.01 | 0.19 | 2.71 | 0.94 | 11.12 | 1.79 | 0.46 | 1.67 | 0.41 | 0.49 | 0.01 | NA | 0.69 | 0.46 | 0.74 | 0.19 | 1.31
Results from laboratory 17 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Mn in OREAS 782 (Indicative Value 0.008 wt.%) | [20 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 3
3
| 1 | 1 | < 0.0039 | 0.008 | < 0.01 | 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0062 | < 0.008 | < 0.008 | 0.00503 | 0.0045379 | < 0.007 | 0.0061562 | 0.01 | 0.01 | 0.01 | 0.0085 |  | 1
| 1 | 2 | < 0.0039 | 0.008 | < 0.01 | 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0062 | < 0.008 | < 0.008 | 0.00581 | 0.0026471 | < 0.007 | 0.0040906 | 0.01 | 0.01 | 0.01 | 0.0085 |  | 3
| 1 | 3 | < 0.0039 | 0.008 | < 0.01 | 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0077 | < 0.008 | < 0.008 | 0.00542 | 0.0034034 | < 0.007 | 0.0040977 | 0.01 | 0.01 | 0.01 | 0.0077 |  | 16
| 1 | 4 | < 0.0039 | 0.015 | < 0.01 | 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.007 | < 0.008 | < 0.008 | 0.00581 | 0.0037816 | < 0.007 | 0.0032932 | 0.01 | 0.01 | 0.01 | 0.0054 |  | 0.0078275
| 1 | 5 | < 0.0039 | 0.008 | < 0.01 | 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.0054 | < 0.008 | < 0.008 | 0.0062 | 0.0030253 | < 0.007 | 0.0054881 | < 0.01 | 0.01 | 0.01 | 0.0077 |  | 9
| 1 | 6 | < 0.0039 | 0.015 | < 0.01 | 0.01 | < 0.008 | < 0.008 | < 0.008 | 0.007 | < 0.008 | < 0.008 | 0.00542 | 0.0045379 | < 0.007 | 0.0045291 | 0.01 | 0.01 | 0.01 | 0.007
| Mean |  |  | 0.010333 |  | 0.01 |  |  |  | 0.0065833 |  |  | 0.005615 | 0.0036555 |  | 0.0046091 | 0.01 | 0.01 | 0.01 | 0.0074667
| Median |  |  | 0.008 |  | 0.01 |  |  |  | 0.0066 |  |  | 0.005615 | 0.0035925 |  | 0.0043134 | 0.01 | 0.01 | 0.01 | 0.0077
| Std Dev. |  |  | 0.0036148 |  | 0 |  |  |  | 0.00081097 |  |  | 0.00040904 | 0.00078112 |  | 0.0010426 | 0 | 0 | 0 | 0.0011605
| Rel.Std.Dev. |  |  | 34.98% |  | 0% |  |  |  | 12.32% |  |  | 7.285% | 21.37% |  | 22.62% | 0% | 0% | 0% | 15.54%
| PDM3 |  |  | 32.01% |  | 27.75% |  |  |  | -15.89% |  |  | -28.27% | -53.3% |  | -41.12% | 27.75% | 27.75% | 27.75% | -4.61%
| Z-Score (Absolute) |  | 1.64 | 3.17 | 0.11 | 2.98 | 0.53 | 0.53 | 0.53 | 1.02 | 0.53 | 0.53 | 0.46 | 0.66 | 0.69 | 0.11 | 2.5 | 2.98 | 2.98 | 1.54
| Analytical results for Mo in OREAS 782 (Indicative Value 53 ppm) | [20 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 50 | 110 | < 100 | 33 | < 70 | < 50 | 110 | 70 | [11 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | 53 | < 70 | < 50 | 110 | 60 | [11 empty] | 4
| 1 | 3 | < 50 | < 50 | < 100 | 67 | < 70 | < 50 | 110 | 50 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | 40 | < 70 | < 50 | 100 | 60 | [11 empty] | 53.1536
| 1 | 5 | < 50 | < 50 | < 100 | 53 | < 70 | < 50 | 110 | 60 | [11 empty] | 10
| 1 | 6 | 50 | < 50 | < 100 | 40 | < 70 | < 50 | 150 | 60
| Mean |  | 50 | 110 |  | 47.6667 |  |  | 115 | 60
| Median |  | 50 | 110 |  | 46.5 |  |  | 110 | 60
| Std Dev. |  |  |  |  | 12.3558 |  |  | 17.6068 | 6.3246
| Rel.Std.Dev. |  |  |  |  | 25.92% |  |  | 15.31% | 10.54%
| PDM3 |  | -5.933% | 106.9% |  | -10.32% |  |  | 116.4% | 12.88%
| Z-Score (Absolute) |  | 0.79 | 0.24 | 0.36 | 0.24 | 0.56 | 1.02 | 3.95 | 0.91
| Analytical results for Na2O in OREAS 782 (Indicative Value 0.138 wt.%) | [20 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 24 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.29 | 0.09 | < 0.1 | 0.03 | 0.18 | 0.15 | 0.12 | 0.18 | 0.103 | 0.09494 | 0.3 | 0.16333 | 0.134 | [6 empty] | 1
| 1 | 2 | 0.29 | 0.09 | < 0.1 | 0.03 | 0.22 | 0.16 | 0.14 | 0.19 | 0.099 | 0.11009 | 0.3 | 0.19203 | 0.129 | [6 empty] | 5
| 1 | 3 | 0.28 | 0.1 | < 0.1 | 0.03 | 0.18 | 0.14 | 0.15 | 0.17 | 0.1055 | 0.0909 | 0.3 | 0.16393 | 0.13 | [6 empty] | 16
| 1 | 4 | 0.3 | 0.08 | < 0.1 | 0.03 | 0.22 | 0.15 | 0.15 | 0.18 | 0.104 | 0.11918 | 0.3 | 0.15922 | 0.136 | [6 empty] | 0.13844
| 1 | 5 | 0.28 | 0.1 | 0.12 | 0.03 | 0.18 | 0.16 | 0.14 | 0.19 | 0.102 | 0.09898 | 0.2 | 0.16403 | 0.132 | [6 empty] | 11
| 1 | 6 | 0.29 | 0.08 | < 0.1 | 0.03 | 0.19 | 0.14 | 0.13 | 0.19 | 0.1045 | 0.1106 | 0.2 | 0.17544 | 0.125
| Mean |  | 0.28833 | 0.09 | 0.12 | 0.03 | 0.195 | 0.15 | 0.13833 | 0.18333 | 0.103 | 0.10411 | 0.26667 | 0.16966 | 0.131
| Median |  | 0.29 | 0.09 | 0.12 | 0.03 | 0.185 | 0.15 | 0.14 | 0.185 | 0.1035 | 0.10454 | 0.3 | 0.16398 | 0.131
| Std Dev. |  | 0.0075277 | 0.0089443 |  | 0 | 0.019748 | 0.0089443 | 0.01169 | 0.008165 | 0.0023022 | 0.010861 | 0.05164 | 0.012227 | 0.0038987
| Rel.Std.Dev. |  | 2.611% | 9.938% |  | 0% | 10.13% | 5.963% | 8.451% | 4.454% | 2.235% | 10.43% | 19.36% | 7.207% | 2.976%
| PDM3 |  | 108.3% | -34.99% | -13.32% | -78.33% | 40.85% | 8.347% | -0.08034% | 32.42% | -25.6% | -24.8% | 92.62% | 22.55% | -5.377%
| Z-Score (Absolute) |  | 2.6 | 0.76 | 1.24 | 1.77 | 1.02 | 0.26 | 0.06 | 0.82 | 0.54 | 0.52 | NA | 0.59 | 0.06
Results from laboratory 17 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Nb in OREAS 782 (Certified Value < 50 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 13 | 17 | 21 | 22 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | #N/A
| 1 | 3 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50 | [12 empty] | 19
| 1 | 6 | < 50 | < 50 | < 70 | < 70 | < 50 | < 50 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Nd in OREAS 782 (Indicative Value 29.3 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 13 | 15 | 17 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [16 empty] | 1
1
| 1 | 1 | < 43 | 25.7 | < 170 | [16 empty] | 1
| 1 | 2 | 51 | 12.9 | < 170 | [16 empty] | 6
| 1 | 3 | 60 | 17.1 | < 170 | [16 empty] | 16
| 1 | 4 | 60 | 25.7 | < 170 | [16 empty] | 29.2924
| 1 | 5 | < 43 | 8.6 | < 170 | [16 empty] | 12
| 1 | 6 | < 43 | 25.7 | < 170
| Mean |  | 57 | 19.2833
| Median |  | 60 | 21.4
| Std Dev. |  | 5.1962 | 7.5255
| Rel.Std.Dev. |  | 9.116% | 39.03%
| PDM3 |  | 94.59% | -34.17%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.57
| Analytical results for Ni in OREAS 782 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 08 | 10 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 15.7 | < 80 | 110 | < 50 | < 50 | < 47 | [7 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 23.6 | < 80 | 180 | 50 | < 50 | < 47 | [7 empty] | 7
| 1 | 3 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 11.8 | < 80 | 270 | 50 | < 50 | < 47 | [7 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 23.6 | < 80 | 90 | < 50 | < 50 | < 47 | [7 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 7.9 | < 80 | 90 | 50 | < 50 | < 47 | [7 empty] | 13
| 1 | 6 | < 50 | < 50 | < 100 | < 24 | < 39 | < 160 | 19.6 | < 80 | 60 | < 50 | < 50 | < 47
| Mean | [7 empty] | 17.0333 |  | 133.3333 | 50
| Median | [7 empty] | 17.65 |  | 100 | 50
| Std Dev. | [7 empty] | 6.4096 |  | 78.1452 | 0
| Rel.Std.Dev. | [7 empty] | 37.63% |  | 58.61% | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 1.65 | 0.87 | 0.35 | 3.53 | 0.53 | 0.94 | 7.14 | 0.82 | 0 | 0.09
| Analytical results for P2O5 in OREAS 782 (Certified Value 0.041 wt.%) | [20 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.034 | 0.04 | 0.03 | 0.04 | 0.03622 | 0.05 | < 0.01 | 0.04 | 0.04 | 0.045 | 0.05 | 0.04339 | 0.04284 | 0.04 | 0.043814 | 0.04 | 0.05 | 0.04 | 0.045 | 1
| 1 | 2 | 0.031 | 0.04 | 0.03 | 0.05 | 0.03575 | 0.05 | < 0.01 | 0.05 | 0.04 | 0.043 | 0.049 | 0.04339 | 0.04182 | 0.04 | 0.04565 | 0.04 | 0.04 | 0.04 | 0.044 | #N/A
| 1 | 3 | 0.031 | 0.05 | 0.03 | 0.05 | 0.038 | 0.04 | < 0.01 | 0.03 | 0.03 | 0.044 | 0.05 | 0.044835 | 0.04182 | 0.04 | 0.044313 | 0.04 | 0.04 | 0.04 | 0.04 | 16
| 1 | 4 | 0.033 | 0.03 | 0.03 | 0.05 | 0.03732 | 0.05 | < 0.01 | 0.04 | 0.03 | 0.044 | 0.051 | 0.04437 | 0.04335 | 0.04 | 0.044885 | 0.04 | 0.05 | 0.04 | 0.044 | 0.041204
| 1 | 5 | 0.032 | 0.05 | 0.04 | 0.05 | 0.03635 | 0.05 | < 0.01 | 0.03 | 0.04 | 0.044 | 0.049 | 0.044345 | 0.0408 | 0.04 | 0.04318 | 0.04 | 0.05 | 0.04 | 0.041 | 20
| 1 | 6 | 0.034 | 0.03 | 0.03 | 0.05 | 0.03927 | 0.04 | < 0.01 | 0.03 | 0.04 | 0.044 | 0.05 | 0.04388 | 0.04488 | 0.03 | 0.044521 | 0.04 | 0.05 | 0.04 | 0.039
| Mean |  | 0.0325 | 0.04 | 0.031667 | 0.048333 | 0.037152 | 0.046667 |  | 0.036667 | 0.036667 | 0.044 | 0.049833 | 0.044035 | 0.042585 | 0.038333 | 0.044394 | 0.04 | 0.046667 | 0.04 | 0.042167
| Median |  | 0.0325 | 0.04 | 0.03 | 0.05 | 0.036835 | 0.05 |  | 0.035 | 0.04 | 0.044 | 0.05 | 0.044112 | 0.04233 | 0.04 | 0.044417 | 0.04 | 0.05 | 0.04 | 0.0425
| Std Dev. |  | 0.0013784 | 0.0089443 | 0.0040825 | 0.0040825 | 0.0013199 | 0.005164 |  | 0.008165 | 0.005164 | 0.00063246 | 0.00075277 | 0.00058385 | 0.0014335 | 0.0040825 | 0.00085468 | 0 | 0.005164 | 0 | 0.0024833
| Rel.Std.Dev. |  | 4.241% | 22.36% | 12.89% | 8.447% | 3.553% | 11.07% |  | 22.27% | 14.08% | 1.437% | 1.511% | 1.326% | 3.366% | 10.65% | 1.925% | 0% | 11.07% | 0% | 5.889%
| PDM3 |  | -21.12% | -2.921% | -23.15% | 17.3% | -9.834% | 13.26% |  | -11.01% | -11.01% | 6.787% | 20.94% | 6.872% | 3.353% | -6.966% | 7.743% | -2.921% | 13.26% | -2.921% | 2.337%
| Z-Score (Absolute) |  | 1.26 | 0 | 1.4 | 1.4 | 0.48 | 1.12 | 5.9 | 0.56 | 0.56 | 0.67 | 1.66 | 0.68 | 0.44 | 0.28 | 0.74 | 0 | 1.12 | 0 | 0.37
| Analytical results for Pb in OREAS 782 (Indicative Value 52 ppm) | [20 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 14 | 17 | 21 | 22 | 23 | 24 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 50 | 120 | < 100 | < 46 | 659.1 | 90 | 50 | 110 | 50 | < 74 | [9 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 46 | 663.8 | 90 | < 50 | 80 | < 50 | < 74 | [9 empty] | 8
| 1 | 3 | < 50 | < 50 | < 100 | < 46 | 645.2 | < 90 | < 50 | 90 | < 50 | < 74 | [9 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 46 | 654.5 | 90 | < 50 | 90 | < 50 | < 74 | [9 empty] | 52.0194
| 1 | 5 | < 50 | < 50 | < 100 | < 46 | 663.8 | 90 | < 50 | 110 | < 50 | < 74 | [9 empty] | 14
| 1 | 6 | < 50 | < 50 | < 100 | < 46 | 663.8 | 90 | < 50 | 120 | < 50 | < 74
| Mean |  |  | 120 |  |  | 658.3667 | 90 | 50 | 100 | 50
| Median |  |  | 120 |  |  | 661.45 | 90 | 50 | 100 | 50
| Std Dev. | [5 empty] | 7.4487 | 0 |  | 15.4919
| Rel.Std.Dev. | [5 empty] | 1.131% | 0% |  | 15.49%
| PDM3 |  |  | 130.7% |  |  | 1166% | 73.01% | -3.882% | 92.24% | -3.882%
| Z-Score (Absolute) |  | 0.75 | 0.1 | 0.59 | 0.85 | 33.41 | 2.49 | 0.53 | 3.29 | 0.53 | 0.1
| Analytical results for Pr in OREAS 782 (Indicative Value 55 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | 58 | < 170 | [17 empty] | 1
| 1 | 2 | 58 | < 170 | [17 empty] | 9
| 1 | 3 | 58 | < 170 | [17 empty] | 16
| 1 | 4 | 41 | < 170 | [17 empty] | 55.1785
| 1 | 5 | 66 | < 170 | [17 empty] | 15
| 1 | 6 | 50 | < 170
| Mean |  | 55.1667
| Median |  | 58
| Std Dev. |  | 8.5888
| Rel.Std.Dev. |  | 15.57%
| PDM3 |  | -0.02145%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 782 (Indicative Value 103 ppm) | [20 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 17 | 21 | 22 | 23 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | 50 | 110 | < 100 | 64 | < 180 | 140 | 100 | 90 | [11 empty] | 1
| 1 | 2 | 50 | 120 | < 100 | 110 | < 180 | 130 | 100 | 110 | [11 empty] | 10
| 1 | 3 | < 50 | 50 | < 100 | 101 | < 180 | 150 | 120 | 100 | [11 empty] | 16
| 1 | 4 | 50 | 80 | < 100 | 101 | < 180 | 140 | 100 | 100 | [11 empty] | 103.2679
| 1 | 5 | 50 | 100 | < 100 | 110 | < 180 | 130 | 100 | 100 | [11 empty] | 16
| 1 | 6 | 50 | 100 | < 100 | 110 | < 180 | 140 | 110 | 100
| Mean |  | 50 | 93.3333 |  | 99.3333 |  | 138.3333 | 105 | 100
| Median |  | 50 | 100 |  | 105.5 |  | 140 | 100 | 100
| Std Dev. |  | 0 | 25.0333 |  | 17.8624 |  | 7.5277 | 8.3666 | 6.3246
| Rel.Std.Dev. |  | 0% | 26.82% |  | 17.98% |  | 5.442% | 7.968% | 6.325%
| PDM3 |  | -51.58% | -9.62% |  | -3.81% |  | 33.96% | 1.677% | -3.165%
| Z-Score (Absolute) |  | 5.01 | 0.28 | 4.59 | 0.28 | 0.47 | 4.19 | 0.88 | 0.38
| Analytical results for S in OREAS 782 (Certified Value 204 ppm) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | [5 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 200 | 100 | 200 | 40 | < 40 | 180 | 236 | 227.055 | 220.57 | < 800 | 200 | 200 | 100 | 308 | [5 empty] | 1
| 1 | 2 | 200 | 100 | 200 | 40 | < 40 | 180 | 216 | 227.055 | 226.69 | < 800 | 100 | 200 | 200 | 272 | [5 empty] | #N/A
| 1 | 3 | 200 | 100 | 200 | 40 | < 40 | 190 | 228 | 218.566 | 222.61 | < 800 | 200 | 200 | 200 | 244 | [5 empty] | 16
| 1 | 4 | 240 | 100 | 200 | 40 | < 40 | 190 | 220 | 224.893 | 230.78 | < 800 | 200 | 200 | 400 | 264 | [5 empty] | 204.2475
| 1 | 5 | 200 | 100 | 200 | < 40 | < 40 | 180 | 248 | 225.053 | 232.82 | < 800 | 100 | 200 | 100 | 232 | [5 empty] | 21
| 1 | 6 | 200 | 100 | 200 | 40 | < 40 | 180 | 220 | 228.897 | 249.16 | < 800 | 200 | 200 | 100 | 248
| Mean |  | 206.6667 | 100 | 200 | 40 |  | 183.3333 | 228 | 225.2532 | 230.4383 |  | 166.6667 | 200 | 183.3333 | 261.3333
| Median |  | 200 | 100 | 200 | 40 |  | 180 | 224 | 226.054 | 228.735 |  | 200 | 200 | 150 | 256
| Std Dev. |  | 16.3299 | 0 | 0 | 0 |  | 5.164 | 12.1326 | 3.5961 | 10.286 |  | 51.6398 | 0 | 116.9045 | 26.9716
| Rel.Std.Dev. |  | 7.902% | 0% | 0% | 0% |  | 2.817% | 5.321% | 1.596% | 4.464% |  | 30.98% | 0% | 63.77% | 10.32%
| PDM3 |  | 1.184% | -51.04% | -2.08% | -80.42% |  | -10.24% | 11.63% | 10.28% | 12.82% |  | -18.4% | -2.08% | -10.24% | 27.95%
| Z-Score (Absolute) |  | 0.16 | 2.3 | 0 | 3.75 | 4.14 | 0.38 | 0.65 | 0.58 | 0.7 | 4.61 | 0.77 | 0 | 0.38 | 1.42
| Analytical results for Sb in OREAS 782 (Indicative Value 116 ppm) | [20 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 08 | 17 | 21 | 22 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | < 80 | 270 | 70 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 80 | 230 | 100 | < 50 | [12 empty] | 11
| 1 | 3 | < 50 | < 50 | < 100 | < 80 | 250 | 80 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 80 | 220 | 60 | < 50 | [12 empty] | 116.1111
| 1 | 5 | < 50 | < 50 | < 100 | < 80 | 240 | 100 | 50 | [12 empty] | 17
| 1 | 6 | < 50 | < 50 | < 100 | < 80 | 190 | 80 | 50
| Mean | [5 empty] | 233.3333 | 81.6667 | 50
| Median | [5 empty] | 235 | 80 | 50
| Std Dev. | [5 empty] | 27.3252 | 16.0208 | 0
| Rel.Std.Dev. | [5 empty] | 11.71% | 19.62% | 0%
| PDM3 | [5 empty] | 101% | -29.67% | -56.94%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.33 | 0 | 7.7 | 1.6 | 0.34
| Analytical results for Sc in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Sc | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 12
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 1
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 782 (Certified Value 51.22 wt.%) | [20 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 50.44 | 53.45 | 51.6 | 52 | 51.7152 | 50.4 | 49.29 | 53.49 | 50.78 | 51.13 | 51.64 | 51.5051 | 50.4928 | 51.1 | 51.3238 | 51.6 | 51.3 | 51.9 | 48.641 | 1
| 1 | 2 | 50.1 | 53.45 | 51 | 52.6 | 51.9192 | 50.4 | 50.04 | 53.2 | 50.84 | 50.98 | 51.7 | 51.3068 | 50.4123 | 51 | 51.4242 | 51.4 | 51.5 | 51.7 | 48.707 | #N/A
| 1 | 3 | 50.16 | 53.44 | 51.6 | 52 | 52.0128 | 51.5 | 49.13 | 53.06 | 49.98 | 51.14 | 51.64 | 51.5147 | 50.1942 | 50.9 | 51.4542 | 51.3 | 51.8 | 51.8 | 48.65 | 16
| 1 | 4 | 50.39 | 53.46 | 51.3 | 52.2 | 51.8734 | 51.1 | 49.24 | 53.37 | 50.01 | 51.13 | 51.39 | 51.5162 | 50.4095 | 51.3 | 51.5922 | 51.1 | 51.3 | 52 | 48.647 | 51.2225
| 1 | 5 | 50.49 | 53.44 | 51.9 | 52 | 51.8594 | 50.9 | 49.44 | 53.2 | 50.3 | 51.09 | 51.65 | 51.4266 | 50.2308 | 50.9 | 51.4676 | 50.7 | 51.7 | 51.8 | 48.541 | 22
| 1 | 6 | 50.18 | 53.46 | 51.4 | 52.1 | 51.8809 | 50.6 | 50.2 | 53.18 | 50.07 | 51.04 | 51.61 | 51.5392 | 50.212 | 51.2 | 51.36 | 50.4 | 51.5 | 51.7 | 48.658
| Mean |  | 50.2933 | 53.45 | 51.4667 | 52.15 | 51.8768 | 50.8167 | 49.5567 | 53.25 | 50.33 | 51.085 | 51.605 | 51.4681 | 50.3253 | 51.0667 | 51.437 | 51.0833 | 51.5167 | 51.8167 | 48.6407
| Median |  | 50.285 | 53.45 | 51.5 | 52.05 | 51.8771 | 50.75 | 49.365 | 53.2 | 50.185 | 51.11 | 51.64 | 51.5099 | 50.3202 | 51.05 | 51.4392 | 51.2 | 51.5 | 51.8 | 48.6485
| Std Dev. |  | 0.16585 | 0.0089443 | 0.30768 | 0.23452 | 0.096685 | 0.43551 | 0.45045 | 0.15362 | 0.38884 | 0.063482 | 0.10932 | 0.087903 | 0.12781 | 0.1633 | 0.094055 | 0.4535 | 0.20412 | 0.1169 | 0.054335
| Rel.Std.Dev. |  | 0.3298% | 0.01673% | 0.5978% | 0.4497% | 0.1864% | 0.857% | 0.909% | 0.2885% | 0.7726% | 0.1243% | 0.2118% | 0.1708% | 0.254% | 0.3198% | 0.1829% | 0.8878% | 0.3962% | 0.2256% | 0.1117%
| PDM3 |  | -1.814% | 4.349% | 0.4767% | 1.811% | 1.277% | -0.7923% | -3.252% | 3.958% | -1.742% | -0.2684% | 0.7468% | 0.4795% | -1.752% | -0.3042% | 0.4188% | -0.2716% | 0.5743% | 1.16% | -5.04%
| Z-Score (Absolute) |  | 1.75 | 3.09 | 0.05 | 1.09 | 0.67 | 0.95 | 2.88 | 2.78 | 1.7 | 0.54 | 0.26 | 0.05 | 1.7 | 0.57 | 0 | 0.54 | 0.12 | 0.58 | 4.29
| Analytical results for Sm in OREAS 782 (Indicative Value < 170 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | [18 empty] | 1
| No | No | 17 | [18 empty] | ppm
|  |  | OxBF*XRF | [18 empty] | 0
0
| 1 | 1 | < 170 | [18 empty] | 1
| 1 | 2 | < 170 | [18 empty] | 28
| 1 | 3 | < 170 | [18 empty] | 16
| 1 | 4 | < 170 | [18 empty] | < 170
| 1 | 5 | < 170 | [18 empty] | 8
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 782 (Indicative Value 47.8 ppm) | [20 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 14 | 17 | 21 | 22 | 23 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 1
1
| 1 | 1 | < 50 | < 50 | < 100 | 71 | < 160 | < 80 | < 50 | 50 | 60 | [10 empty] | 1
| 1 | 2 | < 50 | 50 | < 100 | 47 | < 160 | < 80 | < 50 | 50 | 60 | [10 empty] | 3
| 1 | 3 | < 50 | < 50 | < 100 | 63 | < 160 | < 80 | < 50 | 100 | < 50 | [10 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | 39 | < 160 | < 80 | 60 | 70 | 50 | [10 empty] | 47.8147
| 1 | 5 | < 50 | 50 | < 100 | 71 | < 160 | < 80 | < 50 | 80 | < 50 | [10 empty] | 9
| 1 | 6 | < 50 | < 50 | < 100 | 63 | < 160 | < 80 | < 50 | 90 | 70
| Mean |  |  | 50 |  | 59 |  |  | 60 | 73.3333 | 60
| Median |  |  | 50 |  | 63 |  |  | 60 | 75 | 60
| Std Dev. |  |  | 0 |  | 13.1453 |  |  |  | 20.6559 | 8.165
| Rel.Std.Dev. |  |  | 0% |  | 22.28% |  |  |  | 28.17% | 13.61%
| PDM3 |  |  | 4.57% |  | 23.39% |  |  | 25.48% | 53.37% | 25.48%
| Z-Score (Absolute) |  | 1.05 | 0.67 | 0.07 | 0.48 | 1.37 | 0.4 | 0.79 | 1.12 | 0
| Analytical results for Sr in OREAS 782 (Indicative Value 88 ppm) | [20 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [6 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 42 | 100 | < 100 | < 80 | 80 | < 42 | 76.1 | < 80 | 53.3248 | 100 | 100 | < 100 | < 80 | [6 empty] | 1
| 1 | 2 | 51 | 100 | < 100 | < 80 | 80 | < 42 | 93 | < 80 | 52.479 | 100 | 100 | < 100 | < 80 | [6 empty] | 4
| 1 | 3 | 42 | 100 | < 100 | < 80 | 80 | < 42 | 93 | < 80 | 53.44 | 100 | 100 | < 100 | < 80 | [6 empty] | 16
| 1 | 4 | 42 | 100 | < 100 | < 80 | < 80 | < 42 | 67.6 | < 80 | 53.2474 | 100 | 100 | < 100 | < 80 | [6 empty] | 87.8169
| 1 | 5 | 59 | 100 | < 100 | < 80 | 80 | < 42 | 93 | < 80 | 52.1274 | 100 | 100 | < 100 | < 80 | [6 empty] | 10
| 1 | 6 | 42 | 100 | < 100 | < 80 | 80 | < 42 | 93 | < 80 | 51.6333 | 100 | 100 | < 100 | < 80
| Mean |  | 46.3333 | 100 |  |  | 80 |  | 85.95 |  | 52.7086 | 100 | 100
| Median |  | 42 | 100 |  |  | 80 |  | 93 |  | 52.8632 | 100 | 100
| Std Dev. |  | 7.174 | 0 |  |  | 0 |  | 11.2477 |  | 0.74183 | 0 | 0
| Rel.Std.Dev. |  | 15.48% | 0% |  |  | 0% |  | 13.09% |  | 1.407% | 0% | 0%
| PDM3 |  | -47.24% | 13.87% |  |  | -8.901% |  | -2.126% |  | -39.98% | 13.87% | 13.87%
| Z-Score (Absolute) |  | 0.31 | 4.37 | 0 | 0.67 | 2.4 | 2.52 | 3.14 | 0.67 | 0.24 | 4.37 | 4.37 | 0 | 0.67
| Analytical results for Ta in OREAS 782 (Indicative Value < 80 ppm) | [20 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 15 | 17 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 100 |  | < 80 | [16 empty] | 1
| 1 | 2 | < 100 |  | < 80 | [16 empty] | 5
| 1 | 3 | < 100 | 8 | < 80 | [16 empty] | 16
| 1 | 4 | < 100 |  | < 80 | [16 empty] | < 80
| 1 | 5 | < 100 |  | < 80 | [16 empty] | 11
| 1 | 6 | < 100 |  | < 80
| Mean |  |  | 8
| Median |  |  | 8
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 2.44 | 0
| Analytical results for Te in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 6
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 12
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 782 (Indicative Value 83 ppm) | [20 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 13 | 15 | 17 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 44 | 83.5 | 90 | [16 empty] | 1
| 1 | 2 | < 44 | 79.1 | 180 | [16 empty] | 7
| 1 | 3 | < 44 | 79.1 | 90 | [16 empty] | 16
| 1 | 4 | < 44 | 79.1 | < 90 | [16 empty] | 83.4871
| 1 | 5 | < 44 | 70.3 | 90 | [16 empty] | 13
| 1 | 6 | < 44 | 83.5 | < 90
| Mean |  |  | 79.1 | 112.5
| Median |  |  | 79.1 | 90
| Std Dev. |  |  | 4.82 | 45
| Rel.Std.Dev. |  |  | 6.094% | 40%
| PDM3 |  |  | -5.255% | 34.75%
| Z-Score (Absolute) |  | 4.38 | 0 | 0.67
| Analytical results for TiO2 in OREAS 782 (Certified Value 0.45 wt.%) | [20 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.48 | 0.51 | 0.45 | 0.46 | 0.46427 | 0.45 | 0.12 | 0.834 | 0.45 | 0.44 | 0.44 | 0.42085 | 0.42279 | 0.46 | 0.46653 | 0.44 | 0.46 | 0.45 | 0.409 | 1
| 1 | 2 | 0.47 | 0.51 | 0.44 | 0.46 | 0.45899 | 0.46 | 0.11 | 0.801 | 0.45 | 0.45 | 0.45 | 0.42229 | 0.42177 | 0.45 | 0.46565 | 0.44 | 0.46 | 0.44 | 0.412 | #N/A
| 1 | 3 | 0.48 | 0.5 | 0.44 | 0.46 | 0.46376 | 0.49 | 0.12 | 0.817 | 0.45 | 0.45 | 0.45 | 0.41904 | 0.41922 | 0.46 | 0.46609 | 0.44 | 0.46 | 0.45 | 0.415 | 16
| 1 | 4 | 0.48 | 0.52 | 0.44 | 0.45 | 0.46652 | 0.47 | 0.12 | 0.817 | 0.45 | 0.46 | 0.44 | 0.41733 | 0.42636 | 0.49 | 0.46803 | 0.44 | 0.46 | 0.45 | 0.419 | 0.44998
| 1 | 5 | 0.48 | 0.5 | 0.45 | 0.46 | 0.46856 | 0.46 | 0.11 | 0.817 | 0.45 | 0.44 | 0.45 | 0.41286 | 0.42432 | 0.46 | 0.46978 | 0.44 | 0.47 | 0.45 | 0.417 | 23
| 1 | 6 | 0.48 | 0.52 | 0.45 | 0.46 | 0.46211 | 0.47 | 0.1 | 0.834 | 0.44 | 0.47 | 0.44 | 0.42616 | 0.42024 | 0.49 | 0.46681 | 0.44 | 0.47 | 0.44 | 0.416
| Mean |  | 0.47833 | 0.51 | 0.445 | 0.45833 | 0.46404 | 0.46667 | 0.11333 | 0.82 | 0.44833 | 0.45167 | 0.445 | 0.41975 | 0.42245 | 0.46833 | 0.46715 | 0.44 | 0.46333 | 0.44667 | 0.41467
| Median |  | 0.48 | 0.51 | 0.445 | 0.46 | 0.46402 | 0.465 | 0.115 | 0.817 | 0.45 | 0.45 | 0.445 | 0.41995 | 0.42228 | 0.46 | 0.46667 | 0.44 | 0.46 | 0.45 | 0.4155
| Std Dev. |  | 0.0040825 | 0.0089443 | 0.0054772 | 0.0040825 | 0.0033448 | 0.013663 | 0.008165 | 0.01249 | 0.0040825 | 0.01169 | 0.0054772 | 0.0045295 | 0.0026336 | 0.017224 | 0.0015202 | 0 | 0.005164 | 0.005164 | 0.0036148
| Rel.Std.Dev. |  | 0.8535% | 1.754% | 1.231% | 0.8907% | 0.7208% | 2.928% | 7.204% | 1.523% | 0.9106% | 2.588% | 1.231% | 1.079% | 0.6234% | 3.678% | 0.3254% | 0% | 1.115% | 1.156% | 0.8717%
| PDM3 |  | 6.3% | 13.34% | -1.107% | 1.856% | 3.123% | 3.708% | -74.81% | 82.23% | -0.3665% | 0.3742% | -1.107% | -6.718% | -6.119% | 4.078% | 3.815% | -2.218% | 2.967% | -0.7369% | -7.848%
| Z-Score (Absolute) |  | 1.2 | 2.62 | 0.3 | 0.3 | 0.56 | 0.67 | 15.21 | 16.56 | 0.15 | 0 | 0.3 | 1.43 | 1.31 | 0.75 | 0.7 | 0.52 | 0.52 | 0.22 | 1.66
| Analytical results for Tl in OREAS 782 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Tl | Round | Replicate | Lab | [18 empty] | 1
| No | No | 08 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 8
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 782 (Indicative Value 155 ppm) | [20 empty] | Indicative
U | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 17 | [17 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [17 empty] | 0
0
| 1 | 1 | < 17 | 170 | [17 empty] | 1
| 1 | 2 | < 17 | 80 | [17 empty] | 9
| 1 | 3 | < 17 | 170 | [17 empty] | 16
| 1 | 4 | < 17 | 170 | [17 empty] | 155.4726
| 1 | 5 | < 17 | 170 | [17 empty] | 15
| 1 | 6 | < 17 | 170
| Mean |  |  | 155
| Median |  |  | 170
| Std Dev. |  |  | 36.7423
| Rel.Std.Dev. |  |  | 23.7%
| PDM3 |  |  | -0.304%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V in OREAS 782 (Indicative Value 67 ppm) | [20 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 45 | 100 | 100 | < 100 | < 60 | 78 | 73 | < 110 | 64.4 | 60 | 46.9845 | 100 | 100 | 100 | 62 |  |  |  |  | 1
| 1 | 2 | 67 | 100 | 100 | < 100 | < 60 | 73 | 67 | < 110 | 44.8 | 60 | 53.1769 | 100 | 100 | 100 | 62 |  |  |  |  | 10
| 1 | 3 | 56 | 100 | 100 | < 100 | < 60 | 73 | 78 | < 110 | 56 | 60 | 50.2181 | 100 | 100 | 100 | 62 |  |  |  |  | 16
| 1 | 4 | 39 | < 100 | 100 | < 100 | < 60 | 84 | 67 | < 110 | 58.8 | 110 | 62.9236 | 100 | 100 | 100 | 62 |  |  |  |  | 66.77
| 1 | 5 | 56 | 100 | 100 | < 100 | < 60 | 73 | 73 | < 110 | 67.2 | 110 | 56.9273 | 100 | 100 | 100 | 73 |  |  |  |  | 16
| 1 | 6 | 67 | 100 | 100 | < 100 | < 60 | 78 | 84 | < 110 | 56 | 110 | 58.1983 | 100 | 100 | 100 | 73
| Mean |  | 55 | 100 | 100 |  |  | 76.5 | 73.6667 |  | 57.8667 | 85 | 54.7381 | 100 | 100 | 100 | 65.6667
| Median |  | 56 | 100 | 100 |  |  | 75.5 | 73 |  | 57.4 | 85 | 55.0521 | 100 | 100 | 100 | 62
| Std Dev. |  | 11.3666 | 0 | 0 |  |  | 4.4159 | 6.5625 |  | 7.8533 | 27.3861 | 5.775 | 0 | 0 | 0 | 5.6804
| Rel.Std.Dev. |  | 20.67% | 0% | 0% |  |  | 5.772% | 8.908% |  | 13.57% | 32.22% | 10.55% | 0% | 0% | 0% | 8.65%
| PDM3 |  | -17.63% | 49.77% | 49.77% |  |  | 14.57% | 10.33% |  | -13.33% | 27.3% | -18.02% | 49.77% | 49.77% | 49.77% | -1.652%
| Z-Score (Absolute) |  | 0.16 | NA | NA | 0.61 | 2.54 | 1.72 | 1.48 | 0.08 | 0.08 | 2.38 | 0.19 | NA | NA | NA | 0.74
Results from laboratories 3, 4, 21, 22 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 782 (Certified Value 187 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 13 | 15 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 180 | 180 | 158.59 | < 100 | 230 | 170.5 | 160 | 210 | 210 | 200 | [9 empty] | 1
| 1 | 2 | 190 | 200 | 158.59 | < 100 | 214 | 166.5 | 160 | 180 | 190 | 200 | [9 empty] | 26
| 1 | 3 | 190 | 180 | 237.89 | < 100 | 222 | 170.5 | 160 | 200 | 230 | 200 | [9 empty] | 16
| 1 | 4 | 190 | 160 | 158.59 | < 100 | 214 | 174.5 | 240 | 190 | 190 | 210 | [9 empty] | 187.223
| 1 | 5 | 190 | 200 | 237.89 | < 100 | 222 | 166.5 | 160 | 190 | 220 | 210 | [9 empty] | 24
| 1 | 6 | 190 | 190 | 158.59 | < 100 | 222 | 174.5 | 240 | 190 | 220 | 200
| Mean |  | 188.3333 | 185 | 185.0233 |  | 220.6667 | 170.5 | 186.6667 | 193.3333 | 210 | 203.3333
| Median |  | 190 | 185 | 158.59 |  | 222 | 170.5 | 160 | 190 | 215 | 200
| Std Dev. |  | 4.0825 | 15.1658 | 40.9503 |  | 6.0222 | 3.5777 | 41.3118 | 10.328 | 16.7332 | 5.164
| Rel.Std.Dev. |  | 2.168% | 8.198% | 22.13% |  | 2.729% | 2.098% | 22.13% | 5.342% | 7.968% | 2.54%
| PDM3 |  | 0.593% | -1.187% | -1.175% |  | 17.86% | -8.932% | -0.2972% | 3.264% | 12.17% | 8.605%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 0.1 | 8.07 | 2.01 | 0.96 | 0.1 | 0.39 | 1.38 | 0.98
| Analytical results for Y2O3 in OREAS 782 (Indicative Value 90 ppm) | [20 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 04 | 13 | 17 | 21 | 22 | 23 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | 80 | 70 | < 200 | 130 | 130 | 130 | [12 empty] | 1
| 1 | 2 | 50 | < 50 | 70 | < 200 | 110 | 110 | 120 | [12 empty] | 11
| 1 | 3 | 50 | < 50 | 80 | < 200 | 130 | 110 | 130 | [12 empty] | 16
| 1 | 4 | < 50 | 90 | 90 | < 200 | 120 | 130 | 130 | [12 empty] | 89.8889
| 1 | 5 | 50 | < 50 | 70 | < 200 | 120 | 150 | 140 | [12 empty] | 17
| 1 | 6 | 50 | < 50 | 90 | < 200 | 120 | 140 | 150
| Mean |  | 50 | 85 | 78.3333 |  | 121.6667 | 128.3333 | 133.3333
| Median |  | 50 | 85 | 75 |  | 120 | 130 | 130
| Std Dev. |  | 0 | 7.0711 | 9.8319 |  | 7.5277 | 16.0208 | 10.328
| Rel.Std.Dev. |  | 0% | 8.319% | 12.55% |  | 6.187% | 12.48% | 7.746%
| PDM3 |  | -44.38% | -5.439% | -12.86% |  | 35.35% | 42.77% | 48.33%
| Z-Score (Absolute) |  | 1.39 | 1.31 | 0.52 | 0 | 0.52 | 0.67 | 0.79
| Analytical results for Zn in OREAS 782 (Indicative Value 33.6 ppm) | [20 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 08 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [7 empty] | 1
1
| 1 | 1 | < 50 | < 50 | < 100 | < 40 | < 160 | 28.1 | < 80 | 19.7685 | < 50 | < 50 | 50 | 32 | [7 empty] | 1
| 1 | 2 | < 50 | < 50 | < 100 | < 40 | < 160 | 28.1 | < 80 | 20.0333 | < 50 | < 50 | 60 | 32 | [7 empty] | 12
| 1 | 3 | < 50 | < 50 | < 100 | < 40 | < 160 | 28.1 | < 80 | 20.8242 | < 50 | < 50 | 50 | 16 | [7 empty] | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 40 | < 160 | 28.1 | < 80 | 19.891 | < 50 | < 50 | 60 | 24 | [7 empty] | 33.6057
| 1 | 5 | < 50 | < 50 | < 100 | < 40 | < 160 | 32.1 | < 80 | 21.9961 | < 50 | < 50 | 60 | 40 | [7 empty] | 18
| 1 | 6 | < 50 | < 50 | < 100 | < 40 | < 160 | 32.1 | < 80 | 20.5534 | < 50 | < 50 | 50 | 32
| Mean | [6 empty] | 29.4333 |  | 20.5111 |  |  | 55 | 29.3333
| Median | [6 empty] | 28.1 |  | 20.2933 |  |  | 55 | 32
| Std Dev. | [6 empty] | 2.0656 |  | 0.833 |  |  | 5.4772 | 8.2624
| Rel.Std.Dev. | [6 empty] | 7.018% |  | 4.061% |  |  | 9.959% | 28.17%
| PDM3 | [6 empty] | -12.42% |  | -38.97% |  |  | 63.66% | -12.71%
| Z-Score (Absolute) |  | 0.34 | 0.34 | 3.43 | 1.08 | 8.01 | 0.34 | 1.95 | 1.01 | 0.34 | 0.34 | 4.19 | 0.34
| Analytical results for Zr in OREAS 782 (Certified Value 115 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 08 | 09 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 100 | < 100 | < 100 | 70 | 150 | < 150 | 125.9 | < 150 | 135.815 | 100 | < 100 | 100 | 126 | [6 empty] | 1
| 1 | 2 | 100 | < 100 | < 100 | 70 | 150 | < 150 | 122.2 | < 150 | 137.5923 | 100 | < 100 | 100 | 148 | [6 empty] | #N/A
| 1 | 3 | 100 | < 100 | < 100 | 70 | 150 | < 150 | 122.2 | < 150 | 138.5037 | 100 | < 100 | 100 | 170 | [6 empty] | 16
| 1 | 4 | 100 | < 100 | < 100 | 70 | 70 | < 150 | 133.3 | < 150 | 138.9818 | 200 | < 100 | 100 | 141 | [6 empty] | 114.6922
| 1 | 5 | 100 | 100 | < 100 | 70 | 150 | < 150 | 140.7 | < 150 | 136.6349 | 100 | < 100 | 100 | 148 | [6 empty] | 25
| 1 | 6 | 100 | < 100 | < 100 | 70 | 70 | < 150 | 144.4 | < 150 | 137.101 | 100 | < 100 | < 100 | 163
| Mean |  | 100 | 100 |  | 70 | 123.3333 |  | 131.45 |  | 137.4381 | 116.6667 |  | 100 | 149.3333
| Median |  | 100 | 100 |  | 70 | 150 |  | 129.6 |  | 137.3466 | 100 |  | 100 | 148
| Std Dev. |  | 0 |  |  | 0 | 41.3118 |  | 9.5772 |  | 1.1777 | 40.8248 |  | 0 | 15.6929
| Rel.Std.Dev. |  | 0% |  |  | 0% | 33.5% |  | 7.286% |  | 0.8569% | 34.99% |  | 0% | 10.51%
| PDM3 |  | -12.81% | -12.81% |  | -38.97% | 7.534% |  | 14.61% |  | 19.83% | 1.722% |  | -12.81% | 30.2%
| Z-Score (Absolute) |  | 0.18 | 0.71 | 0.89 | 0.37 | 0.67 | 0.37 | 0.84 | 0.37 | 0.97 | 0.53 | 0.89 | 0 | 1.23
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C634: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C652: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C670: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C688: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C706: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C904: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C922: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C959: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C977: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C995: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 782 (Indicative Value 0.263 wt.%) | [20 empty] | Indicative
H2O- | Round | Replicate | Lab | [18 empty] | 1
| No | No | 20 | [18 empty] | wt.%
|  |  | BF*XRF | [18 empty] | 3
3
| 1 | 1 | 0.2875 | [18 empty] | 1
| 1 | 2 | 0.2807 | [18 empty] | 11
| 1 | 3 | 0.2125 | [18 empty] | 16
| 1 | 4 | 0.2518 | [18 empty] | 0.26265
| 1 | 5 | 0.2706 | [18 empty] | 20
| 1 | 6 | 0.2728
| Mean |  | 0.26265
| Median |  | 0.2717
| Std Dev. |  | 0.027358
| Rel.Std.Dev. |  | 10.42%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 782 (Certified Value 5.14 wt.%) | [20 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | OxBF*XRF | LOI*Furnace | LOI*TGA | OxBF*XRF | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | BF*XRF | LOI*Furnace | LOI*TGA | LOI*TGA | 2
2
| 1 | 1 | 3.94 | 6.04 | 5.47 | 5.03 | 5.349 | I.N.F. | 5.97 | 6.99 | 6.04 | 2.9817 | 3.98 | 4.056 | 6.625 | 4.26 | 7.5192 | 5.3686 | 4.74 | 3.22 | 3.71 | 1
| 1 | 2 | 4.44 | 6.05 | 5.63 | 5.17 | 5.3297 | I.N.F. | 5.88 | 6.79 | 6.18 | 2.9137 | 4.02 | 4.0404 | 6.6956 | 4.27 | 7.9792 | 5.6812 | 4.74 | 3.16 | 3.83 | #N/A
| 1 | 3 | 4.08 | 6.03 | 5.68 | 5.11 | 5.1677 | I.N.F. | 5.97 | 6.6 | 6.16 | 2.7468 | 4.03 | 4.095 | 6.5495 | 4.57 | 8.1826 | 5.5163 | 4.91 | 3.29 | 3.26 | 16
| 1 | 4 | 4.03 | 6.01 | 5.57 | 4.8 | 5.7096 | I.N.F. | 5.81 | 6.83 | 6.04 | 2.9164 | 3.99 | 4.0699 | 6.9325 | 4.47 | 8.046 | 5.506 | 5.08 | 3.41 | 3.72 | 5.1378
| 1 | 5 | 3.91 | 6.05 | 5.76 | 5.34 | 5.0591 | I.N.F. | 5.99 | 6.81 | 5.97 | 2.9107 | 4.08 | 3.9602 | 6.8718 | 4.4 | 7.9252 | 5.5149 | 4.74 | 3.29 | 3.8 | 27
| 1 | 6 | 4.56 | 6.02 | 5.67 | 5.14 | 5.187 | I.N.F. | 6 | 6.87 | 6.22 | 2.96 | 4.05 | 4.0859 | 6.8115 | 4.12 | 7.6439 | 5.6753 | 4.8 | 3.28 | 3.74
| Mean |  | 4.16 | 6.0333 | 5.63 | 5.0983 | 5.3003 |  | 5.9367 | 6.815 | 6.1017 | 2.9049 | 4.025 | 4.0512 | 6.7476 | 4.3483 | 7.8827 | 5.5437 | 4.835 | 3.275 | 3.6767
| Median |  | 4.055 | 6.035 | 5.65 | 5.125 | 5.2584 |  | 5.97 | 6.82 | 6.1 | 2.915 | 4.025 | 4.063 | 6.7535 | 4.335 | 7.9522 | 5.5156 | 4.77 | 3.285 | 3.73
| Std Dev. |  | 0.27298 | 0.01633 | 0.1002 | 0.17837 | 0.22775 |  | 0.075277 | 0.12708 | 0.098472 | 0.082675 | 0.037283 | 0.04877 | 0.14879 | 0.1629 | 0.25174 | 0.11823 | 0.13708 | 0.083606 | 0.20944
| Rel.Std.Dev. |  | 6.562% | 0.2707% | 1.78% | 3.499% | 4.297% |  | 1.268% | 1.865% | 1.614% | 2.846% | 0.9263% | 1.204% | 2.205% | 3.746% | 3.194% | 2.133% | 2.835% | 2.553% | 5.697%
| PDM3 |  | -19.03% | 17.43% | 9.58% | -0.7682% | 3.164% |  | 15.55% | 32.64% | 18.76% | -43.46% | -21.66% | -21.15% | 31.33% | -15.37% | 53.43% | 7.901% | -5.894% | -36.26% | -28.44%
| Z-Score (Absolute) |  | 0.72 | 0.58 | 0.3 | 0.07 | 0.07 | NA | 0.51 | 1.12 | 0.63 | 1.59 | 0.82 | 0.8 | 1.08 | 0.59 | 1.86 | 0.24 | 0.25 | 1.34 | 1.06
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ISE
| Analytical results for F in OREAS 782 (Certified Value 14.57 wt.%) | [16 empty] | Certified
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 18 | 20 | 21 | 24 | wt.%
|  |  | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | 2
2
| 1 | 1 | 9.19 | 18.839 | 13.95 | 16.3935 | 9.1209 | 14 | 11.15 | 15 | > 10 | 15.6289 | > 10 | 12.97 | 12.4746 | 12.9 | 16.8702 | 1
| 1 | 2 | 9.16 | 18.639 | 14.2 | 16.7418 | 8.8548 | 13.7 | 11.53 | 15 | > 10 | 15.5021 | > 10 | 13.17 | 12.4422 | 12.45 | 18.5105 | #N/A
| 1 | 3 | 9.26 | 18.639 | 14.2 | 15.9011 | 9.4088 | 13.6 | 11.25 | 15 | > 10 | 15.2752 | > 10 | 13.07 | 12.3852 | 12.9 | 11.3965 | 16
| 1 | 4 | 9.29 | 18.533 | 14.25 | 15.9817 | 8.994 | 14 | 11.26 | 15.2 | > 10 | 15.6093 | > 10 | 13.07 | 12.131 | 12.95 | 17.5222 | 14.5697
| 1 | 5 | 9.16 | 18.482 | 13.85 | 16.2787 | 8.9323 | 14.249 | 11.47 | 15.3 | > 10 | 15.3285 | > 10 | 12.67 | 12.0635 | 12.95 | 16.3789 | 29
| 1 | 6 | 9.29 | 18.428 | 14.3 | 15.9954 | 8.8501 | 14.6 | 11.4 | 15 | > 10 | 15.5464 | > 10 | 12.67 | 12.703 | 12.7 | 17.1575
| Mean |  | 9.225 | 18.5933 | 14.125 | 16.2154 | 9.0268 | 14.0248 | 11.3433 | 15.0833 |  | 15.4817 |  | 12.9367 | 12.3666 | 12.8083 | 16.306
| Median |  | 9.225 | 18.586 | 14.2 | 16.137 | 8.9631 | 14 | 11.33 | 15 |  | 15.5242 |  | 13.02 | 12.4137 | 12.9 | 17.0138
| Std Dev. |  | 0.062209 | 0.1469 | 0.18097 | 0.32083 | 0.21241 | 0.3656 | 0.14638 | 0.13292 |  | 0.14742 |  | 0.21602 | 0.23591 | 0.19854 | 2.5096
| Rel.Std.Dev. |  | 0.6744% | 0.7901% | 1.281% | 1.979% | 2.353% | 2.607% | 1.29% | 0.8812% |  | 0.9522% |  | 1.67% | 1.908% | 1.55% | 15.39%
| PDM3 |  | -36.68% | 27.62% | -3.052% | 11.3% | -38.04% | -3.739% | -22.14% | 3.526% |  | 6.26% |  | -11.21% | -15.12% | -12.09% | 11.92%
| Z-Score (Absolute) |  | NA | 2.22 | 0 | 1.04 | NA | 0.05 | 1.38 | 0.48 | NA | 0.67 | NA | 0.59 | 0.87 | 0.65 | 1.08
Results from laboratory 1.01 and 1.06 were removed due to their high Z-score and unsatisfactory performance on the control sample.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 782 (Certified Value 0.347 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 3
3
| 1 | 1 | < 0.3 | 0.37 | 0.33 | < 0.5 | 0.37 | 0.4 | 0.4 | 0.34 | < 1 | < 0.5 | 0.38 | 0.3 | 0.34 | 0.34 | 0.32 | 3.8928 |  |  |  | 1
| 1 | 2 | < 0.3 | 0.38 | 0.33 | < 0.5 | 0.35 | 0.5 | 0.3 | 0.36 | < 1 | < 0.5 | 0.34 | 0.4 | 0.34 | 0.32 | 0.34 | 4.5273 |  |  |  | #N/A
| 1 | 3 | < 0.3 | 0.41 | 0.32 | < 0.5 | 0.35 | 0.5 | 0.3 | 0.36 | < 1 | < 0.5 | 0.35 | 0.4 | 0.35 | 0.35 | 0.32 | 5.0832 |  |  |  | 16
| 1 | 4 | < 0.3 | 0.35 | 0.33 | < 0.5 | 0.35 | 0.5 | 0.4 | 0.35 | < 1 | < 0.5 | 0.35 | 0.3 | 0.36 | 0.32 | 0.34 | 10.4292 |  |  |  | 0.3465
| 1 | 5 | < 0.3 | 0.34 | 0.35 | < 0.5 | 0.36 | 0.5 | 0.4 | 0.33 | < 1 | < 0.5 | 0.35 | 0.3 | 0.35 | 0.35 | 0.33 | 4.3782 |  |  |  | 31
| 1 | 6 | < 0.3 | 0.37 | 0.33 | < 0.5 | 0.38 | 0.4 | 0.4 | 0.35 | < 1 | < 0.5 | 0.36 | 0.4 | 0.34 | 0.33 | 0.35 | 4.0074
| Mean |  |  | 0.37 | 0.33167 |  | 0.36 | 0.46667 | 0.36667 | 0.34833 |  |  | 0.355 | 0.35 | 0.34667 | 0.335 | 0.33333 | 5.3864
| Median |  |  | 0.37 | 0.33 |  | 0.355 | 0.5 | 0.4 | 0.35 |  |  | 0.35 | 0.35 | 0.345 | 0.335 | 0.335 | 4.4527
| Std Dev. |  |  | 0.024495 | 0.0098319 |  | 0.012649 | 0.05164 | 0.05164 | 0.01169 |  |  | 0.013784 | 0.054772 | 0.008165 | 0.013784 | 0.012111 | 2.5063
| Rel.Std.Dev. |  |  | 6.62% | 2.964% |  | 3.514% | 11.07% | 14.08% | 3.356% |  |  | 3.883% | 15.65% | 2.355% | 4.115% | 3.633% | 46.53%
| PDM3 |  |  | 6.782% | -4.281% |  | 3.896% | 34.68% | 5.82% | 0.5291% |  |  | 2.453% | 1.01% | 0.0481% | -3.319% | -3.8% | 1455%
| Z-Score (Absolute) |  | 8.84 | 1.05 | 0.67 | 4.35 | 0.6 | NA | NA | 0.07 | 6.89 | 4.35 | 0.37 | NA | 0 | 0.52 | 0.6 | 226.55
Results from laboratories 10, 12 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 782 (Certified Value 4.63 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 4.0293 | 4.9983 | 4.85 | 4.54 | 4.26 | 4.67 | 4.88 | 4.55 | 4.803 | 4.8059 | 4.61 | 4.45 | 4.7948 | 4.22 | 4.64 | 4.68 | 4.62 | 4.3751 |  | 1
| 1 | 2 | 4.0339 | 4.9891 | 4.93 | 4.56 | 4.29 | 4.67 | 4.84 | 4.42 | 4.8244 | 4.7054 | 4.57 | 4.43 | 4.7893 | 4.33 | 4.56 | 4.58 | 4.58 | 4.3755 |  | 19
| 1 | 3 | 3.9592 | 4.9989 | 4.98 | 4.62 | 4.28 | 4.57 | 4.77 | 4.76 | 4.7728 | 4.7293 | 4.65 | 4.57 | 4.7901 | 4.32 | 4.67 | 4.7 | 4.63 | 4.4485 |  | 16
| 1 | 4 | 4.0417 | 4.996 | 4.95 | 4.57 | 4.3 | 4.58 | 4.76 | 4.85 | 4.7359 | 4.559 | 4.55 | 4.45 | 4.8997 | 4.36 | 4.63 | 4.63 | 4.78 | 4.3734 |  | 4.6344
| 1 | 5 | 4.0062 | 4.9962 | 4.96 | 4.59 | 4.27 | 4.59 | 4.59 | 4.52 | 4.7974 | 4.7662 | 4.65 | 4.59 | 4.8403 | 4.38 | 4.58 | 4.51 | 4.59 | 4.412 |  | 32
| 1 | 6 | 4.0418 | 4.9953 | 4.97 | 4.44 | 4.26 | 4.7 | 4.74 | 4.63 | 4.7372 | 4.7924 | 4.48 | 4.63 | 4.7634 | 4.29 | 4.54 | 4.68 | 4.65 | 4.3546
| Mean |  | 4.0187 | 4.9957 | 4.94 | 4.5533 | 4.2767 | 4.63 | 4.7633 | 4.6217 | 4.7784 | 4.7264 | 4.585 | 4.52 | 4.8129 | 4.3167 | 4.6033 | 4.63 | 4.6417 | 4.3898
| Median |  | 4.0316 | 4.9961 | 4.955 | 4.565 | 4.275 | 4.63 | 4.765 | 4.59 | 4.7851 | 4.7478 | 4.59 | 4.51 | 4.7924 | 4.325 | 4.605 | 4.655 | 4.625 | 4.3753
| Std Dev. |  | 0.031945 | 0.0035217 | 0.047329 | 0.06186 | 0.01633 | 0.056214 | 0.10013 | 0.15968 | 0.036379 | 0.090248 | 0.065651 | 0.086487 | 0.049274 | 0.056804 | 0.050859 | 0.073212 | 0.072503 | 0.034255
| Rel.Std.Dev. |  | 0.7949% | 0.0705% | 0.9581% | 1.359% | 0.3818% | 1.214% | 2.102% | 3.455% | 0.7613% | 1.909% | 1.432% | 1.913% | 1.024% | 1.316% | 1.105% | 1.581% | 1.562% | 0.7803%
| PDM3 |  | -13.29% | 7.795% | 6.594% | -1.749% | -7.719% | -0.09509% | 2.782% | -0.2749% | 3.108% | 1.984% | -1.066% | -2.469% | 3.852% | -6.856% | -0.6705% | -0.09509% | 0.1566% | -5.277%
| Z-Score (Absolute) |  | 3.36 | 2.05 | 1.74 | 0.4 | 1.94 | 0.02 | 0.76 | 0.02 | 0.85 | 0.56 | 0.23 | 0.59 | 1.04 | 1.71 | 0.12 | 0.02 | 0.09 | 1.31
| Analytical results for As in OREAS 782 (Certified Value 276 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
0
| 1 | 1 | 266.9 | 304.69 | 288 | 281 | 252.9 | 281 | 304 | 266 | 280.3 | 278.6789 | 273 | 275 | 265.5 | 277 | 271 | 277 | 203.4604 |  |  | 1
| 1 | 2 | 269 | 306.48 | 284 | 280 | 257.3 | 285 | 295 | 261 | 280.3 | 272.8203 | 279 | 283 | 269.3 | 272 | 265 | 264 | 200.1543 |  |  | #N/A
| 1 | 3 | 266.8 | 305.04 | 290 | 285 | 257.9 | 281 | 299 | 270 | 276.4 | 269.6105 | 275 | 288 | 269.4 | 278 | 270 | 269 | 206.3565 |  |  | 16
| 1 | 4 | 267 | 304.67 | 288 | 284 | 258.5 | 274 | 300 | 273 | 279.2 | 269.8145 | 277 | 284 | 298.5 | 283 | 269 | 282 | 197.9976 |  |  | 276.1314
| 1 | 5 | 266.9 | 304.34 | 289 | 283 | 260.1 | 278 | 294 | 270 | 278.3 | 277.7413 | 275 | 283 | 270 | 277 | 263 | 270 | 203.537 |  |  | 33
| 1 | 6 | 270.1 | 304.2 | 288 | 280 | 256 | 278 | 297 | 265 | 272.8 | 283.16 | 270 | 296 | 269.6 | 275 | 264 | 266 | 178.0406
| Mean |  | 267.7833 | 304.9033 | 287.8333 | 282.1667 | 257.1167 | 279.5 | 298.1667 | 267.5 | 277.8833 | 275.3043 | 274.8333 | 284.8333 | 273.7167 | 277 | 267 | 271.3333 | 198.2577
| Median |  | 266.95 | 304.68 | 288 | 282 | 257.6 | 279.5 | 298 | 268 | 278.75 | 275.2808 | 275 | 283.5 | 269.5 | 277 | 267 | 269.5 | 201.8073
| Std Dev. |  | 1.4134 | 0.82662 | 2.0412 | 2.137 | 2.4693 | 3.7283 | 3.656 | 4.3243 | 2.8854 | 5.4363 | 3.1252 | 6.9113 | 12.2526 | 3.6332 | 3.4059 | 6.8605 | 10.324
| Rel.Std.Dev. |  | 0.5278% | 0.2711% | 0.7092% | 0.7573% | 0.9604% | 1.334% | 1.226% | 1.617% | 1.038% | 1.975% | 1.137% | 2.426% | 4.476% | 1.312% | 1.276% | 2.528% | 5.207%
| PDM3 |  | -3.023% | 10.42% | 4.238% | 2.186% | -6.886% | 1.22% | 7.98% | -3.126% | 0.6345% | -0.2995% | -0.4701% | 3.151% | -0.8745% | 0.3146% | -3.307% | -1.738% | -28.2%
| Z-Score (Absolute) |  | 0.67 | 2.65 | 1.12 | 0.62 | 1.63 | 0.38 | 2.05 | 0.7 | 0.23 | 0 | 0.04 | 0.85 | 0.14 | 0.15 | 0.74 | 0.36 | 6.91
| Analytical results for B in OREAS 782 (Indicative Value 9.02 ppm) | [20 empty] | Indicative
B | Round | Replicate | Lab | [18 empty] | 1
| No | No | 09 | [18 empty] | ppm
|  |  | 4A*OES | [18 empty] | 2
2
| 1 | 1 | 8.49 | [18 empty] | 1
| 1 | 2 | 8.89 | [18 empty] | 16
| 1 | 3 | 9.33 | [18 empty] | 16
| 1 | 4 | 9.2 | [18 empty] | 9.0233
| 1 | 5 | 9.24 | [18 empty] | 22
| 1 | 6 | 8.99
| Mean |  | 9.0233
| Median |  | 9.095
| Std Dev. |  | 0.30839
| Rel.Std.Dev. |  | 3.418%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 782 (Certified Value 1065 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 986 | 1137.18 | 1110 | 1080 | 926.3 | 1075 | 1120 | 1070 | 1098.3 | 1046.2458 | 1030 | 1086 | 1085.9662 | > 1000 | 980.5844 | 1060 | 1070 | 1050 | 1013.9481 | 1
| 1 | 2 | 996 | 1130.625 | 1110 | 1080 | 930.6 | 1081 | 1090 | 1060 | 1111.3 | 1046.4892 | 1050 | 1049 | 961.4667 | > 1000 | 993.7227 | 1060 | 1050 | 1040 | 1043.7262 | 20
| 1 | 3 | 980 | 1118.75 | 1130 | 1090 | 935.6 | 1041 | 1080 | 1050 | 1113.1 | 1036.4767 | 1040 | 1071 | 1088.133 | > 1000 | 1015.196 | 1060 | 1080 | 1040 | 1010.1451 | 16
| 1 | 4 | 991 | 1126.92 | 1120 | 1100 | 936.1 | 1079 | 1100 | 1080 | 1099.2 | 1021.7077 | 1050 | 1040 | 1076.1608 | > 1000 | 1022.658 | 1060 | 1070 | 1090 | 1029.0247 | 1065.1963
| 1 | 5 | 985 | 1119.7 | 1120 | 1110 | 930.4 | 1045 | 1060 | 1070 | 1102.8 | 1054.3729 | 1060 | 1086 | 1090.2627 | > 1000 | 1008.988 | 1080 | 1050 | 1040 | 1012.2749 | 34
| 1 | 6 | 999 | 1133.19 | 1140 | 1080 | 932.1 | 1088 | 1090 | 1080 | 1092.9 | 1063.5456 | 1030 | 1078 | 1088.7667 | > 1000 | 1012.858 | 1050 | 1080 | 1040 | 1032.9426
| Mean |  | 989.5 | 1127.7275 | 1121.6667 | 1090 | 931.85 | 1068.1667 | 1090 | 1068.3333 | 1102.9333 | 1044.8063 | 1043.3333 | 1068.3333 | 1065.126 |  | 1005.6679 | 1061.6667 | 1066.6667 | 1050 | 1023.6769
| Median |  | 988.5 | 1128.7725 | 1120 | 1085 | 931.35 | 1077 | 1090 | 1070 | 1101 | 1046.3675 | 1045 | 1074.5 | 1087.0496 |  | 1010.923 | 1060 | 1070 | 1040 | 1021.4864
| Std Dev. |  | 7.1764 | 7.3929 | 11.6905 | 12.6491 | 3.6512 | 19.9842 | 20 | 11.6905 | 7.8671 | 14.4949 | 12.1106 | 19.5004 | 51.0321 |  | 15.5823 | 9.8319 | 13.6626 | 20 | 13.5955
| Rel.Std.Dev. |  | 0.7253% | 0.6556% | 1.042% | 1.16% | 0.3918% | 1.871% | 1.835% | 1.094% | 0.7133% | 1.387% | 1.161% | 1.825% | 4.791% |  | 1.549% | 0.9261% | 1.281% | 1.905% | 1.328%
| PDM3 |  | -7.106% | 5.87% | 5.301% | 2.329% | -12.52% | 0.2789% | 2.329% | 0.2945% | 3.543% | -1.914% | -2.052% | 0.2945% | -0.006595% |  | -5.588% | -0.3314% | 0.138% | -1.427% | -3.898%
| Z-Score (Absolute) |  | 2.21 | 1.79 | 1.61 | 0.7 | 3.87 | 0.07 | 0.7 | 0.07 | 1.07 | 0.61 | 0.65 | 0.07 | 0.02 | NA | 1.74 | 0.12 | 0.02 | 0.46 | 1.22
| Analytical results for Be in OREAS 782 (Certified Value 2.68 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.14 | 2.88 | 2.74 | 2.42 | 2.59 | 2.8 | 2.6 | 2.62 | 2.6952 | 2.8 | 2.72 | 2.8012 | 3 | 2.749 | 2.74 | 2.65 | 2.68 | 2.5777 |  | 1
| 1 | 2 | 2.18 | 2.9 | 2.74 | 2.41 | 2.73 | 2.8 | 2.6 | 2.66 | 2.6996 | 2.8 | 2.51 | 3.3897 | 3 | 2.6949 | 2.67 | 2.63 | 2.6 | 2.4766 |  | #N/A
| 1 | 3 | 2.13 | 2.94 | 2.78 | 2.55 | 2.51 | 2.6 | 2.6 | 2.74 | 2.6935 | 2.9 | 2.63 | 2.8214 | 3 | 2.7663 | 2.75 | 2.65 | 2.66 | 2.5316 |  | 16
| 1 | 4 | 2.16 | 2.9 | 2.8 | 2.45 | 2.58 | 2.5 | 2.6 | 2.59 | 2.653 | 2.8 | 2.67 | 3.3836 | 3 | 2.7555 | 2.72 | 2.64 | 2.73 | 2.5403 |  | 2.6781
| 1 | 5 | 2.16 | 2.93 | 2.83 | 2.34 | 2.59 | 2.9 | 2.6 | 2.67 | 2.7082 | 2.9 | 2.65 | 3.0668 | 3 | 2.7338 | 2.72 | 2.58 | 2.63 | 2.4651 |  | 35
| 1 | 6 | 2.18 | 2.96 | 2.74 | 2.44 | 2.83 | 3.4 | 2.6 | 2.7 | 2.7448 | 2.7 | 2.64 | 2.5088 | 3 | 2.6996 | 2.67 | 2.62 | 2.63 | 2.521
| Mean |  | 2.1583 | 2.9183 | 2.7717 | 2.435 | 2.6383 | 2.8333 | 2.6 | 2.6633 | 2.699 | 2.8167 | 2.6367 | 2.9952 | 3 | 2.7332 | 2.7117 | 2.6283 | 2.655 | 2.5187
| Median |  | 2.16 | 2.915 | 2.76 | 2.43 | 2.59 | 2.8 | 2.6 | 2.665 | 2.6974 | 2.8 | 2.645 | 2.9441 | 3 | 2.7414 | 2.72 | 2.635 | 2.645 | 2.5263
| Std Dev. |  | 0.020412 | 0.029944 | 0.038166 | 0.068337 | 0.11805 | 0.31411 | 0 | 0.053914 | 0.02946 | 0.075277 | 0.069761 | 0.35102 | 0 | 0.029789 | 0.034303 | 0.026394 | 0.045935 | 0.041868
| Rel.Std.Dev. |  | 0.9457% | 1.026% | 1.377% | 2.806% | 4.475% | 11.09% | 0% | 2.024% | 1.092% | 2.673% | 2.646% | 11.72% | 0% | 1.09% | 1.265% | 1.004% | 1.73% | 1.662%
| PDM3 |  | -19.41% | 8.971% | 3.494% | -9.077% | -1.484% | 5.797% | -2.916% | -0.5508% | 0.7822% | 5.175% | -1.547% | 11.84% | 12.02% | 2.057% | 1.254% | -1.858% | -0.862% | -5.951%
| Z-Score (Absolute) |  | 4.88 | 2.46 | 1.05 | 2.2 | 0.24 | 1.64 | 0.61 | 0 | 0.34 | 1.48 | 0.26 | 3.2 | NA | 0.67 | 0.47 | 0.34 | 0.08 | 1.4
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 782 (Certified Value 248 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 271.62 | 173.2092 | 271 | 264 | 216.8 | 252 | 236 | 231 | 246.9 | 251.7493 | 243 | 261 | 242.5169 | 251.6 | 231.3327 | 259 | 245 | 262 | 174.9835 | 1
| 1 | 2 | 240.1 | 173.3375 | 263 | 260 | 217.5 | 262 | 229 | 232 | 238.45 | 246.3331 | 250 | 257 | 235.5669 | 251 | 237.411 | 254 | 245 | 253 | 176.9842 | 1
| 1 | 3 | 254.14 | 173.7668 | 279 | 260 | 238.2 | 256 | 233 | 235 | 242.27 | 246.6434 | 244 | 250 | 245.5569 | 240.3 | 247.188 | 258 | 250 | 253 | 176.8554 | 16
| 1 | 4 | 270.2 | 173.8113 | 274 | 262 | 214.5 | 244 | 230 | 234 | 239.74 | 244.5184 | 242 | 256 | 240.6669 | 240.9 | 231.4376 | 255 | 243 | 269 | 179.348 | 247.7663
| 1 | 5 | 274.9 | 173.4154 | 273 | 265 | 214 | 253 | 231 | 230 | 242.37 | 250.6674 | 250 | 258 | 241.7369 | 243.2 | 234.3617 | 257 | 250 | 255 | 174.7617 | 36
| 1 | 6 | 268.34 | 174.2345 | 280 | 257 | 215 | 253 | 234 | 232 | 241.53 | 251.8017 | 246 | 254 | 244.8869 | 249.4 | 234.8252 | 251 | 246 | 259 | 174.8969
| Mean |  | 263.2167 | 173.6291 | 273.3333 | 261.3333 | 219.3333 | 253.3333 | 232.1667 | 232.3333 | 241.8767 | 248.6189 | 245.8333 | 256 | 241.8219 | 246.0667 | 236.0927 | 255.6667 | 246.5 | 258.5 | 176.3049
| Median |  | 269.27 | 173.5911 | 273.5 | 261 | 215.9 | 253 | 232 | 232 | 241.9 | 248.6554 | 245 | 256.5 | 242.1269 | 246.3 | 234.5935 | 256 | 245.5 | 257 | 175.9194
| Std Dev. |  | 13.4082 | 0.38096 | 6.1536 | 2.9439 | 9.3414 | 5.8538 | 2.6394 | 1.8619 | 2.9003 | 3.1644 | 3.4881 | 3.7417 | 3.5859 | 5.1814 | 5.8953 | 2.9439 | 2.881 | 6.253 | 1.7964
| Rel.Std.Dev. |  | 5.094% | 0.2194% | 2.251% | 1.127% | 4.259% | 2.311% | 1.137% | 0.8014% | 1.199% | 1.273% | 1.419% | 1.462% | 1.483% | 2.106% | 2.497% | 1.151% | 1.169% | 2.419% | 1.019%
| PDM3 |  | 6.236% | -29.92% | 10.32% | 5.476% | -11.48% | 2.247% | -6.296% | -6.229% | -2.377% | 0.3441% | -0.7802% | 3.323% | -2.399% | -0.686% | -4.712% | 3.189% | -0.5111% | 4.332% | -28.84%
| Z-Score (Absolute) |  | 1.16 | 4.9 | 1.84 | 1.03 | 1.81 | 0.49 | 0.94 | 0.93 | 0.28 | 0.17 | 0.02 | 0.67 | 0.29 | 0 | 0.67 | 0.65 | 0.03 | 0.84 | 4.72
| Analytical results for Ca in OREAS 782 (Certified Value 16.42 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 15.3846 | 17.4842 | 17.4 | 16.65 | 16.83 | 16.43 | 16.4 | 16.4 | 16.3972 | 15.5127 | 16.9 | 15.575 | 16.803 | 16.37 | 16.6 | 16.85 | 16.45 | 15.3086 |  | 1
| 1 | 2 | 16.0047 | 17.4528 | 17.35 | 16.7 | 16.59 | 16.49 | 16.3 | 16.8 | 16.0515 | 15.607 | 16.9 | 15.338 | 16.8631 | 16.13 | 16.15 | 16.45 | 16 | 15.2091 |  | 22
| 1 | 3 | 15.627 | 17.4138 | 17.7 | 16.9 | 16.98 | 16.04 | 16.1 | 16.8 | 16.1999 | 15.3662 | 17 | 15.708 | 16.8919 | 16.14 | 16.5 | 17.1 | 16.25 | 15.1953 |  | 16
| 1 | 4 | 15.2836 | 17.6807 | 17.55 | 17 | 16.35 | 16.39 | 16.3 | 16.7 | 16.0599 | 15.6469 | 16.8 | 15.368 | 17 | 16.35 | 16.35 | 16.8 | 17.05 | 15.2987 |  | 16.4202
| 1 | 5 | 15.4037 | 17.5409 | 17.7 | 16.95 | 16.46 | 16.52 | 16 | 16.5 | 16.5318 | 15.7296 | 17.1 | 15.691 | 17.0271 | 16.35 | 16.4 | 16.5 | 16.2 | 15.1564 |  | 37
| 1 | 6 | 16.1738 | 17.3681 | 17.85 | 16.7 | 16.49 | 16.05 | 16.4 | 16.7 | 15.7743 | 16.02 | 16.5 | 15.842 | 16.8091 | 16.33 | 16.25 | 16.8 | 16.3 | 15.2416
| Mean |  | 15.6462 | 17.4901 | 17.5917 | 16.8167 | 16.6167 | 16.32 | 16.25 | 16.65 | 16.1691 | 15.6471 | 16.8667 | 15.587 | 16.899 | 16.2783 | 16.375 | 16.75 | 16.375 | 15.235
| Median |  | 15.5153 | 17.4685 | 17.625 | 16.8 | 16.54 | 16.41 | 16.3 | 16.7 | 16.1299 | 15.6269 | 16.9 | 15.633 | 16.8775 | 16.34 | 16.375 | 16.8 | 16.275 | 15.2254
| Std Dev. |  | 0.36496 | 0.11049 | 0.19343 | 0.15055 | 0.2408 | 0.21781 | 0.16432 | 0.16432 | 0.27061 | 0.2211 | 0.20656 | 0.20029 | 0.095129 | 0.11179 | 0.16355 | 0.24083 | 0.36159 | 0.059905
| Rel.Std.Dev. |  | 2.333% | 0.6317% | 1.1% | 0.8953% | 1.449% | 1.335% | 1.011% | 0.9869% | 1.674% | 1.413% | 1.225% | 1.285% | 0.5629% | 0.6867% | 0.9988% | 1.438% | 2.208% | 0.3932%
| PDM3 |  | -4.713% | 6.516% | 7.134% | 2.415% | 1.197% | -0.6102% | -1.036% | 1.4% | -1.529% | -4.708% | 2.719% | -5.074% | 2.916% | -0.8639% | -0.2752% | 2.009% | -0.2752% | -7.218%
| Z-Score (Absolute) |  | 1.2 | 1.84 | 2.01 | 0.73 | 0.4 | 0.09 | 0.21 | 0.45 | 0.34 | 1.2 | 0.81 | 1.3 | 0.87 | 0.16 | 0 | 0.62 | 0 | 1.88
| Analytical results for Cd in OREAS 782 (Indicative Value 0.028 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.04 | < 0.02 | 0.03 | < 0.2 | 0.05 | < 0.3 | < 0.05 | 0.03 | < 0.5 | < 0.5 | 0.05 | < 0.003 | < 0.1 | 0.03638 | 0.03 | < 0.02 | < 0.02 | 0.7292 |  | 1
| 1 | 2 | 0.03 | < 0.02 | < 0.02 | < 0.2 | 0.05 | < 0.3 | < 0.05 | 0.02 | < 0.5 | < 0.5 | 0.05 | < 0.0492 | < 0.1 | 0.03109 | < 0.02 | 0.03 | < 0.02 | 0.9793 |  | 2
| 1 | 3 | 0.03 | 0.02 | < 0.02 | < 0.2 | 0.05 | < 0.3 | < 0.05 | 0.03 | < 0.5 | < 0.5 | 0.04 | < 0.0294 | < 0.1 | 0.0401 | < 0.02 | < 0.02 | < 0.02 | 0.9918 |  | 16
| 1 | 4 | 0.03 | 0.03 | < 0.02 | < 0.2 | 0.05 | < 0.3 | < 0.05 | < 0.02 | < 0.5 | < 0.5 | 0.04 | < 0.0025 | < 0.1 | 0.03426 | < 0.02 | < 0.02 | < 0.02 | 0.7744 |  | 0.028113
| 1 | 5 | 0.04 | 0.02 | < 0.02 | < 0.2 | 0.05 | < 0.3 | < 0.05 | 0.02 | < 0.5 | < 0.5 | 0.05 | < 0.039 | < 0.1 | 0.03821 | 0.02 | 0.03 | < 0.02 | 0.7178 |  | 23
| 1 | 6 | 0.03 | 0.02 | < 0.02 | < 0.2 | 0.05 | < 0.3 | < 0.05 | 0.03 | < 0.5 | < 0.5 | 0.06 | < 0.0256 | < 0.1 | 0.02804 | < 0.02 | < 0.02 | < 0.02 | 0.8505
| Mean |  | 0.033333 | 0.0225 | 0.03 |  | 0.05 |  |  | 0.026 |  |  | 0.048333 |  |  | 0.03468 | 0.025 | 0.03 |  | 0.8405
| Median |  | 0.03 | 0.02 | 0.03 |  | 0.05 |  |  | 0.03 |  |  | 0.05 |  |  | 0.03532 | 0.025 | 0.03 |  | 0.81245
| Std Dev. |  | 0.005164 | 0.005 |  |  | 7.6012e-18 |  |  | 0.0054772 |  |  | 0.0075277 |  |  | 0.004514 | 0.0070711 | 0 |  | 0.12171
| Rel.Std.Dev. |  | 15.49% | 22.22% |  |  | 1.52e-14% |  |  | 21.07% |  |  | 15.57% |  |  | 13.02% | 28.28% | 0% |  | 14.48%
| PDM3 |  | 18.57% | -19.96% | 6.714% |  | 77.86% |  |  | -7.515% |  |  | 71.93% |  |  | 23.36% | -11.07% | 6.714% |  | 2890%
| Z-Score (Absolute) |  | 0.02 | 0.58 | 0.77 | 2.45 | 0.59 | 4.3 | 0.33 | 0.4 | 8.01 | 8.01 | 0.53 | 0.8 | 0.59 | 0.02 | 0.71 | 0.64 | 0.89 | 29.92
| Analytical results for Ce in OREAS 782 (Certified Value 64 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 48.9 | 64.393 | 70 | 63.3 | 65 | 65.4 | 64.29 | 65.8 | 61.82 | 66.2247 | 66 | 56.0259 | 64.7 | 64.8 | 58 | 60.2055 |  |  |  | 1
| 1 | 2 | 50.3 | 63.5816 | 71.6 | 64.5 | 66 | 63.8 | 63.8 | 67.2 | 62.58 | 65.7796 | 70 | 57.9494 | 65.1 | 64.8 | 58.6 | 60.6914 |  |  |  | 23
| 1 | 3 | 49.6 | 64.4159 | 73.2 | 63.9 | 65 | 64.8 | 63.69 | 66.2 | 63.02 | 66.9326 | 68 | 58.5086 | 65.2 | 67.2 | 57.6 | 61.1276 |  |  |  | 16
| 1 | 4 | 50.2 | 64.4625 | 71.3 | 62.8 | 65 | 64.5 | 64.38 | 66.6 | 59.34 | 67.5564 | 66 | 60.4076 | 66.7 | 63.8 | 61.4 | 57.4616 |  |  |  | 64.1815
| 1 | 5 | 49.7 | 63.8467 | 70.9 | 63 | 67 | 63.3 | 64.39 | 63.8 | 63.03 | 66.3336 | 69 | 58.1571 | 66 | 63.3 | 59.2 | 65.3962 |  |  |  | 38
| 1 | 6 | 51 | 64.4412 | 72.3 | 63 | 66 | 64.1 | 63.29 | 64.8 | 63.12 | 66.9807 | 65 | 58.9294 | 63.7 | 66.4 | 59.6 | 61.3052
| Mean |  | 49.95 | 64.1902 | 71.55 | 63.4167 | 65.6667 | 64.3167 | 63.9733 | 65.7333 | 62.1517 | 66.6346 | 67.3333 | 58.3297 | 65.2333 | 65.05 | 59.0667 | 61.0312
| Median |  | 49.95 | 64.4044 | 71.45 | 63.15 | 65.5 | 64.3 | 64.045 | 66 | 62.8 | 66.6331 | 67 | 58.3329 | 65.15 | 64.8 | 58.9 | 60.9095
| Std Dev. |  | 0.71764 | 0.37881 | 1.1113 | 0.65549 | 0.8165 | 0.74677 | 0.4509 | 1.2437 | 1.4603 | 0.64 | 1.9664 | 1.4277 | 1.0386 | 1.4963 | 1.3604 | 2.5564
| Rel.Std.Dev. |  | 1.437% | 0.5901% | 1.553% | 1.034% | 1.243% | 1.161% | 0.7048% | 1.892% | 2.35% | 0.9605% | 2.92% | 2.448% | 1.592% | 2.3% | 2.303% | 4.189%
| PDM3 |  | -22.17% | 0.01343% | 11.48% | -1.192% | 2.314% | 0.2105% | -0.3244% | 2.418% | -3.163% | 3.822% | 4.911% | -9.118% | 1.639% | 1.353% | -7.969% | -4.908%
| Z-Score (Absolute) |  | 5.39 | 0.02 | 2.75 | 0.32 | 0.53 | 0.02 | 0.11 | 0.56 | 0.79 | 0.9 | 1.16 | 2.23 | 0.37 | 0.3 | 1.95 | 1.21
| Analytical results for Co in OREAS 782 (Certified Value 2.04 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
2
| 1 | 1 | 1.3 | 2.0397 | 2.2 | 2.4 | 4.28 | 1.99 | 2 | 2 | 1.9 | 1.8894 | 2 | 2.3 | 3.6999 | 1.9 | 1.973 | 2.1 | 2.1 | 2.4 | 4.1947 | 1
| 1 | 2 | 1.4 | 1.992 | 2.3 | 2.5 | 4.24 | 1.76 | 1 | 2 | 2 | 1.8454 | 2 | 2.2 | 3.451 | 1.9 | 2.1939 | 2.1 | 2 | 2.2 | 3.9317 | #N/A
| 1 | 3 | 1.4 | 1.9429 | 2.3 | 2.5 | 4.16 | 1.85 | 2 | 2 | 2 | 1.8652 | 2 | 2 | 3.8442 | 1.9 | 1.9696 | 2.1 | 2.2 | 2.2 | 4.0982 | 16
| 1 | 4 | 1.5 | 1.9496 | 2.3 | 2.6 | 4.31 | 1.78 | 1 | 2 | 2 | 1.8788 | 2 | 2.2 | 3.9566 | 1.8 | 1.962 | 2.3 | 2.2 | 2.3 | 4.0986 | 2.0359
| 1 | 5 | 1.6 | 1.9792 | 2.3 | 2.5 | 4.28 | 1.9 | 1 | 2 | 1.9 | 1.9032 | 2 | 2.1 | 3.8623 | 1.9 | 1.9119 | 2.2 | 2.1 | 2.2 | 4.1105 | 39
| 1 | 6 | 1.5 | 1.9735 | 2.3 | 2.5 | 4.13 | 1.72 | 2 | 2 | 1.9 | 1.9469 | 2 | 2.2 | 3.6245 | 2 | 1.8859 | 2.1 | 2.2 | 2.2 | 4.4314
| Mean |  | 1.45 | 1.9795 | 2.2833 | 2.5 | 4.2333 | 1.8333 | 1.5 | 2 | 1.95 | 1.8881 | 2 | 2.1667 | 3.7397 | 1.9 | 1.9827 | 2.15 | 2.1333 | 2.25 | 4.1442
| Median |  | 1.45 | 1.9764 | 2.3 | 2.5 | 4.26 | 1.815 | 1.5 | 2 | 1.95 | 1.8841 | 2 | 2.2 | 3.772 | 1.9 | 1.9658 | 2.1 | 2.15 | 2.2 | 4.1045
| Std Dev. |  | 0.10488 | 0.034784 | 0.040825 | 0.063246 | 0.072572 | 0.10033 | 0.54772 | 0 | 0.054772 | 0.034986 | 0 | 0.10328 | 0.18502 | 0.063246 | 0.10925 | 0.083666 | 0.08165 | 0.083666 | 0.16463
| Rel.Std.Dev. |  | 7.233% | 1.757% | 1.788% | 2.53% | 1.714% | 5.473% | 36.51% | 0% | 2.809% | 1.853% | 0% | 4.767% | 4.948% | 3.329% | 5.51% | 3.891% | 3.827% | 3.718% | 3.973%
| PDM3 |  | -28.78% | -2.772% | 12.15% | 22.79% | 107.9% | -9.951% | -26.32% | -1.764% | -4.22% | -7.259% | -1.764% | 6.422% | 83.69% | -6.676% | -2.614% | 5.603% | 4.785% | 10.51% | 103.6%
| Z-Score (Absolute) |  | 2.15 | 0.5 | 0.44 | 1.12 | 6.51 | 0.96 | NA | NA | 0.6 | 0.79 | NA | 0.08 | 4.97 | 0.75 | 0.49 | 0.03 | 0.03 | 0.34 | 6.23
Results from laboratories 10, 12 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 782 (Certified Value 60 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 61.9 | 15.606 | 61 | 49 | 75.05 | 60 | 48 | 40 | 41 | 56.9502 | 36 | 45.9674 | 65 | 76.9016 | 56 | 50 | 59 | 67.0404 |  | 1
| 1 | 2 | 62.8 | 16.74 | 60 | 52 | 74 | 60 | 53 | 35 | 65 | 59.462 | 39 | 43.234 | 65 | 76.5635 | 56 | 54 | 60 | 69.1835 |  | #N/A
| 1 | 3 | 61.8 | 17.676 | 59 | 51 | 73.65 | 57 | 36 | 39 | 59 | 58.5718 | 38 | 50.701 | 64 | 78.7064 | 53 | 49 | 60 | 70.6162 |  | 16
| 1 | 4 | 61.7 | 16.0625 | 57 | 53 | 72.04 | 61 | 36 | 33 | 64 | 57.3731 | 31 | 50.5937 | 64 | 77.5152 | 54 | 51 | 62 | 97.2933 |  | 59.9281
| 1 | 5 | 61.6 | 16.873 | 61 | 51 | 75.76 | 59 | 38 | 36 | 68 | 57.2903 | 27 | 60.3385 | 64 | 78.2106 | 58 | 55 | 58 | 73.9641 |  | 40
| 1 | 6 | 62.1 | 18.231 | 59 | 50 | 74.25 | 58 | 41 | 40 | 70 | 55.7041 | 37 | 51.4499 | 64 | 78.1057 | 53 | 52 | 59 | 86.8932
| Mean |  | 61.9833 | 16.8647 | 59.5 | 51 | 74.125 | 59.1667 | 42 | 37.1667 | 61.1667 | 57.5586 | 34.6667 | 50.3808 | 64.3333 | 77.6672 | 55 | 51.8333 | 59.6667 | 77.4984
| Median |  | 61.85 | 16.8065 | 59.5 | 51 | 74.125 | 59.5 | 39.5 | 37.5 | 64.5 | 57.3317 | 36.5 | 50.6474 | 64 | 77.8105 | 55 | 51.5 | 59.5 | 72.2902
| Std Dev. |  | 0.43551 | 0.97657 | 1.5166 | 1.4142 | 1.2754 | 1.472 | 7.0143 | 2.9269 | 10.572 | 1.3088 | 4.6762 | 5.8481 | 0.5164 | 0.824 | 2 | 2.3166 | 1.3663 | 11.9872
| Rel.Std.Dev. |  | 0.7026% | 5.791% | 2.549% | 2.773% | 1.721% | 2.488% | 16.7% | 7.875% | 17.28% | 2.274% | 13.49% | 11.61% | 0.8027% | 1.061% | 3.636% | 4.469% | 2.29% | 15.47%
| PDM3 |  | 3.43% | -71.86% | -0.7143% | -14.9% | 23.69% | -1.271% | -29.92% | -37.98% | 2.067% | -3.954% | -42.15% | -15.93% | 7.351% | 29.6% | -8.223% | -13.51% | -0.4362% | 29.32%
| Z-Score (Absolute) |  | 0.35 | 4.03 | 0.11 | 0.71 | 1.53 | 0.08 | 1.59 | 2.06 | 0.27 | 0.08 | 2.3 | 0.77 | 0.58 | 1.87 | 0.33 | 0.63 | 0.13 | 1.86
| Analytical results for Cs in OREAS 782 (Certified Value 6.53 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 7 | 5.6299 | 6.54 | 6.33 | 6.31 | 6.91 | 6.4 | 6.54 | 6.4 | 6.4 | 6.45 | 7 | 6.5615 | 6.87 | 6.91 | 6.34 | 6.1851 |  |  | 1
| 1 | 2 | 6.1 | 5.8337 | 6.79 | 6.4 | 6.43 | 6.61 | 6.4 | 6.65 | 6.2 | 6.23 | 5.8368 | 7.2 | 6.5179 | 6.83 | 6.81 | 6.51 | 6.2701 |  |  | 5
| 1 | 3 | 6.6 | 5.4907 | 6.87 | 6.31 | 6.26 | 6.76 | 6.4 | 6.56 | 6.4 | 6.52 | 6.3986 | 6.9 | 6.4628 | 6.67 | 7.01 | 6.41 | 6.3943 |  |  | 16
| 1 | 4 | 7 | 5.682 | 6.53 | 6.21 | 6.45 | 6.63 | 6.4 | 6.51 | 6.4 | 6.27 | 6.4076 | 6.8 | 6.4677 | 6.74 | 6.69 | 7.05 | 6.1815 |  |  | 6.5309
| 1 | 5 | 7 | 5.7294 | 6.73 | 6.29 | 6.33 | 6.57 | 6.5 | 6.55 | 6.6 | 6.61 | 6.3806 | 7.1 | 6.5792 | 6.96 | 6.55 | 6.6 | 6.4819 |  |  | 41
| 1 | 6 | 6.9 | 5.6013 | 6.78 | 6.3 | 6.29 | 6.56 | 6.4 | 6.43 | 6.2 | 6.75 | 6.3201 | 7.1 | 6.3385 | 6.78 | 6.87 | 6.57 | 6.359
| Mean |  | 6.7667 | 5.6612 | 6.7067 | 6.3067 | 6.345 | 6.6733 | 6.4167 | 6.54 | 6.3667 | 6.4633 | 6.2989 | 7.0167 | 6.4879 | 6.8083 | 6.8067 | 6.58 | 6.312
| Median |  | 6.95 | 5.6559 | 6.755 | 6.305 | 6.32 | 6.62 | 6.4 | 6.545 | 6.4 | 6.46 | 6.3896 | 7.05 | 6.4928 | 6.805 | 6.84 | 6.54 | 6.3146
| Std Dev. |  | 0.36148 | 0.11702 | 0.14038 | 0.061536 | 0.077395 | 0.13633 | 0.040825 | 0.071554 | 0.15055 | 0.20136 | 0.23034 | 0.1472 | 0.087193 | 0.10187 | 0.16464 | 0.25012 | 0.12061
| Rel.Std.Dev. |  | 5.342% | 2.067% | 2.093% | 0.9757% | 1.22% | 2.043% | 0.6362% | 1.094% | 2.365% | 3.115% | 3.657% | 2.098% | 1.344% | 1.496% | 2.419% | 3.801% | 1.911%
| PDM3 |  | 3.61% | -13.32% | 2.692% | -3.433% | -2.846% | 2.181% | -1.749% | 0.1395% | -2.515% | -1.034% | -3.551% | 7.438% | -0.6576% | 4.248% | 4.223% | 0.752% | -3.352%
| Z-Score (Absolute) |  | 1.04 | 3.08 | 0.81 | 0.67 | 0.53 | 0.69 | 0.27 | 0.19 | 0.45 | 0.09 | 0.7 | 1.97 | 0 | 1.19 | 1.19 | 0.34 | 0.65
| Analytical results for Cu in OREAS 782 (Certified Value 80 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 66.9 | 90.04 | 83.3 | 81.5 | 74.37 | 79 | 83 | 81 | 80 | 81.238 | 76 | 80 | 76.0969 | 82.7 | 76.2943 | 83.9 | 78.6 | 81.4 | 73.177 | 1
| 1 | 2 | 66.2 | 88.63 | 83.6 | 84.4 | 75.09 | 82 | 80 | 81 | 81 | 78.4639 | 75 | 77.5 | 76.4269 | 82.7 | 78.4783 | 81.9 | 77.1 | 79 | 75.4545 | #N/A
| 1 | 3 | 66.3 | 88.52 | 85.3 | 86 | 74.59 | 80 | 80 | 82 | 81 | 78.1798 | 76 | 80.4 | 78.1569 | 79.8 | 78.9814 | 82.4 | 80.8 | 79.2 | 78.2264 | 16
| 1 | 4 | 68.4 | 88.29 | 82.6 | 84 | 74.38 | 79 | 80 | 80 | 78 | 80.3602 | 76 | 80.5 | 76.9869 | 76.7 | 78.3388 | 85.4 | 79.3 | 83.2 | 73.1934 | 80.0911
| 1 | 5 | 70.7 | 87.47 | 83.3 | 82.1 | 76.53 | 81 | 78 | 82 | 80 | 80.7855 | 79 | 81 | 77.1469 | 82.1 | 77.6182 | 82.3 | 78.4 | 79.7 | 77.5713 | 42
| 1 | 6 | 67.8 | 88.09 | 84 | 82.7 | 74.81 | 80 | 80 | 82 | 78 | 82.2047 | 74 | 80.9 | 78.2369 | 81.1 | 76.5655 | 81.4 | 79.1 | 82.5 | 73.0428
| Mean |  | 67.7167 | 88.5067 | 83.6833 | 83.45 | 74.9617 | 80.1667 | 80.1667 | 81.3333 | 79.6667 | 80.2053 | 76 | 80.05 | 77.1753 | 80.85 | 77.7127 | 82.8833 | 78.8833 | 80.8333 | 75.1109
| Median |  | 67.35 | 88.405 | 83.45 | 83.35 | 74.7 | 80 | 80 | 81.5 | 80 | 80.5729 | 76 | 80.45 | 77.0669 | 81.6 | 77.9785 | 82.35 | 78.85 | 80.55 | 74.3239
| Std Dev. |  | 1.694 | 0.85577 | 0.91524 | 1.667 | 0.81556 | 1.169 | 1.6021 | 0.8165 | 1.3663 | 1.5849 | 1.6733 | 1.3004 | 0.87748 | 2.315 | 1.0885 | 1.4905 | 1.2156 | 1.7896 | 2.3483
| Rel.Std.Dev. |  | 2.502% | 0.9669% | 1.094% | 1.998% | 1.088% | 1.458% | 1.998% | 1.004% | 1.715% | 1.976% | 2.202% | 1.624% | 1.137% | 2.863% | 1.401% | 1.798% | 1.541% | 2.214% | 3.126%
| PDM3 |  | -15.45% | 10.51% | 4.485% | 4.194% | -6.404% | 0.09439% | 0.09439% | 1.551% | -0.5299% | 0.1427% | -5.108% | -0.05128% | -3.641% | 0.9476% | -2.97% | 3.486% | -1.508% | 0.9268% | -6.218%
| Z-Score (Absolute) |  | 3.42 | 2.29 | 0.97 | 0.9 | 1.43 | 0 | 0 | 0.32 | 0.14 | 0.01 | 1.14 | 0.03 | 0.82 | 0.19 | 0.67 | 0.75 | 0.35 | 0.18 | 1.39
| Analytical results for Dy in OREAS 782 (Certified Value 3.59 ppm) | [20 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 2.56 | 3.606 | 3.27 | 3.7 | 3.68 | 3.6 | 3.43 | 3.8492 | 4.0859 | 3.5616 | [9 empty] | 1
| 1 | 2 | 2.34 | 3.71 | 3.48 | 3.7 | 3.79 | 3.65 | 3.33 | 3.628 | 4.378 | 3.5453 | [9 empty] | 6
| 1 | 3 | 2.43 | 3.5763 | 3.37 | 3.7 | 3.67 | 3.65 | 3.43 | 4.0167 | 4.4584 | 3.4379 | [9 empty] | 16
| 1 | 4 | 2.69 | 3.559 | 3.36 | 3.5 | 3.66 | 3.5 | 3.38 | 3.8249 | 4.6195 | 3.4927 | [9 empty] | 3.5857
| 1 | 5 | 2.75 | 3.7069 | 3.46 | 3.6 | 3.55 | 3.75 | 3.4 | 3.8446 | 4.4131 | 3.5343 | [9 empty] | 43
| 1 | 6 | 2.69 | 3.5645 | 3.47 | 3.6 | 3.69 | 3.5 | 3.47 | 3.7747 | 4.4592 | 3.5419
| Mean |  | 2.5767 | 3.6204 | 3.4017 | 3.6333 | 3.6733 | 3.6083 | 3.4067 | 3.823 | 4.4024 | 3.5189
| Median |  | 2.625 | 3.5912 | 3.415 | 3.65 | 3.675 | 3.625 | 3.415 | 3.8347 | 4.4357 | 3.5381
| Std Dev. |  | 0.16342 | 0.070072 | 0.082805 | 0.08165 | 0.076594 | 0.09704 | 0.048442 | 0.12571 | 0.17574 | 0.0459
| Rel.Std.Dev. |  | 6.342% | 1.935% | 2.434% | 2.247% | 2.085% | 2.689% | 1.422% | 3.288% | 3.992% | 1.304%
| PDM3 |  | -28.14% | 0.9686% | -5.133% | 1.328% | 2.443% | 0.6307% | -4.993% | 6.618% | 22.77% | -1.862%
| Z-Score (Absolute) |  | 4.62 | 0.03 | 0.95 | 0.08 | 0.26 | 0.03 | 0.92 | 0.93 | 3.51 | 0.42
| Analytical results for Er in OREAS 782 (Certified Value 1.89 ppm) | [20 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 1.36 | 1.9026 | 1.82 | 2 | 1.85 | 1.95 | 1.9 | 2.0862 | 2.4326 | 1.7369 | [9 empty] | 1
| 1 | 2 | 1.26 | 1.9711 | 1.64 | 2 | 1.93 | 1.8 | 1.89 | 1.7977 | 2.5099 | 1.796 | [9 empty] | 26
| 1 | 3 | 1.3 | 1.9337 | 1.6 | 1.9 | 1.86 | 1.9 | 2.07 | 2.0266 | 2.6326 | 1.6921 | [9 empty] | 16
| 1 | 4 | 1.46 | 1.9706 | 1.69 | 1.9 | 1.88 | 1.9 | 1.79 | 2.018 | 2.6769 | 1.7377 | [9 empty] | 1.8888
| 1 | 5 | 1.45 | 1.952 | 1.66 | 1.9 | 2.1 | 2 | 1.88 | 2.0404 | 2.5715 | 1.9155 | [9 empty] | 44
| 1 | 6 | 1.44 | 2.0642 | 1.76 | 1.9 | 1.92 | 1.9 | 1.95 | 2.0064 | 2.6416 | 1.7739
| Mean |  | 1.3783 | 1.9657 | 1.695 | 1.9333 | 1.9233 | 1.9083 | 1.9133 | 1.9959 | 2.5775 | 1.7753
| Median |  | 1.4 | 1.9613 | 1.675 | 1.9 | 1.9 | 1.9 | 1.895 | 2.0223 | 2.6021 | 1.7558
| Std Dev. |  | 0.084951 | 0.054671 | 0.081425 | 0.05164 | 0.092232 | 0.066458 | 0.092664 | 0.10098 | 0.092335 | 0.077327
| Rel.Std.Dev. |  | 6.163% | 2.781% | 4.804% | 2.671% | 4.795% | 3.483% | 4.843% | 5.059% | 3.582% | 4.356%
| PDM3 |  | -27.03% | 4.072% | -10.26% | 2.359% | 1.829% | 1.035% | 1.3% | 5.671% | 36.46% | -6.006%
| Z-Score (Absolute) |  | 5.83 | 0.51 | 2.41 | 0.16 | 0.05 | 0.11 | 0.05 | 0.84 | 7.12 | 1.54
| Analytical results for Eu in OREAS 782 (Certified Value 0.98 ppm) | [20 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 0.93 | 2.3857 | 0.96 | 0.99 | 1.02 | 0.95 | 1.04 | 1.0342 | 0.89687 | 1.521 | [9 empty] | 1
| 1 | 2 | 0.82 | 2.1549 | 0.92 | 0.97 | 1.03 | 0.95 | 1.05 | 0.92237 | 0.96286 | 1.5007 | [9 empty] | 7
| 1 | 3 | 0.88 | 2.2301 | 0.93 | 0.97 | 1.02 | 0.95 | 1.06 | 1.0148 | 0.94591 | 1.4427 | [9 empty] | 16
| 1 | 4 | 0.96 | 2.2535 | 0.95 | 1.04 | 0.96 | 0.95 | 1.01 | 1.0565 | 0.97017 | 1.4268 | [9 empty] | 0.97856
| 1 | 5 | 0.97 | 2.1621 | 0.97 | 0.94 | 1.01 | 1 | 1.07 | 1.0259 | 0.96333 | 1.4819 | [9 empty] | 45
| 1 | 6 | 0.94 | 2.2128 | 0.95 | 1.02 | 1.01 | 0.9 | 1.06 | 1.0394 | 0.98872 | 1.5318
| Mean |  | 0.91667 | 2.2332 | 0.94667 | 0.98833 | 1.0083 | 0.95 | 1.0483 | 1.0155 | 0.95464 | 1.4842
| Median |  | 0.935 | 2.2215 | 0.95 | 0.98 | 1.015 | 0.95 | 1.055 | 1.03 | 0.9631 | 1.4913
| Std Dev. |  | 0.056804 | 0.084019 | 0.018619 | 0.03656 | 0.024833 | 0.031623 | 0.02137 | 0.047716 | 0.031494 | 0.042224
| Rel.Std.Dev. |  | 6.197% | 3.762% | 1.967% | 3.699% | 2.463% | 3.329% | 2.038% | 4.699% | 3.299% | 2.845%
| PDM3 |  | -6.325% | 128.2% | -3.26% | 0.9984% | 3.042% | -2.919% | 7.13% | 3.778% | -2.444% | 51.67%
| Z-Score (Absolute) |  | 1.12 | 16.94 | 0.71 | 0.14 | 0.14 | 0.66 | 0.69 | 0.24 | 0.6 | 6.66
| Analytical results for Fe in OREAS 782 (Certified Value 1.82 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.6945 | 1.9633 | 1.87 | 1.82 | 1.79 | 1.72 | 1.87 | 1.78 | 1.79 | 1.8106 | 1.82 | 1.78 | 1.9157 | 1.69 | 1.8 | 1.82 | 1.84 | 1.8 |  | 1
| 1 | 2 | 1.6911 | 1.9651 | 1.87 | 1.8 | 1.85 | 1.71 | 1.83 | 1.79 | 1.82 | 1.7792 | 1.82 | 1.73 | 1.9247 | 1.68 | 1.74 | 1.78 | 1.79 | 1.9 |  | 28
| 1 | 3 | 1.677 | 1.975 | 1.92 | 1.83 | 1.8 | 1.66 | 1.85 | 1.83 | 1.8 | 1.7866 | 1.8 | 1.79 | 1.9144 | 1.68 | 1.79 | 1.82 | 1.81 | 1.8 |  | 16
| 1 | 4 | 1.6888 | 1.9613 | 1.89 | 1.84 | 1.84 | 1.61 | 1.84 | 1.86 | 1.79 | 1.7346 | 1.81 | 1.74 | 1.9302 | 1.67 | 1.77 | 1.82 | 1.9 | 1.8 |  | 1.8248
| 1 | 5 | 1.6657 | 1.9683 | 1.9 | 1.84 | 1.77 | 1.63 | 1.85 | 1.8 | 1.83 | 1.7937 | 1.86 | 1.81 | 1.9525 | 1.69 | 1.78 | 1.77 | 1.85 | 1.8 |  | 46
| 1 | 6 | 1.6937 | 1.9635 | 1.92 | 1.81 | 1.82 | 1.67 | 1.86 | 1.82 | 1.79 | 1.811 | 1.79 | 1.81 | 1.9289 | 1.66 | 1.75 | 1.81 | 1.82 | 1.9
| Mean |  | 1.6851 | 1.9661 | 1.895 | 1.8233 | 1.8117 | 1.6667 | 1.85 | 1.8133 | 1.8033 | 1.786 | 1.8167 | 1.7767 | 1.9277 | 1.6783 | 1.7717 | 1.8033 | 1.835 | 1.8333
| Median |  | 1.69 | 1.9643 | 1.895 | 1.825 | 1.81 | 1.665 | 1.85 | 1.81 | 1.795 | 1.7902 | 1.815 | 1.785 | 1.9268 | 1.68 | 1.775 | 1.815 | 1.83 | 1.8
| Std Dev. |  | 0.011437 | 0.004953 | 0.022583 | 0.01633 | 0.030605 | 0.043205 | 0.014142 | 0.029439 | 0.017512 | 0.028224 | 0.024221 | 0.034448 | 0.013816 | 0.01169 | 0.023166 | 0.022509 | 0.038341 | 0.05164
| Rel.Std.Dev. |  | 0.6787% | 0.2519% | 1.192% | 0.8956% | 1.689% | 2.592% | 0.7644% | 1.623% | 0.9711% | 1.58% | 1.333% | 1.939% | 0.7167% | 0.6966% | 1.308% | 1.248% | 2.089% | 2.817%
| PDM3 |  | -7.653% | 7.743% | 3.848% | -0.07969% | -0.719% | -8.665% | 1.382% | -0.6277% | -1.176% | -2.128% | -0.445% | -2.637% | 5.642% | -8.026% | -2.911% | -1.176% | 0.5597% | 0.4683%
| Z-Score (Absolute) |  | 2.75 | 3.32 | 1.78 | 0.23 | 0.02 | 3.15 | 0.81 | 0.02 | 0.2 | 0.57 | 0.09 | 0.77 | 2.49 | 2.9 | 0.88 | 0.2 | 0.49 | 0.45
| Analytical results for Ga in OREAS 782 (Certified Value 11.8 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 10.68 | 11.777 | 13.05 | 12.45 | 11.84 | 12.3 | 11.9 | 11.67 | 11.4 | 12.35 | 13.5623 | 13 | 11.6901 | 11.25 | 11 | 11.55 | 10.9909 |  |  | 1
| 1 | 2 | 10.55 | 11.7847 | 13 | 12.5 | 12.34 | 12 | 12 | 12.06 | 10.6 | 11.64 | 12.0145 | 13 | 11.7153 | 11.15 | 11.05 | 11.45 | 10.8575 |  |  | 8
| 1 | 3 | 10.59 | 11.2561 | 13.25 | 12.65 | 12.09 | 12.2 | 11.6 | 12.13 | 11.4 | 12.31 | 13.1773 | 13 | 11.6444 | 11.15 | 11.3 | 11.5 | 10.42 |  |  | 16
| 1 | 4 | 10.46 | 11.7962 | 12.95 | 12.45 | 11.52 | 12.5 | 12.1 | 12.24 | 11 | 11.99 | 13.2009 | 13 | 11.7902 | 11.75 | 10.9 | 12.15 | 10.3946 |  |  | 11.834
| 1 | 5 | 10.38 | 11.5922 | 13.3 | 12.2 | 11.32 | 11.4 | 11.8 | 12.02 | 11.8 | 12.26 | 13.3181 | 13 | 11.8115 | 11.4 | 10.9 | 11.85 | 10.8879 |  |  | 47
| 1 | 6 | 10.65 | 11.8237 | 12.8 | 12.3 | 11.8 | 11.9 | 11.8 | 11.98 | 11 | 12.2 | 13.0021 | 13 | 11.3578 | 11.1 | 11.2 | 11.8 | 10.0811
| Mean |  | 10.5517 | 11.6717 | 13.0583 | 12.425 | 11.8183 | 12.05 | 11.8667 | 12.0167 | 11.2 | 12.125 | 13.0459 | 13 | 11.6682 | 11.3 | 11.0583 | 11.7167 | 10.6053
| Median |  | 10.57 | 11.7808 | 13.025 | 12.45 | 11.82 | 12.1 | 11.85 | 12.04 | 11.2 | 12.23 | 13.1891 | 13 | 11.7027 | 11.2 | 11.025 | 11.675 | 10.6387
| Std Dev. |  | 0.11444 | 0.21976 | 0.18819 | 0.15732 | 0.36989 | 0.38341 | 0.17512 | 0.19294 | 0.41952 | 0.26913 | 0.53813 | 0 | 0.16434 | 0.24495 | 0.16253 | 0.26771 | 0.35935
| Rel.Std.Dev. |  | 1.085% | 1.883% | 1.441% | 1.266% | 3.13% | 3.182% | 1.476% | 1.606% | 3.746% | 2.22% | 4.125% | 0% | 1.408% | 2.168% | 1.47% | 2.285% | 3.388%
| PDM3 |  | -10.84% | -1.372% | 10.35% | 4.994% | -0.1323% | 1.825% | 0.2762% | 1.544% | -5.357% | 2.459% | 10.24% | 9.853% | -1.401% | -4.512% | -6.554% | -0.9914% | -10.38%
| Z-Score (Absolute) |  | 1.65 | 0.19 | 1.61 | 0.79 | 0 | 0.3 | 0.06 | 0.26 | 0.8 | 0.4 | 1.6 | 1.54 | 0.2 | 0.67 | 0.99 | 0.13 | 1.58
| Analytical results for Gd in OREAS 782 (Certified Value 4.45 ppm) | [20 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 3.7 | 8.0598 | 4.25 | 4.4 | 4.5 | 4.4 | 4.6 | 4.7292 | 4.3136 | 4.3343 | [9 empty] | 1
| 1 | 2 | 3.36 | 8.0429 | 4.41 | 4.5 | 4.43 | 4.2 | 4.5 | 4.1524 | 4.4276 | 4.2644 | [9 empty] | 9
| 1 | 3 | 3.48 | 7.9678 | 4.31 | 4.3 | 4.41 | 4.4 | 4.7 | 4.7433 | 4.4813 | 4.2588 | [9 empty] | 16
| 1 | 4 | 3.88 | 7.9881 | 4.43 | 4.4 | 4.29 | 4.2 | 4.5 | 4.8073 | 4.7198 | 4.252 | [9 empty] | 4.4464
| 1 | 5 | 3.91 | 7.8155 | 4.23 | 4.3 | 4.47 | 4.4 | 4.7 | 4.7341 | 4.556 | 4.3645 | [9 empty] | 48
| 1 | 6 | 3.82 | 7.8589 | 4.64 | 4.5 | 4.31 | 4 | 4.6 | 4.7383 | 4.5374 | 4.1333
| Mean |  | 3.6917 | 7.9555 | 4.3783 | 4.4 | 4.4017 | 4.2667 | 4.6 | 4.6508 | 4.5059 | 4.2679
| Median |  | 3.76 | 7.9779 | 4.36 | 4.4 | 4.42 | 4.3 | 4.6 | 4.7362 | 4.5093 | 4.2616
| Std Dev. |  | 0.2256 | 0.098669 | 0.15184 | 0.089443 | 0.084951 | 0.1633 | 0.089443 | 0.24583 | 0.13634 | 0.080257
| Rel.Std.Dev. |  | 6.111% | 1.24% | 3.468% | 2.033% | 1.93% | 3.827% | 1.944% | 5.286% | 3.026% | 1.88%
| PDM3 |  | -16.97% | 78.92% | -1.53% | -1.043% | -1.005% | -4.042% | 3.455% | 4.597% | 1.34% | -4.014%
| Z-Score (Absolute) |  | 3.58 | 17.95 | 0.11 | 0 | 0 | 0.68 | 1.01 | 1.26 | 0.53 | 0.67
| Analytical results for Ge in OREAS 782 (Indicative Value 0.19 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 0.36 | 0.12 | 0.12 | 0.2 | 0.3 | 1.9 | 1.6 | < 0.2 | 0.28364 | 3.3 | 0.14 | 0.11 | 0.31 | [6 empty] | 1
| 1 | 2 | 0.24 | 0.12 | 0.13 | 0.19 | 0.3 | 1.9 | 2 | < 0.2 | 0.29836 | 3.2 | 0.11 | 0.15 | 0.23 | [6 empty] | 10
| 1 | 3 | 0.34 | 0.14 | 0.14 | 0.21 | 0.2 | 1.9 | 1.9 | 0.2 | 0.26468 | 3.2 | 0.12 | 0.12 | 0.26 | [6 empty] | 16
| 1 | 4 | 0.18 | 0.11 | 0.15 | 0.19 | 0.2 | 1.9 | 1.9 | < 0.2 | 0.33497 | 3.3 | 0.06 | 0.12 | 0.23 | [6 empty] | 0.19356
| 1 | 5 | 0.13 | 0.11 | 0.13 | 0.21 | 0.2 | 1.9 | 1.9 | < 0.2 | 0.31419 | 3.3 | 0.11 | 0.15 | 0.25 | [6 empty] | 22
| 1 | 6 | 0.11 | 0.12 | 0.12 | 0.21 | 0.3 | 1.9 | 2 | < 0.2 | 0.26768 | 3.3 | 0.12 | 0.15 | 0.27
| Mean |  | 0.22667 | 0.12 | 0.13167 | 0.20167 | 0.25 | 1.9 | 1.8833 | 0.2 | 0.29392 | 3.2667 | 0.11 | 0.13333 | 0.25833
| Median |  | 0.21 | 0.12 | 0.13 | 0.205 | 0.25 | 1.9 | 1.9 | 0.2 | 0.291 | 3.3 | 0.115 | 0.135 | 0.255
| Std Dev. |  | 0.10577 | 0.010954 | 0.01169 | 0.0098319 | 0.054772 | 2.4324e-16 | 0.1472 |  | 0.027439 | 0.05164 | 0.026833 | 0.018619 | 0.029944
| Rel.Std.Dev. |  | 46.66% | 9.129% | 8.879% | 4.875% | 21.91% | 1.28e-14% | 7.816% |  | 9.335% | 1.581% | 24.39% | 13.96% | 11.59%
| PDM3 |  | 17.1% | -38% | -31.98% | 4.189% | 29.16% | 881.6% | 873% | 3.328% | 51.85% | 1588% | -43.17% | -31.11% | 33.47%
| Z-Score (Absolute) |  | 0 | 0.76 | 0.67 | 0.18 | 0.17 | 11.88 | 11.76 | 0.78 | 0.48 | 21.58 | 0.83 | 0.66 | 0.22
| Analytical results for Hf in OREAS 782 (Certified Value 2.53 ppm) | [20 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.14 | 2.7 | 2.6 | 2.16 | 2.7 | 2.6 | 2.54 | 2.6 | 2.61 | 2.6319 | 2.1 | 3.6465 | 2.5 | 2.4 | 2.3 |  |  |  |  | 1
| 1 | 2 | 1.89 | 2.9 | 2.6 | 2.08 | 2.6 | 2.6 | 2.57 | 2.4 | 2.59 | 2.3108 | 2.3 | 3.6849 | 2.5 | 2.4 | 2.3 |  |  |  |  | #N/A
| 1 | 3 | 2.03 | 2.8 | 2.5 | 2.26 | 2.5 | 2.6 | 2.51 | 2.6 | 2.64 | 2.5601 | 2.2 | 3.757 | 2.6 | 2.5 | 2.2 |  |  |  |  | 16
| 1 | 4 | 2.19 | 2.9 | 2.4 | 2.18 | 2.5 | 2.6 | 2.51 | 2.6 | 2.53 | 2.6365 | 2.2 | 3.7689 | 2.4 | 2.6 | 2.4 |  |  |  |  | 2.5261
| 1 | 5 | 2.21 | 2.9 | 2.5 | 2.2 | 2.7 | 2.6 | 2.6 | 2.6 | 2.64 | 2.5435 | 2.3 | 3.7792 | 2.4 | 2.3 | 2.4 |  |  |  |  | 49
| 1 | 6 | 2.14 | 2.9 | 2.4 | 2.32 | 2.7 | 2.5 | 2.68 | 2.4 | 2.7 | 2.5109 | 2.2 | 3.7648 | 2.4 | 2.5 | 2.4
| Mean |  | 2.1 | 2.85 | 2.5 | 2.2 | 2.6167 | 2.5833 | 2.5683 | 2.5333 | 2.6183 | 2.5323 | 2.2167 | 3.7336 | 2.4667 | 2.45 | 2.3333
| Median |  | 2.14 | 2.9 | 2.5 | 2.19 | 2.65 | 2.6 | 2.555 | 2.6 | 2.625 | 2.5518 | 2.2 | 3.7609 | 2.45 | 2.45 | 2.35
| Std Dev. |  | 0.12033 | 0.083666 | 0.089443 | 0.082946 | 0.098319 | 0.040825 | 0.064936 | 0.10328 | 0.057067 | 0.11935 | 0.075277 | 0.054443 | 0.08165 | 0.10488 | 0.08165
| Rel.Std.Dev. |  | 5.73% | 2.936% | 3.578% | 3.77% | 3.757% | 1.58% | 2.528% | 4.077% | 2.18% | 4.713% | 3.396% | 1.458% | 3.31% | 4.281% | 3.499%
| PDM3 |  | -16.87% | 12.82% | -1.033% | -12.91% | 3.585% | 2.265% | 1.672% | 0.2861% | 3.651% | 0.2452% | -12.25% | 47.8% | -2.353% | -3.013% | -7.631%
| Z-Score (Absolute) |  | 3.46 | 2.54 | 0.26 | 2.66 | 0.67 | 0.41 | 0.29 | 0.01 | 0.69 | 0 | 2.52 | 9.6 | 0.52 | 0.66 | 1.59
| Analytical results for Hg in OREAS 782 (Indicative Value 0.12 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 14 | [17 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [17 empty] | 2
2
| 1 | 1 | 0.14 | < 1 | [17 empty] | 1
| 1 | 2 | 0.11 | < 1 | [17 empty] | 12
| 1 | 3 | 0.12 | < 1 | [17 empty] | 16
| 1 | 4 | 0.11 | < 1 | [17 empty] | 0.12333
| 1 | 5 | 0.13 | < 1 | [17 empty] | 23
| 1 | 6 | 0.13 | < 1
| Mean |  | 0.12333
| Median |  | 0.125
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 9.819%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 782 (Certified Value 0.67 ppm) | [20 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 0.42 | 0.70295 | 0.57 | 0.7 | 0.68 | 0.68 | 0.66 | 0.73941 | 0.82898 | 0.665 | [9 empty] | 1
| 1 | 2 | 0.38 | 0.70606 | 0.59 | 0.7 | 0.69 | 0.68 | 0.65 | 0.64207 | 0.83928 | 0.6811 | [9 empty] | 20
| 1 | 3 | 0.4 | 0.68401 | 0.57 | 0.7 | 0.74 | 0.68 | 0.67 | 0.73053 | 0.86871 | 0.6331 | [9 empty] | 16
| 1 | 4 | 0.44 | 0.7112 | 0.59 | 0.7 | 0.68 | 0.66 | 0.65 | 0.72212 | 0.91572 | 0.6755 | [9 empty] | 0.67187
| 1 | 5 | 0.45 | 0.72105 | 0.58 | 0.7 | 0.7 | 0.66 | 0.67 | 0.73385 | 0.88853 | 0.6557 | [9 empty] | 50
| 1 | 6 | 0.44 | 0.73576 | 0.61 | 0.7 | 0.69 | 0.66 | 0.67 | 0.72289 | 0.90189 | 0.6762
| Mean |  | 0.42167 | 0.71017 | 0.585 | 0.7 | 0.69667 | 0.67 | 0.66167 | 0.71514 | 0.87385 | 0.66443
| Median |  | 0.43 | 0.70863 | 0.585 | 0.7 | 0.69 | 0.67 | 0.665 | 0.72671 | 0.87862 | 0.67025
| Std Dev. |  | 0.027142 | 0.017481 | 0.015166 | 1.2162e-16 | 0.022509 | 0.010954 | 0.0098319 | 0.036399 | 0.034611 | 0.017872
| Rel.Std.Dev. |  | 6.437% | 2.461% | 2.592% | 1.737e-14% | 3.231% | 1.635% | 1.486% | 5.09% | 3.961% | 2.69%
| PDM3 |  | -37.24% | 5.701% | -12.93% | 4.187% | 3.691% | -0.2782% | -1.519% | 6.441% | 30.06% | -1.107%
| Z-Score (Absolute) |  | 4.17 | 0.67 | 1.43 | NA | 0.45 | 0 | 0.14 | 0.76 | 3.42 | 0.09
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 782 (Certified Value 0.062 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.096 | 0.064 | 0.068 | 0.08 | < 0.1 | 0.06 | 0.05 | 0.05 | 0.063 | 0.087032 | 0.08 | 0.05868 | 0.062 | 0.056 | 0.065 |  |  |  |  | 1
| 1 | 2 | 0.086 | 0.063 | 0.068 | 0.07 | < 0.1 | 0.05 | 0.07 | 0.05 | 0.065 | 0.080471 | 0.09 | 0.0571 | 0.059 | 0.063 | 0.062 |  |  |  |  | 1
| 1 | 3 | 0.089 | 0.065 | 0.062 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.05 | 0.064 | 0.079982 | 0.09 | 0.05768 | 0.059 | 0.061 | 0.062 |  |  |  |  | 16
| 1 | 4 | 0.087 | 0.064 | 0.067 | 0.08 | < 0.1 | 0.05 | 0.07 | 0.05 | 0.068 | 0.095995 | 0.09 | 0.05772 | 0.075 | 0.058 | 0.064 |  |  |  |  | 0.062192
| 1 | 5 | 0.098 | 0.068 | 0.059 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.05 | 0.071 | 0.086503 | 0.09 | 0.05666 | 0.06 | 0.06 | 0.062 |  |  |  |  | 51
| 1 | 6 | 0.088 | 0.065 | 0.056 | 0.08 | < 0.1 | 0.06 | 0.07 | 0.05 | 0.069 | 0.090199 | 0.09 | 0.056 | 0.059 | 0.067 | 0.063
| Mean |  | 0.090667 | 0.064833 | 0.063333 | 0.075 |  | 0.056667 | 0.066667 | 0.05 | 0.066667 | 0.086697 | 0.088333 | 0.057307 | 0.062333 | 0.060833 | 0.063
| Median |  | 0.0885 | 0.0645 | 0.0645 | 0.075 |  | 0.06 | 0.07 | 0.05 | 0.0665 | 0.086767 | 0.09 | 0.05739 | 0.0595 | 0.0605 | 0.0625
| Std Dev. |  | 0.0050465 | 0.0017224 | 0.0051251 | 0.0054772 |  | 0.005164 | 0.008165 | 7.6012e-18 | 0.0031411 | 0.0060452 | 0.0040825 | 0.00093429 | 0.006314 | 0.0038687 | 0.0012649
| Rel.Std.Dev. |  | 5.566% | 2.657% | 8.092% | 7.303% |  | 9.113% | 12.25% | 1.52e-14% | 4.712% | 6.973% | 4.622% | 1.63% | 10.13% | 6.359% | 2.008%
| PDM3 |  | 45.79% | 4.248% | 1.836% | 20.6% |  | -8.884% | 7.196% | -19.6% | 7.196% | 39.4% | 42.03% | -7.855% | 0.228% | -2.184% | 1.3%
| Z-Score (Absolute) |  | 3.06 | 0.17 | 0 | 1.31 | 1.49 | 0.75 | 0.37 | 1.49 | 0.37 | 2.61 | 2.8 | 0.67 | 0.11 | 0.28 | 0.04
| Analytical results for K in OREAS 782 (Certified Value 1.47 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.4201 | 1.5063 | 1.49 | 1.44 | 1.51 | 1.42 | 1.48 | 1.49 | 1.5305 | 1.5168 | 1.48 | 1.47 | 1.5381 | 1.43 | 1.45 | 1.48 | 1.45 | 1.4523 |  | 1
| 1 | 2 | 1.412 | 1.5029 | 1.49 | 1.45 | 1.51 | 1.43 | 1.48 | 1.53 | 1.5308 | 1.4867 | 1.49 | 1.44 | 1.5028 | 1.47 | 1.41 | 1.45 | 1.42 | 1.455 |  | #N/A
| 1 | 3 | 1.4022 | 1.4999 | 1.51 | 1.46 | 1.5 | 1.42 | 1.47 | 1.5 | 1.5086 | 1.4942 | 1.49 | 1.47 | 1.5117 | 1.48 | 1.46 | 1.49 | 1.43 | 1.4672 |  | 16
| 1 | 4 | 1.4269 | 1.504 | 1.52 | 1.44 | 1.55 | 1.39 | 1.48 | 1.51 | 1.4948 | 1.446 | 1.48 | 1.43 | 1.5165 | 1.47 | 1.46 | 1.46 | 1.49 | 1.4599 |  | 1.4728
| 1 | 5 | 1.405 | 1.4992 | 1.51 | 1.45 | 1.55 | 1.36 | 1.45 | 1.5 | 1.517 | 1.5074 | 1.5 | 1.48 | 1.524 | 1.45 | 1.45 | 1.41 | 1.44 | 1.4653 |  | 52
| 1 | 6 | 1.4236 | 1.5025 | 1.51 | 1.44 | 1.51 | 1.41 | 1.46 | 1.52 | 1.4978 | 1.512 | 1.45 | 1.49 | 1.4949 | 1.5 | 1.43 | 1.45 | 1.43 | 1.4566
| Mean |  | 1.415 | 1.5025 | 1.505 | 1.4467 | 1.5217 | 1.405 | 1.47 | 1.5083 | 1.5132 | 1.4938 | 1.4817 | 1.4633 | 1.5147 | 1.4667 | 1.4433 | 1.4567 | 1.4433 | 1.4594
| Median |  | 1.4161 | 1.5027 | 1.51 | 1.445 | 1.51 | 1.415 | 1.475 | 1.505 | 1.5128 | 1.5008 | 1.485 | 1.47 | 1.5141 | 1.47 | 1.45 | 1.455 | 1.435 | 1.4583
| Std Dev. |  | 0.010143 | 0.0026265 | 0.012247 | 0.008165 | 0.022286 | 0.025884 | 0.012649 | 0.01472 | 0.015623 | 0.026003 | 0.017224 | 0.023381 | 0.015362 | 0.024221 | 0.019664 | 0.028048 | 0.025033 | 0.0058908
| Rel.Std.Dev. |  | 0.7168% | 0.1748% | 0.8138% | 0.5644% | 1.465% | 1.842% | 0.8605% | 0.9759% | 1.032% | 1.741% | 1.162% | 1.598% | 1.014% | 1.651% | 1.362% | 1.925% | 1.734% | 0.4037%
| PDM3 |  | -3.926% | 2.015% | 2.187% | -1.774% | 3.318% | -4.603% | -0.1896% | 2.413% | 2.747% | 1.43% | 0.6025% | -0.6423% | 2.845% | -0.416% | -2% | -1.095% | -2% | -0.9105%
| Z-Score (Absolute) |  | 1.42 | 0.91 | 0.98 | 0.58 | 1.42 | 1.69 | 0.04 | 1.07 | 1.2 | 0.68 | 0.36 | 0.13 | 1.24 | 0.04 | 0.67 | 0.31 | 0.67 | 0.24
| Analytical results for La in OREAS 782 (Certified Value 32.3 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 25.6 | 32.0629 | 34 | 31.4 | 25.77 | 33.4 | 32.5 | 32.81 | 33.272 | 31.8 | 32.1 | 30.7311 | 33 | 28.6828 | 33.1 | 33 | 30.1 | 31.4679 |  | 1
| 1 | 2 | 26.2 | 32.1682 | 35.8 | 31 | 25.95 | 33.2 | 31.5 | 32.8 | 33.264 | 32.9 | 31.3 | 27.1719 | 34 | 29.7274 | 33.1 | 33.4 | 30.9 | 31.0852 |  | 21
| 1 | 3 | 25.9 | 31.6601 | 36.3 | 31.8 | 25.35 | 33.3 | 32.3 | 32.86 | 32.7876 | 32.1 | 31.9 | 30.7375 | 34 | 29.9855 | 32.2 | 34.3 | 29.9 | 31.1379 |  | 16
| 1 | 4 | 26.3 | 32.1609 | 34.2 | 31.2 | 25.36 | 33.5 | 34.1 | 32.95 | 32.6601 | 32.6 | 31.2 | 30.9694 | 33 | 30.7405 | 34.6 | 33.1 | 32.8 | 29.446 |  | 32.3096
| 1 | 5 | 26 | 31.9162 | 35.5 | 31.5 | 25.53 | 33.2 | 31.1 | 33.22 | 33.6014 | 32.8 | 32.2 | 30.7432 | 33 | 29.591 | 34.4 | 32.9 | 30.7 | 34.1613 |  | 53
| 1 | 6 | 26.5 | 31.8647 | 35.5 | 31.2 | 25.3 | 33.4 | 31.7 | 32.19 | 33.8351 | 31.8 | 32.3 | 30.7088 | 34 | 30.3262 | 33.2 | 33.7 | 31 | 30.9625
| Mean |  | 26.0833 | 31.9722 | 35.2167 | 31.35 | 25.5433 | 33.3333 | 32.2 | 32.805 | 33.2367 | 32.3333 | 31.8333 | 30.177 | 33.5 | 29.8422 | 33.4333 | 33.4 | 30.9 | 31.3768
| Median |  | 26.1 | 31.9896 | 35.5 | 31.3 | 25.445 | 33.35 | 32 | 32.835 | 33.268 | 32.35 | 32 | 30.7343 | 33.5 | 29.8564 | 33.15 | 33.25 | 30.8 | 31.1115
| Std Dev. |  | 0.31885 | 0.19713 | 0.91524 | 0.28107 | 0.26349 | 0.12111 | 1.064 | 0.33898 | 0.45328 | 0.49666 | 0.47188 | 1.4753 | 0.54772 | 0.70458 | 0.9048 | 0.52915 | 1.0296 | 1.5365
| Rel.Std.Dev. |  | 1.222% | 0.6166% | 2.599% | 0.8966% | 1.032% | 0.3633% | 3.304% | 1.033% | 1.364% | 1.536% | 1.482% | 4.889% | 1.635% | 2.361% | 2.706% | 1.584% | 3.332% | 4.897%
| PDM3 |  | -19.27% | -1.044% | 8.997% | -2.97% | -20.94% | 3.168% | -0.3393% | 1.533% | 2.869% | 0.07338% | -1.474% | -6.601% | 3.684% | -7.637% | 3.478% | 3.375% | -4.363% | -2.887%
| Z-Score (Absolute) |  | 3.33 | 0.06 | 1.74 | 0.41 | 3.63 | 0.69 | 0.06 | 0.4 | 0.64 | 0.14 | 0.14 | 1.06 | 0.78 | 1.24 | 0.75 | 0.73 | 0.66 | 0.39
| Analytical results for Li in OREAS 782 (Certified Value 39 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 37 | 27.1516 | 43.1 | 39.5 | 39.4 | 39 | 41 | 37.5 | 37.4 | 39.2676 | 39 | 38.8 | 35.45 | 40.6 | 37.6853 | 40.6 | 42 | 41.1 | 51.3356 | 1
| 1 | 2 | 37 | 27.5634 | 42.9 | 39.7 | 39.02 | 40 | 40 | 37.7 | 37.6 | 39.2577 | 37.5 | 36.3 | 35.12 | 40.5 | 37.8028 | 39.7 | 41 | 40 | 54.3988 | 22
| 1 | 3 | 36 | 27.55 | 43.7 | 39.7 | 39.28 | 39 | 40 | 37.2 | 37.7 | 39.2149 | 39.5 | 37.7 | 35.82 | 40 | 36.4729 | 40.6 | 42.3 | 40.5 | 53.2079 | 16
| 1 | 4 | 37 | 27.3487 | 43.4 | 39.9 | 38.41 | 37 | 40 | 37.6 | 37.1 | 38.5402 | 37.5 | 37.3 | 37.04 | 39.2 | 37.0506 | 40.5 | 41.7 | 42.5 | 54.2017 | 39.0289
| 1 | 5 | 37 | 27.6148 | 43.6 | 39.7 | 39.86 | 38 | 38 | 36.5 | 37.9 | 38.68 | 39 | 38.3 | 38.35 | 40.9 | 37.574 | 40.4 | 40.8 | 40.3 | 51.1813 | 54
| 1 | 6 | 37 | 27.8514 | 44 | 39.3 | 39.22 | 39 | 39 | 37.1 | 37.2 | 40.0678 | 37 | 37.5 | 37.16 | 38.3 | 36.8028 | 39.8 | 41.4 | 40.3 | 54.1463
| Mean |  | 36.8333 | 27.5133 | 43.45 | 39.6333 | 39.1983 | 38.6667 | 39.6667 | 37.2667 | 37.4833 | 39.1714 | 38.25 | 37.65 | 36.49 | 39.9167 | 37.2314 | 40.2667 | 41.5333 | 40.7833 | 53.0786
| Median |  | 37 | 27.5567 | 43.5 | 39.7 | 39.25 | 39 | 40 | 37.35 | 37.5 | 39.2363 | 38.25 | 37.6 | 36.43 | 40.25 | 37.3123 | 40.45 | 41.55 | 40.4 | 53.6771
| Std Dev. |  | 0.40825 | 0.23934 | 0.40373 | 0.20656 | 0.47726 | 1.0328 | 1.0328 | 0.44121 | 0.30605 | 0.54075 | 1.0368 | 0.86197 | 1.2344 | 0.99079 | 0.53696 | 0.40825 | 0.5785 | 0.91742 | 1.4696
| Rel.Std.Dev. |  | 1.108% | 0.8699% | 0.9292% | 0.5212% | 1.218% | 2.671% | 2.604% | 1.184% | 0.8165% | 1.38% | 2.711% | 2.289% | 3.383% | 2.482% | 1.442% | 1.014% | 1.393% | 2.25% | 2.769%
| PDM3 |  | -5.625% | -29.51% | 11.33% | 1.549% | 0.4342% | -0.9281% | 1.634% | -4.515% | -3.96% | 0.3651% | -1.996% | -3.533% | -6.505% | 2.275% | -4.606% | 3.171% | 6.417% | 4.495% | 36%
| Z-Score (Absolute) |  | 0.98 | 4.88 | 1.79 | 0.19 | 0.01 | 0.21 | 0.21 | 0.8 | 0.71 | 0 | 0.39 | 0.64 | 1.12 | 0.31 | 0.81 | 0.46 | 0.99 | 0.67 | 5.82
| Analytical results for Lu in OREAS 782 (Certified Value 0.26 ppm) | [20 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [10 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.21 | 0.27 | 0.3 | 0.36 | 0.26 | 0.27 | 0.29937 | 0.32615 | 0.1876 | [10 empty] | 1
| 1 | 2 | 0.19 | 0.27 | 0.3 | 0.29 | 0.28 | 0.27 | 0.25592 | 0.3291 | 0.1851 | [10 empty] | 23
| 1 | 3 | 0.21 | 0.26 | 0.3 | 0.28 | 0.28 | 0.27 | 0.28875 | 0.35333 | 0.1868 | [10 empty] | 16
| 1 | 4 | 0.23 | 0.26 | 0.3 | 0.26 | 0.28 | 0.26 | 0.28502 | 0.35208 | 0.1902 | [10 empty] | 0.26335
| 1 | 5 | 0.23 | 0.25 | 0.3 | 0.25 | 0.3 | 0.27 | 0.28706 | 0.34433 | 0.1912 | [10 empty] | 55
| 1 | 6 | 0.22 | 0.28 | 0.3 | 0.28 | 0.26 | 0.27 | 0.28255 | 0.34738 | 0.1815
| Mean |  | 0.215 | 0.265 | 0.3 | 0.28667 | 0.27667 | 0.26833 | 0.28311 | 0.34206 | 0.18707
| Median |  | 0.215 | 0.265 | 0.3 | 0.28 | 0.28 | 0.27 | 0.28604 | 0.34586 | 0.1872
| Std Dev. |  | 0.015166 | 0.010488 | 0 | 0.038816 | 0.015055 | 0.0040825 | 0.014526 | 0.011679 | 0.0035246
| Rel.Std.Dev. |  | 7.054% | 3.958% | 0% | 13.54% | 5.442% | 1.521% | 5.131% | 3.414% | 1.884%
| PDM3 |  | -18.36% | 0.6259% | 13.92% | 8.853% | 5.056% | 1.892% | 7.503% | 29.89% | -28.97%
| Z-Score (Absolute) |  | 3.13 | 0.41 | NA | 0.77 | 0.23 | 0.23 | 0.58 | 3.79 | 4.65
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 782 (Certified Value 0.195 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.1402 | 0.22376 | 0.2 | 0.19 | 0.19 | 0.16 | 0.19 | 0.199 | 0.1998 | 0.20187 | 0.17 | 0.189 | 0.22719 | 0.15 | 0.19 | 0.2 | 0.2 | 0.1654 |  | 1
| 1 | 2 | 0.1415 | 0.21427 | 0.2 | 0.19 | 0.2 | 0.16 | 0.19 | 0.198 | 0.2047 | 0.19316 | 0.17 | 0.181 | 0.22298 | 0.16 | 0.18 | 0.2 | 0.19 | 0.1618 |  | 24
| 1 | 3 | 0.1396 | 0.22537 | 0.21 | 0.19 | 0.19 | 0.16 | 0.18 | 0.206 | 0.2024 | 0.19403 | 0.17 | 0.189 | 0.22438 | 0.15 | 0.19 | 0.2 | 0.2 | 0.1582 |  | 16
| 1 | 4 | 0.1422 | 0.22521 | 0.2 | 0.2 | 0.2 | 0.15 | 0.18 | 0.203 | 0.2006 | 0.18948 | 0.17 | 0.182 | 0.22732 | 0.16 | 0.19 | 0.2 | 0.2 | 0.1724 |  | 0.1948
| 1 | 5 | 0.1419 | 0.22402 | 0.21 | 0.2 | 0.2 | 0.15 | 0.19 | 0.19 | 0.2045 | 0.19271 | 0.17 | 0.188 | 0.22702 | 0.16 | 0.18 | 0.2 | 0.2 | 0.1578 |  | 56
| 1 | 6 | 0.143 | 0.22246 | 0.21 | 0.19 | 0.19 | 0.15 | 0.19 | 0.2 | 0.2013 | 0.19633 | 0.17 | 0.182 | 0.22352 | 0.16 | 0.18 | 0.2 | 0.2 | 0.1722
| Mean |  | 0.1414 | 0.22252 | 0.205 | 0.19333 | 0.195 | 0.155 | 0.18667 | 0.19933 | 0.20222 | 0.1946 | 0.17 | 0.18517 | 0.2254 | 0.15667 | 0.185 | 0.2 | 0.19833 | 0.16463
| Median |  | 0.1417 | 0.22389 | 0.205 | 0.19 | 0.195 | 0.155 | 0.19 | 0.1995 | 0.20185 | 0.19359 | 0.17 | 0.185 | 0.2257 | 0.16 | 0.185 | 0.2 | 0.2 | 0.1636
| Std Dev. |  | 0.0012759 | 0.0041761 | 0.0054772 | 0.005164 | 0.0054772 | 0.0054772 | 0.005164 | 0.0054283 | 0.0020351 | 0.004196 | 3.0405e-17 | 0.0038687 | 0.0019993 | 0.005164 | 0.0054772 | 3.0405e-17 | 0.0040825 | 0.0065463
| Rel.Std.Dev. |  | 0.9024% | 1.877% | 2.672% | 2.671% | 2.809% | 3.534% | 2.766% | 2.723% | 1.006% | 2.156% | 1.789e-14% | 2.089% | 0.887% | 3.296% | 2.961% | 1.52e-14% | 2.058% | 3.976%
| PDM3 |  | -27.41% | 14.23% | 5.236% | -0.7529% | 0.1027% | -20.43% | -4.175% | 2.327% | 3.807% | -0.104% | -12.73% | -4.945% | 15.71% | -19.58% | -5.031% | 2.669% | 1.814% | -15.49%
| Z-Score (Absolute) |  | 3.94 | 2.22 | 0.89 | 0 | 0.13 | 2.91 | 0.51 | 0.46 | 0.67 | 0.1 | 1.77 | 0.62 | 2.43 | 2.78 | 0.63 | NA | 0.38 | 2.18
Results from laboratory 22 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Mn in OREAS 782 (Certified Value 0.005 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0039 | 0.0051279 | 0.0055 | 0.0056 | 0.005035 | 0.004989 | 0.0053 | 0.0051 | 0.0054 | 0.0053676 | 0.0056 | 0.0054 | 0.005646 | 0.0048 | 0.0052 | 0.0053 | 0.0053 | < 0.1 |  | 1
| 1 | 2 | 0.0043 | 0.0051203 | 0.0055 | 0.0053 | 0.004932 | 0.004984 | 0.0051 | 0.005 | 0.0065 | 0.00524 | 0.0052 | 0.00537 | 0.005574 | 0.0045 | 0.0049 | 0.0051 | 0.0053 | < 0.1 |  | 3
| 1 | 3 | 0.0037 | 0.0051639 | 0.0056 | 0.0054 | 0.005003 | 0.004983 | 0.0052 | 0.0054 | 0.0056 | 0.0051864 | 0.0054 | 0.00516 | 0.005747 | 0.0044 | 0.005 | 0.0053 | 0.0053 | < 0.1 |  | 16
| 1 | 4 | 0.0037 | 0.0051094 | 0.0055 | 0.0056 | 0.005056 | 0.005 | 0.0051 | 0.0053 | 0.0052 | 0.0050145 | 0.0052 | 0.00538 | 0.005597 | 0.0045 | 0.0052 | 0.0053 | 0.0054 | < 0.1 |  | 0.0052702
| 1 | 5 | 0.0037 | 0.0051617 | 0.0056 | 0.0055 | 0.00507 | 0.005 | 0.0052 | 0.0052 | 0.0053 | 0.0051185 | 0.0054 | 0.00524 | 0.005606 | 0.0048 | 0.005 | 0.0052 | 0.0056 | < 0.1 |  | 57
| 1 | 6 | 0.0036 | 0.0051522 | 0.0057 | 0.0055 | 0.004954 | 0.0051 | 0.0052 | 0.0052 | 0.0052 | 0.0053686 | 0.0054 | 0.00539 | 0.005612 | 0.0045 | 0.005 | 0.0051 | 0.0053 | < 0.1
| Mean |  | 0.0038167 | 0.0051393 | 0.0055667 | 0.0054833 | 0.0050083 | 0.0050093 | 0.0051833 | 0.0052 | 0.0055333 | 0.0052159 | 0.0053667 | 0.0053233 | 0.0056303 | 0.0045833 | 0.00505 | 0.0052167 | 0.0053667
| Median |  | 0.0037 | 0.0051401 | 0.00555 | 0.0055 | 0.005019 | 0.0049945 | 0.0052 | 0.0052 | 0.00535 | 0.0052132 | 0.0054 | 0.005375 | 0.005609 | 0.0045 | 0.005 | 0.00525 | 0.0053
| Std Dev. |  | 0.00025626 | 2.3057e-05 | 8.165e-05 | 0.0001169 | 5.5845e-05 | 4.5041e-05 | 7.5277e-05 | 0.00014142 | 0.00049666 | 0.00013988 | 0.00015055 | 9.9331e-05 | 6.1757e-05 | 0.00017224 | 0.00012247 | 9.8319e-05 | 0.00012111
| Rel.Std.Dev. |  | 6.714% | 0.4487% | 1.467% | 2.132% | 1.115% | 0.8991% | 1.452% | 2.72% | 8.976% | 2.682% | 2.805% | 1.866% | 1.097% | 3.758% | 2.425% | 1.885% | 2.257%
| PDM3 |  | -27.58% | -2.485% | 5.625% | 4.044% | -4.969% | -4.95% | -1.649% | -1.332% | 4.993% | -1.03% | 1.83% | 1.008% | 6.833% | -13.03% | -4.178% | -1.016% | 1.83%
| Z-Score (Absolute) |  | 5.06 | 0.28 | 1.27 | 0.96 | 0.75 | 0.75 | 0.12 | 0.06 | 1.15 | 0 | 0.54 | 0.39 | 1.5 | 2.29 | 0.6 | 0 | 0.54 | 161.8
| Analytical results for Mo in OREAS 782 (Certified Value 3.68 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.6 | 2.7988 | 3.95 | 3.55 | 1.9 | 4.15 | 4 | 3.4 | 3.7 | < 50 | 3.9 | 4.1 | 4.2698 | 3.9 | 3.3769 | 3.64 | 3.46 | 3.37 | 29731 | 1
| 1 | 2 | 4.5 | 2.8619 | 3.79 | 3.65 | 1.99 | 3.51 | 4 | 3.4 | 4 | < 50 | 3.8 | 4.2 | 3.4918 | 3.9 | 3.4448 | 3.6 | 3.46 | 3.59 | 29323 | 4
| 1 | 3 | 4.2 | 2.6553 | 3.86 | 3.61 | 2 | 3.51 | 4 | 3.6 | 3.7 | < 50 | 3.6 | 4 | 3.9568 | 3.7 | 3.7639 | 3.4 | 3.73 | 3.38 | 29951 | 16
| 1 | 4 | 4 | 2.901 | 3.82 | 3.72 | 2.08 | 3.62 | 4 | 3.4 | 3.7 | < 50 | 3.7 | 3.9 | 4.0523 | 3.6 | 3.4849 | 3.49 | 3.42 | 3.62 | 30470 | 3.6836
| 1 | 5 | 4.2 | 2.6821 | 3.9 | 3.78 | 1.99 | 3.69 | 4 | 3.4 | 3.7 | < 50 | 3.9 | 3.9 | 3.668 | 3.8 | 3.7695 | 3.57 | 3.53 | 4.24 | 29782 | 58
| 1 | 6 | 4.1 | 2.7125 | 3.92 | 3.52 | 1.98 | 3.71 | 4 | 3.5 | 3.8 | < 50 | 3.7 | 4.2 | 3.6575 | 3.7 | 3.6273 | 3.5 | 3.49 | 3.5 | 29295
| Mean |  | 4.2667 | 2.7686 | 3.8733 | 3.6383 | 1.99 | 3.6983 | 4 | 3.45 | 3.7667 |  | 3.7667 | 4.05 | 3.8494 | 3.7667 | 3.5779 | 3.5333 | 3.515 | 3.6167 | 29758.6667
| Median |  | 4.2 | 2.7556 | 3.88 | 3.63 | 1.99 | 3.655 | 4 | 3.4 | 3.7 |  | 3.75 | 4.05 | 3.8124 | 3.75 | 3.5561 | 3.535 | 3.475 | 3.545 | 29756.5
| Std Dev. |  | 0.23381 | 0.10062 | 0.06121 | 0.099482 | 0.057271 | 0.23718 | 0 | 0.083666 | 0.12111 |  | 0.12111 | 0.13784 | 0.29227 | 0.12111 | 0.16766 | 0.087101 | 0.11149 | 0.32241 | 435.6809
| Rel.Std.Dev. |  | 5.48% | 3.634% | 1.58% | 2.734% | 2.878% | 6.413% | 0% | 2.425% | 3.215% |  | 3.215% | 3.403% | 7.593% | 3.215% | 4.686% | 2.465% | 3.172% | 8.914% | 1.464%
| PDM3 |  | 15.83% | -24.84% | 5.15% | -1.23% | -45.98% | 0.399% | 8.588% | -6.342% | 2.254% |  | 2.254% | 9.946% | 4.5% | 2.254% | -2.871% | -4.08% | -4.578% | -1.818% | 807762%
| Z-Score (Absolute) |  | 2.32 | 3.8 | 0.72 | 0.25 | 6.98 | 0 | NA | 1.01 | 0.28 | 87.05 | 0.28 | 1.44 | 0.62 | 0.28 | 0.49 | 0.67 | 0.75 | 0.33 | NA
Results from laboratory 1.24 were removed due to their systematic bias. Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 782 (Certified Value 0.1 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0679 | 0.08982 | 0.1 | 0.1 | 0.22 | 0.09 | 0.1 | 0.091 | 0.1026 | 0.11621 | 0.11 | 0.101 | 0.51113 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1247 |  | 1
| 1 | 2 | 0.0687 | 0.089868 | 0.11 | 0.1 | 0.22 | 0.09 | 0.09 | 0.092 | 0.1048 | 0.10577 | 0.1 | 0.097 | 0.50104 | 0.11 | 0.1 | 0.11 | 0.1 | 0.1252 |  | #N/A
| 1 | 3 | 0.0678 | 0.090388 | 0.11 | 0.1 | 0.23 | 0.09 | 0.09 | 0.097 | 0.1033 | 0.10661 | 0.1 | 0.1 | 0.50592 | 0.12 | 0.1 | 0.11 | 0.1 | 0.1271 |  | 16
| 1 | 4 | 0.0688 | 0.08976 | 0.11 | 0.1 | 0.22 | 0.09 | 0.09 | 0.094 | 0.1023 | 0.10249 | 0.1 | 0.097 | 0.50893 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1297 |  | 0.099765
| 1 | 5 | 0.0687 | 0.090272 | 0.11 | 0.1 | 0.22 | 0.09 | 0.09 | 0.092 | 0.1044 | 0.10715 | 0.11 | 0.099 | 0.49926 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1231 |  | 59
| 1 | 6 | 0.0692 | 0.090368 | 0.11 | 0.1 | 0.23 | 0.09 | 0.09 | 0.095 | 0.1023 | 0.10773 | 0.11 | 0.099 | 0.49851 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1309
| Mean |  | 0.068517 | 0.090079 | 0.10833 | 0.1 | 0.22333 | 0.09 | 0.091667 | 0.0935 | 0.10328 | 0.10766 | 0.105 | 0.098833 | 0.50413 | 0.105 | 0.1 | 0.11 | 0.1 | 0.12678
| Median |  | 0.0687 | 0.09007 | 0.11 | 0.1 | 0.22 | 0.09 | 0.09 | 0.093 | 0.10295 | 0.10688 | 0.105 | 0.099 | 0.50348 | 0.1 | 0.1 | 0.11 | 0.1 | 0.12615
| Std Dev. |  | 0.00054924 | 0.00029313 | 0.0040825 | 1.5202e-17 | 0.005164 | 1.5202e-17 | 0.0040825 | 0.0022583 | 0.0010907 | 0.0045781 | 0.0054772 | 0.0016021 | 0.0052933 | 0.0083666 | 1.5202e-17 | 0 | 1.5202e-17 | 0.0030321
| Rel.Std.Dev. |  | 0.8016% | 0.3254% | 3.768% | 1.52e-14% | 2.312% | 1.689e-14% | 4.454% | 2.415% | 1.056% | 4.252% | 5.216% | 1.621% | 1.05% | 7.968% | 1.52e-14% | 0% | 1.52e-14% | 2.392%
| PDM3 |  | -31.32% | -9.709% | 8.588% | 0.2354% | 123.9% | -9.788% | -8.118% | -6.28% | 3.526% | 7.915% | 5.247% | -0.934% | 405.3% | 5.247% | 0.2354% | 10.26% | 0.2354% | 27.08%
| Z-Score (Absolute) |  | 2.14 | 0.87 | 0.2 | NA | 6.94 | 0.88 | 0.78 | 0.67 | 0.1 | 0.16 | 0 | 0.36 | 23.4 | 0 | NA | 0.29 | NA | 1.28
Results from laboratories 4, 21 and 23 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb in OREAS 782 (Certified Value 7.83 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 3.93 | 8.8 | 8.3 | 6.45 | 7.8 | 5.7 | 7.48 | 6.8 | 8.6 | 8.2061 | 7.3 | 9.0309 | 8.1 | 7.8 | 7.8 | 5.2572 |  |  |  | 1
| 1 | 2 | 3.69 | 9 | 8 | 6.43 | 8.1 | 5.7 | 8.05 | 6.2 | 7.7 | 7.1654 | 7.6 | 9.1022 | 7.8 | 7.9 | 7.5 | 5.0021 |  |  |  | #N/A
| 1 | 3 | 3.74 | 9 | 7.8 | 6.39 | 6.9 | 5.8 | 8.33 | 6.7 | 8.2 | 7.8501 | 7.4 | 9.1104 | 7.9 | 8 | 7.5 | 5.0334 |  |  |  | 16
| 1 | 4 | 3.82 | 8.7 | 8.2 | 6.96 | 7.2 | 6.8 | 7.92 | 6.4 | 8 | 7.9196 | 7.2 | 9.0363 | 7.7 | 7.7 | 8.1 | 5.7251 |  |  |  | 7.8287
| 1 | 5 | 4.2 | 8.8 | 8.1 | 6.59 | 7.4 | 5.5 | 8.21 | 6.7 | 8.3 | 8.1497 | 7.4 | 9.039 | 7.9 | 7.8 | 8 | 4.9802 |  |  |  | 60
| 1 | 6 | 3.94 | 9.1 | 8 | 6.92 | 7.7 | 5.8 | 8.44 | 7.2 | 7.9 | 7.9752 | 7.7 | 8.8828 | 7.6 | 8 | 7.7 | 5.6167
| Mean |  | 3.8867 | 8.9 | 8.0667 | 6.6233 | 7.5167 | 5.8833 | 8.0717 | 6.6667 | 8.1167 | 7.8777 | 7.4333 | 9.0336 | 7.8333 | 7.8667 | 7.7667 | 5.2691
| Median |  | 3.875 | 8.9 | 8.05 | 6.52 | 7.55 | 5.75 | 8.13 | 6.7 | 8.1 | 7.9474 | 7.4 | 9.0376 | 7.85 | 7.85 | 7.75 | 5.1453
| Std Dev. |  | 0.18305 | 0.15492 | 0.17512 | 0.25469 | 0.43551 | 0.46224 | 0.34499 | 0.34448 | 0.31885 | 0.37449 | 0.18619 | 0.0817 | 0.17512 | 0.12111 | 0.25033 | 0.32839
| Rel.Std.Dev. |  | 4.71% | 1.741% | 2.171% | 3.845% | 5.794% | 7.857% | 4.274% | 5.167% | 3.928% | 4.754% | 2.505% | 0.9044% | 2.236% | 1.539% | 3.223% | 6.232%
| PDM3 |  | -50.35% | 13.68% | 3.04% | -15.4% | -3.986% | -24.85% | 3.104% | -14.84% | 3.678% | 0.6261% | -5.05% | 15.39% | 0.05931% | 0.4851% | -0.7923% | -32.69%
| Z-Score (Absolute) |  | 7.72 | 2.17 | 0.53 | 2.32 | 0.56 | 3.78 | 0.54 | 2.24 | 0.62 | 0.15 | 0.72 | 2.43 | 0.07 | 0.13 | 0.07 | 4.99
| Analytical results for Nd in OREAS 782 (Certified Value 26.7 ppm) | [20 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 24.04 | 26.6622 | 27.4 | 27.2 | 26.97 | 26.9 | 24.8 | 26.0469 | 23.8119 | 29.4797 | [9 empty] | 1
| 1 | 2 | 21.39 | 26.5681 | 27 | 27.5 | 26.77 | 27.8 | 24.5 | 24.4746 | 24.945 | 28.4694 | [9 empty] | 6
| 1 | 3 | 23.04 | 26.7626 | 25.8 | 26.4 | 26.37 | 27.5 | 25.4 | 26.321 | 25.3732 | 28.0761 | [9 empty] | 16
| 1 | 4 | 24.84 | 26.8554 | 26.8 | 27 | 26.92 | 27.4 | 24.3 | 26.5188 | 26.0295 | 27.5712 | [9 empty] | 26.6697
| 1 | 5 | 24.88 | 26.7475 | 26.4 | 26.4 | 26.58 | 27.9 | 25.1 | 26.0804 | 25.2506 | 29.4942 | [9 empty] | 61
| 1 | 6 | 24.46 | 26.9102 | 27.5 | 27.4 | 26.64 | 27.6 | 25.3 | 26.2056 | 25.3822 | 29.2314
| Mean |  | 23.775 | 26.751 | 26.8167 | 26.9833 | 26.7083 | 27.5167 | 24.9 | 25.9412 | 25.1321 | 28.7203
| Median |  | 24.25 | 26.7551 | 26.9 | 27.1 | 26.705 | 27.55 | 24.95 | 26.143 | 25.3119 | 28.8504
| Std Dev. |  | 1.3508 | 0.12468 | 0.64005 | 0.48339 | 0.2248 | 0.35449 | 0.44272 | 0.73883 | 0.73729 | 0.80439
| Rel.Std.Dev. |  | 5.682% | 0.4661% | 2.387% | 1.791% | 0.8417% | 1.288% | 1.778% | 2.848% | 2.934% | 2.801%
| PDM3 |  | -10.85% | 0.305% | 0.5512% | 1.176% | 0.145% | 3.176% | -6.635% | -2.731% | -5.765% | 7.689%
| Z-Score (Absolute) |  | 2.53 | 0.02 | 0.07 | 0.22 | 0.02 | 0.67 | 1.57 | 0.67 | 1.37 | 1.7
| Analytical results for Ni in OREAS 782 (Certified Value 11.3 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 11.4 | 10.2 | 11.8 | 10.9 | 10.84 | 11.21 | 12 | 11 | 14 | 12.0388 | 12 | 15 | 11.95 | 12.4 | 11.6735 | 11.2 | 11.4 | 11.3 | 8.3288 | 1
| 1 | 2 | 11.8 | 10.29 | 12 | 10.6 | 10.41 | 11.57 | 12 | 11 | 13 | 11.2474 | 12 | 15 | 17.5325 | 12.1 | 13.8064 | 11.2 | 11 | 11.2 | 8.4352 | 7
| 1 | 3 | 11.5 | 10.27 | 11.8 | 10.6 | 10.59 | 11.4 | 12 | 11 | 12 | 11.4421 | 12 | 14 | 10.5175 | 12.8 | 11.4994 | 10.9 | 11.9 | 10.9 | 8.4405 | 16
| 1 | 4 | 11.3 | 10.92 | 15.2 | 12.8 | 10.48 | 11.26 | 11 | 11 | 12 | 11.3527 | 12 | 15 | 10.2767 | 11.4 | 12.3917 | 15.6 | 11.9 | 11.7 | 8.1675 | 11.2944
| 1 | 5 | 11.1 | 10.2 | 13.8 | 10.4 | 10.31 | 11.38 | 11 | 11 | 12 | 11.5122 | 14 | 15 | 8.8677 | 12.3 | 11.7013 | 11.2 | 11.3 | 11.2 | 8.3054 | 62
| 1 | 6 | 11.3 | 10.41 | 12.9 | 10.6 | 10.45 | 11.53 | 12 | 11 | 12 | 11.8809 | 12 | 15 | 9.4174 | 12.1 | 12.2156 | 11 | 11.5 | 11.1 | 8.7842
| Mean |  | 11.4 | 10.3817 | 12.9167 | 10.9833 | 10.5133 | 11.3917 | 11.6667 | 11 | 12.5 | 11.579 | 12.3333 | 14.8333 | 11.427 | 12.1833 | 12.2146 | 11.85 | 11.5 | 11.2333 | 8.4103
| Median |  | 11.35 | 10.28 | 12.45 | 10.6 | 10.465 | 11.39 | 12 | 11 | 12 | 11.4771 | 12 | 15 | 10.3971 | 12.2 | 11.9585 | 11.2 | 11.45 | 11.2 | 8.382
| Std Dev. |  | 0.23664 | 0.27477 | 1.3659 | 0.90425 | 0.18425 | 0.14247 | 0.5164 | 0 | 0.83666 | 0.31205 | 0.8165 | 0.40825 | 3.1714 | 0.46224 | 0.85241 | 1.8415 | 0.35214 | 0.26583 | 0.20877
| Rel.Std.Dev. |  | 2.076% | 2.647% | 10.57% | 8.233% | 1.753% | 1.251% | 4.426% | 0% | 6.693% | 2.695% | 6.62% | 2.752% | 27.75% | 3.794% | 6.979% | 15.54% | 3.062% | 2.366% | 2.482%
| PDM3 |  | 0.9345% | -8.082% | 14.36% | -2.755% | -6.916% | 0.8608% | 3.296% | -2.607% | 10.67% | 2.519% | 9.198% | 31.33% | 1.173% | 7.87% | 8.147% | 4.919% | 1.82% | -0.5411% | -25.54%
| Z-Score (Absolute) |  | 0.02 | 1.61 | 2.34 | 0.67 | 1.4 | 0.03 | NA | NA | NA | 0.26 | NA | NA | 0.02 | 1.2 | 1.25 | 0.68 | 0.13 | 0.28 | 4.67
Results from laboratories 10, 12, 13, 16 and 17 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 782 (Certified Value 0.02 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0189 | 0.024064 | 0.021 | 0.021 | 0.02 | 0.0195 | 0.02 | 0.02 | 0.0199 | 0.020219 | 0.021 | 0.02 | 0.019729 | 0.02 | 0.02 | 0.02 | 0.021 | < 0.1 |  | 1
| 1 | 2 | 0.0194 | 0.024709 | 0.021 | 0.02 | 0.02 | 0.0203 | 0.02 | 0.02 | 0.0199 | 0.020013 | 0.021 | 0.021 | 0.019304 | 0.019 | 0.02 | 0.02 | 0.02 | < 0.1 |  | #N/A
| 1 | 3 | 0.0188 | 0.024577 | 0.021 | 0.021 | 0.02 | 0.0189 | 0.019 | 0.019 | 0.0196 | 0.020061 | 0.021 | 0.02 | 0.019621 | 0.019 | 0.021 | 0.02 | 0.02 | < 0.1 |  | 16
| 1 | 4 | 0.0187 | 0.024433 | 0.021 | 0.021 | 0.02 | 0.0194 | 0.019 | 0.02 | 0.0196 | 0.019753 | 0.022 | 0.019 | 0.019819 | 0.019 | 0.02 | 0.02 | 0.021 | < 0.1 |  | 0.020028
| 1 | 5 | 0.0197 | 0.024058 | 0.022 | 0.02 | 0.02 | 0.0204 | 0.019 | 0.02 | 0.0197 | 0.020115 | 0.022 | 0.02 | 0.019164 | 0.02 | 0.02 | 0.019 | 0.02 | < 0.1 |  | 63
| 1 | 6 | 0.0203 | 0.024844 | 0.022 | 0.022 | 0.02 | 0.0192 | 0.02 | 0.02 | 0.0194 | 0.020633 | 0.021 | 0.02 | 0.018624 | 0.019 | 0.02 | 0.02 | 0.019 | < 0.1
| Mean |  | 0.0193 | 0.024448 | 0.021333 | 0.020833 | 0.02 | 0.019617 | 0.0195 | 0.019833 | 0.019683 | 0.020132 | 0.021333 | 0.02 | 0.019377 | 0.019333 | 0.020167 | 0.019833 | 0.020167
| Median |  | 0.01915 | 0.024505 | 0.021 | 0.021 | 0.02 | 0.01945 | 0.0195 | 0.02 | 0.01965 | 0.020088 | 0.021 | 0.02 | 0.019463 | 0.019 | 0.02 | 0.02 | 0.02
| Std Dev. |  | 0.0006229 | 0.00032903 | 0.0005164 | 0.00075277 | 0 | 0.0006047 | 0.00054772 | 0.00040825 | 0.00019408 | 0.00029053 | 0.0005164 | 0.00063246 | 0.00044643 | 0.0005164 | 0.00040825 | 0.00040825 | 0.00075277
| Rel.Std.Dev. |  | 3.227% | 1.346% | 2.421% | 3.613% | 0% | 3.083% | 2.809% | 2.058% | 0.986% | 1.443% | 2.421% | 3.162% | 2.304% | 2.671% | 2.024% | 2.058% | 3.733%
| PDM3 |  | -3.633% | 22.07% | 6.519% | 4.023% | -0.1381% | -2.052% | -2.635% | -0.9703% | -1.719% | 0.5227% | 6.519% | -0.1381% | -3.249% | -3.467% | 0.6941% | -0.9703% | 0.6941%
| Z-Score (Absolute) |  | 1.07 | 6.79 | 2.04 | 1.27 | 0 | 0.59 | 0.76 | 0.25 | 0.48 | 0.2 | 2.04 | 0 | 0.95 | 1.02 | 0.25 | 0.25 | 0.25 | 45.8
| Analytical results for Pb in OREAS 782 (Certified Value 14.5 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 19 | 15.4 | 15.3 | 106.43 | 12.58 | 14 | 16 | 16.6 | 14.7931 | 14 | 14.3 | 16.8538 | 17.3 | 14.7002 | 14.3 | 15.6 | 13.8 | 12.7021 |  | 1
| 1 | 2 | 17 | 15.6 | 15 | 101.55 | 12.45 | 14 | 16 | 15.8 | 14.7196 | 13 | 14.8 | 17.3547 | 14.2 | 14.6692 | 13.6 | 15.1 | 14.2 | 12.7489 |  | 8
| 1 | 3 | 17 | 15.8 | 14.8 | 105.81 | 13.48 | 14 | 16 | 15.5 | 14.4549 | 14 | 14.3 | 17.0096 | 14.5 | 14.824 | 14.5 | 15.6 | 14 | 12.2008 |  | 16
| 1 | 4 | 18 | 15.9 | 14.8 | 104.88 | 12.53 | 14 | 16 | 15.6 | 14.5783 | 14 | 13.9 | 18.2354 | 13.4 | 14.2955 | 13.8 | 15.8 | 14.7 | 12.992 |  | 14.5388
| 1 | 5 | 17 | 15.9 | 14.8 | 100.65 | 13.58 | 14 | 16 | 15.6 | 15.0313 | 15 | 14 | 16.1072 | 13.7 | 14.7515 | 14 | 16 | 14.4 | 12.3984 |  | 64
| 1 | 6 | 18 | 15.8 | 14.4 | 101.21 | 12.7 | 14 | 16 | 15.5 | 15.2772 | 15 | 14.7 | 17.0944 | 13.8 | 14.4813 | 14.2 | 15 | 14.2 | 12.7573
| Mean |  | 17.6667 | 15.7333 | 14.85 | 103.4217 | 12.8867 | 14 | 16 | 15.7667 | 14.8091 | 14.1667 | 14.3333 | 17.1092 | 14.4833 | 14.6203 | 14.0667 | 15.5167 | 14.2167 | 12.6332
| Median |  | 17.5 | 15.8 | 14.8 | 103.215 | 12.64 | 14 | 16 | 15.6 | 14.7564 | 14 | 14.3 | 17.052 | 14 | 14.6847 | 14.1 | 15.6 | 14.2 | 12.7255
| Std Dev. |  | 0.8165 | 0.19664 | 0.29496 | 2.5674 | 0.50587 | 0 | 0 | 0.42269 | 0.30182 | 0.75277 | 0.36148 | 0.69398 | 1.4331 | 0.19612 | 0.33267 | 0.392 | 0.31252 | 0.28438
| Rel.Std.Dev. |  | 4.622% | 1.25% | 1.986% | 2.482% | 3.926% | 0% | 0% | 2.681% | 2.038% | 5.314% | 2.522% | 4.056% | 9.895% | 1.341% | 2.365% | 2.526% | 2.198% | 2.251%
| PDM3 |  | 21.51% | 8.216% | 2.14% | 611.3% | -11.36% | -3.706% | 10.05% | 8.445% | 1.859% | -2.56% | -1.413% | 17.68% | -0.3818% | 0.5602% | -3.248% | 6.726% | -2.216% | -13.11%
| Z-Score (Absolute) |  | 3.31 | 1.21 | 0.25 | NA | 1.88 | 0.67 | 1.5 | 1.25 | 0.21 | 0.49 | 0.31 | 2.71 | 0.15 | 0 | 0.6 | 0.97 | 0.44 | 2.16
Results from laboratory 1.08 were removed due to their systematic bias.
| Analytical results for Pr in OREAS 782 (Certified Value 7.05 ppm) | [20 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 6.57 | 7.2023 | 7.23 | 7.4 | 7.16 | 7.4 | 6.82 | 7.2672 | 6.4439 | 7.0421 | [9 empty] | 1
| 1 | 2 | 5.93 | 7.1927 | 7.51 | 7.2 | 7.02 | 7.6 | 6.8 | 6.4953 | 6.5179 | 6.9648 | [9 empty] | 9
| 1 | 3 | 6.34 | 6.964 | 7.07 | 7.4 | 6.92 | 7.4 | 6.85 | 7.2922 | 6.7101 | 7.0018 | [9 empty] | 16
| 1 | 4 | 6.83 | 7.2293 | 7.33 | 7.2 | 7.07 | 7.45 | 6.69 | 7.3078 | 6.9151 | 6.7286 | [9 empty] | 7.0455
| 1 | 5 | 6.82 | 6.9708 | 6.95 | 7.1 | 7.08 | 7.45 | 6.91 | 7.2476 | 6.7436 | 7.3421 | [9 empty] | 65
| 1 | 6 | 6.74 | 7.006 | 7.52 | 7.2 | 7 | 7.4 | 6.9 | 7.2306 | 6.7966 | 7.1265
| Mean |  | 6.5383 | 7.0942 | 7.2683 | 7.25 | 7.0417 | 7.45 | 6.8283 | 7.1401 | 6.6879 | 7.0343
| Median |  | 6.655 | 7.0994 | 7.28 | 7.2 | 7.045 | 7.425 | 6.835 | 7.2574 | 6.7268 | 7.022
| Std Dev. |  | 0.35085 | 0.12619 | 0.23138 | 0.12247 | 0.081588 | 0.07746 | 0.080353 | 0.31714 | 0.17633 | 0.20132
| Rel.Std.Dev. |  | 5.366% | 1.779% | 3.183% | 1.689% | 1.159% | 1.04% | 1.177% | 4.442% | 2.637% | 2.862%
| PDM3 |  | -7.198% | 0.6914% | 3.163% | 2.903% | -0.05415% | 5.742% | -3.082% | 1.343% | -5.076% | -0.1585%
| Z-Score (Absolute) |  | 1.87 | 0.09 | 0.71 | 0.64 | 0.09 | 1.35 | 0.84 | 0.25 | 1.34 | 0.12
| Analytical results for Rb in OREAS 782 (Certified Value 83 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
1
| 1 | 1 | 83.6 | 80.0357 | 86.8 | 85.5 | 84 | 88.5 | 82.8 | 86.07 | 83.8 | 80.2 | 91.5322 | 92.5 | 81.685 | 84.4 | 83 | 81.6 | 82.4573 |  |  | 1
| 1 | 2 | 77.9 | 80.0271 | 91.4 | 85.8 | 86 | 88.2 | 81 | 87.27 | 84.8 | 78.5 | 74.6331 | 94.2 | 82.1314 | 87.4 | 79.8 | 82.6 | 82.0383 |  |  | 10
| 1 | 3 | 78.9 | 79.9444 | 94.4 | 86.4 | 82 | 87.8 | 82.6 | 84.82 | 84.4 | 82.5 | 81.0102 | 92.7 | 80.996 | 83.8 | 83.7 | 81.4 | 80.7469 |  |  | 16
| 1 | 4 | 79.8 | 80.0755 | 87.5 | 85.5 | 83 | 90.7 | 81.8 | 85.44 | 84.6 | 79.7 | 81.0544 | 90.2 | 81.4524 | 89.5 | 79.3 | 85.5 | 79.6668 |  |  | 83.3108
| 1 | 5 | 82.3 | 79.9308 | 90.1 | 86 | 81 | 85.7 | 83.2 | 86.14 | 84.4 | 82.2 | 81.2119 | 92.8 | 80.3991 | 87.7 | 79.5 | 83.2 | 82.4182 |  |  | 66
| 1 | 6 | 80.2 | 80.083 | 92.3 | 84.4 | 83 | 88 | 81.7 | 83.24 | 83.8 | 83.1 | 81.6123 | 93 | 79.0403 | 86.2 | 80.7 | 83.9 | 82.08
| Mean |  | 80.45 | 80.0161 | 90.4167 | 85.6 | 83.1667 | 88.15 | 82.1833 | 85.4967 | 84.3 | 81.0333 | 81.8423 | 92.5667 | 80.9507 | 86.5 | 81 | 83.0333 | 81.5679
| Median |  | 80 | 80.0314 | 90.75 | 85.65 | 83 | 88.1 | 82.2 | 85.755 | 84.4 | 81.2 | 81.1331 | 92.75 | 81.2242 | 86.8 | 80.25 | 82.9 | 82.0592
| Std Dev. |  | 2.1324 | 0.064712 | 2.9006 | 0.67823 | 1.7224 | 1.5984 | 0.82077 | 1.3741 | 0.41473 | 1.8261 | 5.4338 | 1.3064 | 1.1082 | 2.1466 | 1.8953 | 1.5345 | 1.1212
| Rel.Std.Dev. |  | 2.651% | 0.08087% | 3.208% | 0.7923% | 2.071% | 1.813% | 0.9987% | 1.607% | 0.492% | 2.254% | 6.639% | 1.411% | 1.369% | 2.482% | 2.34% | 1.848% | 1.375%
| PDM3 |  | -3.434% | -3.955% | 8.529% | 2.748% | -0.173% | 5.809% | -1.353% | 2.624% | 1.187% | -2.734% | -1.763% | 11.11% | -2.833% | 3.828% | -2.774% | -0.333% | -2.092%
| Z-Score (Absolute) |  | 0.84 | 0.98 | 2.39 | 0.83 | 0.04 | 1.66 | 0.28 | 0.8 | 0.41 | 0.65 | 0.39 | 3.09 | 0.67 | 1.12 | 0.66 | 0 | 0.47
| Analytical results for Re in OREAS 782 (Indicative Value 0.003 ppm) | [20 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 19 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.005 | 0.003 | 0.005 | 0.002 | 0.003 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | 0.006 | 0.004 | < 0.002 | 0.003 | [6 empty] | 1
| 1 | 2 | 0.005 | 0.003 | 0.005 | 0.002 | 0.002 | < 0.05 | 0.003 | < 0.1 | < 0.1 | 0.006 | 0.004 | < 0.002 | 0.003 | [6 empty] | 16
| 1 | 3 | 0.005 | < 0.002 | 0.005 | 0.002 | 0.003 | < 0.05 | 0.003 | < 0.1 | < 0.1 | < 0.005 | 0.003 | < 0.002 | 0.003 | [6 empty] | 16
| 1 | 4 | 0.006 | 0.003 | 0.004 | 0.002 | 0.003 | < 0.05 | 0.003 | < 0.1 | < 0.1 | 0.007 | 0.006 | < 0.002 | 0.003 | [6 empty] | 0.00335
| 1 | 5 | 0.005 | 0.003 | 0.004 | 0.002 | 0.003 | < 0.05 | < 0.002 | < 0.1 | < 0.1 | < 0.005 | 0.003 | < 0.002 | 0.002 | [6 empty] | 22
| 1 | 6 | 0.005 | 0.002 | 0.003 | 0.002 | 0.003 | < 0.05 | 0.003 | < 0.1 | < 0.1 | 0.008 | 0.003 | < 0.002 | 0.003
| Mean |  | 0.0051667 | 0.0028 | 0.0043333 | 0.002 | 0.0028333 |  | 0.003 |  |  | 0.00675 | 0.0038333 |  | 0.0028333
| Median |  | 0.005 | 0.003 | 0.0045 | 0.002 | 0.003 |  | 0.003 |  |  | 0.0065 | 0.0035 |  | 0.003
| Std Dev. |  | 0.00040825 | 0.00044721 | 0.0008165 | 0 | 0.00040825 |  | 0 |  |  | 0.00095743 | 0.001169 |  | 0.00040825
| Rel.Std.Dev. |  | 7.902% | 15.97% | 18.84% | 0% | 14.41% |  | 0% |  |  | 14.18% | 30.5% |  | 14.41%
| PDM3 |  | 54.23% | -16.42% | 29.35% | -40.3% | -15.42% |  | -10.45% |  |  | 101.5% | 14.43% |  | -15.42%
| Z-Score (Absolute) |  | 0.6 | 0.6 | 0.22 | 0.82 | 0.45 | 9.52 | 0.67 | 20.75 | 20.75 | 0.67 | 0 | 1.27 | 0.45
| Analytical results for S in OREAS 782 (Certified Value 0.025 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 19 | 21 | 22 | 23 | 24 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0221 | 0.023847 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.025 | 0.0255 | < 0.05 | 0.03 | 0.024 | < 0.05 | 0.02 | 0.03 | 0.02 | < 0.1 |  |  | 1
| 1 | 2 | 0.0217 | 0.023227 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.024 | 0.025 | < 0.05 | 0.03 | 0.023 | < 0.05 | 0.03 | 0.03 | 0.03 | < 0.1 |  |  | #N/A
| 1 | 3 | 0.0219 | 0.023365 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.025 | 0.0247 | < 0.05 | 0.03 | 0.024 | < 0.05 | 0.02 | 0.03 | 0.02 | < 0.1 |  |  | 16
| 1 | 4 | 0.0224 | 0.023229 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.026 | 0.0248 | < 0.05 | 0.03 | 0.024 | < 0.05 | 0.02 | 0.03 | 0.02 | < 0.1 |  |  | 0.024754
| 1 | 5 | 0.0226 | 0.023594 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.025 | 0.0249 | < 0.05 | 0.03 | 0.024 | < 0.05 | 0.02 | 0.03 | 0.02 | < 0.1 |  |  | 67
| 1 | 6 | 0.0228 | 0.023758 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.026 | 0.0249 | < 0.05 | 0.03 | 0.025 | < 0.05 | 0.02 | 0.03 | 0.03 | < 0.1
| Mean |  | 0.02225 | 0.023503 | 0.021667 | 0.03 | 0.02 | 0.02 | 0.03 | 0.025167 | 0.024967 |  | 0.03 | 0.024 |  | 0.021667 | 0.03 | 0.023333
| Median |  | 0.02225 | 0.02348 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.025 | 0.0249 |  | 0.03 | 0.024 |  | 0.02 | 0.03 | 0.02
| Std Dev. |  | 0.00042308 | 0.00026905 | 0.0040825 | 0 | 0 | 0 | 0 | 0.00075277 | 0.00028048 |  | 0 | 0.00063246 |  | 0.0040825 | 0 | 0.005164
| Rel.Std.Dev. |  | 1.902% | 1.145% | 18.84% | 0% | 0% | 0% | 0% | 2.991% | 1.123% |  | 0% | 2.635% |  | 18.84% | 0% | 22.13%
| PDM3 |  | -10.11% | -5.052% | -12.47% | 21.19% | -19.2% | -19.2% | 21.19% | 1.668% | 0.8599% |  | 21.19% | -3.045% |  | -12.47% | 21.19% | -5.738%
| Z-Score (Absolute) |  | 0.56 | 0.3 | 0.67 | 1.03 | 1.01 | 1.01 | 1.03 | 0.04 | 0 | 0.01 | 1.03 | 0.2 | 0.01 | 0.67 | 1.03 | 0.33 | 5.12
| Analytical results for Sb in OREAS 782 (Certified Value 5.16 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.99 | 4.6136 | 5.63 | 5.15 | < 5 | 4.58 | 5.4 | 5.4 | 5.51 | 4.4007 | 5.3 | 5.72 | 5.0945 | 5.5 | 5.0151 | 5.21 | 5.22 | 4.92 | 4.74 | 1
| 1 | 2 | 4.64 | 4.5424 | 5.75 | 5.06 | < 5 | 4.57 | 5.4 | 5.3 | 5.7 | 4.5764 | 5.3 | 5.48 | 4.5124 | 5.6 | 5.1136 | 5.13 | 5.3 | 4.94 | 5.4551 | 11
| 1 | 3 | 4.67 | 4.6127 | 5.79 | 5.15 | < 5 | 4.88 | 5.6 | 5.4 | 5.64 | 4.5792 | 5.4 | 5.64 | 5.1195 | 5.6 | 5.0933 | 5.09 | 5.49 | 4.95 | 4.143 | 16
| 1 | 4 | 4.79 | 4.6348 | 5.79 | 5.06 | < 5 | 4.9 | 5.5 | 5.3 | 5.45 | 4.8506 | 5.4 | 5.6 | 4.9698 | 5.3 | 5.1185 | 5.43 | 5.25 | 5.28 | 4.3015 | 5.1611
| 1 | 5 | 4.91 | 4.4333 | 5.83 | 5.14 | < 5 | 4.89 | 5.6 | 5.3 | 5.51 | 4.6724 | 5.5 | 5.8 | 4.9966 | 5.4 | 5.1372 | 5.31 | 5.26 | 5.26 | 4.6146 | 68
| 1 | 6 | 4.86 | 4.6613 | 5.77 | 5.1 | < 5 | 4.96 | 5.3 | 5.4 | 5.51 | 4.8798 | 5.4 | 5.88 | 4.9829 | 5.2 | 5.0094 | 5.23 | 5.39 | 4.95 | 4.8136
| Mean |  | 4.81 | 4.583 | 5.76 | 5.11 |  | 4.7967 | 5.4667 | 5.35 | 5.5533 | 4.6598 | 5.3833 | 5.6867 | 4.946 | 5.4333 | 5.0812 | 5.2333 | 5.3183 | 5.05 | 4.678
| Median |  | 4.825 | 4.6131 | 5.78 | 5.12 |  | 4.885 | 5.45 | 5.35 | 5.51 | 4.6258 | 5.4 | 5.68 | 4.9898 | 5.45 | 5.1035 | 5.22 | 5.28 | 4.95 | 4.6773
| Std Dev. |  | 0.13697 | 0.083291 | 0.068993 | 0.042895 |  | 0.17397 | 0.12111 | 0.054772 | 0.095219 | 0.18196 | 0.075277 | 0.14404 | 0.2212 | 0.1633 | 0.055215 | 0.12356 | 0.10265 | 0.17088 | 0.46002
| Rel.Std.Dev. |  | 2.848% | 1.817% | 1.198% | 0.8394% |  | 3.627% | 2.215% | 1.024% | 1.715% | 3.905% | 1.398% | 2.533% | 4.472% | 3.006% | 1.087% | 2.361% | 1.93% | 3.384% | 9.834%
| PDM3 |  | -6.803% | -11.2% | 11.6% | -0.99% |  | -7.061% | 5.921% | 3.66% | 7.6% | -9.712% | 4.306% | 10.18% | -4.168% | 5.275% | -1.548% | 1.4% | 3.047% | -2.153% | -9.361%
| Z-Score (Absolute) |  | 0.65 | 1.13 | 1.4 | 0 | 5.62 | 0.67 | 0.77 | 0.52 | 0.95 | 0.97 | 0.59 | 1.24 | 0.35 | 0.7 | 0.06 | 0.27 | 0.45 | 0.13 | 0.93
| Analytical results for Sc in OREAS 782 (Certified Value 9.02 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 7.94 | 9.01 | 10.2 | 8.9 | 10 | 8 | 9 | 9.3 | 9.4603 | 9 | 9.5 | 9.1444 | 8.8 | 8.5447 | 9.6 | 8.5 | 8.7 | 8.6875 |  | 1
| 1 | 2 | 7.96 | 9.14 | 10.6 | 9 | 10.16 | 8 | 9 | 9.4 | 9.1719 | 9 | 8.7 | 8.4951 | 8.5 | 8.6809 | 9.4 | 8.5 | 8.8 | 9.2962 |  | 12
| 1 | 3 | 7.86 | 9.04 | 10.6 | 9 | 10.13 | 8 | 9 | 9.3 | 9.1715 | 9 | 9.4 | 9.0456 | 8.4 | 8.6685 | 9.6 | 8.8 | 8.6 | 9.4745 |  | 16
| 1 | 4 | 7.94 | 9.06 | 10.3 | 8.9 | 10.15 | 8 | 9 | 9.1 | 9.1074 | 9 | 9.1 | 9.2862 | 8.4 | 9.1138 | 9.4 | 8.4 | 9.2 | 9.1526 |  | 9.0161
| 1 | 5 | 7.81 | 9.05 | 10.6 | 8.7 | 10.09 | 8 | 9 | 9.3 | 9.1672 | 9 | 9.4 | 9.2677 | 8.6 | 9.0885 | 9.7 | 8.2 | 8.9 | 9.199 |  | 69
| 1 | 6 | 7.96 | 8.97 | 10.7 | 9.2 | 10.66 | 8 | 9 | 9.2 | 8.8859 | 9 | 9.3 | 9.3997 | 8.4 | 8.7469 | 9.6 | 8.7 | 8.9 | 9.3334
| Mean |  | 7.9117 | 9.045 | 10.5 | 8.95 | 10.1983 | 8 | 9 | 9.2667 | 9.1607 | 9 | 9.2333 | 9.1064 | 8.5167 | 8.8072 | 9.55 | 8.5167 | 8.85 | 9.1905
| Median |  | 7.94 | 9.045 | 10.6 | 8.95 | 10.14 | 8 | 9 | 9.3 | 9.1694 | 9 | 9.35 | 9.2061 | 8.45 | 8.7139 | 9.6 | 8.5 | 8.85 | 9.2476
| Std Dev. |  | 0.062102 | 0.056833 | 0.2 | 0.16432 | 0.23353 | 0 | 0 | 0.10328 | 0.18351 | 0 | 0.29439 | 0.32348 | 0.16021 | 0.23701 | 0.12247 | 0.2137 | 0.20736 | 0.27084
| Rel.Std.Dev. |  | 0.7849% | 0.6283% | 1.905% | 1.836% | 2.29% | 0% | 0% | 1.115% | 2.003% | 0% | 3.188% | 3.552% | 1.881% | 2.691% | 1.282% | 2.509% | 2.343% | 2.947%
| PDM3 |  | -12.25% | 0.3205% | 16.46% | -0.7332% | 13.11% | -11.27% | -0.1786% | 2.779% | 1.604% | -0.1786% | 2.409% | 1.002% | -5.539% | -2.317% | 5.922% | -5.539% | -1.842% | 1.935%
| Z-Score (Absolute) |  | 3.14 | 0.16 | 3.66 | 0.41 | 2.87 | NA | NA | 0.42 | 0.14 | NA | 0.33 | 0 | 1.55 | 0.79 | 1.17 | 1.55 | 0.67 | 0.22
Results from laboratories 10, 12 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 782 (Certified Value 1 ppm) | [20 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
2
| 1 | 1 | 1.1 | 1.2046 | 2 | 2 | < 1 | 0.7 | 1 | 0.9 | < 5 | < 5 | < 1 | 2.5941 | 1 | 1 | 1 | < 0.1 |  |  |  | 1
| 1 | 2 | 1 | 1.1185 | 2 | 2 | < 1 | 0.7 | 1 | 1 | < 5 | < 5 | < 1 | 0.93822 | 1 | 1 | 1 | < 0.1 |  |  |  | 1
| 1 | 3 | 1.07 | 1.1854 | 2 | 2 | < 1 | 0.8 | 1 | 0.7 | < 5 | < 5 | 1 | 1.9546 | 1 | 1 | 1 | < 0.1 |  |  |  | 16
| 1 | 4 | 1.02 | 1.1197 | 2 | 2 | < 1 | 1 | 1 | 0.9 | < 5 | < 5 | < 1 | 1.6961 | 1 | 1 | 1 | < 0.1 |  |  |  | 0.99641
| 1 | 5 | 1.22 | 1.0688 | 2 | 2 | < 1 | 1 | 1 | 1.1 | < 5 | < 5 | < 1 | 2.4055 | 1 | 1 | 1 | < 0.1 |  |  |  | 70
| 1 | 6 | 1.02 | 1.1792 | 3 | 1 | < 1 | 0.8 | 1 | 0.9 | < 5 | < 5 | < 1 | 2.0524 | 1 | 1 | 1 | < 0.1
| Mean |  | 1.0717 | 1.146 | 2.1667 | 1.8333 |  | 0.83333 | 1 | 0.91667 |  |  | 1 | 1.9401 | 1 | 1 | 1
| Median |  | 1.045 | 1.1494 | 2 | 2 |  | 0.8 | 1 | 0.9 |  |  | 1 | 2.0035 | 1 | 1 | 1
| Std Dev. |  | 0.081588 | 0.051942 | 0.40825 | 0.40825 |  | 0.13663 | 0 | 0.13292 |  |  |  | 0.58657 | 0 | 0 | 0
| Rel.Std.Dev. |  | 7.613% | 4.532% | 18.84% | 22.27% |  | 16.4% | 0% | 14.5% |  |  |  | 30.23% | 0% | 0% | 0%
| PDM3 |  | 7.553% | 15.02% | 117.4% | 83.99% |  | -16.37% | 0.3603% | -8.003% |  |  | 0.3603% | 94.71% | 0.3603% | 0.3603% | 0.3603%
| Z-Score (Absolute) |  | 0.17 | 0.34 | 2.7 | 1.93 | 1.16 | 0.39 | 0 | 0.19 | 3.47 | 3.47 | 0.96 | 2.17 | 0 | 0 | 0 | 2.2
| Analytical results for Sm in OREAS 782 (Certified Value 5 ppm) | [20 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 4.3 | 5.0239 | 5.13 | 4.9 | 4.98 | 5.05 | 4.6 | 5.1735 | 4.674 | 5.134 | [9 empty] | 1
| 1 | 2 | 3.92 | 5.1552 | 5.01 | 5.1 | 5.08 | 5.05 | 4.6 | 4.5564 | 4.4961 | 4.9778 | [9 empty] | 28
| 1 | 3 | 4.13 | 4.8768 | 5.33 | 5 | 4.93 | 5.1 | 4.7 | 5.0365 | 4.5888 | 4.7978 | [9 empty] | 16
| 1 | 4 | 4.52 | 5.1396 | 5.24 | 4.9 | 5.03 | 5.1 | 4.6 | 5.1737 | 4.932 | 4.7873 | [9 empty] | 4.9983
| 1 | 5 | 4.53 | 4.8708 | 5.06 | 5.1 | 4.91 | 5.15 | 4.7 | 5.1164 | 4.6379 | 5.0044 | [9 empty] | 71
| 1 | 6 | 4.44 | 5.1403 | 5.52 | 5 | 4.98 | 4.95 | 4.7 | 5.0662 | 4.8691 | 5.0882
| Mean |  | 4.3067 | 5.0344 | 5.215 | 5 | 4.985 | 5.0667 | 4.65 | 5.0204 | 4.6997 | 4.9649
| Median |  | 4.37 | 5.0817 | 5.185 | 5 | 4.98 | 5.075 | 4.65 | 5.0913 | 4.656 | 4.9911
| Std Dev. |  | 0.24246 | 0.13311 | 0.18982 | 0.089443 | 0.062849 | 0.068313 | 0.054772 | 0.234 | 0.16789 | 0.14491
| Rel.Std.Dev. |  | 5.63% | 2.644% | 3.64% | 1.789% | 1.261% | 1.348% | 1.178% | 4.661% | 3.572% | 2.919%
| PDM3 |  | -13.84% | 0.7233% | 4.336% | 0.03478% | -0.2653% | 1.369% | -6.968% | 0.4437% | -5.974% | -0.6671%
| Z-Score (Absolute) |  | 7.97 | 0.49 | 2.59 | 0.09 | 0.09 | 0.86 | 3.98 | 0.32 | 3.4 | 0.32
| Analytical results for Sn in OREAS 782 (Certified Value 9.86 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 9.9 | 12.1 | 10.8 | < 5 | 8.72 | 11 | 10 | 10.5 | 10.2001 | 9.3 | 8.3 | 9.206 | 9.7 | 10.3404 | 9.7 | 9.8 | 9.9 | 7.0495 |  | 1
| 1 | 2 | 9.19 | 12.5 | 10.6 | < 5 | 8.72 | 10 | 10.1 | 10.7 | 10.3583 | 9.2 | 9 | 6.7481 | 10.2 | 10.555 | 10 | 9.8 | 9.9 | 7.2167 |  | 3
| 1 | 3 | 9.36 | 12.1 | 10.8 | < 5 | 8.86 | 10 | 10.1 | 10.6 | 10.41 | 9.5 | 9.9 | 7.9561 | 10.6 | 10.4469 | 9.5 | 10 | 9.6 | 7.0307 |  | 16
| 1 | 4 | 9.39 | 11.7 | 10.6 | < 5 | 8.15 | 10 | 10.3 | 10.3 | 10.3939 | 9.6 | 9.6 | 8.7092 | 10 | 10.3345 | 9.9 | 9.7 | 10.2 | 7.3282 |  | 9.858
| 1 | 5 | 9.62 | 12 | 10.8 | < 5 | 8.84 | 10 | 10.1 | 10.4 | 10.5752 | 9.5 | 10 | 9.6573 | 10.6 | 10.5225 | 10 | 9.7 | 10.4 | 6.9959 |  | 72
| 1 | 6 | 9.65 | 12.1 | 10.5 | < 5 | 8.62 | 10 | 10 | 10.6 | 10.7659 | 9.3 | 10.1 | 8.9622 | 9.9 | 10.2592 | 9.6 | 10.1 | 10 | 7.3877
| Mean |  | 9.5183 | 12.0833 | 10.6833 |  | 8.6517 | 10.1667 | 10.1 | 10.5167 | 10.4506 | 9.4 | 9.4833 | 8.5398 | 10.1667 | 10.4098 | 9.7833 | 9.85 | 10 | 7.1681
| Median |  | 9.505 | 12.1 | 10.7 |  | 8.72 | 10 | 10.1 | 10.55 | 10.4019 | 9.4 | 9.75 | 8.8357 | 10.1 | 10.3937 | 9.8 | 9.8 | 9.95 | 7.1331
| Std Dev. |  | 0.25404 | 0.25626 | 0.13292 |  | 0.26111 | 0.40825 | 0.10954 | 0.1472 | 0.19548 | 0.15492 | 0.70261 | 1.0441 | 0.37238 | 0.1169 | 0.2137 | 0.16432 | 0.27568 | 0.16662
| Rel.Std.Dev. |  | 2.669% | 2.121% | 1.244% |  | 3.018% | 4.016% | 1.085% | 1.4% | 1.871% | 1.648% | 7.409% | 12.23% | 3.663% | 1.123% | 2.184% | 1.668% | 2.757% | 2.324%
| PDM3 |  | -3.446% | 22.57% | 8.372% |  | -12.24% | 3.131% | 2.455% | 6.681% | 6.011% | -4.646% | -3.801% | -13.37% | 3.131% | 5.597% | -0.7575% | -0.08126% | 1.44% | -27.29%
| Z-Score (Absolute) |  | 0.4 | 2.69 | 1 | 8.85 | 1.44 | NA | 0.3 | 0.8 | 0.72 | 0.54 | 0.44 | 1.58 | 0.38 | 0.67 | 0.08 | 0 | 0.18 | 3.23
Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 782 (Certified Value 51 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 49 | 47.9095 | 54.9 | 52.8 | 48.76 | 49.4 | 51.4 | 50.9 | 50.94 | 52.987 | 51.5 | 51 | 52.3799 | 55 | 46.3209 | 52.3 | 52.7 | 49.6 | 49.8837 | 1
| 1 | 2 | 49 | 48.0235 | 56.6 | 53.5 | 48.8 | 49.1 | 51.6 | 50.7 | 51.51 | 50.5595 | 51.5 | 50 | 53.824 | 52 | 48.5634 | 52.2 | 51.6 | 49.6 | 47.4147 | 4
| 1 | 3 | 49 | 48.3275 | 57.7 | 53 | 48.42 | 49.1 | 50.7 | 52.5 | 51.24 | 50.9452 | 51 | 51 | 57.1542 | 53 | 47.7136 | 50.7 | 53.8 | 48.9 | 47.5429 | 16
| 1 | 4 | 49 | 47.3015 | 55.7 | 52.5 | 49.09 | 49.4 | 51.9 | 53.4 | 51.04 | 50.6871 | 51.5 | 50 | 49.691 | 51 | 47.7955 | 54.1 | 52.2 | 51.8 | 48.631 | 51.0203
| 1 | 5 | 49 | 48.2515 | 57.6 | 52.4 | 48.33 | 49.5 | 50.1 | 52 | 51.45 | 52.4564 | 52.5 | 52 | 53.6216 | 56 | 47.1452 | 52.4 | 51 | 50.7 | 48.0917 | 73
| 1 | 6 | 49 | 47.8715 | 57.8 | 52.8 | 48.27 | 50 | 51.5 | 51.5 | 50.59 | 52.0191 | 51.5 | 52 | 51.2548 | 52 | 46.3133 | 52 | 52.5 | 50.3 | 48.2956
| Mean |  | 49 | 47.9475 | 56.7167 | 52.8333 | 48.6117 | 49.4167 | 51.2 | 51.8333 | 51.1283 | 51.609 | 51.5833 | 51 | 52.9876 | 53.1667 | 47.3087 | 52.2833 | 52.3 | 50.15 | 48.3099
| Median |  | 49 | 47.9665 | 57.1 | 52.8 | 48.59 | 49.4 | 51.45 | 51.75 | 51.14 | 51.4822 | 51.5 | 51 | 53.0008 | 52.5 | 47.4294 | 52.25 | 52.35 | 49.95 | 48.1937
| Std Dev. |  | 0 | 0.36527 | 1.2057 | 0.39328 | 0.32221 | 0.33116 | 0.66933 | 1.0191 | 0.34487 | 1.0175 | 0.4916 | 0.89443 | 2.5574 | 1.9408 | 0.89091 | 1.087 | 0.96333 | 1.0213 | 0.89637
| Rel.Std.Dev. |  | 0% | 0.7618% | 2.126% | 0.7444% | 0.6628% | 0.6701% | 1.307% | 1.966% | 0.6745% | 1.972% | 0.953% | 1.754% | 4.826% | 3.65% | 1.883% | 2.079% | 1.842% | 2.036% | 1.855%
| PDM3 |  | -3.96% | -6.023% | 11.16% | 3.554% | -4.721% | -3.143% | 0.3522% | 1.594% | 0.2117% | 1.154% | 1.104% | -0.03983% | 3.856% | 4.207% | -7.275% | 2.476% | 2.508% | -1.706% | -5.312%
| Z-Score (Absolute) |  | 0.91 | 1.34 | 2.28 | 0.67 | 1.07 | 0.74 | 0 | 0.26 | 0.03 | 0.17 | 0.16 | 0.08 | 0.74 | 0.81 | 1.61 | 0.45 | 0.45 | 0.43 | 1.19
| Analytical results for Ta in OREAS 782 (Certified Value 0.59 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 0.33 | 0.64 | 0.61 | 0.4 | 0.6 | 0.45 | 0.56 | 0.5 | 0.64 | 0.65466 | 0.6 | 0.66967 | 0.56 | 0.57 | 0.6 |  |  |  |  | 1
| 1 | 2 | 0.33 | 0.65 | 0.6 | 0.43 | 0.6 | 0.48 | 0.62 | 0.55 | 0.54 | 0.54484 | 0.7 | 0.67965 | 0.55 | 0.57 | 0.6 |  |  |  |  | 5
| 1 | 3 | 0.3 | 0.64 | 0.58 | 0.41 | 0.5 | 0.48 | 0.62 | 0.5 | 0.63 | 0.6137 | 0.6 | 0.69006 | 0.56 | 0.67 | 0.59 |  |  |  |  | 16
| 1 | 4 | 0.32 | 0.67 | 0.59 | 0.42 | 0.5 | 0.48 | 0.62 | 0.5 | 0.6 | 0.66239 | 0.6 | 0.69768 | 0.58 | 0.57 | 0.63 |  |  |  |  | 0.59342
| 1 | 5 | 0.34 | 0.68 | 0.6 | 0.42 | 0.6 | 0.47 | 0.63 | 0.55 | 0.61 | 0.62894 | 0.6 | 0.68695 | 0.57 | 0.56 | 0.65 |  |  |  |  | 74
| 1 | 6 | 0.32 | 0.66 | 0.59 | 0.39 | 0.5 | 0.44 | 0.68 | 0.5 | 0.59 | 0.63498 | 0.7 | 0.69221 | 0.53 | 0.6 | 0.6
| Mean |  | 0.32333 | 0.65667 | 0.595 | 0.41167 | 0.55 | 0.46667 | 0.62167 | 0.51667 | 0.60167 | 0.62325 | 0.63333 | 0.68604 | 0.55833 | 0.59 | 0.61167
| Median |  | 0.325 | 0.655 | 0.595 | 0.415 | 0.55 | 0.475 | 0.62 | 0.5 | 0.605 | 0.63196 | 0.6 | 0.6885 | 0.56 | 0.57 | 0.6
| Std Dev. |  | 0.013663 | 0.01633 | 0.010488 | 0.01472 | 0.054772 | 0.017512 | 0.038166 | 0.02582 | 0.035449 | 0.042257 | 0.05164 | 0.0099936 | 0.017224 | 0.041473 | 0.023166
| Rel.Std.Dev. |  | 4.226% | 2.487% | 1.763% | 3.576% | 9.959% | 3.753% | 6.139% | 4.997% | 5.892% | 6.78% | 8.154% | 1.457% | 3.085% | 7.029% | 3.787%
| PDM3 |  | -45.51% | 10.66% | 0.2663% | -30.63% | -7.317% | -21.36% | 4.76% | -12.93% | 1.39% | 5.027% | 6.726% | 15.61% | -5.913% | -0.5763% | 3.075%
| Z-Score (Absolute) |  | 5 | 1.13 | 0 | 3.37 | NA | 2.36 | 0.49 | 1.44 | 0.12 | 0.52 | NA | 1.67 | 0.67 | 0.09 | 0.31
Results from laboratories 10 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 782 (Certified Value 0.65 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | < 0.2 | 0.84801 | 0.63 | 0.7 | 0.66 | 0.62 | 0.61 | 0.66815 | 0.65125 | 0.5258 | [9 empty] | 1
| 1 | 2 | < 0.2 | 0.85384 | 0.67 | 0.7 | 0.65 | 0.62 | 0.61 | 0.59039 | 0.67288 | 0.5123 | [9 empty] | 20
| 1 | 3 | < 0.2 | 0.78362 | 0.61 | 0.7 | 0.74 | 0.62 | 0.63 | 0.6499 | 0.69319 | 0.505 | [9 empty] | 16
| 1 | 4 | < 0.2 | 0.84387 | 0.63 | 0.7 | 0.71 | 0.64 | 0.6 | 0.66469 | 0.7346 | 0.5155 | [9 empty] | 0.6477
| 1 | 5 | < 0.2 | 0.77799 | 0.62 | 0.7 | 0.65 | 0.62 | 0.64 | 0.65529 | 0.67525 | 0.5171 | [9 empty] | 75
| 1 | 6 | < 0.2 | 0.80965 | 0.65 | 0.7 | 0.63 | 0.6 | 0.63 | 0.65474 | 0.71702 | 0.5048
| Mean |  |  | 0.8195 | 0.635 | 0.7 | 0.67333 | 0.62 | 0.62 | 0.64719 | 0.6907 | 0.51342
| Median |  |  | 0.82676 | 0.63 | 0.7 | 0.655 | 0.62 | 0.62 | 0.65502 | 0.68422 | 0.5139
| Std Dev. |  |  | 0.033745 | 0.021679 | 1.2162e-16 | 0.042269 | 0.012649 | 0.015492 | 0.028641 | 0.030785 | 0.0079698
| Rel.Std.Dev. |  |  | 4.118% | 3.414% | 1.737e-14% | 6.278% | 2.04% | 2.499% | 4.425% | 4.457% | 1.552%
| PDM3 |  |  | 26.52% | -1.961% | 8.074% | 3.957% | -4.277% | -4.277% | -0.07901% | 6.638% | -20.73%
| Z-Score (Absolute) |  | 9.41 | 3.25 | 0 | NA | 0.67 | 0.26 | 0.26 | 0.21 | 0.98 | 2.14
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 782 (Certified Value 0.2 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.2 | 0.21 | 0.21 | 0.47 | 0.2 | 0.2 | 0.2 | < 5 | < 0.2 | 0.19 | 0.19372 | 0.2 | 0.21601 | 0.22 | 0.23 | 0.21 |  |  |  | 1
| 1 | 2 | 0.26 | 0.19 | 0.18 | 0.46 | 0.3 | 0.2 | < 0.2 | < 5 | < 0.2 | 0.2 | 0.21044 | 0.2 | 0.19645 | 0.22 | 0.2 | 0.19 |  |  |  | 6
| 1 | 3 | 0.2 | 0.21 | 0.23 | 0.47 | 0.2 | 0.2 | 0.2 | < 5 | < 0.2 | 0.18 | 0.21391 | 0.2 | 0.21446 | 0.2 | 0.21 | 0.17 |  |  |  | 16
| 1 | 4 | 0.14 | 0.23 | 0.18 | 0.47 | 0.2 | 0.2 | 0.2 | < 5 | < 0.2 | 0.21 | 0.23279 | 0.2 | 0.22128 | 0.2 | 0.2 | 0.21 |  |  |  | 0.20099
| 1 | 5 | 0.19 | 0.2 | 0.21 | 0.49 | 0.1 | 0.2 | < 0.2 | < 5 | < 0.2 | 0.19 | 0.21395 | 0.2 | 0.21868 | 0.21 | 0.15 | 0.19 |  |  |  | 76
| 1 | 6 | 0.16 | 0.33 | 0.17 | 0.5 | 0.2 | 0.2 | 0.2 | < 5 | < 0.2 | 0.19 | 0.13442 | 0.3 | 0.21062 | 0.2 | 0.15 | 0.2
| Mean |  | 0.19167 | 0.22833 | 0.19667 | 0.47667 | 0.2 | 0.2 | 0.2 |  |  | 0.19333 | 0.19987 | 0.21667 | 0.21292 | 0.20833 | 0.19 | 0.195
| Median |  | 0.195 | 0.21 | 0.195 | 0.47 | 0.2 | 0.2 | 0.2 |  |  | 0.19 | 0.21217 | 0.2 | 0.21524 | 0.205 | 0.2 | 0.195
| Std Dev. |  | 0.041191 | 0.051543 | 0.023381 | 0.015055 | 0.063246 | 3.0405e-17 | 0 |  |  | 0.010328 | 0.034388 | 0.040825 | 0.0088492 | 0.0098319 | 0.032863 | 0.015166
| Rel.Std.Dev. |  | 21.49% | 22.57% | 11.89% | 3.158% | 31.62% | 1.52e-14% | 0% |  |  | 5.342% | 17.21% | 18.84% | 4.156% | 4.719% | 17.3% | 7.777%
| PDM3 |  | -4.637% | 13.61% | -2.149% | 137.2% | -0.4908% | -0.4908% | -0.4908% |  |  | -3.808% | -0.5552% | 7.802% | 5.936% | 3.655% | -5.466% | -2.979%
| Z-Score (Absolute) |  | 0.51 | 3.2 | 0 | 28.32 | NA | NA | NA | NA | 9.78 | 0.34 | 0.32 | NA | 1.64 | 1.18 | 0.67 | 0.17
Results from laboratory 1.26 were removed due to their systematic bias. Results from laboratories 10, 12, 13 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 782 (Certified Value 9.94 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 4.08 | 10.5 | 9.83 | 9.28 | 10.5 | 9.38 | 10.06 | 9.7 | 10.12 | 9.7254 | 10.1 | 5.4629 | 10.45 | 9.86 | 9.61 | 10.2612 |  |  |  | 1
| 1 | 2 | 3.96 | 10.9 | 9.8 | 9.27 | 10.4 | 9.2 | 9.9 | 9.4 | 9.91 | 8.7324 | 10.4 | 7.3088 | 10.45 | 9.71 | 9.72 | 9.3525 |  |  |  | 7
| 1 | 3 | 3.89 | 10.8 | 9.87 | 9.47 | 10.4 | 9.49 | 9.84 | 9.7 | 10.06 | 9.6953 | 10.2 | 7.4471 | 11 | 10.05 | 9.4 | 10.0679 |  |  |  | 16
| 1 | 4 | 4.53 | 10.8 | 9.59 | 9.49 | 10.3 | 9.44 | 9.93 | 9.6 | 9.95 | 9.9747 | 10 | 6.5575 | 9.4 | 9.58 | 10.05 | 9.1909 |  |  |  | 9.9392
| 1 | 5 | 4.9 | 10.85 | 9.93 | 9.58 | 10.2 | 9.36 | 9.98 | 9.8 | 10.13 | 9.7033 | 10.2 | 6.3446 | 10.6 | 9.64 | 10.15 | 10.7171 |  |  |  | 77
| 1 | 6 | 4.88 | 11.05 | 9.57 | 9.15 | 10.3 | 9.37 | 9.82 | 9.3 | 10.34 | 9.7861 | 10.3 | 7.2654 | 10.5 | 9.8 | 10 | 9.8594
| Mean |  | 4.3733 | 10.8167 | 9.765 | 9.3733 | 10.35 | 9.3733 | 9.9217 | 9.5833 | 10.085 | 9.6029 | 10.2 | 6.7311 | 10.4 | 9.7733 | 9.8217 | 9.9082
| Median |  | 4.305 | 10.825 | 9.815 | 9.375 | 10.35 | 9.375 | 9.915 | 9.65 | 10.09 | 9.7144 | 10.2 | 6.9114 | 10.475 | 9.755 | 9.86 | 9.9637
| Std Dev. |  | 0.45807 | 0.18074 | 0.1499 | 0.16428 | 0.10488 | 0.098319 | 0.089536 | 0.19408 | 0.15346 | 0.43892 | 0.14142 | 0.76416 | 0.53198 | 0.16967 | 0.29144 | 0.57099
| Rel.Std.Dev. |  | 10.47% | 1.671% | 1.535% | 1.753% | 1.013% | 1.049% | 0.9024% | 2.025% | 1.522% | 4.571% | 1.386% | 11.35% | 5.115% | 1.736% | 2.967% | 5.763%
| PDM3 |  | -56% | 8.829% | -1.752% | -5.693% | 4.133% | -5.693% | -0.1762% | -3.58% | 1.467% | -3.384% | 2.624% | -32.28% | 4.636% | -1.669% | -1.182% | -0.312%
| Z-Score (Absolute) |  | 10.6 | 1.99 | 0.06 | 0.83 | 1.08 | 0.83 | 0.24 | 0.42 | 0.56 | 0.38 | 0.79 | 5.99 | 1.18 | 0.05 | 0.05 | 0.22
| Analytical results for Ti in OREAS 782 (Certified Value 0.24 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.1163 | 0.30262 | 0.258 | 0.244 | 0.23 | 0.18 | 0.233 | 0.208 | 0.232 | 0.24182 | 0.223 | 0.258 | 0.26922 | 0.216 | 0.25906 | 0.235 | 0.247 | 0.251 | 0.1641 | 1
| 1 | 2 | 0.1165 | 0.30262 | 0.259 | 0.239 | 0.23 | 0.19 | 0.245 | 0.211 | 0.2499 | 0.23645 | 0.225 | 0.251 | 0.26635 | 0.224 | 0.25579 | 0.23 | 0.244 | 0.245 | 0.2033 | 22
| 1 | 3 | 0.1177 | 0.29668 | 0.261 | 0.235 | 0.23 | 0.18 | 0.216 | 0.217 | 0.2454 | 0.2367 | 0.229 | 0.253 | 0.26698 | 0.21 | 0.26792 | 0.236 | 0.243 | 0.248 | 0.1703 | 16
| 1 | 4 | 0.1193 | 0.30855 | 0.259 | 0.241 | 0.23 | 0.17 | 0.219 | 0.214 | 0.2493 | 0.2315 | 0.225 | 0.247 | 0.26978 | 0.223 | 0.25402 | 0.231 | 0.241 | 0.254 | 0.1823 | 0.24042
| 1 | 5 | 0.1227 | 0.29668 | 0.259 | 0.247 | 0.23 | 0.18 | 0.221 | 0.212 | 0.2561 | 0.23757 | 0.221 | 0.256 | 0.27064 | 0.228 | 0.26262 | 0.234 | 0.238 | 0.247 | 0.1493 | 78
| 1 | 6 | 0.1214 | 0.30855 | 0.26 | 0.24 | 0.23 | 0.17 | 0.228 | 0.215 | 0.2616 | 0.24701 | 0.222 | 0.254 | 0.26551 | 0.226 | 0.24989 | 0.227 | 0.243 | 0.249 | 0.1722
| Mean |  | 0.11898 | 0.30262 | 0.25933 | 0.241 | 0.23 | 0.17833 | 0.227 | 0.21283 | 0.24905 | 0.23851 | 0.22417 | 0.25317 | 0.26808 | 0.22117 | 0.25822 | 0.23217 | 0.24267 | 0.249 | 0.17358
| Median |  | 0.1185 | 0.30262 | 0.259 | 0.2405 | 0.23 | 0.18 | 0.2245 | 0.213 | 0.2496 | 0.23714 | 0.224 | 0.2535 | 0.2681 | 0.2235 | 0.25742 | 0.2325 | 0.243 | 0.2485 | 0.17125
| Std Dev. |  | 0.002637 | 0.0053073 | 0.0010328 | 0.0041473 | 0 | 0.0075277 | 0.010789 | 0.0031885 | 0.010119 | 0.0053049 | 0.0028577 | 0.0038687 | 0.0020761 | 0.006824 | 0.006433 | 0.0034303 | 0.0030111 | 0.0031623 | 0.018173
| Rel.Std.Dev. |  | 2.216% | 1.754% | 0.3983% | 1.721% | 0% | 4.221% | 4.753% | 1.498% | 4.063% | 2.224% | 1.275% | 1.528% | 0.7744% | 3.085% | 2.491% | 1.477% | 1.241% | 1.27% | 10.47%
| PDM3 |  | -50.51% | 25.87% | 7.865% | 0.2396% | -4.336% | -25.83% | -5.583% | -11.48% | 3.588% | -0.7967% | -6.762% | 5.3% | 11.5% | -8.01% | 7.401% | -3.434% | 0.9328% | 3.567% | -27.8%
| Z-Score (Absolute) |  | 5.5 | 2.95 | 0.96 | 0.11 | 0.39 | 2.77 | 0.53 | 1.18 | 0.48 | 0 | 0.66 | 0.67 | 1.36 | 0.8 | 0.91 | 0.29 | 0.19 | 0.48 | 2.99
| Analytical results for Tl in OREAS 782 (Certified Value 0.47 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.481 | 0.5 | 0.48 | 0.4 | 0.48 | 0.5 | 0.5 | 0.5 | 0.56 | 0.42716 | 0.5 | 0.47647 | 0.46 | 0.42 | 0.46 | 0.3131 |  |  |  | 1
| 1 | 2 | 0.43 | 0.5 | 0.49 | 0.41 | 0.48 | 0.48 | 0.5 | 0.5 | 0.56 | 0.38258 | 0.6 | 0.48253 | 0.44 | 0.43 | 0.46 | 0.4953 |  |  |  | 8
| 1 | 3 | 0.455 | 0.51 | 0.47 | 0.44 | 0.48 | 0.5 | 0.49 | 0.5 | 0.53 | 0.42176 | 0.5 | 0.48485 | 0.47 | 0.42 | 0.44 | 0.541 |  |  |  | 16
| 1 | 4 | 0.483 | 0.53 | 0.49 | 0.42 | 0.49 | 0.49 | 0.51 | 0.5 | 0.54 | 0.43248 | 0.5 | 0.47812 | 0.47 | 0.45 | 0.47 | 0.412 |  |  |  | 0.47314
| 1 | 5 | 0.477 | 0.53 | 0.49 | 0.45 | 0.48 | 0.49 | 0.51 | 0.5 | 0.56 | 0.41273 | 0.5 | 0.48639 | 0.45 | 0.41 | 0.47 | 0.4565 |  |  |  | 79
| 1 | 6 | 0.478 | 0.49 | 0.47 | 0.41 | 0.48 | 0.48 | 0.49 | 0.5 | 0.53 | 0.42436 | 0.5 | 0.48224 | 0.45 | 0.43 | 0.46 | 0.6262
| Mean |  | 0.46733 | 0.51 | 0.48167 | 0.42167 | 0.48167 | 0.49 | 0.5 | 0.5 | 0.54667 | 0.41684 | 0.51667 | 0.48177 | 0.45667 | 0.42667 | 0.46 | 0.47402
| Median |  | 0.4775 | 0.505 | 0.485 | 0.415 | 0.48 | 0.49 | 0.5 | 0.5 | 0.55 | 0.42306 | 0.5 | 0.48238 | 0.455 | 0.425 | 0.46 | 0.4759
| Std Dev. |  | 0.020906 | 0.016733 | 0.0098319 | 0.019408 | 0.0040825 | 0.0089443 | 0.0089443 | 0 | 0.015055 | 0.018015 | 0.040825 | 0.0038221 | 0.012111 | 0.013663 | 0.010954 | 0.10784
| Rel.Std.Dev. |  | 4.473% | 3.281% | 2.041% | 4.603% | 0.8476% | 1.825% | 1.789% | 0% | 2.754% | 4.322% | 7.902% | 0.7933% | 2.652% | 3.202% | 2.381% | 22.75%
| PDM3 |  | -1.228% | 7.79% | 1.802% | -10.88% | 1.802% | 3.563% | 5.676% | 5.676% | 15.54% | -11.9% | 9.199% | 1.823% | -3.482% | -9.823% | -2.778% | 0.1847%
| Z-Score (Absolute) |  | 0.36 | 1.11 | 0.13 | 1.94 | 0.13 | 0.42 | 0.77 | NA | 2.38 | 2.11 | NA | 0.14 | 0.73 | 1.77 | 0.62 | 0.13
Results from laboratories 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 782 (Certified Value 0.26 ppm) | [20 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [10 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.17 | 0.27 | 0.3 | 0.26 | 0.26 | 0.28 | 0.28875 | 0.33888 | 0.2432 | [10 empty] | 1
| 1 | 2 | 0.15 | 0.26 | 0.3 | 0.26 | 0.26 | 0.28 | 0.25147 | 0.35023 | 0.249 | [10 empty] | 24
| 1 | 3 | 0.16 | 0.26 | 0.3 | 0.26 | 0.26 | 0.28 | 0.28398 | 0.36494 | 0.2361 | [10 empty] | 16
| 1 | 4 | 0.18 | 0.27 | 0.3 | 0.26 | 0.26 | 0.27 | 0.29443 | 0.37821 | 0.2459 | [10 empty] | 0.26488
| 1 | 5 | 0.18 | 0.26 | 0.3 | 0.27 | 0.27 | 0.28 | 0.28228 | 0.36512 | 0.2435 | [10 empty] | 80
| 1 | 6 | 0.18 | 0.27 | 0.3 | 0.27 | 0.24 | 0.28 | 0.28025 | 0.36841 | 0.2468
| Mean |  | 0.17 | 0.265 | 0.3 | 0.26333 | 0.25833 | 0.27833 | 0.28019 | 0.36096 | 0.24408
| Median |  | 0.175 | 0.265 | 0.3 | 0.26 | 0.26 | 0.28 | 0.28313 | 0.36503 | 0.2447
| Std Dev. |  | 0.012649 | 0.0054772 | 0 | 0.005164 | 0.0098319 | 0.0040825 | 0.014964 | 0.014063 | 0.0044656
| Rel.Std.Dev. |  | 7.441% | 2.067% | 0% | 1.961% | 3.806% | 1.467% | 5.341% | 3.896% | 1.83%
| PDM3 |  | -35.82% | 0.0455% | 13.26% | -0.5837% | -2.471% | 5.079% | 5.782% | 36.28% | -7.851%
| Z-Score (Absolute) |  | 4.21 | 0.04 | NA | 0.04 | 0.26 | 0.63 | 0.72 | 4.32 | 0.9
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 782 (Certified Value 1.94 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.87 | 1.4019 | 2.1 | 2 | 1.87 | 2.1 | 1.95 | 2.06 | < 50 | 1.9 | 1.93 | 1.9338 | 2 | 2.0079 | 2 | 2 | 1.8 | 1.9197 |  | 1
| 1 | 2 | 1.78 | 1.4112 | 2.1 | 2 | 1.95 | 2.1 | 1.83 | 2.06 | < 50 | 1.9 | 2.04 | 1.8103 | 2 | 2.0156 | 1.8 | 2 | 1.8 | 1.8065 |  | 9
| 1 | 3 | 1.78 | 1.3678 | 2.1 | 2 | 1.91 | 2.1 | 1.95 | 2.04 | < 50 | 1.9 | 1.99 | 2.1196 | 2 | 2.0297 | 1.9 | 2 | 1.8 | 1.9414 |  | 16
| 1 | 4 | 1.91 | 1.3521 | 2.1 | 2 | 1.86 | 2.1 | 1.96 | 2.06 | < 50 | 1.9 | 1.91 | 1.9365 | 2 | 1.9214 | 1.9 | 1.9 | 2 | 1.8827 |  | 1.9413
| 1 | 5 | 1.96 | 1.3371 | 2.1 | 2 | 2 | 2.1 | 1.9 | 2.07 | < 50 | 1.9 | 2 | 1.9225 | 2 | 2.0123 | 1.9 | 1.9 | 1.9 | 1.8474 |  | 81
| 1 | 6 | 1.88 | 1.5261 | 2.2 | 2 | 1.92 | 2.1 | 1.91 | 2.04 | < 50 | 1.8 | 2.02 | 1.9537 | 2 | 2.0452 | 1.9 | 1.9 | 1.9 | 1.9099
| Mean |  | 1.8633 | 1.3994 | 2.1167 | 2 | 1.9183 | 2.1 | 1.9167 | 2.055 |  | 1.8833 | 1.9817 | 1.946 | 2 | 2.0054 | 1.9 | 1.95 | 1.8667 | 1.8846
| Median |  | 1.875 | 1.3848 | 2.1 | 2 | 1.915 | 2.1 | 1.93 | 2.06 |  | 1.9 | 1.995 | 1.9351 | 2 | 2.014 | 1.9 | 1.95 | 1.85 | 1.8963
| Std Dev. |  | 0.07174 | 0.068275 | 0.040825 | 0 | 0.051929 | 0 | 0.048854 | 0.012247 |  | 0.040825 | 0.051153 | 0.099382 | 0 | 0.04334 | 0.063246 | 0.054772 | 0.08165 | 0.050171
| Rel.Std.Dev. |  | 3.85% | 4.879% | 1.929% | 0% | 2.707% | 0% | 2.549% | 0.596% |  | 2.168% | 2.581% | 5.107% | 0% | 2.161% | 3.329% | 2.809% | 4.374% | 2.662%
| PDM3 |  | -4.014% | -27.92% | 9.036% | 3.026% | -1.181% | 8.177% | -1.267% | 5.859% |  | -2.984% | 2.082% | 0.2467% | 3.026% | 3.302% | -2.125% | 0.4505% | -3.842% | -2.918%
| Z-Score (Absolute) |  | 0.72 | 6.77 | 2.59 | NA | 0 | 2.37 | 0.02 | 1.78 | NA | 0.46 | 0.83 | 0.36 | NA | 1.14 | 0.24 | 0.41 | 0.67 | 0.44
Results from laboratory 1.14 were removed due to their systematic bias. Results from laboratories 4 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 782 (Certified Value 62 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 56 | 65 | 62 | 53.05 | 60 | 61 | 58 | 56 | 60.5599 | 64 | 60 | 60.9861 | 61 | 64.4229 | 62 | 61 | 63 | 55.679 |  | 1
| 1 | 2 | 56 | 65 | 62 | 53.36 | 61 | 63 | 57 | 67 | 59.3167 | 64 | 61 | 54.0092 | 63 | 64.3021 | 60 | 61 | 60 | 56.5877 |  | 10
| 1 | 3 | 56 | 66 | 63 | 54.99 | 61 | 61 | 59 | 59 | 59.2481 | 64 | 66 | 60.7522 | 61 | 65.4761 | 63 | 61 | 59 | 52.7235 |  | 16
| 1 | 4 | 56 | 65 | 64 | 52.36 | 60 | 62 | 59 | 61 | 59.5856 | 64 | 64 | 60.3531 | 60 | 62.216 | 62 | 61 | 63 | 54.9762 |  | 61.8007
| 1 | 5 | 56 | 65 | 63 | 53.7 | 61 | 61 | 58 | 63 | 60.3376 | 64 | 66 | 60.6231 | 62 | 64.326 | 61 | 60 | 61 | 53.7051 |  | 82
| 1 | 6 | 57 | 66 | 62 | 54.08 | 64 | 62 | 59 | 64 | 62.0038 | 62 | 65 | 60.0174 | 61 | 61.391 | 61 | 61 | 62 | 53.3336
| Mean |  | 56.1667 | 65.3333 | 62.6667 | 53.59 | 61.1667 | 61.6667 | 58.3333 | 61.6667 | 60.1753 | 63.6667 | 63.6667 | 59.4568 | 61.3333 | 63.689 | 61.5 | 60.8333 | 61.3333 | 54.5009
| Median |  | 56 | 65 | 62.5 | 53.53 | 61 | 61.5 | 58.5 | 62 | 59.9616 | 64 | 64.5 | 60.4881 | 61 | 64.3141 | 61.5 | 61 | 61.5 | 54.3406
| Std Dev. |  | 0.40825 | 0.5164 | 0.8165 | 0.90218 | 1.472 | 0.8165 | 0.8165 | 3.8816 | 1.0452 | 0.8165 | 2.582 | 2.6896 | 1.0328 | 1.5469 | 1.0488 | 0.40825 | 1.633 | 1.4915
| Rel.Std.Dev. |  | 0.7269% | 0.7904% | 1.303% | 1.683% | 2.406% | 1.324% | 1.4% | 6.294% | 1.737% | 1.282% | 4.055% | 4.524% | 1.684% | 2.429% | 1.705% | 0.6711% | 2.662% | 2.737%
| PDM3 |  | -9.116% | 5.716% | 1.401% | -13.29% | -1.026% | -0.2169% | -5.611% | -0.2169% | -2.63% | 3.019% | 3.019% | -3.793% | -0.7563% | 3.055% | -0.4866% | -1.565% | -0.7563% | -11.81%
| Z-Score (Absolute) |  | 2.17 | 1.68 | 0.56 | 3.25 | 0.07 | 0.14 | 1.26 | 0.14 | 0.49 | 0.98 | 0.98 | 0.79 | 0 | 0.99 | 0.07 | 0.21 | 0 | 2.87
| Analytical results for W in OREAS 782 (Certified Value 178 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 181.27 | 187.7 | 203 | 175.5 | 188.92 | 186 | 193 | 183 | 178.8 | 183.2724 | 174 | 200 | 175.5273 | 172.5 | 170.6941 | 179 | 174.5 | 165.5 |  | 1
| 1 | 2 | 172.81 | 188.1 | 202 | 174.5 | 188.13 | 185 | 184 | 181 | 178.5 | 181.5076 | 177 | 190 | 161.4319 | 180.3 | 173.5155 | 172 | 174.5 | 168.5 |  | 26
| 1 | 3 | 169.55 | 188 | 206 | 175 | 183.98 | 188 | 183 | 185 | 179.6 | 179.2572 | 175 | 196 | 179.6916 | 177.9 | 176.1635 | 173 | 179.5 | 164.5 |  | 16
| 1 | 4 | 181.52 | 188.2 | 207 | 174.5 | 183.04 | 189 | 185 | 190 | 177.7 | 178.2264 | 179 | 188 | 183.9807 | 169 | 174.1759 | 181 | 173.5 | 178.5 |  | 178.4765
| 1 | 5 | 171.87 | 187.9 | 203 | 176.5 | 183.78 | 184 | 182 | 185 | 179.1 | 182.791 | 174 | 193 | 170.3419 | 171.9 | 174.4386 | 170.5 | 169.5 | 168.5 |  | 83
| 1 | 6 | 176.49 | 187.9 | 204 | 171 | 184.71 | 190 | 184 | 185 | 176 | 187.6184 | 170 | 196 | 177.0849 | 181.5 | 170.2511 | 172 | 174 | 170
| Mean |  | 175.585 | 187.9667 | 204.1667 | 174.5 | 185.4267 | 187 | 185.1667 | 184.8333 | 178.2833 | 182.1122 | 174.8333 | 193.8333 | 174.6764 | 175.5167 | 173.2064 | 174.5833 | 174.25 | 169.25
| Median |  | 174.65 | 187.95 | 203.5 | 174.75 | 184.345 | 187 | 184 | 185 | 178.65 | 182.1493 | 174.5 | 194.5 | 176.3061 | 175.2 | 173.8457 | 172.5 | 174.25 | 168.5
| Std Dev. |  | 5.0256 | 0.17512 | 1.9408 | 1.8708 | 2.4709 | 2.3664 | 3.9707 | 2.9944 | 1.2859 | 3.3389 | 3.0605 | 4.4008 | 7.9051 | 5.0795 | 2.2954 | 4.3176 | 3.1898 | 4.9774
| Rel.Std.Dev. |  | 2.862% | 0.09316% | 0.9506% | 1.072% | 1.333% | 1.265% | 2.144% | 1.62% | 0.7213% | 1.833% | 1.751% | 2.27% | 4.526% | 2.894% | 1.325% | 2.473% | 1.831% | 2.941%
| PDM3 |  | -1.62% | 5.317% | 14.39% | -2.228% | 3.894% | 4.776% | 3.749% | 3.562% | -0.1082% | 2.037% | -2.041% | 8.604% | -2.129% | -1.658% | -2.953% | -2.181% | -2.368% | -5.17%
| Z-Score (Absolute) |  | 0.2 | 1.67 | 4.12 | 0.37 | 1.29 | 1.52 | 1.25 | 1.2 | 0.2 | 0.78 | 0.32 | 2.56 | 0.34 | 0.21 | 0.56 | 0.36 | 0.41 | 1.16
| Analytical results for Y in OREAS 782 (Certified Value 45.5 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 24.26 | 41.9115 | 47.9 | 43.8 | 43.47 | 46.9 | 45.3 | 45.32 | 43.9366 | 45.7 | 47.1 | 44.1156 | 48.2 | 46.2694 | 45.7 | 45.7 | 42.9 | 46.0056 |  | 1
| 1 | 2 | 26.35 | 41.5584 | 49.1 | 44 | 43.85 | 46.4 | 45.9 | 44.85 | 42.8734 | 46.2 | 44.07 | 38.6588 | 49.9 | 49.8923 | 46.2 | 44.4 | 42.2 | 45.1714 |  | #N/A
| 1 | 3 | 24.3 | 41.8287 | 49.5 | 44.7 | 42.54 | 46 | 46.8 | 45.23 | 42.5923 | 45.9 | 46.92 | 43.591 | 48 | 50.7803 | 45.5 | 47.8 | 40.7 | 45.9069 |  | 16
| 1 | 4 | 26.6 | 41.8249 | 48.4 | 44.5 | 43.17 | 46.6 | 45.9 | 45.33 | 43.4346 | 46.9 | 45.45 | 43.3638 | 48 | 50.9259 | 50.4 | 44.3 | 43.9 | 44.1304 |  | 45.4846
| 1 | 5 | 26.72 | 41.6596 | 50.6 | 44.1 | 42.65 | 46.4 | 46.6 | 45.79 | 44.2263 | 46.7 | 46.54 | 43.657 | 48.7 | 48.9 | 46.9 | 45.6 | 42.3 | 45.4931 |  | 84
| 1 | 6 | 27.28 | 41.93 | 50 | 44.3 | 43.25 | 46.3 | 45.5 | 43.99 | 45.0024 | 45.7 | 46.16 | 43.1534 | 48.8 | 50.4507 | 46.2 | 46.3 | 42.3 | 46.3844
| Mean |  | 25.9183 | 41.7855 | 49.25 | 44.2333 | 43.155 | 46.4333 | 46 | 45.085 | 43.6776 | 46.1833 | 46.04 | 42.7566 | 48.6 | 49.5364 | 46.8167 | 45.6833 | 42.3833 | 45.5153
| Median |  | 26.475 | 41.8268 | 49.3 | 44.2 | 43.21 | 46.4 | 45.9 | 45.275 | 43.6856 | 46.05 | 46.35 | 43.4774 | 48.45 | 50.1715 | 46.2 | 45.65 | 42.3 | 45.7
| Std Dev. |  | 1.3051 | 0.14669 | 1.0015 | 0.33267 | 0.49492 | 0.30111 | 0.5933 | 0.61442 | 0.89522 | 0.51543 | 1.1306 | 2.0333 | 0.72388 | 1.7618 | 1.8214 | 1.2983 | 1.0439 | 0.79822
| Rel.Std.Dev. |  | 5.036% | 0.3511% | 2.034% | 0.7521% | 1.147% | 0.6485% | 1.29% | 1.363% | 2.05% | 1.116% | 2.456% | 4.756% | 1.489% | 3.557% | 3.891% | 2.842% | 2.463% | 1.754%
| PDM3 |  | -43.02% | -8.133% | 8.278% | -2.751% | -5.122% | 2.086% | 1.133% | -0.8785% | -3.973% | 1.536% | 1.221% | -5.998% | 6.849% | 8.908% | 2.929% | 0.437% | -6.818% | 0.06756%
| Z-Score (Absolute) |  | 8.07 | 1.56 | 1.5 | 0.56 | 1 | 0.34 | 0.16 | 0.21 | 0.79 | 0.24 | 0.18 | 1.17 | 1.23 | 1.62 | 0.5 | 0.03 | 1.32 | 0.03
| Analytical results for Yb in OREAS 782 (Certified Value 1.86 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 10 | 13 | 16 | 17 | 18 | 20 | 24 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 1.33 | 1.9657 | 1.67 | 1.9 | 1.93 | 1.8 | 1.9 | 2.0283 | 2.2971 | 1.9229 | [9 empty] | 1
| 1 | 2 | 1.21 | 1.9127 | 1.63 | 1.9 | 1.89 | 1.9 | 1.8 | 1.7581 | 2.3165 | 1.9506 | [9 empty] | 27
| 1 | 3 | 1.25 | 1.8244 | 1.64 | 1.8 | 1.9 | 1.85 | 1.8 | 1.9465 | 2.4385 | 1.8907 | [9 empty] | 16
| 1 | 4 | 1.39 | 1.9296 | 1.8 | 1.9 | 1.83 | 1.75 | 1.8 | 2.0262 | 2.5498 | 1.8999 | [9 empty] | 1.8636
| 1 | 5 | 1.41 | 1.956 | 1.7 | 2 | 1.79 | 1.85 | 1.9 | 1.9552 | 2.4421 | 1.9768 | [9 empty] | 85
| 1 | 6 | 1.41 | 1.9692 | 1.72 | 1.9 | 2.17 | 1.8 | 1.8 | 1.9704 | 2.4541 | 1.8534
| Mean |  | 1.3333 | 1.9263 | 1.6933 | 1.9 | 1.9183 | 1.825 | 1.8333 | 1.9475 | 2.4164 | 1.9157
| Median |  | 1.36 | 1.9428 | 1.685 | 1.9 | 1.895 | 1.825 | 1.8 | 1.9628 | 2.4403 | 1.9114
| Std Dev. |  | 0.086178 | 0.054499 | 0.062503 | 0.063246 | 0.13333 | 0.05244 | 0.05164 | 0.099168 | 0.094401 | 0.044179
| Rel.Std.Dev. |  | 6.463% | 2.829% | 3.691% | 3.329% | 6.95% | 2.873% | 2.817% | 5.092% | 3.907% | 2.306%
| PDM3 |  | -28.46% | 3.361% | -9.138% | 1.951% | 2.935% | -2.073% | -1.626% | 4.497% | 29.66% | 2.794%
| Z-Score (Absolute) |  | 6.79 | 0.22 | 2.54 | 0.09 | 0.12 | 0.98 | 0.88 | 0.47 | 6.01 | 0.09
| Analytical results for Zn in OREAS 782 (Certified Value 21.6 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 18 | 18.21 | 24 | 23 | 23.51 | 21 | 22 | 20 | 21 | 23.6077 | 20 | 25 | 21.6549 | 23 | 24.0238 | 22 | 23 | 23 | 55.7241 | 1
| 1 | 2 | 18 | 18.41 | 24 | 22 | 23.99 | 21 | 22 | 20 | 21 | 22.1839 | 18 | 23 | 17.8254 | 19 | 23.6795 | 22 | 23 | 22 | 56.8162 | 12
| 1 | 3 | 18 | 18.4 | 25 | 23 | 23.85 | 21 | 21 | 21 | 20 | 22.3732 | 24 | 22 | 22.0939 | 20 | 22.8942 | 22 | 23 | 22 | 56.2498 | 16
| 1 | 4 | 18 | 18.12 | 24 | 23 | 24.77 | 20 | 22 | 21 | 20 | 22.0821 | 20 | 24 | 34.747 | 21 | 23.1761 | 20 | 29 | 23 | 55.4388 | 21.6232
| 1 | 5 | 18 | 18.43 | 24 | 24 | 24.52 | 20 | 21 | 20 | 21 | 22.2878 | 22 | 22 | 18.6127 | 21 | 26.6916 | 23 | 24 | 23 | 56.2674 | 86
| 1 | 6 | 18 | 18.51 | 24 | 23 | 23.4 | 20 | 21 | 21 | 20 | 22.8443 | 20 | 21 | 22.735 | 20 | 23.8115 | 24 | 23 | 22 | 55.7579
| Mean |  | 18 | 18.3467 | 24.1667 | 23 | 24.0067 | 20.5 | 21.5 | 20.5 | 20.5 | 22.5632 | 20.6667 | 22.8333 | 22.9448 | 20.6667 | 24.0461 | 22.1667 | 24.1667 | 22.5 | 56.0424
| Median |  | 18 | 18.405 | 24 | 23 | 23.92 | 20.5 | 21.5 | 20.5 | 20.5 | 22.3305 | 20 | 22.5 | 21.8744 | 20.5 | 23.7455 | 22 | 23 | 22.5 | 56.0038
| Std Dev. |  | 0 | 0.14868 | 0.40825 | 0.63246 | 0.54511 | 0.54772 | 0.54772 | 0.54772 | 0.54772 | 0.57571 | 2.0656 | 1.472 | 6.1106 | 1.3663 | 1.3618 | 1.3292 | 2.4014 | 0.54772 | 0.49779
| Rel.Std.Dev. |  | 0% | 0.8104% | 1.689% | 2.75% | 2.271% | 2.672% | 2.548% | 2.672% | 2.672% | 2.552% | 9.995% | 6.447% | 26.63% | 6.611% | 5.663% | 5.996% | 9.937% | 2.434% | 0.8882%
| PDM3 |  | -16.76% | -15.15% | 11.76% | 6.367% | 11.02% | -5.195% | -0.5698% | -5.195% | -5.195% | 4.347% | -4.424% | 5.596% | 6.112% | -4.424% | 11.21% | 2.513% | 11.76% | 4.055% | 159.2%
| Z-Score (Absolute) |  | 1.82 | 1.68 | 0.67 | 0.2 | 0.61 | 0.81 | 0.4 | 0.81 | 0.81 | 0.03 | 0.74 | 0.13 | 0.18 | 0.74 | 0.63 | 0.13 | 0.67 | 0 | 13.57
| Analytical results for Zr in OREAS 782 (Certified Value 85 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 69.6 | 93.364 | 98.4 | 88.4 | 98.71 | 80 | 87 | 85.9 | 85.6 | 79.0967 | 84.5 | 90 | 86.2774 | 74 | 133.2531 | 86.2 | 85.2 | 81.4 | 95.9266 | 1
| 1 | 2 | 71 | 93.532 | 101 | 90.3 | 101.6 | 84 | 87 | 85.6 | 87.3 | 78.1432 | 84.5 | 85 | 73.9057 | 75 | 133.2655 | 85.3 | 85.6 | 81.5 | 98.6647 | #N/A
| 1 | 3 | 69.7 | 94.08 | 102.5 | 91.9 | 101.6 | 82 | 84 | 88.6 | 87.1 | 76.4894 | 83.5 | 85 | 83.547 | 75 | 134.2067 | 83.6 | 87.4 | 78.9 | 91.2419 | 16
| 1 | 4 | 70.6 | 93.744 | 102 | 87.5 | 98.76 | 78 | 84 | 84.8 | 87.2 | 77.4801 | 86.5 | 84 | 84.6581 | 71.8 | 132.9494 | 87.7 | 85.1 | 85 | 95.7643 | 85.2807
| 1 | 5 | 70.4 | 93.432 | 100.5 | 88.9 | 101.7 | 83 | 84 | 87.9 | 85.8 | 81.277 | 85.5 | 88 | 83.0565 | 73.1 | 134.1586 | 84.9 | 84.6 | 82.7 | 89.9387 | 87
| 1 | 6 | 71.9 | 93.4 | 102.5 | 88 | 102.94 | 85 | 86 | 85.9 | 88 | 80.5644 | 83.5 | 90 | 83.1445 | 74.4 | 130.7104 | 84.5 | 85.4 | 82.1 | 95.2004
| Mean |  | 70.5333 | 93.592 | 101.15 | 89.1667 | 100.885 | 82 | 85.3333 | 86.45 | 86.8333 | 78.8418 | 84.6667 | 87 | 82.4315 | 73.8833 | 133.0906 | 85.3667 | 85.55 | 81.9333 | 94.4561
| Median |  | 70.5 | 93.482 | 101.5 | 88.65 | 101.6 | 82.5 | 85 | 85.9 | 87.15 | 78.6199 | 84.5 | 86.5 | 83.3457 | 74.2 | 133.2593 | 85.1 | 85.3 | 81.8 | 95.4823
| Std Dev. |  | 0.85713 | 0.27543 | 1.5732 | 1.6464 | 1.741 | 2.6077 | 1.5055 | 1.468 | 0.93524 | 1.8351 | 1.169 | 2.6833 | 4.3493 | 1.2432 | 1.275 | 1.4306 | 0.96695 | 1.9846 | 3.252
| Rel.Std.Dev. |  | 1.215% | 0.2943% | 1.555% | 1.846% | 1.726% | 3.18% | 1.764% | 1.698% | 1.077% | 2.328% | 1.381% | 3.084% | 5.276% | 1.683% | 0.958% | 1.676% | 1.13% | 2.422% | 3.443%
| PDM3 |  | -17.29% | 9.746% | 18.61% | 4.557% | 18.3% | -3.847% | 0.06177% | 1.371% | 1.821% | -7.55% | -0.72% | 2.016% | -3.341% | -13.36% | 56.06% | 0.1009% | 0.3158% | -3.925% | 10.76%
| Z-Score (Absolute) |  | 2.8 | 1.5 | 2.91 | 0.67 | 2.86 | 0.66 | 0.04 | 0.17 | 0.24 | 1.25 | 0.16 | 0.27 | 0.58 | 2.18 | 8.86 | 0.03 | 0 | 0.67 | 1.66
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Al in OREAS 782 (Indicative Value 4.98 wt.%) |  |  |  |  | Indicative
Al | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 2
2
| 1 | 1 | 5.105 |  |  | 1
| 1 | 2 | 4.7981 |  |  | 19
| 1 | 3 | 5.0333 |  |  | 16
| 1 | 4 | 5.0872 |  |  | 4.9824
| 1 | 5 | 4.9526 |  |  | 25
| 1 | 6 | 4.9181
| Mean |  | 4.9824
| Median |  | 4.993
| Std Dev. |  | 0.11627
| Rel.Std.Dev. |  | 2.334%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 782 (Indicative Value 1133 ppm) |  |  |  |  | Indicative
Ba | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 0
0
| 1 | 1 | 1156.1467 |  |  | 1
| 1 | 2 | 1118.5467 |  |  | 20
| 1 | 3 | 1114.4967 |  |  | 16
| 1 | 4 | 1156.6867 |  |  | 1132.94
| 1 | 5 | 1116.4767 |  |  | 26
| 1 | 6 | 1135.2867
| Mean |  | 1132.94
| Median |  | 1126.9167
| Std Dev. |  | 19.6282
| Rel.Std.Dev. |  | 1.732%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 782 (Indicative Value 200 ppm) |  |  |  |  | Indicative
Bi | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 0
0
| 1 | 1 | 202.152 |  |  | 1
| 1 | 2 | 198.8682 |  |  | 1
| 1 | 3 | 202.3827 |  |  | 16
| 1 | 4 | 200.5709 |  |  | 200.2811
| 1 | 5 | 197.0452 |  |  | 27
| 1 | 6 | 200.6675
| Mean |  | 200.2811
| Median |  | 200.6192
| Std Dev. |  | 2.0307
| Rel.Std.Dev. |  | 1.014%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ca in OREAS 782 (Indicative Value 17.01 wt.%) |  |  |  |  | Indicative
Ca | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 2
2
| 1 | 1 | 16.8525 |  |  | 1
| 1 | 2 | 16.86 |  |  | 22
| 1 | 3 | 17.1991 |  |  | 16
| 1 | 4 | 17.1543 |  |  | 17.0122
| 1 | 5 | 16.8843 |  |  | 28
| 1 | 6 | 17.1229
| Mean |  | 17.0122
| Median |  | 17.0036
| Std Dev. |  | 0.16273
| Rel.Std.Dev. |  | 0.9566%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 782 (Indicative Value 63 ppm) |  |  |  |  | Indicative
Ce | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 0
0
| 1 | 1 | 63.0644 |  |  | 1
| 1 | 2 | 63.2567 |  |  | 23
| 1 | 3 | 63.2036 |  |  | 16
| 1 | 4 | 62.1821 |  |  | 63.0155
| 1 | 5 | 63.7889 |  |  | 29
| 1 | 6 | 62.5975
| Mean |  | 63.0155
| Median |  | 63.134
| Std Dev. |  | 0.55917
| Rel.Std.Dev. |  | 0.8874%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 782 (Indicative Value 6.55 ppm) |  |  |  |  | Indicative
Cs | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 6.5244 |  |  | 1
| 1 | 2 | 6.1926 |  |  | 5
| 1 | 3 | 6.7569 |  |  | 16
| 1 | 4 | 6.3764 |  |  | 6.5524
| 1 | 5 | 6.7379 |  |  | 30
| 1 | 6 | 6.7263
| Mean |  | 6.5524
| Median |  | 6.6253
| Std Dev. |  | 0.23136
| Rel.Std.Dev. |  | 3.531%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 782 (Indicative Value 4.11 ppm) |  |  |  |  | Indicative
Dy | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 4.2318 |  |  | 1
| 1 | 2 | 4.0491 |  |  | 6
| 1 | 3 | 4.1259 |  |  | 16
| 1 | 4 | 3.8942 |  |  | 4.1123
| 1 | 5 | 4.1805 |  |  | 31
| 1 | 6 | 4.1923
| Mean |  | 4.1123
| Median |  | 4.1532
| Std Dev. |  | 0.12414
| Rel.Std.Dev. |  | 3.019%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 782 (Indicative Value 2.59 ppm) |  |  |  |  | Indicative
Er | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 2.4975 |  |  | 1
| 1 | 2 | 2.4726 |  |  | 26
| 1 | 3 | 2.59 |  |  | 16
| 1 | 4 | 2.9349 |  |  | 2.5944
| 1 | 5 | 2.5136 |  |  | 32
| 1 | 6 | 2.5574
| Mean |  | 2.5944
| Median |  | 2.5355
| Std Dev. |  | 0.1721
| Rel.Std.Dev. |  | 6.634%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 782 (Indicative Value 1.05 ppm) |  |  |  |  | Indicative
Eu | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 1.0167 |  |  | 1
| 1 | 2 | 1.0119 |  |  | 7
| 1 | 3 | 1.0325 |  |  | 16
| 1 | 4 | 1.1316 |  |  | 1.0531
| 1 | 5 | 1.087 |  |  | 33
| 1 | 6 | 1.0389
| Mean |  | 1.0531
| Median |  | 1.0357
| Std Dev. |  | 0.046812
| Rel.Std.Dev. |  | 4.445%
| PDM3 |  | -4.663e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 782 (Indicative Value 1.56 wt.%) |  |  |  |  | Indicative
Fe | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 2
2
| 1 | 1 | 1.5657 |  |  | 1
| 1 | 2 | 1.5216 |  |  | 28
| 1 | 3 | 1.5944 |  |  | 16
| 1 | 4 | 1.6248 |  |  | 1.5643
| 1 | 5 | 1.5424 |  |  | 34
| 1 | 6 | 1.5371
| Mean |  | 1.5643
| Median |  | 1.5541
| Std Dev. |  | 0.039015
| Rel.Std.Dev. |  | 2.494%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 782 (Indicative Value 4.7 ppm) |  |  |  |  | Indicative
Gd | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 4.7508 |  |  | 1
| 1 | 2 | 4.7157 |  |  | 9
| 1 | 3 | 4.5989 |  |  | 16
| 1 | 4 | 4.6588 |  |  | 4.6956
| 1 | 5 | 4.7555 |  |  | 25
| 1 | 6 | 4.6939
| Mean |  | 4.6956
| Median |  | 4.7048
| Std Dev. |  | 0.059577
| Rel.Std.Dev. |  | 1.269%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 782 (Indicative Value 0.91 ppm) |  |  |  |  | Indicative
Ho | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.89009 |  |  | 1
| 1 | 2 | 0.84246 |  |  | 20
| 1 | 3 | 0.86295 |  |  | 16
| 1 | 4 | 1.0728 |  |  | 0.90697
| 1 | 5 | 0.88041 |  |  | 26
| 1 | 6 | 0.89309
| Mean |  | 0.90697
| Median |  | 0.88525
| Std Dev. |  | 0.083405
| Rel.Std.Dev. |  | 9.196%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 782 (Indicative Value 30.4 ppm) |  |  |  |  | Indicative
La | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 1
1
| 1 | 1 | 30.6678 |  |  | 1
| 1 | 2 | 29.3346 |  |  | 21
| 1 | 3 | 30.7042 |  |  | 16
| 1 | 4 | 30.3355 |  |  | 30.3868
| 1 | 5 | 31.2406 |  |  | 27
| 1 | 6 | 30.0379
| Mean |  | 30.3868
| Median |  | 30.5016
| Std Dev. |  | 0.65423
| Rel.Std.Dev. |  | 2.153%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 782 (Indicative Value 39.2 ppm) |  |  |  |  | Indicative
Li | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 1
1
| 1 | 1 | 40.44 |  |  | 1
| 1 | 2 | 35.63 |  |  | 22
| 1 | 3 | 40.28 |  |  | 16
| 1 | 4 | 41.47 |  |  | 39.1617
| 1 | 5 | 38.2 |  |  | 28
| 1 | 6 | 38.95
| Mean |  | 39.1617
| Median |  | 39.615
| Std Dev. |  | 2.0811
| Rel.Std.Dev. |  | 5.314%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 782 (Indicative Value 0.33 ppm) |  |  |  |  | Indicative
Lu | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.33275 |  |  | 1
| 1 | 2 | 0.32245 |  |  | 23
| 1 | 3 | 0.32822 |  |  | 16
| 1 | 4 | 0.3576 |  |  | 0.33342
| 1 | 5 | 0.33682 |  |  | 29
| 1 | 6 | 0.32267
| Mean |  | 0.33342
| Median |  | 0.33048
| Std Dev. |  | 0.013112
| Rel.Std.Dev. |  | 3.933%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mg in OREAS 782 (Indicative Value 0.242 wt.%) |  |  |  |  | Indicative
Mg | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 3
3
| 1 | 1 | 0.23787 |  |  | 1
| 1 | 2 | 0.23511 |  |  | 24
| 1 | 3 | 0.24423 |  |  | 16
| 1 | 4 | 0.25587 |  |  | 0.24156
| 1 | 5 | 0.23919 |  |  | 30
| 1 | 6 | 0.2371
| Mean |  | 0.24156
| Median |  | 0.23853
| Std Dev. |  | 0.0076517
| Rel.Std.Dev. |  | 3.168%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 782 (Indicative Value 25.1 ppm) |  |  |  |  | Indicative
Nd | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 1
1
| 1 | 1 | 24.6302 |  |  | 1
| 1 | 2 | 24.6655 |  |  | 6
| 1 | 3 | 24.9791 |  |  | 16
| 1 | 4 | 26.1462 |  |  | 25.1293
| 1 | 5 | 26.1838 |  |  | 31
| 1 | 6 | 24.1708
| Mean |  | 25.1293
| Median |  | 24.8223
| Std Dev. |  | 0.84279
| Rel.Std.Dev. |  | 3.354%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 782 (Indicative Value 6.77 ppm) |  |  |  |  | Indicative
Pr | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 6.7755 |  |  | 1
| 1 | 2 | 6.4618 |  |  | 9
| 1 | 3 | 6.7854 |  |  | 16
| 1 | 4 | 6.7751 |  |  | 6.7728
| 1 | 5 | 7.1869 |  |  | 32
| 1 | 6 | 6.6521
| Mean |  | 6.7728
| Median |  | 6.7753
| Std Dev. |  | 0.23788
| Rel.Std.Dev. |  | 3.512%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Si in OREAS 782 (Indicative Value 23.27 wt.%) |  |  |  |  | Indicative
Si | Round | Replicate | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 22 | wt.%
|  |  | PF*OES | PF*OES | BF*OES | 2
2
| 1 | 1 | 20.971 | 24.68 | 23.7 | 1
| 1 | 2 | 21.0229 | 25.38 | 24.17 | 27
| 1 | 3 | 21.2159 | 24.5 | 23.98 | 16
| 1 | 4 | 21.1299 | 24.59 | 23.7 | 23.2654
| 1 | 5 | 22.2577 | 24.68 | 22.63 | 33
| 1 | 6 | 21.4636 | 24.4 | 24.31
| Mean |  | 21.3435 | 24.705 | 23.7483
| Median |  | 21.1729 | 24.635 | 23.84
| Std Dev. |  | 0.48034 | 0.34789 | 0.60038
| Rel.Std.Dev. |  | 2.251% | 1.408% | 2.528%
| PDM3 |  | -8.261% | 6.188% | 2.076%
| Z-Score (Absolute) |  | 1.69 | 0.67 | 0
| Analytical results for Sm in OREAS 782 (Indicative Value 4.73 ppm) |  |  |  |  | Indicative
Sm | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 4.6943 |  |  | 1
| 1 | 2 | 4.6025 |  |  | 28
| 1 | 3 | 4.6799 |  |  | 16
| 1 | 4 | 15.2429 |  |  | 4.726
| 1 | 5 | 4.8481 |  |  | 34
| 1 | 6 | 4.805
| Mean |  | 6.4788
| Median |  | 4.7497
| Std Dev. |  | 4.2944
| Rel.Std.Dev. |  | 66.28%
| PDM3 |  | 37.09%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 782 (Indicative Value 36.6 ppm) |  |  |  |  | Indicative
Sr | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 1
1
| 1 | 1 | 34.5387 |  |  | 1
| 1 | 2 | 36.8887 |  |  | 4
| 1 | 3 | 34.5137 |  |  | 16
| 1 | 4 | 35.2467 |  |  | 36.591
| 1 | 5 | 39.0217 |  |  | 25
| 1 | 6 | 39.3367
| Mean |  | 36.591
| Median |  | 36.0677
| Std Dev. |  | 2.185
| Rel.Std.Dev. |  | 5.971%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 782 (Indicative Value 0.71 ppm) |  |  |  |  | Indicative
Tb | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.71587 |  |  | 1
| 1 | 2 | 0.68271 |  |  | 20
| 1 | 3 | 0.70459 |  |  | 16
| 1 | 4 | 0.74675 |  |  | 0.71244
| 1 | 5 | 0.71775 |  |  | 26
| 1 | 6 | 0.70698
| Mean |  | 0.71244
| Median |  | 0.71143
| Std Dev. |  | 0.02094
| Rel.Std.Dev. |  | 2.939%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 782 (Indicative Value 8.08 ppm) |  |  |  |  | Indicative
Th | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 7.9625 |  |  | 1
| 1 | 2 | 7.8153 |  |  | 7
| 1 | 3 | 7.9225 |  |  | 16
| 1 | 4 | 8.5478 |  |  | 8.0807
| 1 | 5 | 8.1901 |  |  | 27
| 1 | 6 | 8.0457
| Mean |  | 8.0807
| Median |  | 8.0041
| Std Dev. |  | 0.26108
| Rel.Std.Dev. |  | 3.231%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 782 (Indicative Value 0.29 wt.%) |  |  |  |  | Indicative
Ti | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | wt.%
|  |  | PF*OES |  |  | 3
3
| 1 | 1 | 0.29418 |  |  | 1
| 1 | 2 | 0.28054 |  |  | 22
| 1 | 3 | 0.29215 |  |  | 16
| 1 | 4 | 0.29161 |  |  | 0.2904
| 1 | 5 | 0.29331 |  |  | 28
| 1 | 6 | 0.29062
| Mean |  | 0.2904
| Median |  | 0.29188
| Std Dev. |  | 0.0049908
| Rel.Std.Dev. |  | 1.719%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 782 (Indicative Value 0.3 ppm) |  |  |  |  | Indicative
Tl | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.30705 |  |  | 1
| 1 | 2 | 0.30421 |  |  | 8
| 1 | 3 | 0.29352 |  |  | 16
| 1 | 4 | 0.31246 |  |  | 0.30213
| 1 | 5 | 0.2891 |  |  | 29
| 1 | 6 | 0.30643
| Mean |  | 0.30213
| Median |  | 0.30532
| Std Dev. |  | 0.0089186
| Rel.Std.Dev. |  | 2.952%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 782 (Indicative Value 0.36 ppm) |  |  |  |  | Indicative
Tm | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 0.34917 |  |  | 1
| 1 | 2 | 0.33957 |  |  | 24
| 1 | 3 | 0.33905 |  |  | 16
| 1 | 4 | 0.39036 |  |  | 0.35602
| 1 | 5 | 0.35399 |  |  | 30
| 1 | 6 | 0.364
| Mean |  | 0.35602
| Median |  | 0.35158
| Std Dev. |  | 0.019247
| Rel.Std.Dev. |  | 5.406%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 782 (Indicative Value 63 ppm) |  |  |  |  | Indicative
V | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 0
0
| 1 | 1 | 61.379 |  |  | 1
| 1 | 2 | 60.992 |  |  | 10
| 1 | 3 | 64.268 |  |  | 16
| 1 | 4 | 63.602 |  |  | 62.9495
| 1 | 5 | 63.179 |  |  | 31
| 1 | 6 | 64.277
| Mean |  | 62.9495
| Median |  | 63.3905
| Std Dev. |  | 1.4337
| Rel.Std.Dev. |  | 2.278%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 782 (Indicative Value 166 ppm) |  |  |  |  | Indicative
W | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 0
0
| 1 | 1 | 167.7048 |  |  | 1
| 1 | 2 | 166.3428 |  |  | 26
| 1 | 3 | 164.4771 |  |  | 16
| 1 | 4 | 167.0827 |  |  | 166.0043
| 1 | 5 | 163.4397 |  |  | 32
| 1 | 6 | 166.9785
| Mean |  | 166.0043
| Median |  | 166.6606
| Std Dev. |  | 1.675
| Rel.Std.Dev. |  | 1.009%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 782 (Indicative Value 2.31 ppm) |  |  |  |  | Indicative
Yb | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*MS |  |  | 2
2
| 1 | 1 | 2.2589 |  |  | 1
| 1 | 2 | 2.2629 |  |  | 27
| 1 | 3 | 2.2841 |  |  | 16
| 1 | 4 | 2.4055 |  |  | 2.3057
| 1 | 5 | 2.3623 |  |  | 33
| 1 | 6 | 2.2607
| Mean |  | 2.3057
| Median |  | 2.2735
| Std Dev. |  | 0.062744
| Rel.Std.Dev. |  | 2.721%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 782 (Indicative Value 46 ppm) |  |  |  |  | Indicative
Zn | Round | Replicate | Lab |  |  | 1
| No | No | 18 |  |  | ppm
|  |  | PF*OES |  |  | 1
1
| 1 | 1 | 53.8125 |  |  | 1
| 1 | 2 | 60.9 |  |  | 12
| 1 | 3 | 37.8 |  |  | 16
| 1 | 4 | 25.8615 |  |  | 46.0425
| 1 | 5 | 62.0235 |  |  | 34
| 1 | 6 | 35.8575
| Mean |  | 46.0425
| Median |  | 45.8063
| Std Dev. |  | 14.9362
| Rel.Std.Dev. |  | 32.44%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.