## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 340
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
3-Acid Digestion (no HF)
Ag, Silver (ppm) | 14.534 | 13.3519 | 15.7161 | 14.1043 | 14.9637
Borate Fusion XRF
Cu, Copper (ppm) | 51.5556 | 34.4943 | 68.6168 | IND | IND
Fe, Iron (wt.%) | 8.7297 | 8.6186 | 8.8408 | 8.6644 | 8.795
MgO, Magnesium oxide (wt.%) | 1.339 | 1.3054 | 1.3726 | 1.3214 | 1.3565
Mn, Manganese (wt.%) | 1.4173 | 1.3857 | 1.449 | 1.3983 | 1.4364
Pb, Lead (wt.%) | 0.59725 | 0.57904 | 0.61547 | 0.58844 | 0.60607
S, Sulphur (wt.%) | 0.66534 | 0.65074 | 0.67994 | 0.65094 | 0.67974
SiO2, Silicon dioxide (wt.%) | 63.7896 | 63.2281 | 64.351 | 63.5871 | 63.992
Zn, Zinc (wt.%) | 0.37663 | 0.37008 | 0.38318 | 0.37086 | 0.38239
Infrared Combustion
C, Carbon (wt.%) | 0.55291 | 0.53123 | 0.57459 | 0.53939 | 0.56643
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 340
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
3-Acid Digestion (no HF)
As | ppm | 506.9 | Bi | ppm | 0.51 | Co | ppm | 13
Borate Fusion XRF
As | ppm | 460 | Sb | ppm | < 50
Thermogravimetry
LOI1000 | wt.% | 2.5433
Infrared Combustion
C-(Graphite) | wt.% | 0.148 | C-(Organic) | wt.% | 0.126 | S | wt.% | 0.672
Instrumental Neutron Activation Analysis
Ag | ppm | 4.7 | Hf | ppm | < 1 | Se | ppm | < 3
As | ppm | 494.8 | Hg | ppm | < 1 | Sm | ppm | 5.42
Au | ppm | 0.0162 | In | ppm | 0.5 | Sn | ppm | < 200
Ba | ppm | < 50 | Ir | ppm | < 0.005 | Sr | ppm | < 500
Br | ppm | < 0.5 | La | ppm | 35.66 | Ta | ppm | < 0.5
Ca | wt.% | 6.2 | Lu | ppm | 0.101 | Tb | ppm | < 0.5
Ce | ppm | 68.6 | Mo | ppm | 4.5 | Th | ppm | 4.18
Cl | ppm | 500 | Na | wt.% | 0.158 | U | ppm | 4.72
Co | ppm | 11.2 | Nd | ppm | 23 | W | ppm | 13
Cr | ppm | 117 | Ni | ppm | < 20 | Yb | ppm | 1.44
Cs | ppm | < 1 | Rb | ppm | 26.9 | Zn | wt.% | 0.4492
Eu | ppm | 9.86 | Sb | ppm | 46.36
Fe | wt.% | 8.202 | Sc | ppm | 3.94
Ion Selective Electrode
F | wt.% | 1.014
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 340
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
3-Acid Digestion (no HF)
Ag, ppm | 14.534 | 0.55132 | 13.4314 | 15.6366 | 12.8801 | 16.1879 | 3.793% | 7.587% | 11.38% | 13.8073 | 15.2607
Borate Fusion XRF
Cu, ppm | 51.5556 | 13.2537 | 25.0481 | 78.063 | 11.7944 | 91.3167 | 25.71% | 51.42% | 77.12% | 48.9778 | 54.1333
Fe, wt.% | 8.7297 | 0.12151 | 8.4867 | 8.9727 | 8.3652 | 9.0942 | 1.392% | 2.784% | 4.176% | 8.2932 | 9.1662
MgO, wt.% | 1.339 | 0.031408 | 1.2762 | 1.4018 | 1.2447 | 1.4332 | 2.346% | 4.691% | 7.037% | 1.272 | 1.4059
Mn, wt.% | 1.4173 | 0.034792 | 1.3478 | 1.4869 | 1.313 | 1.5217 | 2.455% | 4.909% | 7.364% | 1.3465 | 1.4882
Pb, wt.% | 0.59725 | 0.01808 | 0.5611 | 0.63341 | 0.54302 | 0.65149 | 3.027% | 6.054% | 9.081% | 0.56739 | 0.62712
S, wt.% | 0.66534 | 0.011078 | 0.64318 | 0.6875 | 0.63211 | 0.69857 | 1.665% | 3.33% | 4.995% | 0.63207 | 0.69861
SiO2, wt.% | 63.7896 | 0.46257 | 62.8644 | 64.7147 | 62.4018 | 65.1773 | 0.7252% | 1.45% | 2.175% | 60.6001 | 66.979
Zn, wt.% | 0.37663 | 0.0068703 | 0.36289 | 0.39037 | 0.35602 | 0.39724 | 1.824% | 3.648% | 5.472% | 0.3578 | 0.39546
Infrared Combustion
C, wt.% | 0.55291 | 0.023395 | 0.50612 | 0.5997 | 0.48273 | 0.6231 | 4.231% | 8.462% | 12.69% | 0.52527 | 0.58056
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 340
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma optical emission spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
INAA | instrumental neutron activation analysis
INAA-5S | instrumental neutron activation analysis for short lived isotopes
IRC | infrared combustion
IRC-CS1 | HCl (~50%) digestion to remove carbonates followed by roasting at ~425° C to remove organic carbon; residue determined by infrared combustion
ISE | ion specific electrode
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 340
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Cardiff, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Perth, WA, Australia
Nagrom, Perth, WA, Australia
SGS Australia Mineral Services, Perth, WA, Australia

## 3-Acid
| Analytical results for Ag in OREAS 340 (Certified Value 14.5 ppm) | [11 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | ppm
|  |  | AR*MS | 3A*OES | 3A*MS | 3A*MS | 4A*MS | 3A*OES | 3A*MS | 3A*MS | 3A*MS | AR*MS | 1
|  |  | 01: 10g; 10: 0.5g | 2
| 1 | 1 | 12.95 | 14 | 14 | 15.5 | 14.5 | 15 | 15 | 14.5 | 14.27 | 14.85 | 1
| 1 | 2 | 13.05 | 14 | 14 | 15 | 14 | 15 | 14 | 14.6 | 13.71 | 15.85 | 9
| 1 | 3 | 13 | 14 | 15 | 14.9 | 14 | 15 | 14 | 14.8 | 14.06 | 14 | 16
| 1 | 4 | 12.85 | 15 | 15 | 14.6 | 14 | 14 | 14 | 14.6 | 13.91 | 14.7 | 14.534
| 1 | 5 | 12.9 | 14 | 14 | 15.6 | 14 | 15 | 13 | 14.7 | 14.5 | 14.2 | 7
| Mean |  | 12.95 | 14.2 | 14.4 | 15.12 | 14.1 | 14.8 | 14 | 14.64 | 14.09 | 14.72
| Median |  | 12.95 | 14 | 14 | 15 | 14 | 15 | 14 | 14.6 | 14.06 | 14.7
| Std Dev. |  | 0.079057 | 0.44721 | 0.54772 | 0.42071 | 0.22361 | 0.44721 | 0.70711 | 0.11402 | 0.30749 | 0.72163
| Rel.Std.Dev. |  | 0.6105% | 3.149% | 3.804% | 2.782% | 1.586% | 3.022% | 5.051% | 0.7788% | 2.182% | 4.902%
| PDM3 |  | -10.9% | -2.298% | -0.922% | 4.032% | -2.986% | 1.83% | -3.674% | 0.7293% | -3.055% | 1.28%
| Z-Score (Absolute) |  | 3.04 | NA | NA | 1.61 | 0.58 | NA | NA | 0.58 | 0.6 | 0.75
Results from laboratories 1.01, 1.02, 1.03, 1.06 and 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for As in OREAS 340 (Indicative Value 507 ppm) | [11 empty] | Indicative
As | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 04; 02: 05 | ppm
|  |  | 01: 3A*MS; 02: 4A*MS | 0
0
| 1 | 1 | 01: 535; 02: 515 | 1
| 1 | 2 | 01: 541; 02: 481 | 1
| 1 | 3 | 01: 544; 02: 482 | 16
| 1 | 4 | 01: 538; 02: 455 | 506.9
| 1 | 5 | 01: 530; 02: 448 | 7
| Mean |  | 01: 537.6; 02: 476.2
| Median |  | 01: 538; 02: 481
| Std Dev. |  | 01: 5.4129; 02: 26.4896
| Rel.Std.Dev. |  | 01: 1.007%; 02: 5.563%
| PDM3 |  | 01: 6.056%; 02: -6.056%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
aqua regia digestion, 3-acid digestion and 4-acid digestion method results have been combined for certification.
| Analytical results for Bi in OREAS 340 (Indicative Value 0.51 ppm) | [11 empty] | Indicative
Bi | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 04; 02: 05 | ppm
|  |  | 01: 3A*MS; 02: 4A*MS | 2
2
| 1 | 1 | 01: 0.5; 02: 0.7 | 1
| 1 | 2 | 01: 0.5; 02: 0.4 | 1
| 1 | 3 | 01: 0.5; 02: 0.7 | 16
| 1 | 4 | 01: 0.5; 02: 0.4 | 0.51
| 1 | 5 | 01: 0.5; 02: 0.4 | 7
| Mean |  | 01: 0.5; 02: 0.52
| Median |  | 01: 0.5; 02: 0.4
| Std Dev. |  | 01: 0; 02: 0.16432
| Rel.Std.Dev. |  | 01: 0%; 02: 31.6%
| PDM3 |  | 01: -1.961%; 02: 1.961%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
aqua regia digestion, 3-acid digestion and 4-acid digestion method results have been combined for certification.
| Analytical results for Co in OREAS 340 (Indicative Value 13 ppm) | [11 empty] | Indicative
Co | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 04; 02: 05 | ppm
|  |  | 01: 3A*MS; 02: 4A*MS | 1
1
| 1 | 1 | 01: 13; 02: 13 | 1
| 1 | 2 | 01: 13; 02: 14 | 17
| 1 | 3 | 01: 13; 02: 13 | 16
| 1 | 4 | 01: 12.5; 02: 13 | 13
| 1 | 5 | 01: 12.5; 02: 13 | 7
| Mean |  | 01: 12.8; 02: 13.2
| Median |  | 01: 13; 02: 13
| Std Dev. |  | 01: 0.27386; 02: 0.44721
| Rel.Std.Dev. |  | 01: 2.14%; 02: 3.388%
| PDM3 |  | 01: -1.538%; 02: 1.538%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
aqua regia digestion, 3-acid digestion and 4-acid digestion method results have been combined for certification.
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C34: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C52: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C70: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for As in OREAS 340 (Indicative Value 460 ppm) | [11 empty] | Indicative
As | Round | Replicate | Lab | [9 empty] | 1
| No | No | 01 | [9 empty] | ppm
|  |  | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 500 | [9 empty] | 1
| 1 | 2 | 400 | [9 empty] | 1
| 1 | 3 | 500 | [9 empty] | 16
| 1 | 4 | 400 | [9 empty] | 460
| 1 | 5 | 500 | [9 empty] | 9
| Mean |  | 460
| Median |  | 500
| Std Dev. |  | 54.7723
| Rel.Std.Dev. |  | 11.91%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 340 (Certified Value 52 ppm) | [11 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 0
0
| 1 | 1 | 60 | < 50 | 50 | 40 | 50 | 40 | < 100 | 72 | 50 | < 100 | 1
| 1 | 2 | 90 | < 50 | 50 | 50 | 40 | 30 | < 100 | 33 | 60 | < 100 | #N/A
| 1 | 3 | 50 | < 50 | < 50 | 50 | 60 | 40 | 300 | 30 | 50 | < 100 | 16
| 1 | 4 | 60 | < 50 | < 50 | 40 | 60 | 30 | 200 | 20 | 60 | < 100 | 51.5556
| 1 | 5 | 80 | < 50 | 60 | 50 | 40 | 40 | 1300 | 22 | 60 | < 100 | 9
| Mean |  | 68 |  | 53.3333 | 46 | 50 | 36 | 600 | 35.4 | 56
| Median |  | 60 |  | 50 | 50 | 50 | 40 | 300 | 30 | 60
| Std Dev. |  | 16.4317 |  | 5.7735 | 5.4772 | 10 | 5.4772 | 608.2763 | 21.1613 | 5.4772
| Rel.Std.Dev. |  | 24.16% |  | 10.83% | 11.91% | 20% | 15.21% | 101.4% | 59.78% | 9.781%
| PDM3 |  | 31.9% |  | 3.448% | -10.78% | -3.017% | -30.17% | 1064% | -31.34% | 8.621%
| Z-Score (Absolute) |  | 1.48 | 1.42 | 0.27 | 0 | 0.27 | 0.67 | NA | 0.71 | 0.67 | 0.27
Results from laboratory 7 were removed due to their 100 ppm reading resolution.
| Analytical results for Fe in OREAS 340 (Certified Value 8.73 wt.%) | [11 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 8.67 | 8.6 | 8.812 | 8.72 | 8.85 | 8.629 | 8.68 | 8.7 | 8.64 | 8.9872 | 1
| 1 | 2 | 8.64 | 8.64 | 8.762 | 8.74 | 8.82 | 8.63 | 8.74 | 8.77 | 8.6 | 8.9448 | 21
| 1 | 3 | 8.6 | 8.53 | 8.69 | 8.79 | 8.85 | 8.612 | 8.96 | 8.76 | 8.6 | 8.96 | 16
| 1 | 4 | 8.6 | 8.57 | 8.724 | 8.76 | 8.89 | 8.632 | 8.63 | 8.8 | 8.67 | 8.9983 | 8.7297
| 1 | 5 | 8.67 | 8.67 | 8.786 | 8.73 | 8.82 | 8.645 | 8.67 | 8.71 | 8.6 | 8.983 | 10
| Mean |  | 8.636 | 8.602 | 8.7548 | 8.748 | 8.846 | 8.6296 | 8.736 | 8.748 | 8.622 | 8.9747
| Median |  | 8.64 | 8.6 | 8.762 | 8.74 | 8.85 | 8.63 | 8.68 | 8.76 | 8.6 | 8.983
| Std Dev. |  | 0.035071 | 0.055408 | 0.048592 | 0.027749 | 0.02881 | 0.01176 | 0.13126 | 0.042071 | 0.031937 | 0.02175
| Rel.Std.Dev. |  | 0.4061% | 0.6441% | 0.555% | 0.3172% | 0.3257% | 0.1363% | 1.503% | 0.4809% | 0.3704% | 0.2423%
| PDM3 |  | -1.073% | -1.463% | 0.2875% | 0.2096% | 1.332% | -1.147% | 0.0721% | 0.2096% | -1.234% | 2.806%
| Z-Score (Absolute) |  | 0.68 | 0.9 | 0.08 | 0.04 | 0.67 | 0.72 | 0.04 | 0.04 | 0.77 | 1.49
| Analytical results for MgO in OREAS 340 (Certified Value 1.34 wt.%) | [11 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.34 | 1.4 | 1.309 | 1.35 | 1.33 | 1.32 | 1.25 | 1.31 | 1.36 | 1.3246 | 1
| 1 | 2 | 1.36 | 1.41 | 1.32 | 1.36 | 1.35 | 1.314 | 1.25 | 1.31 | 1.36 | 1.3137 | #N/A
| 1 | 3 | 1.35 | 1.4 | 1.287 | 1.36 | 1.33 | 1.324 | 1.2 | 1.32 | 1.34 | 1.3182 | 16
| 1 | 4 | 1.36 | 1.38 | 1.269 | 1.35 | 1.36 | 1.322 | 1.2 | 1.31 | 1.36 | 1.3095 | 1.339
| 1 | 5 | 1.36 | 1.42 | 1.303 | 1.35 | 1.33 | 1.314 | 1.21 | 1.32 | 1.34 | 1.3257 | 11
| Mean |  | 1.354 | 1.402 | 1.2976 | 1.354 | 1.34 | 1.3188 | 1.222 | 1.314 | 1.352 | 1.3183
| Median |  | 1.36 | 1.4 | 1.303 | 1.35 | 1.33 | 1.32 | 1.21 | 1.31 | 1.36 | 1.3182
| Std Dev. |  | 0.0089443 | 0.014832 | 0.019945 | 0.0054772 | 0.014142 | 0.0046043 | 0.025884 | 0.0054772 | 0.010954 | 0.0069472
| Rel.Std.Dev. |  | 0.6606% | 1.058% | 1.537% | 0.4045% | 1.055% | 0.3491% | 2.118% | 0.4168% | 0.8102% | 0.527%
| PDM3 |  | 1.122% | 4.707% | -3.09% | 1.122% | 0.07684% | -1.506% | -8.736% | -1.865% | 0.9731% | -1.541%
| Z-Score (Absolute) |  | 0.7 | 2.07 | 0.91 | 0.7 | 0.3 | 0.3 | 3.07 | 0.44 | 0.65 | 0.32
| Analytical results for Mn in OREAS 340 (Certified Value 1.42 wt.%) | [11 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.41 | 1.38 | 1.406 | 1.44 | 1.46 | 1.464 | 1.43 | 1.44 | 1.4 | 1.3589 | 1
| 1 | 2 | 1.4 | 1.38 | 1.39 | 1.44 | 1.44 | 1.469 | 1.43 | 1.45 | 1.4 | 1.357 | #N/A
| 1 | 3 | 1.4 | 1.36 | 1.374 | 1.45 | 1.46 | 1.472 | 1.41 | 1.44 | 1.4 | 1.3636 | 16
| 1 | 4 | 1.4 | 1.38 | 1.392 | 1.45 | 1.46 | 1.462 | 1.44 | 1.46 | 1.4 | 1.3677 | 1.4173
| 1 | 5 | 1.42 | 1.4 | 1.399 | 1.44 | 1.46 | 1.455 | 1.43 | 1.44 | 1.38 | 1.3569 | 12
| Mean |  | 1.406 | 1.38 | 1.3922 | 1.444 | 1.456 | 1.4644 | 1.428 | 1.446 | 1.396 | 1.3608
| Median |  | 1.4 | 1.38 | 1.392 | 1.44 | 1.46 | 1.464 | 1.43 | 1.44 | 1.4 | 1.3589
| Std Dev. |  | 0.0089443 | 0.014142 | 0.011967 | 0.0054772 | 0.0089443 | 0.0065803 | 0.010954 | 0.0089443 | 0.0089443 | 0.0047092
| Rel.Std.Dev. |  | 0.6362% | 1.025% | 0.8595% | 0.3793% | 0.6143% | 0.4493% | 0.7671% | 0.6186% | 0.6407% | 0.3461%
| PDM3 |  | -0.8002% | -2.635% | -1.774% | 1.881% | 2.727% | 3.32% | 0.752% | 2.022% | -1.506% | -3.988%
| Z-Score (Absolute) |  | 0.26 | 0.89 | 0.6 | 0.65 | 0.94 | 1.14 | 0.26 | 0.7 | 0.51 | 1.35
| Analytical results for Pb in OREAS 340 (Certified Value 0.597 wt.%) | [11 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.594 | 0.614 | 0.553 | 0.587 | 0.629 | 0.645 | 0.57 | 0.6 | 0.596 | 0.5932 | 1
| 1 | 2 | 0.596 | 0.616 | 0.532 | 0.583 | 0.633 | 0.647 | 0.58 | 0.6 | 0.594 | 0.5894 | #N/A
| 1 | 3 | 0.588 | 0.605 | 0.516 | 0.601 | 0.624 | 0.644 | 0.56 | 0.61 | 0.595 | 0.5921 | 16
| 1 | 4 | 0.594 | 0.613 | 0.542 | 0.594 | 0.623 | 0.642 | 0.56 | 0.6 | 0.598 | 0.5968 | 0.59725
| 1 | 5 | 0.602 | 0.614 | 0.517 | 0.585 | 0.626 | 0.645 | 0.55 | 0.6 | 0.59 | 0.5947 | 13
| Mean |  | 0.5948 | 0.6124 | 0.532 | 0.59 | 0.627 | 0.6446 | 0.564 | 0.602 | 0.5946 | 0.59324
| Median |  | 0.594 | 0.614 | 0.532 | 0.587 | 0.626 | 0.645 | 0.56 | 0.6 | 0.595 | 0.5932
| Std Dev. |  | 0.00502 | 0.0042778 | 0.015984 | 0.0074162 | 0.004062 | 0.0018166 | 0.011402 | 0.0044721 | 0.0029665 | 0.0027772
| Rel.Std.Dev. |  | 0.844% | 0.6985% | 3.005% | 1.257% | 0.6478% | 0.2818% | 2.022% | 0.7429% | 0.4989% | 0.4681%
| PDM3 |  | -0.411% | 2.536% | -10.93% | -1.215% | 4.98% | 7.927% | -5.568% | 0.7945% | -0.4445% | -0.6722%
| Z-Score (Absolute) |  | 0.01 | 0.95 | 3.38 | 0.25 | 1.74 | 2.69 | 1.66 | 0.39 | 0.01 | 0.08
| Analytical results for S in OREAS 340 (Certified Value 0.665 wt.%) | [11 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 10 | 11 |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.67 | 0.68 | 0.67 | 0.67 | 0.717 | 0.682 | 0.66 | 0.67 | 0.6523 |  | 1
| 1 | 2 | 0.66 | 0.68 | 0.657 | 0.67 | 0.702 | 0.686 | 0.65 | 0.67 | 0.6472 |  | 7
| 1 | 3 | 0.67 | 0.68 | 0.653 | 0.68 | 0.69 | 0.655 | 0.66 | 0.67 | 0.6519 |  | 16
| 1 | 4 | 0.66 | 0.68 | 0.671 | 0.67 | 0.685 | 0.658 | 0.65 | 0.67 | 0.6559 |  | 0.66534
| 1 | 5 | 0.66 | 0.67 | 0.667 | 0.68 | 0.668 | 0.659 | 0.64 | 0.67 | 0.6583 |  | 14
| Mean |  | 0.664 | 0.678 | 0.6636 | 0.674 | 0.6924 | 0.668 | 0.652 | 0.67 | 0.65312
| Median |  | 0.66 | 0.68 | 0.667 | 0.67 | 0.69 | 0.659 | 0.65 | 0.67 | 0.6523
| Std Dev. |  | 0.0054772 | 0.0044721 | 0.0081117 | 0.0054772 | 0.018393 | 0.014748 | 0.0083666 | 0 | 0.0042358
| Rel.Std.Dev. |  | 0.8249% | 0.6596% | 1.222% | 0.8126% | 2.656% | 2.208% | 1.283% | 0% | 0.6485%
| PDM3 |  | -0.2014% | 1.903% | -0.2615% | 1.302% | 4.067% | 0.3998% | -2.005% | 0.7004% | -1.837%
| Z-Score (Absolute) |  | 0.45 | 1.12 | 0.49 | 0.67 | 2.74 | 0 | 1.8 | 0.22 | 1.67
| Analytical results for Sb in OREAS 340 (Indicative Value < 50 ppm) | [11 empty] | Indicative
Sb | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 01 |  | ppm
|  |  | 01: OxBF*XRF |  | 0
0
| 1 | 1 | 01: < 50 |  | 1
| 1 | 2 | 01: 50 |  | 3
| 1 | 3 | 01: < 50 |  | 16
| 1 | 4 | 01: < 50 |  | < 50
| 1 | 5 | 01: < 50 |  | 9
| Mean |  | 01: 50
| Median |  | 01: 50
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for SiO2 in OREAS 340 (Certified Value 63.79 wt.%) | [11 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 2
3
| 1 | 1 | 63.7 | 64.4 | 64.04 | 64.24 | 64.38 | 63.152 | 63.3 | 63.13 | 63.8 | 63.869 | 1
| 1 | 2 | 64 | 64.9 | 63.589 | 64.39 | 64.31 | 63.224 | 63.4 | 63.4 | 63.6 | 63.495 | #N/A
| 1 | 3 | 63.2 | 64.4 | 62.991 | 64.23 | 64.26 | 63.189 | 62.3 | 63.51 | 63.4 | 63.675 | 16
| 1 | 4 | 63.6 | 64.4 | 63.288 | 64.25 | 64.35 | 63.33 | 63.4 | 63.69 | 63.9 | 63.924 | 63.7896
| 1 | 5 | 64.2 | 65.4 | 63.781 | 64.27 | 64.33 | 63.252 | 63.5 | 63.31 | 63.7 | 63.904 | 15
| Mean |  | 63.74 | 64.7 | 63.5378 | 64.276 | 64.326 | 63.2294 | 63.18 | 63.408 | 63.68 | 63.7734
| Median |  | 63.7 | 64.4 | 63.589 | 64.25 | 64.33 | 63.224 | 63.4 | 63.4 | 63.7 | 63.869
| Std Dev. |  | 0.38471 | 0.44721 | 0.4109 | 0.065422 | 0.045056 | 0.067608 | 0.49699 | 0.21029 | 0.19235 | 0.18444
| Rel.Std.Dev. |  | 0.6036% | 0.6912% | 0.6467% | 0.1018% | 0.07004% | 0.1069% | 0.7866% | 0.3316% | 0.3021% | 0.2892%
| PDM3 |  | -0.07769% | 1.427% | -0.3947% | 0.7626% | 0.841% | -0.8781% | -0.9556% | -0.5982% | -0.1718% | -0.02533%
| Z-Score (Absolute) |  | 0.05 | 1.71 | 0.3 | 0.98 | 1.06 | 0.83 | 0.91 | 0.52 | 0.05 | 0.11
| Analytical results for Zn in OREAS 340 (Certified Value 0.377 wt.%) | [11 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.374 | 0.365 | 0.386 | 0.385 | 0.379 | 0.432 | 0.38 | 0.38 | 0.371 | 0.3799 | 1
| 1 | 2 | 0.375 | 0.367 | 0.382 | 0.386 | 0.378 | 0.434 | 0.38 | 0.38 | 0.37 | 0.3754 | 19
| 1 | 3 | 0.372 | 0.363 | 0.381 | 0.39 | 0.381 | 0.406 | 0.37 | 0.38 | 0.369 | 0.3796 | 16
| 1 | 4 | 0.375 | 0.364 | 0.382 | 0.388 | 0.381 | 0.388 | 0.37 | 0.37 | 0.375 | 0.3799 | 0.37663
| 1 | 5 | 0.378 | 0.366 | 0.386 | 0.386 | 0.38 | 0.388 | 0.37 | 0.37 | 0.37 | 0.3784 | 16
| Mean |  | 0.3748 | 0.365 | 0.3834 | 0.387 | 0.3798 | 0.4096 | 0.374 | 0.376 | 0.371 | 0.37864
| Median |  | 0.375 | 0.365 | 0.382 | 0.386 | 0.38 | 0.406 | 0.37 | 0.38 | 0.37 | 0.3796
| Std Dev. |  | 0.0021679 | 0.0015811 | 0.0024083 | 0.002 | 0.0013038 | 0.022601 | 0.0054772 | 0.0054772 | 0.0023452 | 0.0019139
| Rel.Std.Dev. |  | 0.5784% | 0.4332% | 0.6281% | 0.5168% | 0.3433% | 5.518% | 1.464% | 1.457% | 0.6321% | 0.5055%
| PDM3 |  | -0.485% | -3.087% | 1.798% | 2.754% | 0.8426% | 8.755% | -0.6974% | -0.1664% | -1.494% | 0.5346%
| Z-Score (Absolute) |  | 0.36 | 1.77 | 0.87 | 1.39 | 0.36 | 4.63 | 0.48 | 0.19 | 0.91 | 0.19
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C33: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C51: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C68: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C102: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C119: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 340 (Indicative Value 2.54 wt.%) |  |  |  |  | Indicative
LOI1000 | Round | Replicate | Lab | Lab | Lab | 1
| No | No | 04 | 06 | 07 | wt.%
|  |  | LOI*TGA | LOI*TGA | BF*XRF | 2
2
| 1 | 1 | 2.54 | 2.34 | 2.92 | 1
| 1 | 2 | 2.5 | 2.31 | 2.79 | 5
| 1 | 3 | 2.48 | 2.22 | 2.81 | 16
| 1 | 4 | 2.57 | 2.25 | 2.77 | 2.5433
| 1 | 5 | 2.56 | 2.32 | 2.77 | 11
| Mean |  | 2.53 | 2.288 | 2.812
| Median |  | 2.54 | 2.31 | 2.79
| Std Dev. |  | 0.03873 | 0.050695 | 0.06261
| Rel.Std.Dev. |  | 1.531% | 2.216% | 2.227%
| PDM3 |  | -0.5242% | -10.04% | 10.56%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.79
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 340 (Certified Value 0.553 wt.%) | [11 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.52 | 0.56 | 0.54 | 0.56 | 0.59 | 0.5 | 0.55 | 0.53 | 0.57 | 0.54 | 1
| 1 | 2 | 0.53 | 0.57 | 0.53 | 0.56 | 0.58 | 0.5 | 0.585 | 0.55 | 0.57 | 0.53 | #N/A
| 1 | 3 | 0.52 | 0.56 | 0.54 | 0.56 | 0.58 | 0.6 | 0.573 | 0.53 | 0.58 | 0.53 | 16
| 1 | 4 | 0.52 | 0.58 | 0.54 | 0.56 | 0.59 | 0.6 | 0.578 | 0.54 | 0.57 | 0.53 | 0.55291
| 1 | 5 | 0.52 | 0.57 | 0.54 | 0.54 | 0.58 | 0.6 | 0.585 | 0.5 | 0.57 | 0.53 | 18
| Mean |  | 0.522 | 0.568 | 0.538 | 0.556 | 0.584 | 0.56 | 0.5742 | 0.53 | 0.572 | 0.532
| Median |  | 0.52 | 0.57 | 0.54 | 0.56 | 0.58 | 0.6 | 0.578 | 0.53 | 0.57 | 0.53
| Std Dev. |  | 0.0044721 | 0.0083666 | 0.0044721 | 0.0089443 | 0.0054772 | 0.054772 | 0.014446 | 0.018708 | 0.0044721 | 0.0044721
| Rel.Std.Dev. |  | 0.8567% | 1.473% | 0.8313% | 1.609% | 0.9379% | 9.781% | 2.516% | 3.53% | 0.7818% | 0.8406%
| PDM3 |  | -5.591% | 2.729% | -2.697% | 0.5587% | 5.623% | 1.282% | 3.85% | -4.144% | 3.452% | -3.782%
| Z-Score (Absolute) |  | 1.26 | 0.44 | 0.67 | 0 | 1.04 | NA | 0.67 | 0.96 | 0.59 | 0.89
Results from laboratory 1.06 were removed due to their 0.1 ppm reading resolution.
| Analytical results for C-(Graphite) in OREAS 340 (Indicative Value 0.148 wt.%) | [11 empty] | Indicative
C-(Graphite) | Round | Replicate | 01: Lab | 1
| No | No | 01: 10 | wt.%
|  |  | 01: IRC-CS1 | 3
3
| 1 | 1 | 01: 0.15 | 1
| 1 | 2 | 01: 0.14 | 7
| 1 | 3 | 01: 0.15 | 16
| 1 | 4 | 01: 0.14 | 0.148
| 1 | 5 | 01: 0.16 | 13
| Mean |  | 01: 0.148
| Median |  | 01: 0.15
| Std Dev. |  | 01: 0.0083666
| Rel.Std.Dev. |  | 01: 5.653%
| PDM3 |  | 01: 2.22e-14%
Z-Score (Absolute) | bar: 01
| Analytical results for C-(Organic) in OREAS 340 (Indicative Value 0.126 wt.%) | [11 empty] | Indicative
C-(Organic) | Round | Replicate | 01: Lab | 1
| No | No | 01: 10 | wt.%
|  |  | 01: IRC | 3
3
| 1 | 1 | 01: 0.13 | 1
| 1 | 2 | 01: 0.12 | 7
| 1 | 3 | 01: 0.12 | 16
| 1 | 4 | 01: 0.12 | 0.126
| 1 | 5 | 01: 0.14 | 13
| Mean |  | 01: 0.126
| Median |  | 01: 0.12
| Std Dev. |  | 01: 0.0089443
| Rel.Std.Dev. |  | 01: 7.099%
PDM3 | bar: 01
Z-Score (Absolute) | bar: 01
| Analytical results for S in OREAS 340 (Indicative Value 0.672 wt.%) | [11 empty] | Indicative
S | Round | Replicate | 01: Lab | 1
| No | No | 01: 10 | wt.%
|  |  | 01: IRC | 3
3
| 1 | 1 | 01: 0.66 | 1
| 1 | 2 | 01: 0.67 | 7
| 1 | 3 | 01: 0.67 | 16
| 1 | 4 | 01: 0.68 | 0.672
| 1 | 5 | 01: 0.68 | 13
| Mean |  | 01: 0.672
| Median |  | 01: 0.67
| Std Dev. |  | 01: 0.0083666
| Rel.Std.Dev. |  | 01: 1.245%
PDM3 | bar: 01
Z-Score (Absolute) | bar: 01
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C34: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C51: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C68: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## INAA
| Analytical results for Ag in OREAS 340 (Indicative Value 4.7 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 7 | 1
| 1 | 2 | < 5 | 9
| 1 | 3 | < 5 | 16
| 1 | 4 | < 5 | 4.7
| 1 | 5 | 9 | 15
| Mean |  | 8
| Median |  | 8
| Std Dev. |  | 1.4142
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 70.21%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 340 (Indicative Value 495 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 0
0
| 1 | 1 | 501 | 1
| 1 | 2 | 501 | 1
| 1 | 3 | 495 | 16
| 1 | 4 | 487 | 494.8
| 1 | 5 | 490 | 16
| Mean |  | 494.8
| Median |  | 495
| Std Dev. |  | 6.3403
| Rel.Std.Dev. |  | 1.281%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Au in OREAS 340 (Indicative Value 0.016 ppm) |  |  | Indicative
Au | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 3
3
| 1 | 1 | 0.018 | 1
| 1 | 2 | 0.017 | 11
| 1 | 3 | 0.014 | 16
| 1 | 4 | 0.013 | 0.0162
| 1 | 5 | 0.019 | 17
| Mean |  | 0.0162
| Median |  | 0.017
| Std Dev. |  | 0.0025884
| Rel.Std.Dev. |  | 15.98%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 340 (Indicative Value < 50 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 0
0
| 1 | 1 | < 50 | 1
| 1 | 2 | < 50 | 12
| 1 | 3 | < 50 | 16
| 1 | 4 | < 50 | < 50
| 1 | 5 | < 50 | 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Br in OREAS 340 (Indicative Value < 0.5 ppm) |  |  | Indicative
Br | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 0.5 | 1
| 1 | 2 | < 0.5 | 13
| 1 | 3 | < 0.5 | 16
| 1 | 4 | < 0.5 | < 0.5
| 1 | 5 | < 0.5 | 19
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ca in OREAS 340 (Indicative Value 6.2 wt.%) |  |  | Indicative
Ca | Round | Replicate | Lab | 1
| No | No | 03 | wt.%
|  |  | INAA | 2
2
| 1 | 1 | 9 | 1
| 1 | 2 | 8 | 14
| 1 | 3 | 3 | 16
| 1 | 4 | 5 | 6.2
| 1 | 5 | 6 | 20
| Mean |  | 6.2
| Median |  | 6
| Std Dev. |  | 2.3875
| Rel.Std.Dev. |  | 38.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 340 (Indicative Value 69 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 0
0
| 1 | 1 | 68 | 1
| 1 | 2 | 67 | 15
| 1 | 3 | 77 | 16
| 1 | 4 | 60 | 68.6
| 1 | 5 | 71 | 21
| Mean |  | 68.6
| Median |  | 68
| Std Dev. |  | 6.1887
| Rel.Std.Dev. |  | 9.021%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 340 (Indicative Value 500 ppm) |  |  | Indicative
Cl | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA-5S | 0
0
| 1 | 1 | 600 | 1
| 1 | 2 | 800 | 16
| 1 | 3 | < 500 | 16
| 1 | 4 | 600 | 500
| 1 | 5 | < 500 | 22
| Mean |  | 666.6667
| Median |  | 600
| Std Dev. |  | 115.4701
| Rel.Std.Dev. |  | 17.32%
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 340 (Indicative Value 11.2 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | 11 | 1
| 1 | 2 | 11 | 17
| 1 | 3 | 11 | 16
| 1 | 4 | 11 | 11.2
| 1 | 5 | 12 | 23
| Mean |  | 11.2
| Median |  | 11
| Std Dev. |  | 0.44721
| Rel.Std.Dev. |  | 3.993%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 340 (Indicative Value 117 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 0
0
| 1 | 1 | 121 | 1
| 1 | 2 | 123 | 18
| 1 | 3 | 115 | 16
| 1 | 4 | 116 | 117
| 1 | 5 | 110 | 24
| Mean |  | 117
| Median |  | 116
| Std Dev. |  | 5.1478
| Rel.Std.Dev. |  | 4.4%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 340 (Indicative Value < 1 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 1 | 1
| 1 | 2 | < 1 | 19
| 1 | 3 | < 1 | 16
| 1 | 4 | < 1 | < 1
| 1 | 5 | < 1 | 25
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 340 (Indicative Value 9.86 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 9.9 | 1
| 1 | 2 | 9.7 | 20
| 1 | 3 | 10 | 16
| 1 | 4 | 9.8 | 9.86
| 1 | 5 | 9.9 | 26
| Mean |  | 9.86
| Median |  | 9.9
| Std Dev. |  | 0.11402
| Rel.Std.Dev. |  | 1.156%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 340 (Indicative Value 8.2 wt.%) |  |  | Indicative
Fe | Round | Replicate | Lab | 1
| No | No | 03 | wt.%
|  |  | INAA | 2
2
| 1 | 1 | 8.45 | 1
| 1 | 2 | 8.3 | 21
| 1 | 3 | 8.06 | 16
| 1 | 4 | 7.91 | 8.202
| 1 | 5 | 8.29 | 27
| Mean |  | 8.202
| Median |  | 8.29
| Std Dev. |  | 0.21464
| Rel.Std.Dev. |  | 2.617%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 340 (Indicative Value < 1 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 1 | 1
| 1 | 2 | < 1 | 9
| 1 | 3 | < 1 | 16
| 1 | 4 | < 1 | < 1
| 1 | 5 | < 1 | 15
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Hg in OREAS 340 (Indicative Value < 1 ppm) |  |  | Indicative
Hg | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 1 | 1
| 1 | 2 | < 1 | 10
| 1 | 3 | < 1 | 16
| 1 | 4 | < 1 | < 1
| 1 | 5 | < 1 | 16
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 340 (Indicative Value 0.5 ppm) |  |  | Indicative
In | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA-5S | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 11
| 1 | 3 | < 0.5 | 16
| 1 | 4 | < 0.5 | 0.5
| 1 | 5 | 0.8 | 17
| Mean |  | 0.66667
| Median |  | 0.6
| Std Dev. |  | 0.11547
| Rel.Std.Dev. |  | 17.32%
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Ir in OREAS 340 (Indicative Value < 0.005 ppm) |  |  | Indicative
Ir | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 3
3
| 1 | 1 | < 0.005 | 1
| 1 | 2 | < 0.005 | 12
| 1 | 3 | < 0.005 | 16
| 1 | 4 | < 0.005 | < 0.005
| 1 | 5 | < 0.005 | 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 340 (Indicative Value 35.7 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | 35.8 | 1
| 1 | 2 | 35.8 | 13
| 1 | 3 | 36.2 | 16
| 1 | 4 | 35.4 | 35.66
| 1 | 5 | 35.1 | 19
| Mean |  | 35.66
| Median |  | 35.8
| Std Dev. |  | 0.4219
| Rel.Std.Dev. |  | 1.183%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 340 (Indicative Value 0.1 ppm) |  |  | Indicative
Lu | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 0.18 | 1
| 1 | 2 | 0.06 | 14
| 1 | 3 | 0.1 | 16
| 1 | 4 | < 0.05 | 0.101
| 1 | 5 | 0.14 | 20
| Mean |  | 0.12
| Median |  | 0.12
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 43.03%
| PDM3 |  | 18.81%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 340 (Indicative Value 4.5 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 10 | 1
| 1 | 2 | 4 | 15
| 1 | 3 | < 1 | 16
| 1 | 4 | 3 | 4.5
| 1 | 5 | 5 | 21
| Mean |  | 5.5
| Median |  | 4.5
| Std Dev. |  | 3.1091
| Rel.Std.Dev. |  | 56.53%
| PDM3 |  | 22.22%
| Z-Score (Absolute) |  | NA
| Analytical results for Na in OREAS 340 (Indicative Value 0.158 wt.%) |  |  | Indicative
Na | Round | Replicate | Lab | 1
| No | No | 03 | wt.%
|  |  | INAA | 3
3
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.17 | 16
| 1 | 3 | 0.16 | 16
| 1 | 4 | 0.16 | 0.158
| 1 | 5 | 0.15 | 22
| Mean |  | 0.158
| Median |  | 0.16
| Std Dev. |  | 0.0083666
| Rel.Std.Dev. |  | 5.295%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 340 (Indicative Value 23 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | 25 | 1
| 1 | 2 | 18 | 17
| 1 | 3 | 24 | 16
| 1 | 4 | 27 | 23
| 1 | 5 | 21 | 23
| Mean |  | 23
| Median |  | 24
| Std Dev. |  | 3.5355
| Rel.Std.Dev. |  | 15.37%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 340 (Indicative Value < 20 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | < 20 | 1
| 1 | 2 | < 20 | 18
| 1 | 3 | < 20 | 16
| 1 | 4 | < 20 | < 20
| 1 | 5 | < 20 | 24
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 340 (Indicative Value 26.9 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | 21 | 1
| 1 | 2 | < 15 | 19
| 1 | 3 | 54 | 16
| 1 | 4 | 29 | 26.9
| 1 | 5 | 23 | 25
| Mean |  | 31.75
| Median |  | 26
| Std Dev. |  | 15.2179
| Rel.Std.Dev. |  | 47.93%
| PDM3 |  | 18.03%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 340 (Indicative Value 46.4 ppm) |  |  | Indicative
Sb | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 47 | 3
| 1 | 3 | 47.3 | 16
| 1 | 4 | 44.6 | 46.36
| 1 | 5 | 46.9 | 26
| Mean |  | 46.36
| Median |  | 46.9
| Std Dev. |  | 1.0968
| Rel.Std.Dev. |  | 2.366%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 340 (Indicative Value 3.94 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 3.9 | 1
| 1 | 2 | 3.9 | 21
| 1 | 3 | 4 | 16
| 1 | 4 | 3.9 | 3.94
| 1 | 5 | 4 | 27
| Mean |  | 3.94
| Median |  | 3.9
| Std Dev. |  | 0.054772
| Rel.Std.Dev. |  | 1.39%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 340 (Indicative Value < 3 ppm) |  |  | Indicative
Se | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 3 | 1
| 1 | 2 | < 3 | 9
| 1 | 3 | < 3 | 16
| 1 | 4 | < 3 | < 3
| 1 | 5 | < 3 | 15
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 340 (Indicative Value 5.42 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 5.4 | 1
| 1 | 2 | 5.7 | 10
| 1 | 3 | 5.5 | 16
| 1 | 4 | 5.2 | 5.42
| 1 | 5 | 5.3 | 16
| Mean |  | 5.42
| Median |  | 5.4
| Std Dev. |  | 0.19235
| Rel.Std.Dev. |  | 3.549%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 340 (Indicative Value < 200 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 0
0
| 1 | 1 | < 200 | 1
| 1 | 2 | < 200 | 11
| 1 | 3 | 900 | 16
| 1 | 4 | < 200 | < 200
| 1 | 5 | < 200 | 17
| Mean |  | 900
| Median |  | 900
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 340 (Indicative Value < 500 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 0
0
| 1 | 1 | < 500 | 1
| 1 | 2 | < 500 | 12
| 1 | 3 | < 500 | 16
| 1 | 4 | < 500 | < 500
| 1 | 5 | < 500 | 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 340 (Indicative Value < 0.5 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 0.5 | 1
| 1 | 2 | < 0.5 | 13
| 1 | 3 | < 0.5 | 16
| 1 | 4 | < 0.5 | < 0.5
| 1 | 5 | < 0.5 | 19
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 340 (Indicative Value < 0.5 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 0.5 | 1
| 1 | 2 | < 0.5 | 14
| 1 | 3 | < 0.5 | 16
| 1 | 4 | < 0.5 | < 0.5
| 1 | 5 | < 0.5 | 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 340 (Indicative Value 4.18 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 3.7 | 1
| 1 | 2 | 5.5 | 15
| 1 | 3 | 4.3 | 16
| 1 | 4 | 4.3 | 4.18
| 1 | 5 | 3.1 | 21
| Mean |  | 4.18
| Median |  | 4.3
| Std Dev. |  | 0.88994
| Rel.Std.Dev. |  | 21.29%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 340 (Indicative Value 4.72 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 5.3 | 1
| 1 | 2 | 6.2 | 16
| 1 | 3 | 4.5 | 16
| 1 | 4 | 3 | 4.72
| 1 | 5 | 4.6 | 22
| Mean |  | 4.72
| Median |  | 4.6
| Std Dev. |  | 1.1777
| Rel.Std.Dev. |  | 24.95%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 340 (Indicative Value 13 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 1
1
| 1 | 1 | 11 | 1
| 1 | 2 | 14 | 17
| 1 | 3 | 13 | 16
| 1 | 4 | 13 | 13
| 1 | 5 | 14 | 23
| Mean |  | 13
| Median |  | 13
| Std Dev. |  | 1.2247
| Rel.Std.Dev. |  | 9.421%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 340 (Indicative Value 1.44 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 03 | ppm
|  |  | INAA | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.2 | 18
| 1 | 3 | 1.2 | 16
| 1 | 4 | 1.6 | 1.44
| 1 | 5 | 1.6 | 24
| Mean |  | 1.44
| Median |  | 1.6
| Std Dev. |  | 0.21909
| Rel.Std.Dev. |  | 15.21%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 340 (Indicative Value 0.449 wt.%) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 03 | wt.%
|  |  | INAA | 3
3
| 1 | 1 | 0.461 | 1
| 1 | 2 | 0.463 | 19
| 1 | 3 | 0.422 | 16
| 1 | 4 | 0.432 | 0.4492
| 1 | 5 | 0.468 | 25
| Mean |  | 0.4492
| Median |  | 0.461
| Std Dev. |  | 0.020729
| Rel.Std.Dev. |  | 4.615%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C33: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C50: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C84: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C322: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C339: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C356: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C373: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C390: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C407: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C424: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C628: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ISE
| Analytical results for F in OREAS 340 (Indicative Value 1.01 wt.%) |  |  | Indicative
F | Round | Replicate | Lab | 1
| No | No | 10 | wt.%
|  |  | ISE | 2
2
| 1 | 1 | 1.015 | 1
| 1 | 2 | 1.01 | 23
| 1 | 3 | 1.02 | 16
| 1 | 4 | 1.005 | 1.014
| 1 | 5 | 1.02 | 29
| Mean |  | 1.014
| Median |  | 1.015
| Std Dev. |  | 0.0065192
| Rel.Std.Dev. |  | 0.6429%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.