## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 612
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 62.2892 | 59.6338 | 64.9447 | 60.5044 | 64.0741
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppb) | 58.3257 | 56.0975 | 60.5539 | 56.4949 | 60.1565
4-Acid Digestion
Ag, Silver (ppm) | 48.1767 | 46.759 | 49.5945 | 47.1895 | 49.164
Al, Aluminium (wt.%) | 7.2195 | 6.9551 | 7.4839 | 7.0487 | 7.3903
As, Arsenic (ppm) | 331.0075 | 321.2365 | 340.7785 | 324.4081 | 337.6069
Ba, Barium (ppm) | 2163.2808 | 2100.9607 | 2225.6009 | 2106.2023 | 2220.3593
Be, Beryllium (ppm) | 2.4848 | 2.3692 | 2.6004 | 2.4076 | 2.562
Bi, Bismuth (ppm) | 0.85257 | 0.78666 | 0.91849 | 0.81334 | 0.89181
Ca, Calcium (wt.%) | 0.70471 | 0.68007 | 0.72935 | 0.69136 | 0.71806
Cd, Cadmium (ppm) | 13.838 | 13.3163 | 14.3597 | 13.5457 | 14.1303
Ce, Cerium (ppm) | 76.536 | 72.8181 | 80.2539 | 73.9714 | 79.1006
Co, Cobalt (ppm) | 4.1101 | 3.8778 | 4.3424 | 3.9179 | 4.3023
Cr, Chromium (ppm) | 21.7867 | 19.1259 | 24.4474 | 20.2532 | 23.3201
Cs, Caesium (ppm) | 6.6282 | 6.3526 | 6.9038 | 6.4799 | 6.7764
Cu, Copper (ppm) | 93.7422 | 89.1968 | 98.2876 | 90.896 | 96.5883
Dy, Dysprosium (ppm) | 3.4074 | 3.2323 | 3.5824 | 3.3036 | 3.5112
Er, Erbium (ppm) | 1.2689 | 1.1819 | 1.3559 | 1.1889 | 1.3489
Eu, Europium (ppm) | 1.3946 | 1.2987 | 1.4906 | 1.3567 | 1.4326
Fe, Iron (wt.%) | 2.2278 | 2.1751 | 2.2805 | 2.1782 | 2.2775
Ga, Gallium (ppm) | 20.1664 | 19.2656 | 21.0672 | 19.6654 | 20.6673
Gd, Gadolinium (ppm) | 5.4494 | 5.1205 | 5.7782 | 5.2823 | 5.6164
Hf, Hafnium (ppm) | 4.2916 | 4.0446 | 4.5386 | 4.0856 | 4.4975
Ho, Holmium (ppm) | 0.51625 | 0.47808 | 0.55442 | 0.49488 | 0.53763
In, Indium (ppm) | 0.14001 | 0.12488 | 0.15513 | 0.13174 | 0.14828
K, Potassium (wt.%) | 4.3267 | 4.1456 | 4.5079 | 4.2528 | 4.4007
La, Lanthanum (ppm) | 35.9836 | 33.6165 | 38.3506 | 34.7381 | 37.229
Li, Lithium (ppm) | 32.2771 | 30.5765 | 33.9778 | 31.2333 | 33.321
Lu, Lutetium (ppm) | 0.14819 | 0.13209 | 0.1643 | IND | IND
Mg, Magnesium (wt.%) | 0.25884 | 0.24923 | 0.26844 | 0.251 | 0.26668
Mn, Manganese (wt.%) | 0.044806 | 0.043396 | 0.046215 | 0.043785 | 0.045826
Mo, Molybdenum (ppm) | 7.0681 | 6.6547 | 7.4815 | 6.8238 | 7.3125
Na, Sodium (wt.%) | 1.6761 | 1.6253 | 1.7269 | 1.6548 | 1.6974
Nb, Niobium (ppm) | 14.5358 | 13.9952 | 15.0764 | 14.101 | 14.9706
Nd, Neodymium (ppm) | 33.6431 | 31.7648 | 35.5214 | 33.0729 | 34.2133
Ni, Nickel (ppm) | 10.826 | 9.8182 | 11.8337 | 10.147 | 11.505
P, Phosphorus (wt.%) | 0.046812 | 0.044905 | 0.048718 | 0.045642 | 0.047981
Pb, Lead (wt.%) | 0.2535 | 0.24567 | 0.26133 | 0.2484 | 0.25859
Pr, Praseodymium (ppm) | 8.9792 | 8.4335 | 9.5249 | 8.7301 | 9.2284
Rb, Rubidium (ppm) | 211.394 | 200.7739 | 222.014 | 206.3364 | 216.4515
S, Sulphur (wt.%) | 0.51826 | 0.50181 | 0.53471 | 0.50774 | 0.52877
Sb, Antimony (ppm) | 14.1426 | 13.4237 | 14.8615 | 13.7615 | 14.5236
Sc, Scandium (ppm) | 5.2761 | 5.0191 | 5.5331 | 5.1074 | 5.4448
Sm, Samarium (ppm) | 6.7511 | 6.3707 | 7.1315 | 6.4995 | 7.0027
Sn, Tin (ppm) | 3.6198 | 3.4073 | 3.8323 | 3.4224 | 3.8172
Sr, Strontium (ppm) | 228.1358 | 219.5418 | 236.7298 | 221.5453 | 234.7263
Ta, Tantalum (ppm) | 1.0809 | 0.9983 | 1.1635 | 1.0319 | 1.1299
Tb, Terbium (ppm) | 0.69827 | 0.6563 | 0.74023 | 0.66631 | 0.73023
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 12.8998 | 12.3493 | 13.4502 | 12.4836 | 13.3159
Ti, Titanium (wt.%) | 0.20057 | 0.19221 | 0.20894 | 0.19693 | 0.20422
Tl, Thallium (ppm) | 3.7683 | 3.5974 | 3.9392 | 3.6271 | 3.9095
Tm, Thulium (ppm) | 0.16548 | 0.15177 | 0.17919 | IND | IND
U, Uranium (ppm) | 4.3589 | 4.1891 | 4.5288 | 4.2064 | 4.5115
V, Vanadium (ppm) | 28.9217 | 27.8514 | 29.9921 | 28.1364 | 29.7071
W, Tungsten (ppm) | 3.0657 | 2.8054 | 3.3261 | 2.9373 | 3.1942
Y, Yttrium (ppm) | 14.2579 | 13.6204 | 14.8955 | 13.8779 | 14.638
Yb, Ytterbium (ppm) | 1.0154 | 0.92757 | 1.1032 | 0.95775 | 1.073
Zn, Zinc (wt.%) | 0.36662 | 0.35746 | 0.37579 | 0.36119 | 0.37206
Zr, Zirconium (ppm) | 149.4674 | 143.9943 | 154.9404 | 146.6593 | 152.2754
Aqua Regia Digestion
Ag, Silver (ppm) | 47.2395 | 45.7519 | 48.7272 | 46.5046 | 47.9745
Al, Aluminium (wt.%) | 0.87352 | 0.83642 | 0.91062 | 0.85457 | 0.89247
As, Arsenic (ppm) | 321.6841 | 311.4981 | 331.8701 | 316.9076 | 326.4607
Ba, Barium (ppm) | 323.9871 | 304.2804 | 343.6938 | 313.6807 | 334.2935
Be, Beryllium (ppm) | 0.63379 | 0.60372 | 0.66386 | 0.61481 | 0.65277
Bi, Bismuth (ppm) | 0.84803 | 0.78428 | 0.91177 | 0.81036 | 0.88569
Ca, Calcium (wt.%) | 0.50119 | 0.48679 | 0.5156 | 0.48872 | 0.51367
Cd, Cadmium (ppm) | 13.9156 | 13.4408 | 14.3904 | 13.6191 | 14.2121
Ce, Cerium (ppm) | 48.6151 | 45.7842 | 51.4459 | 47.1269 | 50.1032
Co, Cobalt (ppm) | 3.6881 | 3.5309 | 3.8453 | 3.5639 | 3.8123
Cr, Chromium (ppm) | 19.5023 | 17.5934 | 21.4112 | 18.2712 | 20.7334
Cs, Caesium (ppm) | 2.2466 | 2.1354 | 2.3577 | 2.1733 | 2.3198
Cu, Copper (ppm) | 92.2795 | 89.0828 | 95.4762 | 90.327 | 94.232
Dy, Dysprosium (ppm) | 1.9117 | 1.7083 | 2.115 | 1.8079 | 2.0155
Er, Erbium (ppm) | 0.59355 | 0.52545 | 0.66164 | 0.56724 | 0.61985
Eu, Europium (ppm) | 0.84199 | 0.71046 | 0.97351 | 0.8077 | 0.87627
Fe, Iron (wt.%) | 1.8355 | 1.7829 | 1.8882 | 1.7981 | 1.8729
Ga, Gallium (ppm) | 4.3318 | 4.0177 | 4.6459 | 4.1929 | 4.4708
Gd, Gadolinium (ppm) | 3.5008 | 3.0809 | 3.9207 | 3.3462 | 3.6554
Hf, Hafnium (ppm) | 0.81543 | 0.75452 | 0.87634 | 0.77373 | 0.85714
Hg, Mercury (ppm) | 0.19458 | 0.17476 | 0.2144 | 0.17791 | 0.21126
Ho, Holmium (ppm) | 0.27442 | 0.23399 | 0.31485 | 0.25884 | 0.29
In, Indium (ppm) | 0.064303 | 0.054741 | 0.073865 | 0.055234 | 0.073372
K, Potassium (wt.%) | 0.43563 | 0.4144 | 0.45686 | 0.42225 | 0.44901
La, Lanthanum (ppm) | 23.5825 | 22.2262 | 24.9387 | 22.9887 | 24.1762
Li, Lithium (ppm) | 10.3974 | 9.7904 | 11.0044 | 10.1103 | 10.6845
Lu, Lutetium (ppm) | 0.057411 | 0.05155 | 0.063272 | IND | IND
Mg, Magnesium (wt.%) | 0.16474 | 0.15573 | 0.17375 | 0.1615 | 0.16798
Mn, Manganese (wt.%) | 0.040858 | 0.039818 | 0.041899 | 0.040247 | 0.04147
Mo, Molybdenum (ppm) | 6.488 | 6.1832 | 6.7928 | 6.2643 | 6.7118
Na, Sodium (wt.%) | 0.060282 | 0.052103 | 0.068461 | 0.057353 | 0.063211
Nd, Neodymium (ppm) | 22.5337 | 19.8156 | 25.2519 | 21.9489 | 23.1185
Ni, Nickel (ppm) | 10.2688 | 9.358 | 11.1796 | 9.7679 | 10.7697
P, Phosphorus (wt.%) | 0.036671 | 0.035146 | 0.038196 | 0.035673 | 0.037669
Pb, Lead (wt.%) | 0.24768 | 0.24073 | 0.25463 | 0.24218 | 0.25317
Pr, Praseodymium (ppm) | 5.779 | 5.103 | 6.455 | 5.5959 | 5.9622
Rb, Rubidium (ppm) | 23.544 | 22.2921 | 24.7959 | 22.7921 | 24.296
S, Sulphur (wt.%) | 0.5146 | 0.49929 | 0.52991 | 0.50351 | 0.52569
Sb, Antimony (ppm) | 10.4893 | 9.7084 | 11.2701 | 10.1056 | 10.8729
Sc, Scandium (ppm) | 1.6053 | 1.4449 | 1.7656 | 1.5465 | 1.6641
Sm, Samarium (ppm) | 4.2782 | 3.8435 | 4.7129 | 4.0923 | 4.4641
Sn, Tin (ppm) | 0.97504 | 0.85154 | 1.0985 | 0.91514 | 1.0349
Sr, Strontium (ppm) | 49.2121 | 45.2503 | 53.1739 | 47.5048 | 50.9193
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.411 | 0.38271 | 0.43929 | 0.39444 | 0.42755
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 7.787 | 7.1328 | 8.4411 | 7.4959 | 8.078
Ti, Titanium (wt.%) | 0.016049 | 0.014433 | 0.017666 | 0.015207 | 0.016892
Tl, Thallium (ppm) | 0.85888 | 0.80804 | 0.90972 | 0.8358 | 0.88196
Tm, Thulium (ppm) | 0.072999 | 0.060192 | 0.085805 | IND | IND
U, Uranium (ppm) | 2.2153 | 2.0803 | 2.3502 | 2.1133 | 2.3173
V, Vanadium (ppm) | 9.4108 | 8.298 | 10.5235 | IND | IND
W, Tungsten (ppm) | 1.0313 | 0.83279 | 1.2299 | 0.94291 | 1.1197
Y, Yttrium (ppm) | 7.2186 | 6.7965 | 7.6407 | 7.0073 | 7.4299
Yb, Ytterbium (ppm) | 0.4046 | 0.37322 | 0.43599 | 0.39001 | 0.4192
Zn, Zinc (wt.%) | 0.36219 | 0.35359 | 0.37079 | 0.35507 | 0.3693
Zr, Zirconium (ppm) | 27.9699 | 26.1829 | 29.7569 | 26.9641 | 28.9757
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 612
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay (Grav)
Ag | ppm | 54.4577
4-Acid Digestion
B | ppm | 8.8333 | Hg | ppm | 0.096667 | Se | ppm | 0.72381
Ge | ppm | 0.19962 | Re | ppm | 0.0023333
Aqua Regia Digestion
B | ppm | < 10 | Pd | ppb | < 10 | Se | ppm | 0.80583
Ge | ppm | 0.094583 | Pt | ppb | < 5
Nb | ppm | 0.38863 | Re | ppm | 0.00087917
Infrared Combustion
C | wt.% | 0.15 | S | wt.% | 0.48
Borate Fusion XRF
Al2O3 | wt.% | 14.38 | MgO | wt.% | 0.46 | S | wt.% | 0.51858
CaO | wt.% | 0.99 | MnO | wt.% | 0.06 | SiO2 | wt.% | 69.755
Fe2O3 | wt.% | 3.295 | Na2O | wt.% | 2.275 | TiO2 | wt.% | 0.35
K2O | wt.% | 5.36 | P2O5 | wt.% | 0.1075
Thermogravimetry
LOI1000 | wt.% | 2.035
Laser Ablation ICP-MS
Ag | ppm | 57.45 | Hf | ppm | 6.435 | Sn | ppm | 3.8
As | ppm | 328.5 | Ho | ppm | 0.575 | Sr | ppm | 229.5
Ba | ppm | 2165 | In | ppm | 0.125 | Ta | ppm | 1.13
Be | ppm | 2.9 | La | ppm | 39.55 | Tb | ppm | 0.75
Bi | ppm | 0.97 | Lu | ppm | 0.15 | Te | ppm | 0.15
Cd | ppm | 15.55 | Mn | wt.% | 0.0466 | Th | ppm | 13.5
Ce | ppm | 74.65 | Mo | ppm | 7 | Ti | wt.% | 0.2095
Co | ppm | 4.3 | Nb | ppm | 14.45 | Tl | ppm | 2.8
Cr | ppm | 29 | Nd | ppm | 35.95 | Tm | ppm | 0.175
Cs | ppm | 6.395 | Ni | ppm | 11 | U | ppm | 4.4
Cu | ppm | 104 | Pb | wt.% | 0.278 | V | ppm | 30.7
Dy | ppm | 3.505 | Pr | ppm | 9.415 | W | ppm | 3.25
Er | ppm | 1.33 | Rb | ppm | 210 | Y | ppm | 16.15
Eu | ppm | 1.395 | Re | ppm | < 0.01 | Yb | ppm | 1.045
Ga | ppm | 19.4 | Sb | ppm | 15.75 | Zn | wt.% | 0.3705
Gd | ppm | 5.5 | Sc | ppm | 5.4 | Zr | ppm | 219.5
Ge | ppm | 1.325 | Sm | ppm | 6.775
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 612
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 62.2892 | 6.0123 | 50.2645 | 74.3139 | 44.2522 | 80.3263 | 9.652% | 19.3% | 28.96% | 59.1748 | 65.4037
Aqua Regia Digestion (sample weights 10-50g)
Au, ppb | 58.3257 | 3.3748 | 51.576 | 65.0754 | 48.2011 | 68.4502 | 5.786% | 11.57% | 17.36% | 55.4094 | 61.242
4-Acid Digestion
Ag, ppm | 48.1767 | 1.4288 | 45.3191 | 51.0343 | 43.8903 | 52.4632 | 2.966% | 5.932% | 8.897% | 45.7679 | 50.5856
Al, wt.% | 7.2195 | 0.24784 | 6.7238 | 7.7152 | 6.476 | 7.963 | 3.433% | 6.866% | 10.3% | 6.8585 | 7.5805
As, ppm | 331.0075 | 10.5576 | 309.8923 | 352.1228 | 299.3346 | 362.6804 | 3.19% | 6.379% | 9.569% | 314.4571 | 347.5579
Ba, ppm | 2163.2808 | 59.2343 | 2044.8122 | 2281.7494 | 1985.5779 | 2340.9837 | 2.738% | 5.476% | 8.215% | 2055.1168 | 2271.4448
Be, ppm | 2.4848 | 0.10856 | 2.2677 | 2.7019 | 2.1592 | 2.8105 | 4.369% | 8.738% | 13.11% | 2.3606 | 2.6091
Bi, ppm | 0.85257 | 0.078766 | 0.69504 | 1.0101 | 0.61628 | 1.0889 | 9.239% | 18.48% | 27.72% | 0.80995 | 0.8952
Ca, wt.% | 0.70471 | 0.026412 | 0.65189 | 0.75753 | 0.62547 | 0.78395 | 3.748% | 7.496% | 11.24% | 0.66947 | 0.73994
Cd, ppm | 13.838 | 0.38221 | 13.0736 | 14.6024 | 12.6914 | 14.9846 | 2.762% | 5.524% | 8.286% | 13.1461 | 14.5299
Ce, ppm | 76.536 | 4.0788 | 68.3783 | 84.6937 | 64.2994 | 88.7725 | 5.329% | 10.66% | 15.99% | 72.7092 | 80.3628
Co, ppm | 4.1101 | 0.20134 | 3.7074 | 4.5128 | 3.5061 | 4.7141 | 4.899% | 9.797% | 14.7% | 3.9046 | 4.3156
Cr, ppm | 21.7867 | 2.7952 | 16.1963 | 27.377 | 13.4012 | 30.1721 | 12.83% | 25.66% | 38.49% | 20.6973 | 22.876
Cs, ppm | 6.6282 | 0.22566 | 6.1769 | 7.0795 | 5.9512 | 7.3052 | 3.405% | 6.809% | 10.21% | 6.2968 | 6.9596
Cu, ppm | 93.7422 | 4.5424 | 84.6573 | 102.827 | 80.1149 | 107.3694 | 4.846% | 9.691% | 14.54% | 89.0551 | 98.4293
Dy, ppm | 3.4074 | 0.097369 | 3.2126 | 3.6021 | 3.1153 | 3.6995 | 2.858% | 5.715% | 8.573% | 3.237 | 3.5777
Er, ppm | 1.2689 | 0.058957 | 1.151 | 1.3868 | 1.092 | 1.4458 | 4.646% | 9.293% | 13.94% | 1.2054 | 1.3323
Eu, ppm | 1.3946 | 0.082148 | 1.2304 | 1.5589 | 1.1482 | 1.6411 | 5.89% | 11.78% | 17.67% | 1.3249 | 1.4644
Fe, wt.% | 2.2278 | 0.05262 | 2.1226 | 2.333 | 2.0699 | 2.3857 | 2.362% | 4.724% | 7.086% | 2.1164 | 2.3392
Ga, ppm | 20.1664 | 0.86006 | 18.4463 | 21.8865 | 17.5862 | 22.7466 | 4.265% | 8.53% | 12.79% | 19.1581 | 21.1747
Gd, ppm | 5.4494 | 0.26117 | 4.927 | 5.9717 | 4.6658 | 6.2329 | 4.793% | 9.585% | 14.38% | 5.1769 | 5.7218
Hf, ppm | 4.2916 | 0.17607 | 3.9394 | 4.6437 | 3.7634 | 4.8198 | 4.103% | 8.206% | 12.31% | 4.077 | 4.5062
Ho, ppm | 0.51625 | 0.024638 | 0.46698 | 0.56553 | 0.44234 | 0.59017 | 4.772% | 9.545% | 14.32% | 0.49044 | 0.54206
In, ppm | 0.14001 | 0.0093306 | 0.12135 | 0.15867 | 0.11202 | 0.168 | 6.664% | 13.33% | 19.99% | 0.13301 | 0.14701
K, wt.% | 4.3267 | 0.26498 | 3.7968 | 4.8567 | 3.5318 | 5.1217 | 6.124% | 12.25% | 18.37% | 4.1104 | 4.5431
La, ppm | 35.9836 | 3.8824 | 28.2188 | 43.7483 | 24.3365 | 47.6307 | 10.79% | 21.58% | 32.37% | 34.1844 | 37.7827
Li, ppm | 32.2771 | 1.4083 | 29.4605 | 35.0938 | 28.0522 | 36.5021 | 4.363% | 8.726% | 13.09% | 30.6633 | 33.891
Lu, ppm | 0.14819 | 0.011906 | 0.12438 | 0.17201 | 0.11248 | 0.18391 | 8.034% | 16.07% | 24.1% | 0.14079 | 0.1556
Mg, wt.% | 0.25884 | 0.0084717 | 0.2419 | 0.27578 | 0.23342 | 0.28425 | 3.273% | 6.546% | 9.819% | 0.2459 | 0.27178
Mn, wt.% | 0.044806 | 0.0016088 | 0.041588 | 0.048023 | 0.039979 | 0.049632 | 3.591% | 7.181% | 10.77% | 0.042565 | 0.047046
Mo, ppm | 7.0681 | 0.35987 | 6.3484 | 7.7879 | 5.9885 | 8.1477 | 5.091% | 10.18% | 15.27% | 6.7147 | 7.4215
Na, wt.% | 1.6761 | 0.058924 | 1.5583 | 1.794 | 1.4994 | 1.8529 | 3.515% | 7.031% | 10.55% | 1.5923 | 1.7599
Nb, ppm | 14.5358 | 0.54922 | 13.4374 | 15.6343 | 12.8882 | 16.1835 | 3.778% | 7.557% | 11.34% | 13.809 | 15.2626
Nd, ppm | 33.6431 | 1.4593 | 30.7245 | 36.5618 | 29.2651 | 38.0211 | 4.338% | 8.675% | 13.01% | 31.961 | 35.3253
Ni, ppm | 10.826 | 0.91606 | 8.9939 | 12.6581 | 8.0778 | 13.5742 | 8.462% | 16.92% | 25.38% | 10.2847 | 11.3673
P, wt.% | 0.046812 | 0.0018811 | 0.04305 | 0.050574 | 0.041169 | 0.052455 | 4.018% | 8.037% | 12.06% | 0.044471 | 0.049153
Pb, wt.% | 0.2535 | 0.0094706 | 0.23456 | 0.27244 | 0.22509 | 0.28191 | 3.736% | 7.472% | 11.21% | 0.24082 | 0.26617
Pr, ppm | 8.9792 | 0.44816 | 8.0829 | 9.8756 | 7.6348 | 10.3237 | 4.991% | 9.982% | 14.97% | 8.5303 | 9.4282
Rb, ppm | 211.394 | 12.4724 | 186.4491 | 236.3388 | 173.9767 | 248.8112 | 5.9% | 11.8% | 17.7% | 200.8243 | 221.9636
S, wt.% | 0.51826 | 0.015127 | 0.48801 | 0.54851 | 0.47288 | 0.56364 | 2.919% | 5.837% | 8.756% | 0.49235 | 0.54417
Sb, ppm | 14.1426 | 0.73333 | 12.6759 | 15.6093 | 11.9426 | 16.3426 | 5.185% | 10.37% | 15.56% | 13.4355 | 14.8497
Sc, ppm | 5.2761 | 0.37837 | 4.5194 | 6.0328 | 4.141 | 6.4112 | 7.171% | 14.34% | 21.51% | 5.0123 | 5.5399
Sm, ppm | 6.7511 | 0.23044 | 6.2902 | 7.212 | 6.0598 | 7.4424 | 3.413% | 6.827% | 10.24% | 6.4135 | 7.0886
Sn, ppm | 3.6198 | 0.12963 | 3.3606 | 3.8791 | 3.2309 | 4.0087 | 3.581% | 7.162% | 10.74% | 3.4388 | 3.8008
Sr, ppm | 228.1358 | 6.6069 | 214.9221 | 241.3495 | 208.3152 | 247.9563 | 2.896% | 5.792% | 8.688% | 216.729 | 239.5426
Ta, ppm | 1.0809 | 0.069879 | 0.94117 | 1.2207 | 0.87129 | 1.2906 | 6.465% | 12.93% | 19.39% | 1.0269 | 1.135
Tb, ppm | 0.69827 | 0.040413 | 0.61744 | 0.77909 | 0.57703 | 0.81951 | 5.788% | 11.58% | 17.36% | 0.66335 | 0.73318
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 12.8998 | 0.60786 | 11.684 | 14.1155 | 11.0762 | 14.7233 | 4.712% | 9.424% | 14.14% | 12.2548 | 13.5447
Ti, wt.% | 0.20057 | 0.0069356 | 0.1867 | 0.21445 | 0.17977 | 0.22138 | 3.458% | 6.916% | 10.37% | 0.19055 | 0.2106
Tl, ppm | 3.7683 | 0.15908 | 3.4502 | 4.0865 | 3.2911 | 4.2456 | 4.221% | 8.443% | 12.66% | 3.5799 | 3.9568
Tm, ppm | 0.16548 | 0.0056037 | 0.15427 | 0.17669 | 0.14867 | 0.18229 | 3.386% | 6.773% | 10.16% | 0.15721 | 0.17376
U, ppm | 4.3589 | 0.20913 | 3.9407 | 4.7772 | 3.7316 | 4.9863 | 4.798% | 9.595% | 14.39% | 4.141 | 4.5769
V, ppm | 28.9217 | 0.98114 | 26.9595 | 30.884 | 25.9783 | 31.8652 | 3.392% | 6.785% | 10.18% | 27.4757 | 30.3678
W, ppm | 3.0657 | 0.23426 | 2.5972 | 3.5343 | 2.363 | 3.7685 | 7.641% | 15.28% | 22.92% | 2.9125 | 3.219
Y, ppm | 14.2579 | 0.7565 | 12.7449 | 15.7709 | 11.9884 | 16.5274 | 5.306% | 10.61% | 15.92% | 13.545 | 14.9708
Yb, ppm | 1.0154 | 0.066918 | 0.88153 | 1.1492 | 0.81461 | 1.2161 | 6.591% | 13.18% | 19.77% | 0.9646 | 1.0661
Zn, wt.% | 0.36662 | 0.010246 | 0.34613 | 0.38712 | 0.33589 | 0.39736 | 2.795% | 5.589% | 8.384% | 0.34829 | 0.38495
Zr, ppm | 149.4674 | 6.3687 | 136.73 | 162.2047 | 130.3613 | 168.5734 | 4.261% | 8.522% | 12.78% | 141.994 | 156.9407
Aqua Regia Digestion
Ag, ppm | 47.2395 | 1.6889 | 43.8618 | 50.6173 | 42.1729 | 52.3062 | 3.575% | 7.15% | 10.73% | 44.8776 | 49.6015
Al, wt.% | 0.87352 | 0.040689 | 0.79214 | 0.9549 | 0.75145 | 0.99559 | 4.658% | 9.316% | 13.97% | 0.82984 | 0.91719
As, ppm | 321.6841 | 10.1585 | 301.3671 | 342.0011 | 291.2086 | 352.1596 | 3.158% | 6.316% | 9.474% | 305.5999 | 337.7683
Ba, ppm | 323.9871 | 28.4883 | 267.0105 | 380.9636 | 238.5223 | 409.4519 | 8.793% | 17.59% | 26.38% | 307.7877 | 340.1864
Be, ppm | 0.63379 | 0.024302 | 0.58518 | 0.68239 | 0.56088 | 0.7067 | 3.834% | 7.669% | 11.5% | 0.6021 | 0.66548
Bi, ppm | 0.84803 | 0.043909 | 0.76021 | 0.93584 | 0.7163 | 0.97975 | 5.178% | 10.36% | 15.53% | 0.80563 | 0.89043
Ca, wt.% | 0.50119 | 0.014716 | 0.47176 | 0.53062 | 0.45704 | 0.54534 | 2.936% | 5.872% | 8.809% | 0.47613 | 0.52625
Cd, ppm | 13.9156 | 0.34859 | 13.2184 | 14.6128 | 12.8698 | 14.9614 | 2.505% | 5.01% | 7.515% | 13.2198 | 14.6114
Ce, ppm | 48.6151 | 3.0936 | 42.4279 | 54.8022 | 39.3343 | 57.8958 | 6.363% | 12.73% | 19.09% | 46.1843 | 51.0458
Co, ppm | 3.6881 | 0.15471 | 3.3787 | 3.9975 | 3.224 | 4.1522 | 4.195% | 8.39% | 12.58% | 3.5037 | 3.8725
Cr, ppm | 19.5023 | 1.3169 | 16.8686 | 22.1361 | 15.5517 | 23.4529 | 6.752% | 13.5% | 20.26% | 18.5272 | 20.4774
Cs, ppm | 2.2466 | 0.094721 | 2.0571 | 2.436 | 1.9624 | 2.5307 | 4.216% | 8.432% | 12.65% | 2.1342 | 2.3589
Cu, ppm | 92.2795 | 3.4242 | 85.4311 | 99.1279 | 82.0069 | 102.5522 | 3.711% | 7.421% | 11.13% | 87.6655 | 96.8935
Dy, ppm | 1.9117 | 0.19341 | 1.5249 | 2.2985 | 1.3315 | 2.4919 | 10.12% | 20.23% | 30.35% | 1.8161 | 2.0073
Er, ppm | 0.59355 | 0.055891 | 0.48176 | 0.70533 | 0.42587 | 0.76122 | 9.416% | 18.83% | 28.25% | 0.56387 | 0.62322
Eu, ppm | 0.84199 | 0.11431 | 0.61336 | 1.0706 | 0.49904 | 1.1849 | 13.58% | 27.15% | 40.73% | 0.79989 | 0.88409
Fe, wt.% | 1.8355 | 0.06399 | 1.7075 | 1.9635 | 1.6435 | 2.0275 | 3.486% | 6.972% | 10.46% | 1.7437 | 1.9273
Ga, ppm | 4.3318 | 0.33697 | 3.6579 | 5.0058 | 3.3209 | 5.3427 | 7.779% | 15.56% | 23.34% | 4.1152 | 4.5484
Gd, ppm | 3.5008 | 0.33567 | 2.8295 | 4.1721 | 2.4938 | 4.5078 | 9.588% | 19.18% | 28.77% | 3.3258 | 3.6758
Hf, ppm | 0.81543 | 0.04988 | 0.71567 | 0.91519 | 0.66579 | 0.96507 | 6.117% | 12.23% | 18.35% | 0.77466 | 0.8562
Hg, ppm | 0.19458 | 0.012876 | 0.16883 | 0.22033 | 0.15596 | 0.23321 | 6.617% | 13.23% | 19.85% | 0.18485 | 0.20431
Ho, ppm | 0.27442 | 0.032638 | 0.20914 | 0.3397 | 0.17651 | 0.37234 | 11.89% | 23.79% | 35.68% | 0.2607 | 0.28814
In, ppm | 0.064303 | 0.007646 | 0.049011 | 0.079595 | 0.041365 | 0.087241 | 11.89% | 23.78% | 35.67% | 0.061088 | 0.067518
K, wt.% | 0.43563 | 0.021863 | 0.3919 | 0.47936 | 0.37004 | 0.50122 | 5.019% | 10.04% | 15.06% | 0.41385 | 0.45741
La, ppm | 23.5825 | 1.6459 | 20.2906 | 26.8743 | 18.6447 | 28.5202 | 6.979% | 13.96% | 20.94% | 22.4033 | 24.7616
Li, ppm | 10.3974 | 0.75095 | 8.8955 | 11.8993 | 8.1446 | 12.6503 | 7.222% | 14.44% | 21.67% | 9.8775 | 10.9173
Lu, ppm | 0.057411 | 0.0065098 | 0.044392 | 0.070431 | 0.037882 | 0.076941 | 11.34% | 22.68% | 34.02% | 0.054541 | 0.060282
Mg, wt.% | 0.16474 | 0.0089169 | 0.14691 | 0.18257 | 0.13799 | 0.19149 | 5.413% | 10.83% | 16.24% | 0.1565 | 0.17298
Mn, wt.% | 0.040858 | 0.0010197 | 0.038819 | 0.042898 | 0.037799 | 0.043918 | 2.496% | 4.992% | 7.487% | 0.038815 | 0.042901
Mo, ppm | 6.488 | 0.28299 | 5.922 | 7.054 | 5.6391 | 7.337 | 4.362% | 8.723% | 13.09% | 6.1636 | 6.8124
Na, wt.% | 0.060282 | 0.0052355 | 0.049811 | 0.070753 | 0.044576 | 0.075989 | 8.685% | 17.37% | 26.06% | 0.057268 | 0.063296
Nd, ppm | 22.5337 | 2.1442 | 18.2454 | 26.8221 | 16.1012 | 28.9663 | 9.515% | 19.03% | 28.55% | 21.4071 | 23.6604
Ni, ppm | 10.2688 | 0.74901 | 8.7708 | 11.7668 | 8.0218 | 12.5158 | 7.294% | 14.59% | 21.88% | 9.7554 | 10.7822
P, wt.% | 0.036671 | 0.0017918 | 0.033087 | 0.040254 | 0.031296 | 0.042046 | 4.886% | 9.772% | 14.66% | 0.034837 | 0.038504
Pb, wt.% | 0.24768 | 0.0068973 | 0.23388 | 0.26147 | 0.22698 | 0.26837 | 2.785% | 5.57% | 8.354% | 0.23529 | 0.26006
Pr, ppm | 5.779 | 0.54142 | 4.6962 | 6.8619 | 4.1548 | 7.4033 | 9.369% | 18.74% | 28.11% | 5.4901 | 6.068
Rb, ppm | 23.544 | 1.0947 | 21.3546 | 25.7334 | 20.2599 | 26.8281 | 4.65% | 9.299% | 13.95% | 22.3668 | 24.7212
S, wt.% | 0.5146 | 0.013042 | 0.48852 | 0.54069 | 0.47547 | 0.55373 | 2.534% | 5.069% | 7.603% | 0.48887 | 0.54033
Sb, ppm | 10.4893 | 1.1916 | 8.1061 | 12.8725 | 6.9144 | 14.0641 | 11.36% | 22.72% | 34.08% | 9.9648 | 11.0137
Sc, ppm | 1.6053 | 0.21623 | 1.1728 | 2.0377 | 0.95658 | 2.254 | 13.47% | 26.94% | 40.41% | 1.525 | 1.6855
Sm, ppm | 4.2782 | 0.35095 | 3.5763 | 4.9801 | 3.2254 | 5.3311 | 8.203% | 16.41% | 24.61% | 4.0643 | 4.4921
Sn, ppm | 0.97504 | 0.15509 | 0.66486 | 1.2852 | 0.50977 | 1.4403 | 15.91% | 31.81% | 47.72% | 0.92628 | 1.0238
Sr, ppm | 49.2121 | 6.9358 | 35.3405 | 63.0837 | 28.4047 | 70.0195 | 14.09% | 28.19% | 42.28% | 46.7515 | 51.6727
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.411 | 0.022812 | 0.36537 | 0.45662 | 0.34256 | 0.47943 | 5.55% | 11.1% | 16.65% | 0.39045 | 0.43155
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 7.787 | 0.80007 | 6.1868 | 9.3871 | 5.3868 | 10.1872 | 10.27% | 20.55% | 30.82% | 7.3976 | 8.1763
Ti, wt.% | 0.016049 | 0.0020203 | 0.012009 | 0.02009 | 0.0099887 | 0.02211 | 12.59% | 25.18% | 37.76% | 0.015247 | 0.016852
Tl, ppm | 0.85888 | 0.059695 | 0.73949 | 0.97827 | 0.6798 | 1.038 | 6.95% | 13.9% | 20.85% | 0.81594 | 0.90183
Tm, ppm | 0.072999 | 0.0085986 | 0.055801 | 0.090196 | 0.047203 | 0.098794 | 11.78% | 23.56% | 35.34% | 0.069349 | 0.076648
U, ppm | 2.2153 | 0.13214 | 1.951 | 2.4796 | 1.8189 | 2.6117 | 5.965% | 11.93% | 17.9% | 2.1045 | 2.326
V, ppm | 9.4108 | 1.3544 | 6.702 | 12.1195 | 5.3476 | 13.4739 | 14.39% | 28.78% | 43.18% | 8.9402 | 9.8813
W, ppm | 1.0313 | 0.26357 | 0.50418 | 1.5585 | 0.2406 | 1.822 | 25.56% | 51.11% | 76.67% | 0.97976 | 1.0829
Y, ppm | 7.2186 | 0.54256 | 6.1335 | 8.3037 | 5.5909 | 8.8463 | 7.516% | 15.03% | 22.55% | 6.8577 | 7.5795
Yb, ppm | 0.4046 | 0.028134 | 0.34834 | 0.46087 | 0.3202 | 0.48901 | 6.953% | 13.91% | 20.86% | 0.38437 | 0.42484
Zn, wt.% | 0.36219 | 0.0059583 | 0.35027 | 0.3741 | 0.34431 | 0.38006 | 1.645% | 3.29% | 4.935% | 0.34408 | 0.3803
Zr, ppm | 27.9699 | 1.7894 | 24.391 | 31.5488 | 22.6016 | 33.3382 | 6.398% | 12.8% | 19.19% | 26.5714 | 29.3684
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 612
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Text Values:
-- | Element Not Determined (Lab 1.20)
--- | Not Assayable reliablly for AAM (Lab 1.19)
N.A. | Not Applicable (Lab 1.02)
N.A. | Not Applicable (Lab 1.03)
N.A. | Not Applicable (Lab 1.04)
NA | Not Applicable (Lab 1.16)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 612
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Laboratorio Tecnológico de Metalurgia LTM SA de CV, Hermosillo, Sonora, Mexico
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Minerals, Santiago, Chile
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 612 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppb) | Upscaled Value (ppb)
2 | 1 | ANSLu | 1 | 9 | 9 | 307378 | 0.082417 | 66.8678 | 68.326
2 | 1 | ANSLu | 1 | 13 | 7 | 307379 | 0.083205 | 68.775 | 68.4275
2 | 1 | ANSLu | 1 | 17 | 5 | 307380 | 0.082081 | 71.9992 | 68.599
2 | 1 | ANSLu | 1 | 6 | 11 | 307381 | 0.085055 | 67.9254 | 68.3823
2 | 1 | ANSLu | 1 | 1 | 1 | 307382 | 0.086258 | 60.7609 | 68.0012
2 | 1 | ANSLu | 1 | 18 | 5 | 307383 | 0.087558 | 74.7451 | 68.7451
2 | 1 | ANSLu | 1 | 7 | 4 | 307384 | 0.085945 | 65.9277 | 68.276
2 | 1 | ANSLu | 1 | 5 | 11 | 307385 | 0.08392 | 67.407 | 68.3547
2 | 1 | ANSLu | 1 | 16 | 12 | 307386 | 0.086496 | 60.9438 | 68.0109
2 | 1 | ANSLu | 1 | 11 | 2 | 307387 | 0.086519 | 55.5988 | 67.7265
2 | 1 | ANSLu | 1 | 10 | 9 | 307388 | 0.084709 | 63.1371 | 68.1276
2 | 1 | ANSLu | 1 | 2 | 1 | 307389 | 0.086254 | 112.6013 | 70.759
2 | 1 | ANSLu | 1 | 15 | 12 | 307390 | 0.08291 | 65.423 | 68.2492
2 | 1 | ANSLu | 1 | 14 | 7 | 307391 | 0.083377 | 65.4108 | 68.2485
2 | 1 | ANSLu | 1 | 19 | 10 | 307392 | 0.087294 | 73.898 | 68.7
2 | 1 | ANSLu | 1 | 20 | 10 | 307393 | 0.082512 | 62.5741 | 68.0976
2 | 1 | ANSLu | 1 | 4 | 6 | 307394 | 0.087978 | 70.5533 | 68.5221
2 | 1 | ANSLu | 1 | 12 | 2 | 307395 | 0.083241 | 73.1818 | 68.6619
2 | 1 | ANSLu | 1 | 8 | 4 | 307396 | 0.085891 | 61.0142 | 68.0146
2 | 1 | ANSLu | 1 | 3 | 6 | 307397 | 0.084431 | 59.4151 | 67.9296
Mean | [6 empty] | 0.084903 | 68.408 | 68.408
Median | [7 empty] | 66.3978 | 68.301
SD | [7 empty] | 11.6236 | 0.61836
RSD | [7 empty] | 16.99% | 0.9039%
RSD Upscale Factor | [7 empty] | 0.053199
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 612 (Certified Value 62.3 ppb) | [26 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | ppb
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | 1
|  |  | 0.085g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 30g | 30g | 2
| 1 | 1 | 60.7609 | 50 | 55 | 70 | 64 | 60 | 70 | 70 | 60 | 65 | 59 | 61 | 69.2631 | 58 | 70 | 50 | 60 | 79 | 74 | 63 | 61 | 60 | 64 | 62 | 51.8333 | 1
| 1 | 2 | 112.6013 | 50 | 55 | 70 | 65 | 80 | 70 | 70 | 70 | 64 | 59 | 64 | 67.5943 | 58 | 70 | 60 | 60 | 78 | 64 | 63 | 58 | 60 | 62 | 61 | 55.1667 | #N/A
| 1 | 3 | 59.4151 | 50 | 56 | 70 | 63 | 60 | 70 | 60 | 70 | 61 | 58 | 59 | 67.6565 | 61 | 70 | 60 | 60 | 75 | 66 | 60 | 59 | 60 | 64 | 61 | 52 | 16
| 1 | 4 | 70.5533 | 50 | 56 | 70 | 67 | 60 | 70 | 70 | 70 | 61 | 60 | 62 | 67.3035 | 64 | 70 | 60 | 60 | 77 | 68 | 63 | 58 | 60 | 63 | 63 | 52.6667 | 62.2892
| 1 | 5 | 67.407 | 50 | 60 | 70 | 62 | 60 | 70 | 70 | 70 | 61 | 59 | 62 | 68.8417 | 53 | 70 | 60 | 70 | 77 | 68 | 60 | 60 | 60 | 66 | 60 | 52.1667 | 7
| 1 | 6 | 67.9254 | 50 | 50 | 70 | 65 | 60 | 70 | 70 | 60 | 61 | 59 | 66 | 69.0877 | 56 | 70 | 60 | 60 | 73 | 56 | 63 | 58 | 60 | 62 | 59 | 54.3333
|  | 7 | 65.9277
|  | 8 | 61.0142
|  | 9 | 66.8678
|  | 10 | 63.1371
|  | 11 | 55.5988
|  | 12 | 73.1818
|  | 13 | 68.775
|  | 14 | 65.4108
|  | 15 | 65.423
|  | 16 | 60.9438
|  | 17 | 71.9992
|  | 18 | 74.7451
|  | 19 | 73.898
|  | 20 | 62.5741
| Mean |  | 68.408 | 50 | 55.3333 | 70 | 64.3333 | 63.3333 | 70 | 68.3333 | 66.6667 | 62.1667 | 59 | 62.3333 | 68.2911 | 58.3333 | 70 | 58.3333 | 61.6667 | 76.5 | 66 | 62 | 59 | 60 | 63.5 | 61 | 53.0278
| Median |  | 66.3978 | 50 | 55.5 | 70 | 64.5 | 60 | 70 | 70 | 70 | 61 | 59 | 62 | 68.2491 | 58 | 70 | 60 | 60 | 77 | 67 | 63 | 58.5 | 60 | 63.5 | 61 | 52.4167
| Std Dev. |  | 11.6236 | 0 | 3.2042 | 0 | 1.7512 | 8.165 | 0 | 4.0825 | 5.164 | 1.8348 | 0.63246 | 2.4221 | 0.86558 | 3.8297 | 0 | 4.0825 | 4.0825 | 2.1679 | 5.933 | 1.5492 | 1.2649 | 0 | 1.5166 | 1.4142 | 1.3881
| Rel.Std.Dev. |  | 16.99% | 0% | 5.791% | 0% | 2.722% | 12.89% | 0% | 5.974% | 7.746% | 2.951% | 1.072% | 3.886% | 1.267% | 6.565% | 0% | 6.999% | 6.62% | 2.834% | 8.989% | 2.499% | 2.144% | 0% | 2.388% | 2.318% | 2.618%
| PDM3 |  | 9.823% | -19.73% | -11.17% | 12.38% | 3.282% | 1.676% | 12.38% | 9.703% | 7.028% | -0.1968% | -5.281% | 0.07081% | 9.636% | -6.351% | 12.38% | -6.351% | -0.9995% | 22.81% | 5.957% | -0.4643% | -5.281% | -3.675% | 1.944% | -2.07% | -14.87%
| Z-Score (Absolute) |  | NA | 2.15 | 1.22 | 1.36 | 0.37 | 0.19 | 1.36 | 1.07 | 0.78 | 0.01 | 0.57 | 0.01 | 1.06 | 0.69 | 1.36 | 0.69 | 0.1 | 2.51 | 0.66 | 0.04 | 0.57 | 0.4 | 0.22 | 0.22 | 1.62
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fire Assay (Grav)
| Analytical results for Ag in OREAS 612 (Indicative Value 54.5 ppm) | [8 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.13 | 1.16 | 1.18 | 1.19 | 1.20 | ppm
|  |  | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | 1
|  |  | 30g | 30g | 30g | 10g | 30g | 30g | 30g | 1
| 1 | 1 | N.A. | N.A. | < 20 | NA | 54.4059 | --- | -- | 1
| 1 | 2 | N.A. | N.A. | < 20 | NA | 54.7248 | --- | -- | 1
| 1 | 3 | N.A. | N.A. | < 20 | NA | 53.7409 | --- | -- | 16
| 1 | 4 | N.A. | N.A. | < 20 | NA | 54.2336 | --- | -- | 54.4577
| 1 | 5 | N.A. | N.A. | < 20 | NA | 55.1499 | --- | -- | 7
| 1 | 6 | N.A. | N.A. | < 20 | NA | 54.4911 | --- | --
| Mean | [5 empty] | 54.4577
| Median | [5 empty] | 54.4485
| Std Dev. | [5 empty] | 0.47292
| Rel.Std.Dev. | [5 empty] | 0.8684%
| PDM3 | [5 empty] | 0%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 612 (Certified Value 58.3 ppb) | [17 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.24 | ppb
|  |  | INAA | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | 1
|  |  | 0.085g | 25g | 25g | 25g | 25g | 25g | 30g | 10g | 25g | 30g | 30g | 20g | 30g | 25g | 50g | 10g | 2
| 1 | 1 | 60.7609 | 60 | 40 | 60 | 60 | 59 | 63.3 | < 20 | 50 | < 50 | 50 | 64 |  | 60 | 60 | 54.9992 | 1
| 1 | 2 | 112.6013 | 60 | 40 | 50 | 60 | 58 | 60.6 | < 20 | 50 | < 50 | 50 | 67 |  | 57 | 50 | 59.9973 | #N/A
| 1 | 3 | 59.4151 | 60 | 40 | 60 | 60 | 58 | 59.4 | < 20 | 50 | < 50 | 60 | 73 |  | 59 | 60 | 44.9986 | 16
| 1 | 4 | 70.5533 | 60 | 40 | 60 | 60 | 60 | 58.9 | < 20 | 50 | < 50 | 60 | 80 |  | 57 | 60 | 44.9973 | 58.3257
| 1 | 5 | 67.407 | 60 | 50 | 60 | 60 | 61 | 62.3 | < 20 | 50 | < 50 | 60 | 71 |  | 60 | 60 | 49.9978 | 9
| 1 | 6 | 67.9254 | 60 | 50 | 60 | 60 | 58 | 60.1 | < 20 | 50 | < 50 | 50 | 75 |  | 58 | 60 | 54.9967
|  | 7 | 65.9277
|  | 8 | 61.0142
|  | 9 | 66.8678
|  | 10 | 63.1371
|  | 11 | 55.5988
|  | 12 | 73.1818
|  | 13 | 68.775
|  | 14 | 65.4108
|  | 15 | 65.423
|  | 16 | 60.9438
|  | 17 | 71.9992
|  | 18 | 74.7451
|  | 19 | 73.898
|  | 20 | 62.5741
| Mean |  | 68.408 | 60 | 43.3333 | 58.3333 | 60 | 59 | 60.7667 |  | 50 |  | 55 | 71.6667 |  | 58.5 | 58.3333 | 51.6645
| Median |  | 66.3978 | 60 | 40 | 60 | 60 | 58.5 | 60.35 |  | 50 |  | 55 | 72 |  | 58.5 | 60 | 52.4972
| Std Dev. |  | 11.6236 | 0 | 5.164 | 4.0825 | 0 | 1.2649 | 1.7084 |  | 0 |  | 5.4772 | 5.7155 |  | 1.3784 | 4.0825 | 6.0551
| Rel.Std.Dev. |  | 16.99% | 0% | 11.92% | 6.999% | 0% | 2.144% | 2.811% |  | 0% |  | 9.959% | 7.975% |  | 2.356% | 6.999% | 11.72%
| PDM3 |  | 17.29% | 2.871% | -25.7% | 0.01313% | 2.871% | 1.156% | 4.185% |  | -14.27% |  | -5.702% | 22.87% |  | 0.2989% | 0.01313% | -11.42%
| Z-Score (Absolute) |  | NA | 0.39 | 3.51 | 0 | 0.39 | 0.16 | 0.57 | 11.3 | 1.95 | 7.8 | 0.78 | 3.12 | NA | 0.04 | 0 | 1.56
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 612 (Certified Value 48.2 ppm) | [23 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 48.5 | 46.39 | 50.25 | 48 | 49.1 | 47 | 48 | 48.19 | 49 | 45.4 | 48.1295 | 47.62 | 49.1 | 49.3 | 49.5 | 48.4 | 47.6251 | 46.4 | 48.98 | 46.43 | 50.7727 |  | 1
| 1 | 2 | 48.58 | 46.23 | 50.04 | 49 | 47.9 | 48 | 48 | 46.99 | 47 | 46.5 | 48.7735 | 48.53 | 50.3 | 50 | 50 | 48 | 47.9043 | 45.6 | 48.86 | 46.53 | 50.4347 |  | 1
| 1 | 3 | 48.83 | 48.26 | 50.06 | 48 | 48.9 | 47 | 50 | 47.63 | 47 | 45.8 | 49.1374 | 48.11 | 47.9 | 48.4 | 50.3 | 46.6 | 47.043 | 46.1 | 49.2 | 46.96 | 50.258 |  | 16
| 1 | 4 | 49.49 | 49.35 | 49.61 | 49 | 49 | 46 | 48 | 46.7 | 46 | 47.5 | 48.0833 | 48.63 | 47.8 | 49.9 | 50.4 | 46.1 | 47.4743 | 45.7 | 49.2 | 46.72 | 50.767 |  | 48.1767
| 1 | 5 | 48.52 | 48.56 | 50.66 | 47 | 47.9 | 48 | 47 | 46.62 | 47 | 46.7 | 47.8779 | 48.2 | 49.8 | 46.9 | 50.4 | 46.4 | 47.401 | 46 | 48.07 | 47.18 | 51.188 |  | 11
| 1 | 6 | 48.39 | 49.39 | 51.16 | 49 | 48.3 | 47 | 48 | 47.29 | 46 | 46.8 | 48.0022 | 48.77 | 48.7 | 51 | 50.4 | 45.4 | 47.6997 | 46.9 | 48.53 | 47.3 | 50.6877
| Mean |  | 48.7183 | 48.03 | 50.2967 | 48.3333 | 48.5167 | 47.1667 | 48.1667 | 47.2367 | 47 | 46.45 | 48.334 | 48.31 | 48.9333 | 49.25 | 50.1667 | 46.8167 | 47.5246 | 46.1167 | 48.8067 | 46.8533 | 50.6847
| Median |  | 48.55 | 48.41 | 50.155 | 48.5 | 48.6 | 47 | 48 | 47.14 | 47 | 46.6 | 48.1064 | 48.365 | 48.9 | 49.6 | 50.35 | 46.5 | 47.5497 | 46.05 | 48.92 | 46.84 | 50.7273
| Std Dev. |  | 0.40543 | 1.404 | 0.54276 | 0.8165 | 0.55287 | 0.75277 | 0.98319 | 0.59926 | 1.0954 | 0.75033 | 0.50228 | 0.42204 | 1.0053 | 1.4349 | 0.36148 | 1.1531 | 0.29485 | 0.47924 | 0.43853 | 0.35189 | 0.32013
| Rel.Std.Dev. |  | 0.8322% | 2.923% | 1.079% | 1.689% | 1.14% | 1.596% | 2.041% | 1.269% | 2.331% | 1.615% | 1.039% | 0.8736% | 2.054% | 2.914% | 0.7206% | 2.463% | 0.6204% | 1.039% | 0.8985% | 0.751% | 0.6316%
| PDM3 |  | 1.124% | -0.3046% | 4.4% | 0.325% | 0.7056% | -2.097% | -0.02091% | -1.951% | -2.443% | -3.584% | 0.3264% | 0.2766% | 1.57% | 2.228% | 4.13% | -2.823% | -1.354% | -4.276% | 1.308% | -2.747% | 5.206%
| Z-Score (Absolute) |  | 0.29 | 0.2 | 1.43 | 0.02 | 0.15 | 0.82 | 0.1 | 0.77 | 0.94 | 1.33 | 0.02 | 0 | 0.45 | 0.67 | 1.33 | 1.07 | 0.56 | 1.57 | 0.36 | 1.04 | 1.7
| Analytical results for Al in OREAS 612 (Certified Value 7.22 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.44 | 2.02 | 6.87 | 7.49 | 3.6 | 7.2 | 6.67 | 4.91 | 7.3577 | 7.3 | 7.66 | 7.018 | 7.0364 | 7.58 | 7.44 | 7.13 | 6.11 | 7.1806 | 7.01 | 7.3133 | 7.26 | 6.8819 | 1
| 1 | 2 | 7.11 | 3.95 | 7.309 | 7.47 | 3.37 | 7.33 | 6.93 | 5.2 | 7.1269 | 7.04 | 7.78 | 7.0462 | 7.2064 | 7.69 | 7.54 | 7 | 6.48 | 7.1662 | 6.94 | 7.2823 | 7.23 | 6.7112 | #N/A
| 1 | 3 | 7.54 | 2.99 | 7.028 | 7.51 | 4.28 | 7.39 | 6.96 | 5.41 | 7.3469 | 7.11 | 7.48 | 6.8955 | 7.1918 | 7.36 | 7.53 | 7.14 | 6.32 | 7.1659 | 6.97 | 7.3161 | 7.27 | 6.5109 | 16
| 1 | 4 | 7.32 | 2.81 | 7.134 | 7.42 | 3.41 | 7.44 | 7.5 | 5.15 | 7.2018 | 6.97 | 7.25 | 7.0191 | 7.2104 | 7.21 | 7.52 | 7.14 | 6.69 | 7.1926 | 6.9 | 7.361 | 7.19 | 6.8106 | 7.2195
| 1 | 5 | 6.97 | 4.18 | 7.076 | 7.54 | 3.8 | 7.33 | 7.12 | 6.02 | 7.1929 | 7 | 7.91 | 7.0513 | 7.1628 | 7.52 | 7.54 | 7.17 | 6.21 | 7.1761 | 7.04 | 7.2922 | 7.28 | 6.729 | 12
| 1 | 6 | 6.92 | 4.12 | 7.288 | 7.66 | 3.01 | 7.42 | 7.1 | 5.27 | 7.304 | 6.91 | 8.02 | 7.1193 | 7.2164 | 7.53 | 7.54 | 6.99 | 6.16 | 7.1867 | 7.08 | 7.3684 | 7.26 | 6.8449
| Mean |  | 7.2167 | 3.345 | 7.1175 | 7.515 | 3.5783 | 7.3517 | 7.0467 | 5.3267 | 7.255 | 7.055 | 7.6833 | 7.0249 | 7.1707 | 7.4817 | 7.5183 | 7.095 | 6.3283 | 7.178 | 6.99 | 7.3222 | 7.2483 | 6.7481
| Median |  | 7.215 | 3.47 | 7.105 | 7.5 | 3.505 | 7.36 | 7.03 | 5.235 | 7.2529 | 7.02 | 7.72 | 7.0326 | 7.1991 | 7.525 | 7.535 | 7.135 | 6.265 | 7.1783 | 6.99 | 7.3147 | 7.26 | 6.7698
| Std Dev. |  | 0.25508 | 0.87541 | 0.16553 | 0.08167 | 0.43282 | 0.087044 | 0.27449 | 0.37729 | 0.094329 | 0.13751 | 0.28416 | 0.073323 | 0.068532 | 0.17058 | 0.0392 | 0.078677 | 0.22085 | 0.010802 | 0.066332 | 0.035351 | 0.033116 | 0.13356
| Rel.Std.Dev. |  | 3.535% | 26.17% | 2.326% | 1.087% | 12.1% | 1.184% | 3.895% | 7.083% | 1.3% | 1.949% | 3.698% | 1.044% | 0.9557% | 2.28% | 0.5214% | 1.109% | 3.49% | 0.1505% | 0.949% | 0.4828% | 0.4569% | 1.979%
| PDM3 |  | -0.039% | -53.67% | -1.413% | 4.093% | -50.44% | 1.831% | -2.394% | -26.22% | 0.4924% | -2.278% | 6.425% | -2.695% | -0.6757% | 3.632% | 4.14% | -1.724% | -12.34% | -0.5744% | -3.179% | 1.423% | 0.3996% | -6.53%
| Z-Score (Absolute) |  | 0.29 | 15.42 | 0.11 | 1.51 | 14.47 | 0.84 | 0.4 | 7.38 | 0.45 | 0.36 | 2.19 | 0.48 | 0.11 | 1.37 | 1.52 | 0.2 | 3.31 | 0.14 | 0.63 | 0.72 | 0.42 | 1.61
| Analytical results for As in OREAS 612 (Certified Value 331 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 330 | 336 | 307.47 | 329 | 300 | 333 | 320 | 330 | 329 | 336 | 351 | 338.7823 | 317.2 | 338 | 329 | 308 | 173.9 | 321 | 332.9 | 327 | 344.1237 |  | 1
| 1 | 2 | 330 | 324 | 323.6 | 330 | 300 | 325 | 330 | 330 | 334.9 | 332 | 369 | 340.1945 | 325.3 | 349 | 331 | 304 | 170.8 | 322 | 334.5 | 328 | 345.5427 |  | 20
| 1 | 3 | 330 | 339 | 315.85 | 331 | 300 | 339 | 350 | 340 | 333.7 | 334 | 366 | 340.0346 | 301.5 | 339 | 326 | 314 | 165.5 | 317 | 333.7 | 328 | 343.2507 |  | 16
| 1 | 4 | 330 | 350 | 324.91 | 325 | 300 | 330 | 320 | 330 | 332.9 | 327 | 378 | 340.7997 | 324.9 | 333 | 332 | 314 | 187.5 | 320 | 338.3 | 324 | 353.2137 |  | 331.0075
| 1 | 5 | 320 | 359 | 324.46 | 327 | 290 | 328 | 340 | 330 | 337.3 | 321 | 363 | 339.0928 | 324 | 349 | 332 | 310 | 163.5 | 314 | 331.1 | 329 | 354.5835 |  | 13
| 1 | 6 | 330 | 349 | 331.7 | 331 | 320 | 332 | 340 | 340 | 333.7 | 329 | 365 | 340.9501 | 328.9 | 349 | 334 | 305 | 165.3 | 320 | 335.6 | 329 | 347.2927
| Mean |  | 328.3333 | 342.8333 | 321.3317 | 328.8333 | 301.6667 | 331.1667 | 333.3333 | 333.3333 | 333.5833 | 329.8333 | 365.3333 | 339.9757 | 320.3 | 342.8333 | 330.6667 | 309.1667 | 171.0833 | 319 | 334.35 | 327.5 | 348.0011
| Median |  | 330 | 344 | 324.03 | 329.5 | 300 | 331 | 335 | 330 | 333.7 | 330.5 | 365.5 | 340.1145 | 324.45 | 344 | 331.5 | 309 | 168.15 | 320 | 334.1 | 328 | 346.4177
| Std Dev. |  | 4.0825 | 12.3841 | 8.455 | 2.4014 | 9.8319 | 4.7924 | 12.1106 | 5.164 | 2.7221 | 5.4191 | 8.7788 | 0.88127 | 9.9686 | 7.0545 | 2.8048 | 4.3089 | 8.9403 | 2.9665 | 2.4607 | 1.8708 | 4.7884
| Rel.Std.Dev. |  | 1.243% | 3.612% | 2.631% | 0.7303% | 3.259% | 1.447% | 3.633% | 1.549% | 0.816% | 1.643% | 2.403% | 0.2592% | 3.112% | 2.058% | 0.8482% | 1.394% | 5.226% | 0.9299% | 0.736% | 0.5712% | 1.376%
| PDM3 |  | -0.8079% | 3.573% | -2.923% | -0.6568% | -8.864% | 0.04808% | 0.7027% | 0.7027% | 0.7782% | -0.3547% | 10.37% | 2.709% | -3.235% | 3.573% | -0.103% | -6.598% | -48.31% | -3.628% | 1.01% | -1.06% | 5.134%
| Z-Score (Absolute) |  | 0.17 | 0.88 | 0.68 | 0.13 | 2.1 | 0.04 | 0.19 | 0.19 | 0.21 | 0.06 | 2.51 | 0.67 | 0.75 | 0.88 | 0 | 1.56 | 11.56 | 0.85 | 0.27 | 0.23 | 1.26
| Analytical results for B in OREAS 612 (Indicative Value 8.83 ppm) | [23 empty] | Indicative
B | Round | Replicate | Lab | [21 empty] | 1
| No | No | 1.15 | [21 empty] | ppm
|  |  | 4A*OES | [21 empty] | 2
2
| 1 | 1 | 9 | [21 empty] | 1
| 1 | 2 | 9 | [21 empty] | 3
| 1 | 3 | 9 | [21 empty] | 16
| 1 | 4 | 9 | [21 empty] | 8.8333
| 1 | 5 | 8 | [21 empty] | 9
| 1 | 6 | 9
| Mean |  | 8.8333
| Median |  | 9
| Std Dev. |  | 0.40825
| Rel.Std.Dev. |  | 4.622%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 612 (Certified Value 2163 ppm) | [23 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 2144 | 2079 | 3702.72 | 2177 | 2030 | 2150 | 2080 | 2080 | 2173.6 | 2067 | 2220 | 2176.0437 | 2129 | 2230 | 2160 | 2210 | > 1000 | 2104 | 2202 |  |  |  | 1
| 1 | 2 | 2123 | 2084 | 3917.36 | 2165 | 2030 | 2120 | 2150 | 2070 | 2156.2 | 1809 | 2250 | 2206.5346 | 2163 | 2290 | 2199 | 2220 | > 1000 | 2157 | 2196 |  |  |  | 21
| 1 | 3 | 2146 | 2144 | 3586.07 | 2178 | 2060 | 2190 | 2120 | 2120 | 2129.9 | 1896 | 2230 | 2228.1521 | 2155 | 2330 | 2156 | 2270 | > 1000 | 2185 | 2184 |  |  |  | 16
| 1 | 4 | 2195 | 2215 | 3922.22 | 2162 | 2050 | 2160 | 2090 | 2130 | 2104.5 | 1902 | 2310 | 2159.6018 | 2160 | 2280 | 2247 | 2270 | > 1000 | 2177 | 2224 |  |  |  | 2163.2808
| 1 | 5 | 2124 | 2239 | 3683.49 | 2169 | 2040 | 2150 | 2150 | 2140 | 2155 | 1943 | 2250 | 2177.6788 | 2159 | 2370 | 2146 | 2270 | > 1000 | 2145 | 2226 |  |  |  | 14
| 1 | 6 | 2111 | 2165 | 3782.57 | 2229 | 2040 | 2170 | 2180 | 2110 | 2152.2 | 1917 | 2280 | 2219.8618 | 2181 | 2390 | 2191 | 2220 | > 1000 | 2117 | 2234
| Mean |  | 2140.5 | 2154.3333 | 3765.7383 | 2180 | 2041.6667 | 2156.6667 | 2128.3333 | 2108.3333 | 2145.2333 | 1922.3333 | 2256.6667 | 2194.6455 | 2157.8333 | 2315 | 2183.1667 | 2243.3333 |  | 2147.5 | 2211
| Median |  | 2134 | 2154.5 | 3742.645 | 2173 | 2040 | 2155 | 2135 | 2115 | 2153.6 | 1909.5 | 2250 | 2192.1067 | 2159.5 | 2310 | 2175.5 | 2245 |  | 2151 | 2213
| Std Dev. |  | 29.8714 | 65.8655 | 134.7411 | 24.8355 | 11.6905 | 23.3809 | 38.6868 | 27.8687 | 24.3441 | 84.0658 | 33.2666 | 27.4294 | 16.7859 | 59.9166 | 37.5309 | 29.4392 |  | 32.2351 | 19.7889
| Rel.Std.Dev. |  | 1.396% | 3.057% | 3.578% | 1.139% | 0.5726% | 1.084% | 1.818% | 1.322% | 1.135% | 4.373% | 1.474% | 1.25% | 0.7779% | 2.588% | 1.719% | 1.312% |  | 1.501% | 0.895%
| PDM3 |  | -1.053% | -0.4136% | 74.08% | 0.7729% | -5.622% | -0.3057% | -1.615% | -2.54% | -0.8343% | -11.14% | 4.317% | 1.45% | -0.2518% | 7.013% | 0.9192% | 3.701% |  | -0.7295% | 2.206%
| Z-Score (Absolute) |  | 0.34 | 0.06 | 32.71 | 0.46 | 2.35 | 0.01 | 0.59 | 0.99 | 0.24 | 4.78 | 2.02 | 0.76 | 0.01 | 3.21 | 0.53 | 1.75 | NA | 0.2 | 1.09
| Analytical results for Be in OREAS 612 (Certified Value 2.48 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.59 | 2.7 | 2.33 | 2.58 | < 10 | 2.45 | < 10 | < 10 | 2.48 | 2.4 | 2.8 | 2.4607 | 2.27 | 2.46 | < 5 | 2.5 | 3 | 2.5245 | 2.6 | 2.58 | 2.61 | 2.4087 | 1
| 1 | 2 | 2.79 | 2.8 | 2.49 | 2.58 | < 10 | 2.41 | < 10 | < 10 | 2.44 | 2.3 | 2.5 | 2.4877 | 2.31 | 2.55 | < 5 | 2.3 | 3 | 2.5269 | 2.4 | 2.47 | 2.6 | 2.405 | 22
| 1 | 3 | 2.58 | 2.9 | 2.41 | 2.73 | < 10 | 2.48 | < 10 | < 10 | 2.52 | 2.3 | 2.4 | 2.4914 | 2.3 | 2.44 | < 5 | 2.5 | 3 | 2.5363 | 2.5 | 2.5 | 2.59 | 2.4303 | 16
| 1 | 4 | 2.6 | 2.9 | 2.47 | 2.67 | < 10 | 2.46 | < 10 | < 10 | 2.64 | 2.3 | 2.6 | 2.4446 | 2.3 | 2.4 | < 5 | 2.4 | 3 | 2.6166 | 2.5 | 2.43 | 2.57 | 2.4463 | 2.4848
| 1 | 5 | 2.52 | 3.1 | 2.48 | 2.62 | < 10 | 2.45 | < 10 | < 10 | 2.55 | 2.3 | 2.9 | 2.4327 | 2.3 | 2.49 | < 5 | 2.5 | 3 | 2.4994 | 2.6 | 2.48 | 2.59 | 2.4415 | 15
| 1 | 6 | 2.52 | 2.9 | 2.48 | 2.7 | < 10 | 2.48 | < 10 | < 10 | 2.49 | 2.2 | 3 | 2.4646 | 2.32 | 2.52 | < 5 | 2.4 | 3 | 2.4532 | 2.6 | 2.5 | 2.62 | 2.4607
| Mean |  | 2.6 | 2.8833 | 2.4433 | 2.6467 |  | 2.455 |  |  | 2.52 | 2.3 | 2.7 | 2.4636 | 2.3 | 2.4767 |  | 2.4333 | 3 | 2.5261 | 2.5333 | 2.4933 | 2.5967 | 2.4321
| Median |  | 2.585 | 2.9 | 2.475 | 2.645 |  | 2.455 |  |  | 2.505 | 2.3 | 2.7 | 2.4626 | 2.3 | 2.475 |  | 2.45 | 3 | 2.5257 | 2.55 | 2.49 | 2.595 | 2.4359
| Std Dev. |  | 0.099398 | 0.13292 | 0.062503 | 0.06314 |  | 0.025884 |  |  | 0.06957 | 0.063246 | 0.23664 | 0.023155 | 0.016733 | 0.05465 |  | 0.08165 | 0 | 0.053545 | 0.08165 | 0.049666 | 0.017512 | 0.02188
| Rel.Std.Dev. |  | 3.823% | 4.61% | 2.558% | 2.386% |  | 1.054% |  |  | 2.761% | 2.75% | 8.765% | 0.9399% | 0.7275% | 2.207% |  | 3.355% | 0% | 2.12% | 3.223% | 1.992% | 0.6744% | 0.8996%
| PDM3 |  | 4.635% | 16.04% | -1.67% | 6.513% |  | -1.2% |  |  | 1.416% | -7.438% | 8.66% | -0.8532% | -7.438% | -0.3283% |  | -2.072% | 20.73% | 1.662% | 1.952% | 0.3425% | 4.501% | -2.122%
| Z-Score (Absolute) |  | 0.67 | 3.06 | 0.65 | 1.07 | 20.9 | 0.55 | 20.9 | 20.9 | 0 | 1.85 | 1.52 | 0.48 | 1.85 | 0.37 | 0.17 | 0.73 | NA | 0.05 | 0.11 | 0.22 | 0.65 | 0.74
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 612 (Certified Value 0.85 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 0.87 | 0.75 | < 1 | 0.86 | < 20 | 0.85 | < 20 | < 20 | 0.86 | < 5 | 0.8 | 0.90926 | 0.93 | 0.87 | 1.45 | 0.95 | 0.7 | 0.69821 | 0.72 | 0.94 | 0.83 | < 1 | 1
| 1 | 2 | 0.95 | 0.8 | < 1 | 0.83 | < 20 | 0.83 | < 20 | < 20 | 0.98 | < 5 | 0.84 | 0.91591 | 0.96 | 0.92 | 1.51 | 0.95 | 0.7 | 0.71371 | 0.7 | 0.88 | 0.85 | < 1 | 23
| 1 | 3 | 0.88 | 0.73 | 1.74 | 0.87 | < 20 | 0.81 | < 20 | < 20 | 0.9 | < 5 | 0.86 | 0.88621 | 0.94 | 0.94 | 1.51 | 0.99 | 0.7 | 0.71585 | 0.74 | 0.89 | 0.83 | < 1 | 16
| 1 | 4 | 0.88 | 0.73 | < 1 | 0.91 | < 20 | 0.86 | < 20 | < 20 | 0.84 | < 5 | 0.84 | 0.86511 | 0.92 | 0.87 | 1.5 | 0.99 | 0.8 | 0.70887 | 0.71 | 0.89 | 0.83 | < 1 | 0.85257
| 1 | 5 | 0.83 | 0.81 | < 1 | 0.86 | < 20 | 0.84 | < 20 | < 20 | 0.89 | < 5 | 0.81 | 0.91057 | 0.93 | 0.94 | 1.47 | 0.97 | 0.7 | 0.68626 | 0.76 | 0.89 | 0.82 | < 1 | 16
| 1 | 6 | 0.91 | 0.75 | < 1 | 0.87 | < 20 | 0.83 | < 20 | < 20 | 0.86 | < 5 | 0.81 | 0.85864 | 0.94 | 0.89 | 1.54 | 0.98 | 0.7 | 0.71767 | 0.72 | 0.9 | 0.85 | < 1
| Mean |  | 0.88667 | 0.76167 | 1.74 | 0.86667 |  | 0.83667 |  |  | 0.88833 |  | 0.82667 | 0.89095 | 0.93667 | 0.905 | 1.4967 | 0.97167 | 0.71667 | 0.70676 | 0.725 | 0.89833 | 0.835
| Median |  | 0.88 | 0.75 | 1.74 | 0.865 |  | 0.835 |  |  | 0.875 |  | 0.825 | 0.89774 | 0.935 | 0.905 | 1.505 | 0.975 | 0.7 | 0.71129 | 0.72 | 0.89 | 0.83
| Std Dev. |  | 0.040332 | 0.034881 |  | 0.02582 |  | 0.017512 |  |  | 0.049967 |  | 0.023381 | 0.024807 | 0.013663 | 0.032711 | 0.032042 | 0.018348 | 0.040825 | 0.012228 | 0.021679 | 0.02137 | 0.012247
| Rel.Std.Dev. |  | 4.549% | 4.58% |  | 2.979% |  | 2.093% |  |  | 5.625% |  | 2.828% | 2.784% | 1.459% | 3.614% | 2.141% | 1.888% | 5.696% | 1.73% | 2.99% | 2.379% | 1.467%
| PDM3 |  | 3.999% | -10.66% | 104.1% | 1.653% |  | -1.866% |  |  | 4.194% |  | -3.039% | 4.501% | 9.863% | 6.149% | 75.55% | 13.97% | -15.94% | -17.1% | -14.96% | 5.367% | -2.061%
| Z-Score (Absolute) |  | 0.01 | 1.02 | 1.47 | 0.18 | 73.73 | 0.42 | 73.73 | 73.73 | 0 | 13.04 | 0.5 | 0.02 | 0.39 | 0.13 | 4.92 | 0.67 | NA | 1.47 | 1.32 | 0.08 | 0.43 | 3.14
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 612 (Certified Value 0.705 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.7 | 0.22 | 0.78 | 0.721 | 0.49 | 0.7 | 0.69 | 0.54 | 0.726 | 0.701 | 0.7 | 0.70187 | 0.7349 | 0.72 | 0.69 | 0.677 | 0.66 | 0.72423 | 0.66 | 0.7277 | 0.7 | 0.83219 | 1
| 1 | 2 | 0.67 | 0.44 | 0.83 | 0.717 | 0.48 | 0.7 | 0.7 | 0.54 | 0.7054 | 0.673 | 0.72 | 0.70993 | 0.7503 | 0.75 | 0.7 | 0.673 | 0.67 | 0.72795 | 0.68 | 0.718 | 0.7 | 0.83049 | #N/A
| 1 | 3 | 0.7 | 0.32 | 0.79 | 0.719 | 0.51 | 0.71 | 0.67 | 0.56 | 0.7259 | 0.68 | 0.72 | 0.71643 | 0.7457 | 0.74 | 0.7 | 0.688 | 0.65 | 0.72942 | 0.67 | 0.7168 | 0.71 | 0.8255 | 16
| 1 | 4 | 0.73 | 0.28 | 0.81 | 0.713 | 0.48 | 0.71 | 0.71 | 0.53 | 0.7132 | 0.667 | 0.74 | 0.7004 | 0.7493 | 0.72 | 0.72 | 0.687 | 0.67 | 0.72536 | 0.67 | 0.7263 | 0.7 | 0.83261 | 0.70471
| 1 | 5 | 0.67 | 0.46 | 0.79 | 0.723 | 0.51 | 0.7 | 0.71 | 0.57 | 0.7114 | 0.67 | 0.71 | 0.69031 | 0.7458 | 0.75 | 0.69 | 0.691 | 0.64 | 0.72311 | 0.69 | 0.7304 | 0.71 | 0.82758 | 17
| 1 | 6 | 0.67 | 0.48 | 0.81 | 0.735 | 0.49 | 0.71 | 0.72 | 0.54 | 0.7217 | 0.663 | 0.71 | 0.70422 | 0.7504 | 0.76 | 0.7 | 0.673 | 0.63 | 0.72231 | 0.67 | 0.7346 | 0.71 | 0.85742
| Mean |  | 0.69 | 0.36667 | 0.80167 | 0.72133 | 0.49333 | 0.705 | 0.7 | 0.54667 | 0.71727 | 0.67567 | 0.71667 | 0.70386 | 0.74607 | 0.74 | 0.7 | 0.6815 | 0.65333 | 0.7254 | 0.67333 | 0.72563 | 0.705 | 0.8343
| Median |  | 0.685 | 0.38 | 0.8 | 0.72 | 0.49 | 0.705 | 0.705 | 0.54 | 0.71745 | 0.6715 | 0.715 | 0.70305 | 0.74755 | 0.745 | 0.7 | 0.682 | 0.655 | 0.7248 | 0.67 | 0.727 | 0.705 | 0.83134
| Std Dev. |  | 0.024495 | 0.10783 | 0.018348 | 0.0075277 | 0.013663 | 0.0054772 | 0.017889 | 0.015055 | 0.0085116 | 0.013677 | 0.013663 | 0.0088832 | 0.0058661 | 0.016733 | 0.010954 | 0.0080932 | 0.01633 | 0.0027872 | 0.010328 | 0.0069876 | 0.0054772 | 0.011653
| Rel.Std.Dev. |  | 3.55% | 29.41% | 2.289% | 1.044% | 2.769% | 0.7769% | 2.556% | 2.754% | 1.187% | 2.024% | 1.906% | 1.262% | 0.7863% | 2.261% | 1.565% | 1.188% | 2.499% | 0.3842% | 1.534% | 0.963% | 0.7769% | 1.397%
| PDM3 |  | -2.087% | -47.97% | 13.76% | 2.359% | -29.99% | 0.04124% | -0.6683% | -22.43% | 1.782% | -4.121% | 1.697% | -0.1203% | 5.869% | 5.008% | -0.6683% | -3.293% | -7.29% | 2.936% | -4.452% | 2.969% | 0.04124% | 18.39%
| Z-Score (Absolute) |  | 0.44 | 10.32 | 2.97 | 0.52 | 6.45 | 0.02 | 0.14 | 4.82 | 0.39 | 0.88 | 0.37 | 0.02 | 1.27 | 1.09 | 0.14 | 0.7 | 1.56 | 0.64 | 0.95 | 0.65 | 0.02 | 3.97
| Analytical results for Cd in OREAS 612 (Certified Value 13.8 ppm) | [23 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 14.2 | 13.56 | 12.72 | 14.38 | 10 | 15.9 | 10 | 10 | 14.23 | 14 | 14.9 | 14.068 | 12.47 | 14 | 13.59 | 13.7 | 13.2 | 13.6577 | 14.1 | 13.7 | 13.79 | 13.8657 | 1
| 1 | 2 | 14.23 | 13.86 | 13.09 | 14.48 | 10 | 15.1 | 10 | 10 | 13.89 | 13 | 15.4 | 14.0198 | 12.9 | 13.9 | 14.44 | 13.7 | 12.8 | 13.7876 | 13.9 | 13.8 | 13.76 | 13.8823 | 24
| 1 | 3 | 14.37 | 13.87 | 12.92 | 14.46 | 10 | 15.5 | 10 | 10 | 13.63 | 14 | 14.9 | 14.0143 | 12.94 | 13.25 | 13.87 | 14.1 | 11.9 | 13.8604 | 13.9 | 13.6 | 13.54 | 13.974 | 16
| 1 | 4 | 14.32 | 13.93 | 13.21 | 14.23 | 10 | 15.6 | 10 | 10 | 13.65 | 14 | 14.5 | 13.8531 | 12.99 | 13.5 | 13.99 | 14.1 | 13.5 | 13.8619 | 13.9 | 13.9 | 13.27 | 14.123 | 13.838
| 1 | 5 | 14.14 | 14.25 | 13.24 | 14.55 | 10 | 15.1 | 10 | 10 | 13.85 | 14 | 15.3 | 13.7791 | 12.9 | 14.1 | 13.54 | 13.7 | 12.1 | 13.8085 | 14.2 | 13.8 | 13.05 | 14.108 | 18
| 1 | 6 | 14.09 | 14.13 | 13.31 | 14.56 | 10 | 15.2 | 10 | 10 | 13.56 | 13 | 15.2 | 13.9839 | 13.29 | 13.2 | 14.01 | 13.5 | 13.2 | 13.9179 | 14.1 | 13.9 | 13.38 | 14.165
| Mean |  | 14.225 | 13.9333 | 13.0817 | 14.4433 | 10 | 15.4 | 10 | 10 | 13.8017 | 13.6667 | 15.0333 | 13.953 | 12.915 | 13.6583 | 13.9067 | 13.8 | 12.7833 | 13.8157 | 14.0167 | 13.7833 | 13.465 | 14.0197
| Median |  | 14.215 | 13.9 | 13.15 | 14.47 | 10 | 15.35 | 10 | 10 | 13.75 | 14 | 15.05 | 13.9991 | 12.92 | 13.7 | 13.93 | 13.7 | 13 | 13.8344 | 14 | 13.8 | 13.46 | 14.041
| Std Dev. |  | 0.10597 | 0.23989 | 0.22391 | 0.1234 | 0 | 0.32249 | 0 | 0 | 0.24677 | 0.5164 | 0.33267 | 0.11192 | 0.26281 | 0.39296 | 0.32807 | 0.24495 | 0.64936 | 0.089872 | 0.13292 | 0.1169 | 0.28836 | 0.12982
| Rel.Std.Dev. |  | 0.745% | 1.722% | 1.712% | 0.8543% | 0% | 2.094% | 0% | 0% | 1.788% | 3.779% | 2.213% | 0.8021% | 2.035% | 2.877% | 2.359% | 1.775% | 5.08% | 0.6505% | 0.9483% | 0.8482% | 2.142% | 0.926%
| PDM3 |  | 2.797% | 0.6889% | -5.466% | 4.374% | -27.74% | 11.29% | -27.74% | -27.74% | -0.2626% | -1.238% | 8.638% | 0.8312% | -6.67% | -1.298% | 0.4962% | -0.2746% | -7.622% | -0.1615% | 1.291% | -0.395% | -2.695% | 1.313%
| Z-Score (Absolute) |  | 1.37 | 0.39 | 2.46 | 2.11 | NA | 5.31 | NA | NA | 0.05 | 0.5 | 4.08 | 0.46 | 3.02 | 0.53 | 0.31 | 0.05 | 3.46 | 0 | 0.67 | 0.11 | 1.18 | 0.68
Results from laboratories 1.05, 1.07 and 1.08 were removed due to their 10 ppm reading resolution.
| Analytical results for Ce in OREAS 612 (Certified Value 77 ppm) | [23 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | [7 empty] | 0
1
| 1 | 1 | 79.45 | 48.1 | 77.6 | 88.3 | 80.09 | 76.7 | 78.4576 | 64.3 | 80.4 | 73 | 63 | 71.9259 | 72.02 | 71 | 75.8 | [7 empty] | 1
| 1 | 2 | 78.53 | 54.22 | 76.55 | 84.2 | 77.22 | 77.7 | 76.455 | 66.2 | 82.9 | 73 | 50 | 71.7092 | 68.87 | 75 | 78.3 | [7 empty] | 25
| 1 | 3 | 79.26 | 62.46 | 75.37 | 87.4 | 80.06 | 77.6 | 76.6958 | 65 | 75.1 | 73 | 56 | 71.7319 | 69.95 | 74 | 77.2 | [7 empty] | 16
| 1 | 4 | 77.79 | 58.34 | 76.05 | 85.4 | 78.76 | 79.5 | 76.9877 | 65.5 | 77.3 | 75 | 66 | 71.7357 | 68.9 | 76 | 76.3 | [7 empty] | 76.536
| 1 | 5 | 75.89 | 71.8 | 76.2 | 83.6 | 78.26 | 78.6 | 75.576 | 65.8 | 80.5 | 71 | 55 | 71.714 | 71.38 | 75 | 74.1 | [7 empty] | 19
| 1 | 6 | 76.96 | 69.15 | 76.35 | 84.9 | 77.82 | 77.7 | 76.6625 | 65.8 | 77.3 | 77 | 56 | 71.7793 | 70.43 | 74 | 75.6
| Mean |  | 77.98 | 60.6783 | 76.3533 | 85.6333 | 78.7017 | 77.9667 | 76.8057 | 65.4333 | 78.9167 | 73.6667 | 57.6667 | 71.766 | 70.2583 | 74.1667 | 76.2167
| Median |  | 78.16 | 60.4 | 76.275 | 85.15 | 78.51 | 77.7 | 76.6791 | 65.65 | 78.85 | 73 | 56 | 71.7338 | 70.19 | 74.5 | 76.05
| Std Dev. |  | 1.3818 | 8.9892 | 0.73148 | 1.8446 | 1.1784 | 0.96264 | 0.94126 | 0.68313 | 2.8414 | 2.0656 | 5.8195 | 0.082149 | 1.2852 | 1.7224 | 1.4386
| Rel.Std.Dev. |  | 1.772% | 14.81% | 0.958% | 2.154% | 1.497% | 1.235% | 1.226% | 1.044% | 3.601% | 2.804% | 10.09% | 0.1145% | 1.829% | 2.322% | 1.888%
| PDM3 |  | 1.887% | -20.72% | -0.2386% | 11.89% | 2.83% | 1.869% | 0.3525% | -14.51% | 3.111% | -3.749% | -24.65% | -6.232% | -8.202% | -3.096% | -0.4172%
| Z-Score (Absolute) |  | 0.47 | 4.11 | 0.04 | 2.49 | 0.66 | 0.46 | 0.16 | 2.85 | 0.71 | 0.67 | 4.91 | 1.18 | 1.58 | 0.54 | 0
| Analytical results for Co in OREAS 612 (Certified Value 4.11 ppm) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 3.8 | 4.2 | 2.79 | 3.9 | < 10 | 4.3 | < 10 | 10 | 4.2 | 3 | 5 | 4.0443 | 4.1 | 4.2 | 3.9 | 4 | 4.2 | 4.2 | 4.2 | 4.06 | 5.462 |  | 1
| 1 | 2 | 3.9 | 4.3 | 3.12 | 4.2 | < 10 | 4.1 | < 10 | 10 | 4 | 3 | 4 | 4.1034 | 4.1 | 4.3 | 4.1 | 4 | 3.7 | 3.9 | 4.3 | 4.13 | 5.193 |  | 26
| 1 | 3 | 3.9 | 4.5 | 2.85 | 4 | < 10 | 4.3 | < 10 | 10 | 4.3 | 3 | 4 | 4.0688 | 4.3 | 4.3 | 3.9 | 4 | 3.4 | 4 | 4.1 | 4.07 | 5.12 |  | 16
| 1 | 4 | 4 | 4.5 | 3.1 | 4.1 | < 10 | 4.3 | < 10 | < 10 | 4.2 | 3 | 4 | 4.04 | 4.1 | 4.3 | 3.9 | 4 | 3.7 | 4 | 4.3 | 3.97 | 5.2663 |  | 4.1101
| 1 | 5 | 3.9 | 4.6 | 3.45 | 4 | < 10 | 4.2 | < 10 | < 10 | 4.1 | 3 | 4 | 4.0039 | 4.3 | 4.7 | 4 | 4 | 3.5 | 4.2 | 4.3 | 3.88 | 5.6695 |  | 20
| 1 | 6 | 3.7 | 4.5 | 3.24 | 4.1 | < 10 | 4.2 | < 10 | < 10 | 4.3 | 3 | 4 | 4.0269 | 4.3 | 4.4 | 4.1 | 4 | 3.8 | 4.2 | 4.2 | 4.11 | 5.333
| Mean |  | 3.8667 | 4.4333 | 3.0917 | 4.05 |  | 4.2333 |  | 10 | 4.1833 | 3 | 4.1667 | 4.0479 | 4.2 | 4.3667 | 3.9833 | 4 | 3.7167 | 4.0833 | 4.2333 | 4.0367 | 5.3406
| Median |  | 3.9 | 4.5 | 3.11 | 4.05 |  | 4.25 |  | 10 | 4.2 | 3 | 4 | 4.0422 | 4.2 | 4.3 | 3.95 | 4 | 3.7 | 4.1 | 4.25 | 4.065 | 5.2997
| Std Dev. |  | 0.10328 | 0.15055 | 0.24523 | 0.10488 |  | 0.08165 |  | 0 | 0.1169 | 0 | 0.40825 | 0.034556 | 0.10954 | 0.17512 | 0.098319 | 0 | 0.27869 | 0.13292 | 0.08165 | 0.094587 | 0.19947
| Rel.Std.Dev. |  | 2.671% | 3.396% | 7.932% | 2.59% |  | 1.929% |  | 0% | 2.795% | 0% | 9.798% | 0.8537% | 2.608% | 4.01% | 2.468% | 0% | 7.498% | 3.255% | 1.929% | 2.343% | 3.735%
| PDM3 |  | -5.923% | 7.865% | -24.78% | -1.462% |  | 2.998% |  | 143.3% | 1.782% | -27.01% | 1.376% | -1.514% | 2.187% | 6.243% | -3.084% | -2.679% | -9.572% | -0.6511% | 2.998% | -1.787% | 29.94%
| Z-Score (Absolute) |  | 1.16 | 0.92 | 4.02 | 0.49 | 3 | 0.18 | 3 | NA | 0 | NA | NA | 0.5 | 0.06 | 0.67 | 0.74 | NA | 1.72 | 0.37 | 0.18 | 0.54 | 4.26
Results from laboratories 1.10, 1.11 and 1.16 were removed due to their 1 ppm reading resolution. Results from laboratory 1.08 were removed due to their 10 ppm reading resolution.
| Analytical results for Cr in OREAS 612 (Certified Value 21.8 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
1
| 1 | 1 | 18 | 12 | 24.04 | 22 | 20 | 20 | 20 | 20 | 22 | 21 | 20 | 20.6229 | 22 | 20 | 24 | 25 | 21 | 22.1512 | 25 | 27 | 22 | 13.628 | 1
| 1 | 2 | 20 | 14 | 26.99 | 20 | 20 | 20 | 20 | 20 | 22 | 21 | 18 | 21.1777 | 22.7 | 31 | 22 | 24 | 19 | 23.6737 | 25 | 26 | 22 | 16.5005 | 27
| 1 | 3 | 18 | 14 | 28.38 | 23 | 20 | 21 | 20 | 20 | 23 | 23 | 17 | 21.0646 | 21.5 | 21 | 22 | 26 | 18 | 24.4425 | 23 | 26 | 20 | 16.107 | 16
| 1 | 4 | 21 | 13 | 26.33 | 23 | 20 | 19 | 20 | 20 | 24 | 23 | 20 | 21.073 | 23.2 | 21 | 22 | 24 | 18 | 22.9976 | 25 | 26 | 22 | 13.802 | 21.7867
| 1 | 5 | 20 | 11 | 31.07 | 22 | 20 | 22 | 20 | 20 | 21 | 21 | 19 | 21.0045 | 21.5 | 20 | 25 | 23 | 20 | 24.2203 | 24 | 24 | 19 | 15.671 | 21
| 1 | 6 | 19 | 13 | 27.14 | 19 | 20 | 22 | 20 | 20 | 23 | 20 | 20 | 20.7875 | 22.5 | 21 | 23 | 23 | 19 | 23.1857 | 26 | 26 | 23 | 18.393
| Mean |  | 19.3333 | 12.8333 | 27.325 | 21.5 | 20 | 20.6667 | 20 | 20 | 22.5 | 21.5 | 19 | 20.955 | 22.2333 | 22.3333 | 23 | 24.1667 | 19.1667 | 23.4452 | 24.6667 | 25.8333 | 21.3333 | 15.6836
| Median |  | 19.5 | 13 | 27.065 | 22 | 20 | 20.5 | 20 | 20 | 22.5 | 21 | 19.5 | 21.0346 | 22.25 | 21 | 22.5 | 24 | 19 | 23.4297 | 25 | 26 | 22 | 15.889
| Std Dev. |  | 1.2111 | 1.169 | 2.3266 | 1.6432 | 0 | 1.2111 | 0 | 0 | 1.0488 | 1.2247 | 1.2649 | 0.208 | 0.68605 | 4.274 | 1.2649 | 1.169 | 1.169 | 0.84734 | 1.0328 | 0.98319 | 1.5055 | 1.7863
| Rel.Std.Dev. |  | 6.264% | 9.109% | 8.515% | 7.643% | 0% | 5.86% | 0% | 0% | 4.661% | 5.696% | 6.657% | 0.9926% | 3.086% | 19.14% | 5.5% | 4.837% | 6.099% | 3.614% | 4.187% | 3.806% | 7.057% | 11.39%
| PDM3 |  | -11.26% | -41.1% | 25.42% | -1.316% | -8.201% | -5.141% | -8.201% | -8.201% | 3.274% | -1.316% | -12.79% | -3.817% | 2.05% | 2.509% | 5.569% | 10.92% | -12.03% | 7.613% | 13.22% | 18.57% | -2.081% | -28.01%
| Z-Score (Absolute) |  | 0.75 | 3 | 2.02 | 0 | NA | 0.29 | NA | NA | 0.35 | 0 | 0.87 | 0.19 | 0.25 | 0.29 | 0.52 | 0.92 | 0.81 | 0.67 | 1.1 | 1.5 | 0.06 | 2.02
Results from laboratories 1.05, 1.07 and 1.08 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 612 (Certified Value 6.63 ppm) | [23 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | [6 empty] | 2
3
| 1 | 1 | 6.81 | 6 | 6.78 | 7.23 | 6.87 | 6.8 | 6.5432 | 6.2 | 6.88 | 6.34 | 6.3 | 6.4 | 6.4883 | 6.54 | 6.88 | 6.68 | [6 empty] | 1
| 1 | 2 | 6.9 | 6 | 6.79 | 6.81 | 6.58 | 6.63 | 6.522 | 6.4 | 6.94 | 6.54 | 6.4 | 6.5 | 6.4623 | 6.31 | 6.89 | 6.66 | [6 empty] | 28
| 1 | 3 | 6.84 | 7 | 6.72 | 7.07 | 6.75 | 6.89 | 6.6088 | 6.5 | 6.36 | 6.39 | 6.4 | 5.9 | 6.482 | 6.15 | 6.97 | 6.66 | [6 empty] | 16
| 1 | 4 | 6.93 | 6 | 6.72 | 6.85 | 6.73 | 6.75 | 6.522 | 6.4 | 6.49 | 6.55 | 6.5 | 7.1 | 6.5213 | 6.36 | 6.92 | 6.7 | [6 empty] | 6.6282
| 1 | 5 | 6.55 | 7 | 6.84 | 6.79 | 6.69 | 6.83 | 6.5273 | 6.5 | 6.67 | 6.18 | 6.5 | 6.1 | 6.4637 | 6.66 | 6.96 | 6.5 | [6 empty] | 22
| 1 | 6 | 6.88 | 7 | 6.84 | 6.82 | 6.79 | 6.95 | 6.5487 | 6.7 | 6.38 | 6.7 | 6.3 | 6.6 | 6.4985 | 6.19 | 6.91 | 6.58
| Mean |  | 6.8183 | 6.5 | 6.7817 | 6.9283 | 6.735 | 6.8083 | 6.5453 | 6.45 | 6.62 | 6.45 | 6.4 | 6.4333 | 6.486 | 6.3683 | 6.9217 | 6.63
| Median |  | 6.86 | 6.5 | 6.785 | 6.835 | 6.74 | 6.815 | 6.5353 | 6.45 | 6.58 | 6.465 | 6.4 | 6.45 | 6.4851 | 6.335 | 6.915 | 6.66
| Std Dev. |  | 0.13819 | 0.54772 | 0.053821 | 0.18005 | 0.097519 | 0.11179 | 0.033043 | 0.16432 | 0.25084 | 0.18396 | 0.089443 | 0.41793 | 0.02229 | 0.19874 | 0.03656 | 0.075631
| Rel.Std.Dev. |  | 2.027% | 8.427% | 0.7936% | 2.599% | 1.448% | 1.642% | 0.5048% | 2.548% | 3.789% | 2.852% | 1.398% | 6.496% | 0.3437% | 3.121% | 0.5282% | 1.141%
| PDM3 |  | 2.869% | -1.934% | 2.316% | 4.528% | 1.612% | 2.718% | -1.25% | -2.688% | -0.1234% | -2.688% | -3.443% | -2.94% | -2.145% | -3.92% | 4.428% | 0.02748%
| Z-Score (Absolute) |  | 0.79 | NA | 0.64 | 1.22 | 0.46 | 0.75 | 0.3 | 0.67 | 0 | 0.67 | 0.87 | 0.74 | 0.53 | 1 | 1.2 | 0.04
Results from laboratory 1.02 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 612 (Certified Value 94 ppm) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 95.8 | 90.2 | 110 | 90.9 | 90 | 98.3 | 100 | 90 | 88 | 92 | 101 | 93.0824 | 92.2 | 91.1 | 98 | 84 | 88.6 | 94.6067 | 95.1 | 93 | 86.8 | 86.8043 | 1
| 1 | 2 | 95.4 | 91.8 | 90 | 92.5 | 90 | 94.5 | 100 | 100 | 87 | 90.6 | 97 | 92.2971 | 94.7 | 92.5 | 100 | 84 | 88.8 | 92.8079 | 97.8 | 92 | 87.8 | 85.9373 | 29
| 1 | 3 | 99.4 | 95.5 | 90 | 95.8 | 90 | 97.4 | 100 | 100 | 88 | 92.5 | 95 | 95.6899 | 96.5 | 98.1 | 99 | 88 | 92.5 | 94.3342 | 97.6 | 93 | 88.1 | 86.9297 | 16
| 1 | 4 | 98.9 | 98.3 | 90 | 92.6 | 90 | 97.8 | 100 | 100 | 86 | 93 | 92 | 93.1144 | 96.3 | 101.5 | 101 | 88 | 92.5 | 93.2289 | 95.7 | 99 | 88.8 | 88.6677 | 93.7422
| 1 | 5 | 93.5 | 97.9 | 90 | 91.9 | 90 | 95.5 | 100 | 100 | 87 | 91.4 | 101 | 93.6464 | 96.3 | 103 | 99 | 88 | 92.7 | 94.008 | 97.2 | 94 | 88.8 | 89.6485 | 23
| 1 | 6 | 97.6 | 95.5 | 90 | 91.1 | 100 | 97.1 | 100 | 100 | 88 | 93.6 | 101 | 93.5647 | 96.3 | 95.5 | 102 | 85 | 88.9 | 94.6256 | 98.7 | 93 | 89.8 | 91.472
| Mean |  | 96.7667 | 94.8667 | 93.3333 | 92.4667 | 91.6667 | 96.7667 | 100 | 98.3333 | 87.3333 | 92.1833 | 97.8333 | 93.5658 | 95.3833 | 96.95 | 99.8333 | 86.1667 | 90.6667 | 93.9352 | 97.0167 | 94 | 88.35 | 88.2433
| Median |  | 96.7 | 95.5 | 90 | 92.2 | 90 | 97.25 | 100 | 100 | 87.5 | 92.25 | 99 | 93.3395 | 96.3 | 96.8 | 99.5 | 86.5 | 90.7 | 94.1711 | 97.4 | 93 | 88.45 | 87.7987
| Std Dev. |  | 2.2651 | 3.2549 | 8.165 | 1.7761 | 4.0825 | 1.461 | 0 | 4.0825 | 0.8165 | 1.0889 | 3.8166 | 1.1459 | 1.6952 | 4.7907 | 1.472 | 2.0412 | 2.0849 | 0.75651 | 1.3586 | 2.5298 | 1.0271 | 2.0829
| Rel.Std.Dev. |  | 2.341% | 3.431% | 8.748% | 1.921% | 4.454% | 1.51% | 0% | 4.152% | 0.9349% | 1.181% | 3.901% | 1.225% | 1.777% | 4.941% | 1.474% | 2.369% | 2.299% | 0.8054% | 1.4% | 2.691% | 1.163% | 2.36%
| PDM3 |  | 3.226% | 1.2% | -0.4361% | -1.361% | -2.214% | 3.226% | 6.676% | 4.898% | -6.837% | -1.663% | 4.364% | -0.1881% | 1.751% | 3.422% | 6.498% | -8.081% | -3.281% | 0.2059% | 3.493% | 0.275% | -5.752% | -5.866%
| Z-Score (Absolute) |  | 0.65 | 0.21 | 0.15 | 0.35 | 0.54 | 0.65 | 1.41 | 1.02 | 1.55 | 0.42 | 0.9 | 0.09 | 0.33 | 0.7 | 1.37 | 1.82 | 0.77 | 0.01 | 0.71 | 0.01 | 1.31 | 1.34
| Analytical results for Dy in OREAS 612 (Certified Value 3.41 ppm) | [23 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 3.41 | 3.4 | 3.2 | 3.4337 | 3.26 | 3.57 | 3.5121 | 3.48 | [14 empty] | 1
| 1 | 2 | 3.51 | 3.29 | 3.4 | 3.365 | 3.38 | 3.36 | 3.4949 | 4.56 | [14 empty] | 30
| 1 | 3 | 3.45 | 3.32 | 3.4 | 3.2769 | 3.32 | 3.38 | 3.5144 | 3.5 | [14 empty] | 16
| 1 | 4 | 3.6 | 3.3 | 3.5 | 3.3528 | 3.32 | 3.42 | 3.4472 | 3.42 | [14 empty] | 3.4074
| 1 | 5 | 3.34 | 3.26 | 3.6 | 3.2708 | 3.33 | 3.4 | 3.4659 | 3.42 | [14 empty] | 24
| 1 | 6 | 3.49 | 3.31 | 3.5 | 3.2706 | 3.35 | 3.54 | 3.4789 | 3.48
| Mean |  | 3.4667 | 3.3133 | 3.4333 | 3.3283 | 3.3267 | 3.445 | 3.4856 | 3.6433
| Median |  | 3.47 | 3.305 | 3.45 | 3.3149 | 3.325 | 3.41 | 3.4869 | 3.48
| Std Dev. |  | 0.089144 | 0.047188 | 0.13663 | 0.066823 | 0.039833 | 0.088034 | 0.026578 | 0.45032
| Rel.Std.Dev. |  | 2.571% | 1.424% | 3.979% | 2.008% | 1.197% | 2.555% | 0.7625% | 12.36%
| PDM3 |  | 1.741% | -2.759% | 0.7623% | -2.32% | -2.368% | 1.105% | 2.295% | 6.925%
| Z-Score (Absolute) |  | 0.24 | 1.08 | 0.05 | 0.95 | 0.96 | 0.05 | 0.4 | 1.75
| Analytical results for Er in OREAS 612 (Certified Value 1.27 ppm) | [23 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 1.31 | 1.29 | 1.2 | 1.2693 | 1.23 | 1.3 | 1.225 | 1.33 | [14 empty] | 1
| 1 | 2 | 1.39 | 1.23 | 1.1 | 1.211 | 1.28 | 1.3 | 1.2322 | 1.26 | [14 empty] | 31
| 1 | 3 | 1.36 | 1.25 | 1.2 | 1.1946 | 1.25 | 1.26 | 1.3109 | 1.38 | [14 empty] | 16
| 1 | 4 | 1.39 | 1.2 | 1.3 | 1.2106 | 1.26 | 1.32 | 1.2836 | 1.32 | [14 empty] | 1.2689
| 1 | 5 | 1.29 | 1.24 | 1.2 | 1.1822 | 1.26 | 1.36 | 1.284 | 1.28 | [14 empty] | 25
| 1 | 6 | 1.28 | 1.25 | 1.2 | 1.2367 | 1.28 | 1.33 | 1.2667 | 1.32
| Mean |  | 1.3367 | 1.2433 | 1.2 | 1.2174 | 1.26 | 1.3117 | 1.2671 | 1.315
| Median |  | 1.335 | 1.245 | 1.2 | 1.2108 | 1.26 | 1.31 | 1.2751 | 1.32
| Std Dev. |  | 0.049666 | 0.029439 | 0.063246 | 0.031325 | 0.018974 | 0.033714 | 0.033067 | 0.041833
| Rel.Std.Dev. |  | 3.716% | 2.368% | 5.27% | 2.573% | 1.506% | 2.57% | 2.61% | 3.181%
| PDM3 |  | 5.341% | -2.014% | -5.429% | -4.058% | -0.7006% | 3.371% | -0.1446% | 3.634%
| Z-Score (Absolute) |  | 1.05 | 0.29 | 0.91 | 0.66 | 0.05 | 0.69 | 0.05 | 0.74
| Analytical results for Eu in OREAS 612 (Certified Value 1.39 ppm) | [23 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 1.43 | 1.36 | 1.31 | 1.4235 | 1.51 | 1.36 | 2.5476 | 1.36 | [14 empty] | 1
| 1 | 2 | 1.44 | 1.29 | 1.36 | 1.3858 | 1.55 | 1.43 | 2.553 | 1.34 | [14 empty] | 32
| 1 | 3 | 1.44 | 1.35 | 1.27 | 1.3504 | 1.57 | 1.44 | 2.6665 | 1.36 | [14 empty] | 16
| 1 | 4 | 1.44 | 1.31 | 1.3 | 1.3849 | 1.55 | 1.4 | 2.5607 | 1.35 | [14 empty] | 1.3946
| 1 | 5 | 1.37 | 1.29 | 1.29 | 1.3613 | 1.57 | 1.46 | 2.6229 | 1.35 | [14 empty] | 26
| 1 | 6 | 1.41 | 1.32 | 1.33 | 1.3694 | 1.56 | 1.48 | 2.5066 | 1.35
| Mean |  | 1.4217 | 1.32 | 1.31 | 1.3792 | 1.5517 | 1.4283 | 2.5762 | 1.3517
| Median |  | 1.435 | 1.315 | 1.305 | 1.3771 | 1.555 | 1.435 | 2.5568 | 1.35
| Std Dev. |  | 0.027869 | 0.029665 | 0.031623 | 0.02564 | 0.022286 | 0.043089 | 0.057922 | 0.0075277
| Rel.Std.Dev. |  | 1.96% | 2.247% | 2.414% | 1.859% | 1.436% | 3.017% | 2.248% | 0.5569%
| PDM3 |  | 1.937% | -5.353% | -6.07% | -1.107% | 11.26% | 2.415% | 84.72% | -3.082%
| Z-Score (Absolute) |  | 0.22 | 0.84 | 0.94 | 0.22 | 1.58 | 0.29 | 12.27 | 0.51
| Analytical results for Fe in OREAS 612 (Certified Value 2.23 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.26 | 1.8 | 2.25 | 2.29 | 1.86 | 2.19 | 2.17 | 1.95 | 2.27 | 2.27 | 2.22 | 2.21 | 2.2069 | 2.24 | 2.27 | 2.25 | 2.19 | 2.1473 | 2.13 | 2.16 | 2.18 | 2.5906 | 1
| 1 | 2 | 2.26 | 2.01 | 2.23 | 2.28 | 1.87 | 2.16 | 2.19 | 1.96 | 2.2 | 2.19 | 2.32 | 2.2429 | 2.2607 | 2.31 | 2.25 | 2.21 | 2.23 | 2.1379 | 2.19 | 2.16 | 2.18 | 2.6064 | #N/A
| 1 | 3 | 2.32 | 1.93 | 2.22 | 2.29 | 1.93 | 2.22 | 2.2 | 2 | 2.26 | 2.22 | 2.29 | 2.2528 | 2.2432 | 2.27 | 2.26 | 2.25 | 2.19 | 2.1477 | 2.16 | 2.18 | 2.2 | 2.445 | 16
| 1 | 4 | 2.29 | 1.99 | 2.23 | 2.27 | 1.86 | 2.21 | 2.2 | 1.94 | 2.21 | 2.17 | 2.3 | 2.1967 | 2.2514 | 2.22 | 2.28 | 2.24 | 2.25 | 2.1225 | 2.14 | 2.19 | 2.19 | 2.6985 | 2.2278
| 1 | 5 | 2.26 | 2.09 | 2.23 | 2.3 | 1.9 | 2.2 | 2.24 | 1.99 | 2.23 | 2.18 | 2.38 | 2.1882 | 2.2454 | 2.31 | 2.28 | 2.25 | 2.2 | 2.1559 | 2.18 | 2.19 | 2.21 | 2.8041 | 27
| 1 | 6 | 2.23 | 2.05 | 2.27 | 2.34 | 1.88 | 2.22 | 2.24 | 1.96 | 2.26 | 2.16 | 2.38 | 2.2296 | 2.2802 | 2.34 | 2.27 | 2.2 | 2.18 | 2.1439 | 2.17 | 2.21 | 2.21 | 2.5205
| Mean |  | 2.27 | 1.9783 | 2.2383 | 2.295 | 1.8833 | 2.2 | 2.2067 | 1.9667 | 2.2383 | 2.1983 | 2.315 | 2.22 | 2.248 | 2.2817 | 2.2683 | 2.2333 | 2.2067 | 2.1425 | 2.1617 | 2.1817 | 2.195 | 2.6108
| Median |  | 2.26 | 2 | 2.23 | 2.29 | 1.875 | 2.205 | 2.2 | 1.96 | 2.245 | 2.185 | 2.31 | 2.2198 | 2.2484 | 2.29 | 2.27 | 2.245 | 2.195 | 2.1456 | 2.165 | 2.185 | 2.195 | 2.5985
| Std Dev. |  | 0.030984 | 0.10284 | 0.018348 | 0.02429 | 0.027325 | 0.022804 | 0.028048 | 0.023381 | 0.029269 | 0.040702 | 0.060581 | 0.02588 | 0.024202 | 0.046224 | 0.01169 | 0.022509 | 0.027325 | 0.011421 | 0.023166 | 0.019408 | 0.013784 | 0.1274
| Rel.Std.Dev. |  | 1.365% | 5.198% | 0.8197% | 1.058% | 1.451% | 1.037% | 1.271% | 1.189% | 1.308% | 1.852% | 2.617% | 1.166% | 1.077% | 2.026% | 0.5154% | 1.008% | 1.238% | 0.5331% | 1.072% | 0.8896% | 0.628% | 4.88%
| PDM3 |  | 1.894% | -11.2% | 0.4725% | 3.016% | -15.46% | -1.248% | -0.9489% | -11.72% | 0.4725% | -1.323% | 3.914% | -0.3485% | 0.9049% | 2.418% | 1.819% | 0.2481% | -0.9489% | -3.829% | -2.969% | -2.071% | -1.473% | 17.19%
| Z-Score (Absolute) |  | 0.89 | 3.67 | 0.39 | 1.28 | 5.16 | 0.21 | 0.1 | 3.86 | 0.39 | 0.23 | 1.59 | 0.1 | 0.54 | 1.07 | 0.86 | 0.31 | 0.1 | 1.11 | 0.81 | 0.5 | 0.29 | 6.21
| Analytical results for Ga in OREAS 612 (Certified Value 20.2 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 19.83 | 20.2 | 19.54 | 19.47 | < 50 | 21.7 | < 50 | < 50 | 20.23 | 19.8 | 19.7775 | 20.47 | 21.3 | 18.4 | 19.9 | 20 | 19.1749 | 19.6 | 20 | 20.32 | 20.9543 |  | 1
| 1 | 2 | 20.21 | 21.3 | 20.82 | 19.78 | < 50 | 20.5 | < 50 | < 50 | 20.17 | 20.4 | 19.068 | 20.93 | 21.7 | 19.3 | 20 | 19 | 19.0689 | 19 | 21 | 20.36 | 21.1163 |  | 33
| 1 | 3 | 20.98 | 21.6 | 20.24 | 19.73 | < 50 | 21.2 | < 50 | < 50 | 19.84 | 19.6 | 19.6791 | 21.07 | 19.1 | 18.8 | 20.3 | 18 | 19.2035 | 19.4 | 20 | 19.9 | 20.9033 |  | 16
| 1 | 4 | 21.1 | 22 | 20.16 | 19.82 | < 50 | 21.7 | < 50 | < 50 | 19.78 | 20.7 | 20.2818 | 20.84 | 19.35 | 19.1 | 20.7 | 21 | 19.1966 | 19.4 | 21 | 19.44 | 21.3113 |  | 20.1664
| 1 | 5 | 20.42 | 22.7 | 20.15 | 19.82 | < 50 | 20.6 | < 50 | < 50 | 20.05 | 20.1 | 19.2201 | 20.96 | 20.1 | 18.7 | 20.4 | 19 | 19.1422 | 20.1 | 20 | 19.89 | 21.0355 |  | 28
| 1 | 6 | 21.23 | 22 | 19.94 | 19.9 | < 50 | 21.2 | < 50 | < 50 | 19.46 | 20.2 | 19.4619 | 21.13 | 19.3 | 19.4 | 20.2 | 20 | 19.1505 | 20.2 | 21 | 19.52 | 21.4737
| Mean |  | 20.6283 | 21.6333 | 20.1417 | 19.7533 |  | 21.15 |  |  | 19.9217 | 20.1333 | 19.5814 | 20.9 | 20.1417 | 18.95 | 20.25 | 19.5 | 19.1561 | 19.6167 | 20.5 | 19.905 | 21.1324
| Median |  | 20.7 | 21.8 | 20.155 | 19.8 |  | 21.2 |  |  | 19.945 | 20.15 | 19.5705 | 20.945 | 19.725 | 18.95 | 20.25 | 19.5 | 19.1627 | 19.5 | 20.5 | 19.895 | 21.0759
| Std Dev. |  | 0.55926 | 0.84538 | 0.41782 | 0.14962 |  | 0.51672 |  |  | 0.28743 | 0.39833 | 0.43521 | 0.23444 | 1.1128 | 0.38341 | 0.2881 | 1.0488 | 0.049119 | 0.45789 | 0.54772 | 0.38573 | 0.22014
| Rel.Std.Dev. |  | 2.711% | 3.908% | 2.074% | 0.7575% |  | 2.443% |  |  | 1.443% | 1.978% | 2.223% | 1.122% | 5.525% | 2.023% | 1.423% | 5.379% | 0.2564% | 2.334% | 2.672% | 1.938% | 1.042%
| PDM3 |  | 2.291% | 7.274% | -0.1226% | -2.048% |  | 4.877% |  |  | -1.213% | -0.1639% | -2.901% | 3.638% | -0.1226% | -6.032% | 0.4146% | -3.304% | -5.01% | -2.726% | 1.654% | -1.296% | 4.79%
| Z-Score (Absolute) |  | 0.59 | 1.8 | 0 | 0.47 | 5.85 | 1.21 | 5.85 | 5.85 | 0.26 | 0.01 | 0.67 | 0.91 | 0 | 1.43 | 0.13 | 0.77 | 1.19 | 0.63 | 0.43 | 0.28 | 1.19
| Analytical results for Gd in OREAS 612 (Certified Value 5.45 ppm) | [23 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 5.56 | 5.51 | 5.8 | 5.4954 | 5.65 | 5.33 | 5.2071 | 5.45 | [14 empty] | 1
| 1 | 2 | 5.48 | 5.39 | 5.7 | 5.2815 | 5.95 | 5.19 | 5.1751 | 5.36 | [14 empty] | 34
| 1 | 3 | 5.58 | 5.47 | 5.6 | 5.2552 | 5.95 | 5.17 | 5.1345 | 5.45 | [14 empty] | 16
| 1 | 4 | 5.62 | 5.46 | 5.3 | 5.287 | 5.84 | 4.98 | 5.1257 | 5.59 | [14 empty] | 5.4494
| 1 | 5 | 5.31 | 5.39 | 5.8 | 5.2197 | 5.96 | 4.91 | 5.1531 | 5.38 | [14 empty] | 29
| 1 | 6 | 5.48 | 5.72 | 5.6 | 5.2227 | 5.98 | 5.43 | 5.1419 | 5.53
| Mean |  | 5.505 | 5.49 | 5.6333 | 5.2936 | 5.8883 | 5.1683 | 5.1562 | 5.46
| Median |  | 5.52 | 5.465 | 5.65 | 5.2683 | 5.95 | 5.18 | 5.1475 | 5.45
| Std Dev. |  | 0.11059 | 0.12215 | 0.18619 | 0.10286 | 0.12671 | 0.19864 | 0.030213 | 0.087636
| Rel.Std.Dev. |  | 2.009% | 2.225% | 3.305% | 1.943% | 2.152% | 3.843% | 0.5859% | 1.605%
| PDM3 |  | 1.021% | 0.7459% | 3.376% | -2.859% | 8.056% | -5.157% | -5.379% | 0.1954%
| Z-Score (Absolute) |  | 0.12 | 0.06 | 0.63 | 0.72 | 1.64 | 1.22 | 1.27 | 0.06
| Analytical results for Ge in OREAS 612 (Indicative Value 0.2 ppm) | [23 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.06 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.07 | 0.2 | 0.14 | 1.5 | < 0.1 | 0.2 | 0.24 | 0.36 | 1.2 | 3 | 0.3 | 1.6 | [10 empty] | 1
| 1 | 2 | 0.05 | 0.2 | 0.15 | 1.5 | 0.1 | 0.19 | 0.25 | 0.36 | 1.3 | 3.1 | 0.3 | 1.5 | [10 empty] | 4
| 1 | 3 | 0.06 | 0.2 | 0.17 | 1.5 | 0.1 | 0.2 | 0.23 | 0.37 | 1.2 | 2.9 | 0.3 | 1.5 | [10 empty] | 16
| 1 | 4 | 0.08 | 0.2 | 0.18 | 1.5 | 0.1 | 0.19 | 0.22 | 0.36 | 1.3 | 3.4 | 0.2 | 1.5 | [10 empty] | 0.19962
| 1 | 5 | 0.07 | 0.2 | 0.17 | 1.5 | 0.1 | 0.21 | 0.21 | 0.39 | 1.4 | 3.1 | 0.3 | 1.6 | [10 empty] | 10
| 1 | 6 | 0.08 | 0.2 | 0.18 | 1.5 | 0.1 | 0.2 | 0.17 | 0.36 | 1.4 | 3.3 | 0.3 | 1.5
| Mean |  | 0.068333 | 0.2 | 0.165 | 1.5 | 0.1 | 0.19833 | 0.22 | 0.36667 | 1.3 | 3.1333 | 0.28333 | 1.5333
| Median |  | 0.07 | 0.2 | 0.17 | 1.5 | 0.1 | 0.2 | 0.225 | 0.36 | 1.3 | 3.1 | 0.3 | 1.5
| Std Dev. |  | 0.01169 | 3.0405e-17 | 0.016432 | 0 | 0 | 0.0075277 | 0.028284 | 0.012111 | 0.089443 | 0.18619 | 0.040825 | 0.05164
| Rel.Std.Dev. |  | 17.11% | 1.52e-14% | 9.959% | 0% | 0% | 3.795% | 12.86% | 3.303% | 6.88% | 5.942% | 14.41% | 3.368%
| PDM3 |  | -65.77% | 0.1879% | -17.35% | 651.4% | -49.91% | -0.647% | 10.21% | 83.68% | 551.2% | 1470% | 41.93% | 668.1%
| Z-Score (Absolute) |  | 0.9 | 0.25 | 0.43 | 6.12 | 0.78 | 0.26 | 0.16 | 0.56 | 5.14 | 14.13 | 0.16 | 6.29
| Analytical results for Hf in OREAS 612 (Certified Value 4.29 ppm) | [23 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
3
| 1 | 1 | 4.2 | 4.49 | 4.35 | 4.3 | 4.51 | 4.1 | 4.4139 | 4.06 | 4.6 | 5.3 | 4.2 | 4.2 | 4.43 | 4.53 | 4.36 | [7 empty] | 1
| 1 | 2 | 4.3 | 4.43 | 4.39 | 4.1 | 4.18 | 4.1 | 4.1958 | 4.17 | 4.7 | 5.3 | 4.2 | 4.1 | 4.21 | 4.46 | 4.26 | [7 empty] | 19
| 1 | 3 | 4.4 | 4.68 | 4.25 | 4.2 | 4.39 | 4.2 | 4.1125 | 4.17 | 4.3 | 5.5 | 4.4 | 3.7 | 4.26 | 4.39 | 4.22 | [7 empty] | 16
| 1 | 4 | 4.4 | 4.76 | 4.18 | 4.3 | 4.13 | 5.1 | 4.1985 | 4.06 | 4.3 | 5.5 | 4.3 | 4.2 | 4.27 | 4.35 | 4.25 | [7 empty] | 4.2916
| 1 | 5 | 4.3 | 4.82 | 4.3 | 4.1 | 4.19 | 4.2 | 4.0437 | 4.12 | 4.5 | 5.4 | 4.3 | 3.8 | 4.44 | 4.24 | 4.12 | [7 empty] | 30
| 1 | 6 | 4.4 | 4.64 | 4.4 | 4.1 | 4.43 | 4.4 | 4.1279 | 4.19 | 4.4 | 5.7 | 4.3 | 4.1 | 4.5 | 4.34 | 4.2
| Mean |  | 4.3333 | 4.6367 | 4.3117 | 4.1833 | 4.305 | 4.35 | 4.1821 | 4.1283 | 4.4667 | 5.45 | 4.2833 | 4.0167 | 4.3517 | 4.385 | 4.235
| Median |  | 4.35 | 4.66 | 4.325 | 4.15 | 4.29 | 4.2 | 4.1619 | 4.145 | 4.45 | 5.45 | 4.3 | 4.1 | 4.35 | 4.37 | 4.235
| Std Dev. |  | 0.08165 | 0.15161 | 0.085654 | 0.098319 | 0.1577 | 0.38341 | 0.12739 | 0.057764 | 0.1633 | 0.15166 | 0.075277 | 0.2137 | 0.11923 | 0.10095 | 0.07893
| Rel.Std.Dev. |  | 1.884% | 3.27% | 1.987% | 2.35% | 3.663% | 8.814% | 3.046% | 1.399% | 3.656% | 2.783% | 1.757% | 5.32% | 2.74% | 2.302% | 1.864%
| PDM3 |  | 0.973% | 8.041% | 0.4682% | -2.522% | 0.3128% | 1.361% | -2.552% | -3.804% | 4.08% | 26.99% | -0.1921% | -6.406% | 1.4% | 2.177% | -1.318%
| Z-Score (Absolute) |  | 0.19 | 2.86 | 0 | 1.13 | 0.06 | 0.34 | 1.14 | 1.61 | 1.36 | 10.01 | 0.25 | 2.59 | 0.35 | 0.64 | 0.67
| Analytical results for Hg in OREAS 612 (Indicative Value 0.097 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.12 | 1.13 | [19 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | [19 empty] | 3
3
| 1 | 1 | < 20 | < 1 | 0.12 | [19 empty] | 1
| 1 | 2 | < 20 | < 1 | 0.11 | [19 empty] | 3
| 1 | 3 | < 20 | < 1 | 0.06 | [19 empty] | 16
| 1 | 4 | < 20 | < 1 | 0.11 | [19 empty] | 0.096667
| 1 | 5 | < 20 | < 1 | 0.09 | [19 empty] | 9
| 1 | 6 | < 20 | < 1 | 0.09
| Mean |  |  |  | 0.096667
| Median |  |  |  | 0.1
| Std Dev. |  |  |  | 0.021602
| Rel.Std.Dev. |  |  |  | 22.35%
| PDM3 |  |  |  | -4.441e-14%
| Z-Score (Absolute) |  | 15.88 | 0 | 0.67
| Analytical results for Ho in OREAS 612 (Certified Value 0.52 ppm) | [23 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.22 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.53 | 0.54 | 0.5 | 0.51906 | 0.48 | 0.54 | 0.55 | [15 empty] | 1
| 1 | 2 | 0.54 | 0.51 | 0.5 | 0.49433 | 0.49 | 0.55 | 0.53 | [15 empty] | 20
| 1 | 3 | 0.53 | 0.51 | 0.6 | 0.4871 | 0.49 | 0.55 | 0.53 | [15 empty] | 16
| 1 | 4 | 0.54 | 0.51 | 0.5 | 0.49688 | 0.48 | 0.54 | 0.53 | [15 empty] | 0.51625
| 1 | 5 | 0.5 | 0.51 | 0.5 | 0.47831 | 0.48 | 0.57 | 0.52 | [15 empty] | 31
| 1 | 6 | 0.52 | 0.5 | 0.6 | 0.48939 | 0.49 | 0.54 | 0.52
| Mean |  | 0.52667 | 0.51333 | 0.53333 | 0.49418 | 0.485 | 0.54833 | 0.53
| Median |  | 0.53 | 0.51 | 0.5 | 0.49186 | 0.485 | 0.545 | 0.53
| Std Dev. |  | 0.015055 | 0.013663 | 0.05164 | 0.013793 | 0.0054772 | 0.01169 | 0.010954
| Rel.Std.Dev. |  | 2.859% | 2.662% | 9.682% | 2.791% | 1.129% | 2.132% | 2.067%
| PDM3 |  | 2.017% | -0.5654% | 3.309% | -4.276% | -6.054% | 6.214% | 2.663%
| Z-Score (Absolute) |  | 0.25 | 0.25 | NA | 0.97 | 1.32 | 1.07 | 0.38
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 612 (Certified Value 0.14 ppm) | [23 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
3
| 1 | 1 | 0.169 | 0.13 | 0.13 | 0.145 | 0.14 | 0.1 | 0.14811 | 0.14 | 0.147 | 0.23 | 0.14 | 0.12 | 0.13 | 0.17 | 0.13 | [7 empty] | 1
| 1 | 2 | 0.167 | 0.13 | 0.133 | 0.141 | 0.14 | 0.1 | 0.15309 | 0.146 | 0.139 | 0.22 | 0.14 | 0.12 | 0.13 | 0.14 | 0.18 | [7 empty] | 21
| 1 | 3 | 0.178 | 0.12 | 0.133 | 0.147 | 0.14 | 0.1 | 0.14776 | 0.147 | 0.126 | 0.23 | 0.14 | 0.1 | 0.13 | 0.15 | 0.12 | [7 empty] | 16
| 1 | 4 | 0.163 | 0.14 | 0.135 | 0.147 | 0.13 | 0.2 | 0.139 | 0.147 | 0.133 | 0.23 | 0.15 | 0.12 | 0.13 | 0.16 | 0.14 | [7 empty] | 0.14001
| 1 | 5 | 0.175 | 0.14 | 0.138 | 0.141 | 0.14 | 0.2 | 0.14776 | 0.143 | 0.155 | 0.22 | 0.15 | 0.11 | 0.13 | 0.15 | 0.15 | [7 empty] | 32
| 1 | 6 | 0.169 | 0.14 | 0.134 | 0.145 | 0.13 | 0.1 | 0.14984 | 0.144 | 0.125 | 0.24 | 0.14 | 0.12 | 0.14 | 0.15 | 0.13
| Mean |  | 0.17017 | 0.13333 | 0.13383 | 0.14433 | 0.13667 | 0.13333 | 0.1476 | 0.1445 | 0.1375 | 0.22833 | 0.14333 | 0.115 | 0.13167 | 0.15333 | 0.14167
| Median |  | 0.169 | 0.135 | 0.1335 | 0.145 | 0.14 | 0.1 | 0.14794 | 0.145 | 0.136 | 0.23 | 0.14 | 0.12 | 0.13 | 0.15 | 0.135
| Std Dev. |  | 0.0054559 | 0.008165 | 0.0026394 | 0.0027325 | 0.005164 | 0.05164 | 0.0046783 | 0.0027386 | 0.011895 | 0.0075277 | 0.005164 | 0.0083666 | 0.0040825 | 0.010328 | 0.02137
| Rel.Std.Dev. |  | 3.206% | 6.124% | 1.972% | 1.893% | 3.779% | 38.73% | 3.17% | 1.895% | 8.651% | 3.297% | 3.603% | 7.275% | 3.101% | 6.736% | 15.08%
| PDM3 |  | 21.54% | -4.768% | -4.411% | 3.089% | -2.387% | -4.768% | 5.419% | 3.208% | -1.792% | 63.09% | 2.375% | -17.86% | -5.958% | 9.517% | 1.184%
| Z-Score (Absolute) |  | 2.57 | 0.85 | 0.81 | 0.17 | 0.54 | NA | 0.47 | 0.19 | 0.47 | 7.98 | 0.08 | 2.56 | 1.01 | 1.01 | 0.08
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 612 (Certified Value 4.33 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 4.34 | 4.44 | 4.16 | 4.49 | 3.5 | 4.41 | 4.2 | 3.9 | 4.4622 | 4.01 | 4.59 | 4.2505 | 3.9288 | 4.45 | 4.19 | 4.72 | 4.39 | 4.1566 | 4.16 | 4.4399 | 4.57 | 3.8047 | 1
| 1 | 2 | 4.31 | 4.35 | 4.4 | 4.48 | 3.6 | 4.43 | 4.1 | 3.9 | 4.3392 | 4.04 | 4.68 | 4.2824 | 4.018 | 4.61 | 4.34 | 4.76 | 4.39 | 4.1438 | 4.32 | 4.4385 | 4.57 | 3.7083 | #N/A
| 1 | 3 | 4.32 | 4.51 | 4.27 | 4.48 | 3.2 | 4.48 | 4.3 | 3.9 | 4.4728 | 4.08 | 4.65 | 4.1242 | 4.0147 | 4.43 | 4.2 | 4.94 | 4.38 | 4.1554 | 4.22 | 4.4026 | 4.6 | 3.6061 | 16
| 1 | 4 | 4.37 | 4.64 | 4.33 | 4.48 | 3.5 | 4.5 | 4.2 | 3.2 | 4.3642 | 4.05 | 4.82 | 4.2649 | 3.9985 | 4.43 | 4.46 | 4.87 | 4.37 | 4.14 | 4.22 | 4.4851 | 4.58 | 3.7608 | 4.3267
| 1 | 5 | 4.29 | 4.66 | 4.28 | 4.5 | 3.3 | 4.44 | 4.1 | 4.1 | 4.3825 | 3.86 | 4.73 | 4.2789 | 3.9837 | 4.5 | 4.28 | 4.73 | 4.37 | 4.1848 | 4.31 | 4.4817 | 4.65 | 3.7436 | 33
| 1 | 6 | 4.31 | 4.67 | 4.37 | 4.61 | 3.4 | 4.43 | 4.2 | 4 | 4.4388 | 4.02 | 4.73 | 4.3004 | 3.9998 | 4.61 | 4.36 | 4.97 | 4.37 | 4.1494 | 4.31 | 4.5173 | 4.66 | 3.785
| Mean |  | 4.3233 | 4.545 | 4.3017 | 4.5067 | 3.4167 | 4.4483 | 4.1833 | 3.8333 | 4.41 | 4.01 | 4.7 | 4.2502 | 3.9906 | 4.505 | 4.305 | 4.8317 | 4.3783 | 4.155 | 4.2567 | 4.4608 | 4.605 | 3.7347
| Median |  | 4.315 | 4.575 | 4.305 | 4.485 | 3.45 | 4.435 | 4.2 | 3.9 | 4.4107 | 4.03 | 4.705 | 4.2719 | 3.9992 | 4.475 | 4.31 | 4.815 | 4.375 | 4.1524 | 4.265 | 4.4608 | 4.59 | 3.7522
| Std Dev. |  | 0.028048 | 0.13278 | 0.085654 | 0.051251 | 0.1472 | 0.034303 | 0.075277 | 0.32042 | 0.055434 | 0.07746 | 0.078994 | 0.063998 | 0.032697 | 0.085264 | 0.1031 | 0.1098 | 0.0098319 | 0.015959 | 0.065929 | 0.04133 | 0.040373 | 0.071329
| Rel.Std.Dev. |  | 0.6487% | 2.921% | 1.991% | 1.137% | 4.308% | 0.7711% | 1.799% | 8.359% | 1.257% | 1.932% | 1.681% | 1.506% | 0.8194% | 1.893% | 2.395% | 2.273% | 0.2246% | 0.3841% | 1.549% | 0.9265% | 0.8767% | 1.91%
| PDM3 |  | -0.07852% | 5.045% | -0.5793% | 4.159% | -21.03% | 2.811% | -3.314% | -11.4% | 1.923% | -7.32% | 8.627% | -1.769% | -7.769% | 4.12% | -0.5022% | 11.67% | 1.193% | -3.969% | -1.619% | 3.1% | 6.431% | -13.68%
| Z-Score (Absolute) |  | 0.04 | 0.89 | 0.05 | 0.74 | 3.46 | 0.52 | 0.5 | 1.85 | 0.37 | 1.17 | 1.49 | 0.25 | 1.25 | 0.74 | 0.04 | 1.99 | 0.25 | 0.61 | 0.22 | 0.57 | 1.12 | 2.23
| Analytical results for La in OREAS 612 (Certified Value 36 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 1
1
| 1 | 1 | 40 | 23.6 | < 1 | 38 | < 50 | 45.3 | < 50 | < 50 | 39.45 | 37.1 | 37.3 | 38.387 | 32.2 | 40 | 36 | 30 | 34.3147 | 30.1 | 36 | 37.5 | 31.7173 |  | 1
| 1 | 2 | 39.1 | 24.2 | < 1 | 37.79 | < 50 | 43.3 | < 50 | < 50 | 38.21 | 35.6 | 37.6 | 38.5488 | 33.2 | 41.5 | 36 | 34 | 34.32 | 29.9 | 36 | 38 | 31.277 |  | 22
| 1 | 3 | 39.9 | 27.7 | < 1 | 37.43 | < 50 | 43.9 | < 50 | < 50 | 39.14 | 36.5 | 37.4 | 38.3185 | 32.5 | 36.8 | 36 | 29 | 34.545 | 30.8 | 35 | 38.2 | 31.7123 |  | 16
| 1 | 4 | 41 | 25.6 | < 1 | 37.27 | < 50 | 44.6 | < 50 | < 50 | 38.91 | 35.7 | 38.7 | 37.8654 | 32.9 | 38.1 | 37 | 38 | 34.5351 | 31.3 | 37 | 38 | 32.1127 |  | 35.9836
| 1 | 5 | 39.1 | 32.4 | < 1 | 37.13 | < 50 | 41.8 | < 50 | < 50 | 38.55 | 35.6 | 38.2 | 37.6131 | 32.9 | 39.5 | 35 | 33 | 34.5027 | 30.5 | 37 | 36 | 32.834 |  | 34
| 1 | 6 | 40.1 | 30.4 | < 1 | 36.98 | < 50 | 43.7 | < 50 | < 50 | 39.2 | 35.4 | 37.6 | 38.2671 | 32.8 | 37.3 | 37 | 36 | 34.7023 | 28.8 | 37 | 37.5 | 32.531
| Mean |  | 39.8667 | 27.3167 |  | 37.4333 |  | 43.7667 |  |  | 38.91 | 35.9833 | 37.8 | 38.1667 | 32.75 | 38.8667 | 36.1667 | 33.3333 | 34.4866 | 30.2333 | 36.3333 | 37.5333 | 32.0307
| Median |  | 39.95 | 26.65 |  | 37.35 |  | 43.8 |  |  | 39.025 | 35.65 | 37.6 | 38.2928 | 32.85 | 38.8 | 36 | 33.5 | 34.5189 | 30.3 | 36.5 | 37.75 | 31.915
| Std Dev. |  | 0.71181 | 3.5205 |  | 0.39307 |  | 1.1961 |  |  | 0.45786 | 0.66758 | 0.54037 | 0.35354 | 0.35071 | 1.7851 | 0.75277 | 3.4448 | 0.14822 | 0.86178 | 0.8165 | 0.80416 | 0.57792
| Rel.Std.Dev. |  | 1.785% | 12.89% |  | 1.05% |  | 2.733% |  |  | 1.177% | 1.855% | 1.43% | 0.9263% | 1.071% | 4.593% | 2.081% | 10.33% | 0.4298% | 2.85% | 2.247% | 2.143% | 1.804%
| PDM3 |  | 10.79% | -24.09% |  | 4.029% |  | 21.63% |  |  | 8.133% | -0.0006551% | 5.048% | 6.067% | -8.986% | 8.012% | 0.5088% | -7.365% | -4.16% | -15.98% | 0.972% | 4.307% | -10.99%
| Z-Score (Absolute) |  | 0.89 | 2 | 8.18 | 0.33 | 2.53 | 1.79 | 2.53 | 2.53 | 0.67 | 0 | 0.42 | 0.5 | 0.74 | 0.66 | 0.04 | 0.61 | 0.34 | 1.32 | 0.08 | 0.36 | 0.91
| Analytical results for Li in OREAS 612 (Certified Value 32.3 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 1
2
| 1 | 1 | 31.8 | 37 | 89.2 | 35 | 32.3 | 31.3 | 36 | 34 | 32.2376 | 31 | 31.8 | 30.7 | 32 | 31.2 | 30.5784 | 30 | 32 | 34.8 | 32.1127 |  |  |  | 1
| 1 | 2 | 31.3 | 36 | 32.78 | 34 | 31.6 | 32.5 | 35 | 35 | 33.2747 | 31 | 32.7 | 32.8 | 31.3 | 31.6 | 30.6032 | 30 | 32 | 34.7 | 31.7073 |  |  |  | #N/A
| 1 | 3 | 31.6 | 38 | 35.15 | 35 | 32.4 | 33.1 | 35 | 33 | 32.9437 | 31 | 32.8 | 33.4 | 31.3 | 32 | 30.564 | 30 | 32 | 32.8 | 32.3887 |  |  |  | 16
| 1 | 4 | 32 | 40 | 33.77 | 34 | 32.2 | 32.2 | 35 | 32 | 31.8366 | 31 | 32 | 32.5 | 30.7 | 32 | 30.372 | 30 | 33 | 32.2 | 33.4617 |  |  |  | 32.2771
| 1 | 5 | 31.2 | 39 | 31.04 | 35 | 32.9 | 32.2 | 36 | 35 | 31.6794 | 31 | 33.4 | 30.9 | 30.4 | 30.9 | 30.7168 | 30 | 33 | 32.9 | 34.471 |  |  |  | 35
| 1 | 6 | 31.8 | 38 | 32.52 | 36 | 32.3 | 33 | 35 | 35 | 32.1083 | 31 | 33.9 | 32.3 | 32 | 32.1 | 30.2616 | 30 | 33 | 32 | 34.8385
| Mean |  | 31.6167 | 38 | 42.41 | 34.8333 | 32.2833 | 32.3833 | 35.3333 | 34 | 32.3467 | 31 | 32.7667 | 32.1 | 31.2833 | 31.6333 | 30.516 | 30 | 32.5 | 33.2333 | 33.1633
| Median |  | 31.7 | 38 | 33.275 | 35 | 32.3 | 32.35 | 35 | 34.5 | 32.1729 | 31 | 32.75 | 32.4 | 31.3 | 31.8 | 30.5712 | 30 | 32.5 | 32.85 | 32.9252
| Std Dev. |  | 0.31252 | 1.4142 | 22.963 | 0.75277 | 0.41673 | 0.65549 | 0.5164 | 1.2649 | 0.63115 | 0 | 0.80166 | 1.0752 | 0.65549 | 0.49261 | 0.16708 | 0 | 0.54772 | 1.2242 | 1.2984
| Rel.Std.Dev. |  | 0.9885% | 3.722% | 54.15% | 2.161% | 1.291% | 2.024% | 1.462% | 3.72% | 1.951% | 0% | 2.447% | 3.349% | 2.095% | 1.557% | 0.5475% | 0% | 1.685% | 3.684% | 3.915%
| PDM3 |  | -2.046% | 17.73% | 31.39% | 7.92% | 0.01919% | 0.329% | 9.469% | 5.338% | 0.2156% | -3.957% | 1.517% | -0.5488% | -3.079% | -1.995% | -5.456% | -7.055% | 0.6905% | 2.962% | 2.745%
| Z-Score (Absolute) |  | 0.61 | 4.46 | 7.95 | 1.94 | 0.08 | 0 | 2.34 | 1.28 | 0.03 | 1.1 | 0.3 | 0.22 | 0.87 | 0.59 | 1.48 | 1.89 | 0.09 | 0.67 | 0.62
| Analytical results for Lu in OREAS 612 (Certified Value 0.15 ppm) | [23 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.20 | 1.22 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.15 | 0.09 | 0.15 | 0.14 | 0.1 | 0.14335 | 0.17 | 0.19 | 0.14 | 0.15 | [12 empty] | 1
| 1 | 2 | 0.15 | 0.11 | 0.15 | 0.13 | 0.2 | 0.13367 | 0.17 | 0.18 | 0.13 | 0.15 | [12 empty] | 23
| 1 | 3 | 0.15 | 0.11 | 0.16 | 0.14 | 0.2 | 0.13452 | 0.17 | 0.18 | 0.14 | 0.15 | [12 empty] | 16
| 1 | 4 | 0.15 | 0.11 | 0.16 | 0.14 | 0.1 | 0.14533 | 0.17 | 0.18 | 0.14 | 0.15 | [12 empty] | 0.14819
| 1 | 5 | 0.14 | 0.13 | 0.16 | 0.14 | 0.1 | 0.13828 | 0.16 | 0.2 | 0.15 | 0.15 | [12 empty] | 36
| 1 | 6 | 0.14 | 0.12 | 0.16 | 0.14 | 0.1 | 0.12904 | 0.17 | 0.18 | 0.13 | 0.16
| Mean |  | 0.14667 | 0.11167 | 0.15667 | 0.13833 | 0.13333 | 0.13736 | 0.16833 | 0.185 | 0.13833 | 0.15167
| Median |  | 0.15 | 0.11 | 0.16 | 0.14 | 0.1 | 0.1364 | 0.17 | 0.18 | 0.14 | 0.15
| Std Dev. |  | 0.005164 | 0.013292 | 0.005164 | 0.0040825 | 0.05164 | 0.0061832 | 0.0040825 | 0.0083666 | 0.0075277 | 0.0040825
| Rel.Std.Dev. |  | 3.521% | 11.9% | 3.296% | 2.951% | 38.73% | 4.501% | 2.425% | 4.522% | 5.442% | 2.692%
| PDM3 |  | -1.031% | -24.65% | 5.717% | -6.654% | -10.03% | -7.308% | 13.59% | 24.84% | -6.654% | 2.343%
| Z-Score (Absolute) |  | 0 | 2.54 | 0.72 | 0.6 | NA | 0.67 | 1.57 | 2.78 | 0.6 | 0.36
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 612 (Certified Value 0.259 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.26 | 0.21 | 0.25 | 0.262 | 0.15 | 0.25 | 0.2 | 0.19 | 0.264 | 0.263 | 0.26 | 0.25429 | 0.2647 | 0.24 | 0.25 | 0.273 | 0.26 | 0.25648 | 0.24 | 0.2621 | 0.26 | 0.31872 | 1
| 1 | 2 | 0.26 | 0.21 | 0.27 | 0.26 | 0.15 | 0.24 | 0.22 | 0.2 | 0.2573 | 0.253 | 0.26 | 0.25737 | 0.2699 | 0.25 | 0.26 | 0.27 | 0.27 | 0.25572 | 0.25 | 0.26 | 0.26 | 0.30195 | #N/A
| 1 | 3 | 0.26 | 0.23 | 0.27 | 0.261 | 0.17 | 0.25 | 0.23 | 0.21 | 0.265 | 0.255 | 0.26 | 0.25778 | 0.2697 | 0.24 | 0.25 | 0.27 | 0.26 | 0.25244 | 0.25 | 0.2603 | 0.26 | 0.30298 | 16
| 1 | 4 | 0.27 | 0.24 | 0.26 | 0.259 | 0.16 | 0.25 | 0.23 | 0.2 | 0.2594 | 0.25 | 0.27 | 0.25211 | 0.2695 | 0.24 | 0.26 | 0.272 | 0.27 | 0.25456 | 0.25 | 0.2637 | 0.26 | 0.30315 | 0.25884
| 1 | 5 | 0.25 | 0.25 | 0.26 | 0.262 | 0.16 | 0.25 | 0.22 | 0.22 | 0.2598 | 0.25 | 0.26 | 0.25081 | 0.2702 | 0.25 | 0.25 | 0.287 | 0.27 | 0.25831 | 0.25 | 0.2639 | 0.26 | 0.31183 | 37
| 1 | 6 | 0.26 | 0.24 | 0.27 | 0.267 | 0.13 | 0.25 | 0.22 | 0.2 | 0.2637 | 0.248 | 0.26 | 0.25421 | 0.2726 | 0.25 | 0.26 | 0.284 | 0.27 | 0.25366 | 0.25 | 0.2652 | 0.26 | 0.33503
| Mean |  | 0.26 | 0.23 | 0.26333 | 0.26183 | 0.15333 | 0.24833 | 0.22 | 0.20333 | 0.26153 | 0.25317 | 0.26167 | 0.25443 | 0.26943 | 0.245 | 0.255 | 0.276 | 0.26667 | 0.25519 | 0.24833 | 0.26253 | 0.26 | 0.31228
| Median |  | 0.26 | 0.235 | 0.265 | 0.2615 | 0.155 | 0.25 | 0.22 | 0.2 | 0.26175 | 0.2515 | 0.26 | 0.25425 | 0.2698 | 0.245 | 0.255 | 0.2725 | 0.27 | 0.25514 | 0.25 | 0.2629 | 0.26 | 0.30749
| Std Dev. |  | 0.0063246 | 0.016733 | 0.008165 | 0.0027869 | 0.013663 | 0.0040825 | 0.010954 | 0.010328 | 0.0031072 | 0.0054191 | 0.0040825 | 0.0027703 | 0.0025812 | 0.0054772 | 0.0054772 | 0.00751 | 0.005164 | 0.0020966 | 0.0040825 | 0.0020944 | 0 | 0.012926
| Rel.Std.Dev. |  | 2.433% | 7.275% | 3.101% | 1.064% | 8.91% | 1.644% | 4.979% | 5.079% | 1.188% | 2.141% | 1.56% | 1.089% | 0.958% | 2.236% | 2.148% | 2.721% | 1.936% | 0.8216% | 1.644% | 0.7978% | 0% | 4.139%
| PDM3 |  | 0.4487% | -11.14% | 1.737% | 1.157% | -40.76% | -4.059% | -15% | -21.44% | 1.041% | -2.191% | 1.093% | -1.704% | 4.093% | -5.346% | -1.483% | 6.63% | 3.024% | -1.408% | -4.059% | 1.427% | 0.4487% | 20.65%
| Z-Score (Absolute) |  | 0.22 | 2.51 | 0.52 | 0.39 | 9.5 | 0.84 | 3.42 | 4.94 | 0.36 | 0.4 | 0.37 | 0.29 | 1.08 | 1.15 | 0.24 | 1.68 | 0.83 | 0.22 | 0.84 | 0.45 | 0.22 | 4.98
| Analytical results for Mn in OREAS 612 (Certified Value 0.045 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.0443 | 0.03 | 0.042914 | 0.0455 | 0.04 | 0.0446 | 0.045 | 0.041 | 0.045 | 0.0439 | 0.0447 | 0.044566 | 0.0465 | 0.0457 | 0.0428 | 0.0443 | 0.046 | 0.039696 | 0.0444 | 0.0448 | 0.0441 | 0.047199 | 1
| 1 | 2 | 0.044 | 0.04 | 0.046693 | 0.0452 | 0.04 | 0.045 | 0.045 | 0.042 | 0.0438 | 0.0422 | 0.045 | 0.044196 | 0.047 | 0.0472 | 0.0435 | 0.0431 | 0.0448 | 0.038882 | 0.0444 | 0.0449 | 0.0441 | 0.046254 | 24
| 1 | 3 | 0.0447 | 0.03 | 0.044132 | 0.0456 | 0.04 | 0.0454 | 0.045 | 0.042 | 0.0452 | 0.0429 | 0.0455 | 0.04446 | 0.0472 | 0.0472 | 0.0432 | 0.0437 | 0.045 | 0.039554 | 0.0442 | 0.0451 | 0.0447 | 0.04621 | 16
| 1 | 4 | 0.0453 | 0.03 | 0.04522 | 0.045 | 0.039 | 0.0455 | 0.044 | 0.041 | 0.0441 | 0.042 | 0.0466 | 0.044824 | 0.0475 | 0.0475 | 0.0445 | 0.0437 | 0.0416 | 0.039439 | 0.0442 | 0.0455 | 0.0439 | 0.045803 | 0.044806
| 1 | 5 | 0.0433 | 0.04 | 0.044363 | 0.0454 | 0.04 | 0.0449 | 0.045 | 0.042 | 0.0444 | 0.0422 | 0.0463 | 0.043603 | 0.0473 | 0.0481 | 0.0427 | 0.0438 | 0.0455 | 0.039957 | 0.0458 | 0.0455 | 0.0447 | 0.048165 | 38
| 1 | 6 | 0.0437 | 0.04 | 0.045852 | 0.0464 | 0.04 | 0.0453 | 0.045 | 0.041 | 0.045 | 0.0417 | 0.047 | 0.044404 | 0.0479 | 0.0494 | 0.0454 | 0.0442 | 0.0447 | 0.039689 | 0.0455 | 0.0458 | 0.0445 | 0.047284
| Mean |  | 0.044217 | 0.035 | 0.044862 | 0.045517 | 0.039833 | 0.045117 | 0.044833 | 0.0415 | 0.044583 | 0.042483 | 0.04585 | 0.044342 | 0.047233 | 0.047517 | 0.043683 | 0.0438 | 0.0446 | 0.039536 | 0.04475 | 0.045267 | 0.044333 | 0.046819
| Median |  | 0.04415 | 0.035 | 0.044791 | 0.04545 | 0.04 | 0.04515 | 0.045 | 0.0415 | 0.0447 | 0.0422 | 0.0459 | 0.044432 | 0.04725 | 0.04735 | 0.04335 | 0.04375 | 0.0449 | 0.039622 | 0.0444 | 0.0453 | 0.0443 | 0.046726
| Std Dev. |  | 0.00071671 | 0.0054772 | 0.0013444 | 0.00048339 | 0.00040825 | 0.00034303 | 0.00040825 | 0.00054772 | 0.00056716 | 0.00079854 | 0.0009225 | 0.00041642 | 0.00047188 | 0.0012156 | 0.001061 | 0.00042895 | 0.0015479 | 0.00036423 | 0.00070922 | 0.00039328 | 0.00034448 | 0.0008827
| Rel.Std.Dev. |  | 1.621% | 15.65% | 2.997% | 1.062% | 1.025% | 0.7603% | 0.9106% | 1.32% | 1.272% | 1.88% | 2.012% | 0.9391% | 0.999% | 2.558% | 2.429% | 0.9793% | 3.471% | 0.9213% | 1.585% | 0.8688% | 0.777% | 1.885%
| PDM3 |  | -1.314% | -21.88% | 0.1266% | 1.587% | -11.1% | 0.6942% | 0.06186% | -7.378% | -0.4961% | -5.183% | 2.331% | -1.035% | 5.418% | 6.051% | -2.505% | -2.244% | -0.4589% | -11.76% | -0.1241% | 1.029% | -1.054% | 4.494%
| Z-Score (Absolute) |  | 0.3 | 7.61 | 0.21 | 0.73 | 3.77 | 0.42 | 0.19 | 2.45 | 0.01 | 1.67 | 1 | 0.2 | 2.1 | 2.32 | 0.72 | 0.63 | 0.01 | 4.01 | 0.13 | 0.54 | 0.2 | 1.77
| Analytical results for Mo in OREAS 612 (Certified Value 7.07 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.25 | 7.31 | 6.6 | 6.95 | 10 | 7.48 | 10 | 10 | 7.2 | 6 | 9 | < 20 | 6.6 | 7.27 | 7.25 | 6.7 | 6.8 | 7.413 | 7.13 | 7.4 | 6.8 | 5.6747 | 1
| 1 | 2 | 7.47 | 7.23 | 6.44 | 7.22 | 10 | 7.11 | 10 | 10 | 7.1 | 6 | 8 | < 20 | 7.2 | 7.67 | 7.19 | 6.7 | 6 | 7.3459 | 6.96 | 7.3 | 6.8 | 5.4767 | 25
| 1 | 3 | 7.45 | 7.57 | 6.68 | 7.03 | 10 | 7.37 | 10 | 10 | 7.2 | 6 | 7 | < 20 | 7 | 7.02 | 7.04 | 6.8 | 6.4 | 7.3687 | 6.97 | 7.4 | 6.98 | 5.6713 | 16
| 1 | 4 | 7.23 | 7.44 | 6.67 | 7.11 | 10 | 7.33 | < 10 | 10 | 6.9 | 6 | 7 | < 20 | 7.2 | 7 | 7.52 | 6.8 | 6.4 | 7.4137 | 7.03 | 7.8 | 6.69 | 5.587 | 7.0681
| 1 | 5 | 7.07 | 7.6 | 6.75 | 7.41 | 10 | 7.35 | 10 | 10 | 6.9 | 7 | 8 | < 20 | 7.2 | 6.97 | 7.23 | 6.6 | 6.1 | 7.3328 | 7.29 | 7 | 6.41 | 5.8205 | 39
| 1 | 6 | 7.07 | 7.45 | 6.53 | 7.18 | 10 | 7.48 | 10 | 10 | 7.1 | 6 | 8 | < 20 | 7.2 | 6.9 | 7.61 | 6.6 | 6.3 | 7.4178 | 7.27 | 7.2 | 6.91 | 5.447
| Mean |  | 7.2567 | 7.4333 | 6.6117 | 7.15 | 10 | 7.3533 | 10 | 10 | 7.0667 | 6.1667 | 7.8333 |  | 7.0667 | 7.1383 | 7.3067 | 6.7 | 6.3333 | 7.382 | 7.1083 | 7.35 | 6.765 | 5.6129
| Median |  | 7.24 | 7.445 | 6.635 | 7.145 | 10 | 7.36 | 10 | 10 | 7.1 | 6 | 8 |  | 7.2 | 7.01 | 7.24 | 6.7 | 6.35 | 7.3908 | 7.08 | 7.35 | 6.8 | 5.6292
| Std Dev. |  | 0.17512 | 0.14376 | 0.11268 | 0.16087 | 0 | 0.13574 | 0 | 0 | 0.13663 | 0.40825 | 0.75277 |  | 0.24221 | 0.28923 | 0.2151 | 0.089443 | 0.28048 | 0.037815 | 0.14621 | 0.26646 | 0.20067 | 0.13934
| Rel.Std.Dev. |  | 2.413% | 1.934% | 1.704% | 2.25% | 0% | 1.846% | 0% | 0% | 1.933% | 6.62% | 9.61% |  | 3.428% | 4.052% | 2.944% | 1.335% | 4.429% | 0.5123% | 2.057% | 3.625% | 2.966% | 2.483%
| PDM3 |  | 2.667% | 5.167% | -6.458% | 1.158% | 41.48% | 4.035% | 41.48% | 41.48% | -0.02072% | -12.75% | 10.83% |  | -0.02072% | 0.9932% | 3.375% | -5.208% | -10.4% | 4.44% | 0.5688% | 3.988% | -4.289% | -20.59%
| Z-Score (Absolute) |  | 0.37 | 0.93 | 1.65 | 0.04 | NA | 0.67 | NA | NA | 0.22 | NA | NA | 8.98 | 0.22 | 0 | 0.53 | 1.37 | 2.52 | 0.76 | 0.09 | 0.66 | 1.17 | 4.78
Results from laboratories 1.10 and 1.11 were removed due to their 1 ppm reading resolution. Results from laboratories 1.05, 1.07 and 1.08 were removed due to their 10 ppm reading resolution.
| Analytical results for Na in OREAS 612 (Certified Value 1.68 wt.%) | [23 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.61 | 1.62 | 1.93 | 1.609 | 1.62 | 1.67 | 1.63 | 1.73 | 1.6671 | 1.648 | 1.74 | 1.6385 | 1.9272 | 1.72 | 1.64 | 1.66 | 1.68 | 1.7139 | 1.72 | 1.6774 | 1.82 | 1.6762 | 1
| 1 | 2 | 1.61 | 1.59 | 2.07 | 1.612 | 1.62 | 1.68 | 1.63 | 1.73 | 1.6128 | 1.608 | 1.76 | 1.6469 | 1.9664 | 1.72 | 1.65 | 1.69 | 1.7 | 1.7111 | 1.75 | 1.6705 | 1.82 | 1.6116 | #N/A
| 1 | 3 | 1.61 | 1.66 | 1.97 | 1.613 | 1.62 | 1.71 | 1.62 | 1.73 | 1.6612 | 1.626 | 1.75 | 1.6548 | 1.9635 | 1.69 | 1.65 | 1.65 | 1.69 | 1.7093 | 1.71 | 1.6712 | 1.83 | 1.6117 | 16
| 1 | 4 | 1.69 | 1.68 | 2.02 | 1.609 | 1.61 | 1.71 | 1.61 | 1.71 | 1.6294 | 1.604 | 1.81 | 1.6322 | 1.9617 | 1.65 | 1.69 | 1.64 | 1.7 | 1.7238 | 1.68 | 1.679 | 1.82 | 1.6182 | 1.6761
| 1 | 5 | 1.61 | 1.71 | 1.96 | 1.616 | 1.62 | 1.67 | 1.64 | 1.71 | 1.6264 | 1.608 | 1.8 | 1.6275 | 1.9626 | 1.73 | 1.63 | 1.66 | 1.69 | 1.7233 | 1.71 | 1.6719 | 1.83 | 1.6392 | 40
| 1 | 6 | 1.61 | 1.72 | 2 | 1.65 | 1.61 | 1.7 | 1.63 | 1.73 | 1.6523 | 1.582 | 1.79 | 1.645 | 1.9809 | 1.73 | 1.66 | 1.66 | 1.7 | 1.7135 | 1.75 | 1.6709 | 1.84 | 1.6932
| Mean |  | 1.6233 | 1.6633 | 1.9917 | 1.6182 | 1.6167 | 1.69 | 1.6267 | 1.7233 | 1.6415 | 1.6127 | 1.775 | 1.6408 | 1.9604 | 1.7067 | 1.6533 | 1.66 | 1.6933 | 1.7158 | 1.72 | 1.6735 | 1.8267 | 1.6417
| Median |  | 1.61 | 1.67 | 1.985 | 1.6125 | 1.62 | 1.69 | 1.63 | 1.73 | 1.6408 | 1.608 | 1.775 | 1.6418 | 1.963 | 1.72 | 1.65 | 1.66 | 1.695 | 1.7137 | 1.715 | 1.6716 | 1.825 | 1.6287
| Std Dev. |  | 0.03266 | 0.050859 | 0.049565 | 0.015817 | 0.005164 | 0.018974 | 0.010328 | 0.010328 | 0.021717 | 0.022295 | 0.02881 | 0.010073 | 0.017746 | 0.031411 | 0.020656 | 0.016733 | 0.008165 | 0.0062264 | 0.026833 | 0.0037167 | 0.008165 | 0.035242
| Rel.Std.Dev. |  | 2.012% | 3.058% | 2.489% | 0.9774% | 0.3194% | 1.123% | 0.6349% | 0.5993% | 1.323% | 1.382% | 1.623% | 0.6139% | 0.9052% | 1.841% | 1.249% | 1.008% | 0.4822% | 0.3629% | 1.56% | 0.2221% | 0.447% | 2.147%
| PDM3 |  | -3.15% | -0.7632% | 18.83% | -3.458% | -3.547% | 0.8278% | -2.951% | 2.816% | -2.064% | -3.786% | 5.899% | -2.106% | 16.96% | 1.822% | -1.36% | -0.9621% | 1.027% | 2.369% | 2.618% | -0.1576% | 8.981% | -2.055%
| Z-Score (Absolute) |  | 0.7 | 0.08 | 5.02 | 0.78 | 0.8 | 0.34 | 0.65 | 0.85 | 0.42 | 0.87 | 1.66 | 0.43 | 4.54 | 0.59 | 0.23 | 0.13 | 0.39 | 0.74 | 0.8 | 0.08 | 2.46 | 0.42
| Analytical results for Nb in OREAS 612 (Certified Value 14.5 ppm) | [23 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  |  |  | 1
2
| 1 | 1 | 14.9 | 15 | 10.22 | 14.4 | 14.6 | 14.91 | 14.1 | 14.3039 | 11.98 | 15.3 | 11.8 | 12.5 | 14.4 | 13.6105 | 14.7 | 15.34 | 14.75 | 18.8605 |  |  |  |  | 1
| 1 | 2 | 14.7 | 15.1 | 11.25 | 14.5 | 13.8 | 14.5 | 14.4 | 14.0784 | 12.44 | 15.6 | 11.6 | 12.6 | 13.9 | 13.7543 | 13.9 | 14.92 | 14.9 | 19.122 |  |  |  |  | 8
| 1 | 3 | 15 | 15.7 | 10.56 | 14.3 | 14.2 | 14.42 | 14.1 | 14.293 | 12.03 | 13.8 | 11.3 | 12.8 | 12.6 | 13.6459 | 14.3 | 14.99 | 14.66 | 18.7903 |  |  |  |  | 16
| 1 | 4 | 14.6 | 16 | 10.73 | 14.5 | 14.5 | 14.51 | 14.8 | 14.3665 | 12.19 | 14 | 11.1 | 12.4 | 14.4 | 13.9791 | 14.1 | 14.99 | 14.5 | 20.4935 |  |  |  |  | 14.5358
| 1 | 5 | 14.9 | 16 | 11.19 | 14.5 | 13.9 | 14.45 | 14.4 | 14.2727 | 12.21 | 14.7 | 12 | 12.6 | 12.9 | 13.9601 | 14.9 | 14.8 | 13.96 | 21.75 |  |  |  |  | 41
| 1 | 6 | 14.8 | 15.6 | 11.12 | 14.7 | 14.6 | 14.58 | 14.5 | 14.305 | 11.94 | 13.9 | 12.5 | 12.7 | 14.6 | 13.8827 | 15.3 | 14.88 | 14.7 | 22.216
| Mean |  | 14.8167 | 15.5667 | 10.845 | 14.4833 | 14.2667 | 14.5617 | 14.3833 | 14.2699 | 12.1317 | 14.55 | 11.7167 | 12.6 | 13.8 | 13.8054 | 14.5333 | 14.9867 | 14.5783 | 20.2054
| Median |  | 14.85 | 15.65 | 10.925 | 14.5 | 14.35 | 14.505 | 14.4 | 14.2985 | 12.11 | 14.35 | 11.7 | 12.6 | 14.15 | 13.8185 | 14.5 | 14.955 | 14.68 | 19.8077
| Std Dev. |  | 0.1472 | 0.43205 | 0.41079 | 0.13292 | 0.3559 | 0.17927 | 0.26394 | 0.098925 | 0.18691 | 0.77136 | 0.50365 | 0.14142 | 0.85088 | 0.15881 | 0.52789 | 0.18737 | 0.32951 | 1.5163
| Rel.Std.Dev. |  | 0.9934% | 2.775% | 3.788% | 0.9177% | 2.495% | 1.231% | 1.835% | 0.6932% | 1.541% | 5.301% | 4.299% | 1.122% | 6.166% | 1.15% | 3.632% | 1.25% | 2.26% | 7.505%
| PDM3 |  | 1.932% | 7.092% | -25.39% | -0.3611% | -1.852% | 0.1778% | -1.049% | -1.829% | -16.54% | 0.09754% | -19.39% | -13.32% | -5.062% | -5.025% | -0.01712% | 3.102% | 0.2925% | 39%
| Z-Score (Absolute) |  | 0.55 | 1.63 | 5.17 | 0.07 | 0.24 | 0.18 | 0.07 | 0.24 | 3.31 | 0.17 | 3.91 | 2.64 | 0.91 | 0.9 | 0.14 | 0.8 | 0.21 | 8.31
| Analytical results for Nd in OREAS 612 (Certified Value 33.6 ppm) | [23 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 1
2
| 1 | 1 | 36.3 | 34.8 | 34.4 | 35.2804 | 31.3 | 31.2 | 32.6515 | 34.6 | [14 empty] | 1
| 1 | 2 | 34.7 | 32.87 | 33.9 | 34.4008 | 32.58 | 31.2 | 32.5977 | 34.39 | [14 empty] | 27
| 1 | 3 | 36.1 | 33.42 | 34.2 | 34.1592 | 32.42 | 30.5 | 32.6808 | 34.7 | [14 empty] | 16
| 1 | 4 | 36.4 | 33.2 | 35.7 | 34.7473 | 32.46 | 32.2 | 32.7308 | 34.89 | [14 empty] | 33.6431
| 1 | 5 | 34.4 | 32.5 | 34.3 | 33.9115 | 32.76 | 30.2 | 32.8526 | 34.52 | [14 empty] | 42
| 1 | 6 | 35.4 | 33.33 | 34.9 | 34.2841 | 33 | 33.2 | 32.6928 | 34.94
| Mean |  | 35.55 | 33.3533 | 34.5667 | 34.4639 | 32.42 | 31.4167 | 32.7011 | 34.6733
| Median |  | 35.75 | 33.265 | 34.35 | 34.3424 | 32.52 | 31.2 | 32.6868 | 34.65
| Std Dev. |  | 0.85499 | 0.78526 | 0.64395 | 0.48607 | 0.58883 | 1.1143 | 0.086554 | 0.21351
| Rel.Std.Dev. |  | 2.405% | 2.354% | 1.863% | 1.41% | 1.816% | 3.547% | 0.2647% | 0.6158%
| PDM3 |  | 5.668% | -0.8613% | 2.745% | 2.44% | -3.636% | -6.618% | -2.8% | 3.062%
| Z-Score (Absolute) |  | 1.12 | 0.38 | 0.45 | 0.38 | 1.02 | 1.7 | 0.83 | 0.52
| Analytical results for Ni in OREAS 612 (Certified Value 10.8 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 9.9 | 10.2 | 1.59 | 11 | 10 | 11 | 10 | 10 | 11 | 11 | 20 | 10.92 | 11.7 | 11.4 | 11 | 11 | 9.7 | 11 | 12 | 11.8 | 15.9883 |  | 1
| 1 | 2 | 10.2 | 9.5 | 1.73 | 11 | 10 | 10.8 | 10 | 20 | 10 | 10 | 12 | 11.2763 | 12.2 | 16.4 | 11 | 12 | 9 | 10.6 | 11 | 11.1 | 15.1577 |  | 28
| 1 | 3 | 9.8 | 11.8 | 1.65 | 11 | 10 | 11.4 | 10 | 10 | 12 | 11 | 12 | 11.3147 | 12 | 10.8 | 11 | 11 | 8.9 | 10.4 | 12 | 11 | 15.4303 |  | 16
| 1 | 4 | 10.3 | 10.5 | 1.69 | 10 | 10 | 10.7 | 10 | 10 | 10 | 11 | 13 | 10.9182 | 12.4 | 11.2 | 11 | 12 | 9.3 | 11.2 | 11 | 11.4 | 15.0337 |  | 10.826
| 1 | 5 | 10.3 | 9.5 | 1.39 | 11 | 10 | 11.8 | 10 | 10 | 10 | 11 | 13 | 10.8973 | 11.9 | 11.2 | 11 | 12 | 9.2 | 10.3 | 10 | 11.1 | 16.183 |  | 43
| 1 | 6 | 9.9 | 10.3 | 1.67 | 11 | 10 | 11.8 | 10 | 10 | 11 | 10 | 14 | 10.8949 | 12.3 | 10.6 | 10 | 12 | 8.9 | 10.8 | 11 | 12 | 15.5853
| Mean |  | 10.0667 | 10.3 | 1.62 | 10.8333 | 10 | 11.25 | 10 | 11.6667 | 10.6667 | 10.6667 | 14 | 11.0369 | 12.0833 | 11.9333 | 10.8333 | 11.6667 | 9.1667 | 10.7167 | 11.1667 | 11.4 | 15.5631
| Median |  | 10.05 | 10.25 | 1.66 | 11 | 10 | 11.2 | 10 | 10 | 10.5 | 11 | 13 | 10.9191 | 12.1 | 11.2 | 11 | 12 | 9.1 | 10.7 | 11 | 11.25 | 15.5078
| Std Dev. |  | 0.22509 | 0.84617 | 0.12182 | 0.40825 | 0 | 0.48888 | 0 | 4.0825 | 0.8165 | 0.5164 | 3.0332 | 0.20095 | 0.26394 | 2.2079 | 0.40825 | 0.5164 | 0.30768 | 0.34881 | 0.75277 | 0.41473 | 0.45341
| Rel.Std.Dev. |  | 2.236% | 8.215% | 7.52% | 3.768% | 0% | 4.346% | 0% | 34.99% | 7.655% | 4.841% | 21.67% | 1.821% | 2.184% | 18.5% | 3.768% | 4.426% | 3.357% | 3.255% | 6.741% | 3.638% | 2.913%
| PDM3 |  | -7.014% | -4.858% | -85.04% | 0.06795% | -7.63% | 3.917% | -7.63% | 7.765% | -1.472% | -1.472% | 29.32% | 1.948% | 11.61% | 10.23% | 0.06795% | 7.765% | -15.33% | -1.01% | 3.147% | 5.302% | 43.76%
| Z-Score (Absolute) |  | 0.67 | 0.47 | 8.1 | 0 | 0.73 | 0.37 | 0.73 | 0.73 | 0.15 | 0.15 | 2.79 | 0.18 | 1.1 | 0.97 | 0 | 0.73 | 1.47 | 0.1 | 0.29 | 0.5 | 4.16
| Analytical results for P in OREAS 612 (Certified Value 0.047 wt.%) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.0477 | 0.045 | 0.05 | 0.049 | 0.04 | 0.048 | 0.044 | 0.044 | 0.0467 | 0.049 | 0.049 | 0.047262 | 0.0451 | 0.047 | 0.0439 | 0.047 | 0.044 | 0.031339 | 0.0471 | 0.0474 | 0.0466 | 0.056739 | 1
| 1 | 2 | 0.0471 | 0.0439 | 0.05 | 0.049 | 0.041 | 0.045 | 0.044 | 0.045 | 0.0453 | 0.048 | 0.051 | 0.04782 | 0.0459 | 0.049 | 0.0444 | 0.047 | 0.045 | 0.031395 | 0.0463 | 0.0473 | 0.0467 | 0.056851 | #N/A
| 1 | 3 | 0.0477 | 0.0463 | 0.05 | 0.049 | 0.042 | 0.046 | 0.044 | 0.046 | 0.0472 | 0.049 | 0.05 | 0.047878 | 0.0454 | 0.047 | 0.0449 | 0.047 | 0.045 | 0.0311 | 0.0441 | 0.049 | 0.0465 | 0.057053 | 16
| 1 | 4 | 0.0487 | 0.0477 | 0.05 | 0.049 | 0.041 | 0.046 | 0.043 | 0.045 | 0.0455 | 0.048 | 0.048 | 0.046724 | 0.0461 | 0.046 | 0.0454 | 0.047 | 0.045 | 0.031391 | 0.0451 | 0.0488 | 0.0463 | 0.057759 | 0.046812
| 1 | 5 | 0.0471 | 0.048 | 0.05 | 0.049 | 0.041 | 0.045 | 0.045 | 0.044 | 0.0461 | 0.046 | 0.051 | 0.046583 | 0.0458 | 0.048 | 0.0434 | 0.047 | 0.045 | 0.031744 | 0.0464 | 0.0481 | 0.0472 | 0.058617 | 44
| 1 | 6 | 0.0468 | 0.0479 | 0.05 | 0.05 | 0.039 | 0.046 | 0.044 | 0.044 | 0.0466 | 0.046 | 0.051 | 0.04699 | 0.0465 | 0.049 | 0.0445 | 0.047 | 0.045 | 0.031632 | 0.0477 | 0.0471 | 0.047 | 0.058573
| Mean |  | 0.047517 | 0.046467 | 0.05 | 0.049167 | 0.040667 | 0.046 | 0.044 | 0.044667 | 0.046233 | 0.047667 | 0.05 | 0.04721 | 0.0458 | 0.047667 | 0.044417 | 0.047 | 0.044833 | 0.031434 | 0.046117 | 0.04795 | 0.046717 | 0.057599
| Median |  | 0.0474 | 0.047 | 0.05 | 0.049 | 0.041 | 0.046 | 0.044 | 0.0445 | 0.04635 | 0.048 | 0.0505 | 0.047126 | 0.04585 | 0.0475 | 0.04445 | 0.047 | 0.045 | 0.031393 | 0.04635 | 0.04775 | 0.04665 | 0.057406
| Std Dev. |  | 0.0006824 | 0.0017143 | 7.6012e-18 | 0.00040825 | 0.0010328 | 0.0010954 | 0.00063246 | 0.0008165 | 0.00073666 | 0.0013663 | 0.0012649 | 0.00054766 | 0.000498 | 0.0012111 | 0.00070828 | 7.6012e-18 | 0.00040825 | 0.00022778 | 0.0013182 | 0.00081179 | 0.00033116 | 0.00084963
| Rel.Std.Dev. |  | 1.436% | 3.689% | 1.52e-14% | 0.8303% | 2.54% | 2.381% | 1.437% | 1.828% | 1.593% | 2.866% | 2.53% | 1.16% | 1.087% | 2.541% | 1.595% | 1.617e-14% | 0.9106% | 0.7246% | 2.858% | 1.693% | 0.7089% | 1.475%
| PDM3 |  | 1.506% | -0.7375% | 6.81% | 5.03% | -13.13% | -1.734% | -6.007% | -4.583% | -1.236% | 1.826% | 6.81% | 0.8494% | -2.162% | 1.826% | -5.117% | 0.4018% | -4.227% | -32.85% | -1.485% | 2.431% | -0.2034% | 23.04%
| Z-Score (Absolute) |  | 0.51 | 0.07 | 1.89 | 1.43 | 3.28 | 0.33 | 1.44 | 1.07 | 0.2 | 0.6 | 1.89 | 0.34 | 0.44 | 0.6 | 1.21 | 0.23 | 0.97 | 8.4 | 0.26 | 0.75 | 0.07 | 6.1
| Analytical results for Pb in OREAS 612 (Certified Value 0.253 wt.%) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.26 | 0.23572 | 0.24781 | 0.2435 | 0.236 | 0.251 | 0.256 | 0.247 | 0.24685 | 0.2684 | 0.283 | 0.25903 | 0.2583 | 0.26 | 0.288 | 0.25 | 0.273 | 0.25643 | 0.2487 | 0.2518 | 0.2511 | 0.3132 | 1
| 1 | 2 | 0.25 | 0.23966 | 0.25818 | 0.2416 | 0.237 | 0.249 | 0.261 | 0.245 | 0.23864 | 0.2596 | 0.29 | 0.25123 | 0.2622 | 0.268 | 0.2883 | 0.249 | 0.276 | 0.25342 | 0.2401 | 0.2536 | 0.2527 | 0.33467 | 29
| 1 | 3 | 0.26 | 0.24395 | 0.25241 | 0.2364 | 0.24 | 0.256 | 0.261 | 0.25 | 0.24102 | 0.2642 | 0.282 | 0.25102 | 0.2624 | 0.265 | 0.2815 | 0.255 | 0.276 | 0.25538 | 0.2462 | 0.2508 | 0.2525 | 0.31732 | 16
| 1 | 4 | 0.25 | 0.25061 | 0.25791 | 0.241 | 0.238 | 0.254 | 0.255 | 0.245 | 0.24498 | 0.2582 | 0.277 | 0.25175 | 0.2635 | 0.26 | 0.2842 | 0.252 | 0.277 | 0.25329 | 0.2459 | 0.2571 | 0.2492 | 0.33246 | 0.2535
| 1 | 5 | 0.25 | 0.2606 | 0.25853 | 0.2453 | 0.238 | 0.253 | 0.263 | 0.244 | 0.24444 | 0.2597 | 0.29 | 0.24429 | 0.2625 | 0.269 | 0.2813 | 0.251 | 0.276 | 0.25572 | 0.2474 | 0.2547 | 0.2521 | 0.33687 | 45
| 1 | 6 | 0.26 | 0.25436 | 0.25999 | 0.2496 | 0.24 | 0.256 | 0.263 | 0.247 | 0.24505 | 0.256 | 0.29 | 0.24908 | 0.2656 | 0.271 | 0.2815 | 0.25 | 0.278 | 0.2566 | 0.2475 | 0.2572 | 0.2533 | 0.33367
| Mean |  | 0.255 | 0.24748 | 0.2558 | 0.2429 | 0.23817 | 0.25317 | 0.25983 | 0.24633 | 0.2435 | 0.26102 | 0.28533 | 0.25107 | 0.26242 | 0.2655 | 0.28413 | 0.25117 | 0.276 | 0.25514 | 0.24597 | 0.2542 | 0.25182 | 0.32803
| Median |  | 0.255 | 0.24728 | 0.25805 | 0.24255 | 0.238 | 0.2535 | 0.261 | 0.246 | 0.24471 | 0.25965 | 0.2865 | 0.25113 | 0.26245 | 0.2665 | 0.28285 | 0.2505 | 0.276 | 0.25555 | 0.2468 | 0.25415 | 0.2523 | 0.33307
| Std Dev. |  | 0.0054772 | 0.0093934 | 0.0046994 | 0.0044398 | 0.0016021 | 0.0027869 | 0.0034881 | 0.0021602 | 0.0030486 | 0.0045053 | 0.0055015 | 0.0047704 | 0.0023794 | 0.0046797 | 0.0032928 | 0.002137 | 0.0016733 | 0.0014557 | 0.0030448 | 0.0026586 | 0.0014757 | 0.010083
| Rel.Std.Dev. |  | 2.148% | 3.796% | 1.837% | 1.828% | 0.6727% | 1.101% | 1.342% | 0.877% | 1.252% | 1.726% | 1.928% | 1.9% | 0.9067% | 1.763% | 1.159% | 0.8508% | 0.6063% | 0.5706% | 1.238% | 1.046% | 0.586% | 3.074%
| PDM3 |  | 0.5922% | -2.373% | 0.9097% | -4.181% | -6.048% | -0.131% | 2.499% | -2.827% | -3.946% | 2.966% | 12.56% | -0.9593% | 3.518% | 4.734% | 12.08% | -0.9199% | 8.876% | 0.6475% | -2.971% | 0.2767% | -0.6635% | 29.4%
| Z-Score (Absolute) |  | 0.04 | 0.64 | 0.11 | 1.06 | 1.48 | 0.13 | 0.47 | 0.75 | 1 | 0.58 | 2.78 | 0.32 | 0.71 | 0.98 | 2.67 | 0.31 | 1.93 | 0.05 | 0.78 | 0.04 | 0.25 | 6.63
| Analytical results for Pr in OREAS 612 (Certified Value 8.98 ppm) | [23 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 9.36 | 9.28 | 9 | 9.0806 | 7.69 | 8.03 | 8.7676 | 9.87 | [14 empty] | 1
| 1 | 2 | 9.04 | 8.73 | 9.1 | 8.8497 | 7.98 | 8.18 | 8.7106 | 9.63 | [14 empty] | 30
| 1 | 3 | 9.33 | 8.97 | 8.8 | 8.8736 | 7.92 | 8.25 | 8.7869 | 9.69 | [14 empty] | 16
| 1 | 4 | 9.39 | 8.96 | 9.2 | 8.8535 | 7.91 | 8.46 | 8.7217 | 9.92 | [14 empty] | 8.9792
| 1 | 5 | 8.9 | 8.87 | 9.1 | 8.7741 | 7.98 | 8.15 | 8.65 | 9.73 | [14 empty] | 46
| 1 | 6 | 9.17 | 8.85 | 9.2 | 8.8947 | 8.08 | 8.6 | 8.6852 | 9.72
| Mean |  | 9.1983 | 8.9433 | 9.0667 | 8.8877 | 7.9267 | 8.2783 | 8.7203 | 9.76
| Median |  | 9.25 | 8.915 | 9.1 | 8.8635 | 7.95 | 8.215 | 8.7162 | 9.725
| Std Dev. |  | 0.19753 | 0.18651 | 0.15055 | 0.10294 | 0.13079 | 0.21217 | 0.050859 | 0.11136
| Rel.Std.Dev. |  | 2.147% | 2.085% | 1.661% | 1.158% | 1.65% | 2.563% | 0.5832% | 1.141%
| PDM3 |  | 2.44% | -0.3999% | 0.9736% | -1.02% | -11.72% | -7.806% | -2.883% | 8.695%
| Z-Score (Absolute) |  | 0.8 | 0.08 | 0.43 | 0.08 | 2.79 | 1.8 | 0.55 | 2.38
| Analytical results for Rb in OREAS 612 (Certified Value 211 ppm) | [23 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | [6 empty] | 0
0
| 1 | 1 | 236.2 | 223.5 | 202 | 234 | 216.29 | 205 | 205.8135 | 190 | 217 | 199 | 201 | 229.4 | 199.881 | 196.5 | 219.6 | 220 | [6 empty] | 1
| 1 | 2 | 245.2 | 209.3 | 205 | 226 | 212.72 | 211 | 205.1324 | 196.7 | 220 | 205 | 203 | 226.4 | 199.1072 | 195.2 | 221.37 | 222 | [6 empty] | 31
| 1 | 3 | 239.6 | 232 | 199 | 234 | 213.32 | 211 | 205.3644 | 197.9 | 198.5 | 204 | 204 | 230.4 | 200.4517 | 202 | 221.55 | 217 | [6 empty] | 16
| 1 | 4 | 242.8 | 231.2 | 204 | 235 | 209.35 | 215 | 203.2447 | 196.1 | 197.5 | 209 | 208 | 242.5 | 200.2757 | 205.3 | 222.48 | 216 | [6 empty] | 211.394
| 1 | 5 | 238.7 | 239.2 | 189 | 225 | 210.11 | 212 | 204.3246 | 198.4 | 207 | 198 | 205 | 218.5 | 200.6693 | 199.3 | 220.47 | 212 | [6 empty] | 47
| 1 | 6 | 243.9 | 235.6 | 209 | 233 | 209.99 | 212 | 206.0798 | 201.9 | 198.5 | 211 | 208 | 238.1 | 199.9313 | 200.9 | 222.13 | 218
| Mean |  | 241.0667 | 228.4667 | 201.3333 | 231.1667 | 211.9633 | 211 | 204.9932 | 196.8333 | 206.4167 | 204.3333 | 204.8333 | 230.8833 | 200.0527 | 199.8667 | 221.2667 | 217.5
| Median |  | 241.2 | 231.6 | 203 | 233.5 | 211.415 | 211.5 | 205.2484 | 197.3 | 202.75 | 204.5 | 204.5 | 229.9 | 200.1035 | 200.1 | 221.46 | 217.5
| Std Dev. |  | 3.4512 | 10.7524 | 6.8896 | 4.446 | 2.6567 | 3.2863 | 1.0502 | 3.9119 | 10.0171 | 5.2026 | 2.7869 | 8.5195 | 0.55271 | 3.7033 | 1.0697 | 3.4496
| Rel.Std.Dev. |  | 1.432% | 4.706% | 3.422% | 1.923% | 1.253% | 1.558% | 0.5123% | 1.987% | 4.853% | 2.546% | 1.361% | 3.69% | 0.2763% | 1.853% | 0.4834% | 1.586%
| PDM3 |  | 14.04% | 8.076% | -4.759% | 9.353% | 0.2693% | -0.1864% | -3.028% | -6.888% | -2.355% | -3.34% | -3.104% | 9.219% | -5.365% | -5.453% | 4.67% | 2.888%
| Z-Score (Absolute) |  | 2.5 | 1.53 | 0.57 | 1.74 | 0.25 | 0.18 | 0.29 | 0.92 | 0.18 | 0.34 | 0.3 | 1.71 | 0.67 | 0.68 | 0.97 | 0.68
| Analytical results for Re in OREAS 612 (Indicative Value 0.002 ppm) | [23 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.22 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 3
3
| 1 | 1 | 0.006 | 0.002 | < 0.002 | 0.001 | < 0.003 | < 0.002 | 0.009 | < 0.05 | 0.006 | < 0.002 | [12 empty] | 1
| 1 | 2 | 0.004 | < 0.002 | < 0.002 | < 0.001 | < 0.003 | < 0.002 | 0.01 | < 0.05 | < 0.005 | < 0.002 | [12 empty] | 4
| 1 | 3 | 0.005 | 0.002 | < 0.002 | 0.001 | < 0.003 | < 0.002 | 0.01 | < 0.05 | < 0.005 | < 0.002 | [12 empty] | 16
| 1 | 4 | 0.004 | < 0.002 | < 0.002 | < 0.001 | < 0.003 | 0.002 | 0.01 | < 0.05 | < 0.005 | < 0.002 | [12 empty] | 0.0023333
| 1 | 5 | 0.005 | < 0.002 | < 0.002 | < 0.001 | < 0.003 | < 0.002 | 0.01 | < 0.05 | 0.006 | < 0.002 | [12 empty] | 10
| 1 | 6 | 0.005 | < 0.002 | < 0.002 | < 0.001 | < 0.003 | < 0.002 | 0.009 | < 0.05 | < 0.005 | < 0.002
| Mean |  | 0.0048333 | 0.002 |  | 0.001 |  | 0.002 | 0.0096667 |  | 0.006
| Median |  | 0.005 | 0.002 |  | 0.001 |  | 0.002 | 0.01 |  | 0.006
| Std Dev. |  | 0.00075277 | 0 |  | 0 |  |  | 0.0005164 |  | 0
| Rel.Std.Dev. |  | 15.57% | 0% |  | 0% |  |  | 5.342% |  | 0%
| PDM3 |  | 107.1% | -14.29% |  | -57.14% |  | -14.29% | 314.3% |  | 157.1%
| Z-Score (Absolute) |  | 3.95 | 0.1 | 0.48 | 0.87 | 0.1 | 0.29 | 9.54 | 27.26 | 2.6 | 0.48
| Analytical results for S in OREAS 612 (Certified Value 0.518 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.53 | 0.45 | 0.53 | 0.544 | 0.46 | 0.5 | 0.55 | 0.5 | 0.4997 | 0.52 | 0.5 | 0.51251 | 0.4989 | 0.54 | 0.51 | 0.517 | 0.47 | 0.5369 | 0.53 | 0.5116 | 0.51 | 0.532 | 1
| 1 | 2 | 0.52 | 0.49 | 0.56 | 0.54 | 0.46 | 0.51 | 0.55 | 0.51 | 0.4862 | 0.501 | 0.52 | 0.51842 | 0.509 | 0.55 | 0.53 | 0.515 | 0.47 | 0.53353 | 0.54 | 0.5122 | 0.51 | 0.527 | 10
| 1 | 3 | 0.53 | 0.48 | 0.54 | 0.542 | 0.47 | 0.52 | 0.55 | 0.51 | 0.5034 | 0.508 | 0.52 | 0.52237 | 0.5155 | 0.51 | 0.52 | 0.516 | 0.46 | 0.53377 | 0.52 | 0.5047 | 0.51 | 0.519 | 16
| 1 | 4 | 0.53 | 0.51 | 0.53 | 0.542 | 0.45 | 0.52 | 0.53 | 0.51 | 0.4904 | 0.499 | 0.53 | 0.51468 | 0.511 | 0.5 | 0.54 | 0.517 | 0.48 | 0.53474 | 0.52 | 0.5149 | 0.51 | 0.515 | 0.51826
| 1 | 5 | 0.52 | 0.49 | 0.52 | 0.544 | 0.46 | 0.51 | 0.56 | 0.51 | 0.4933 | 0.501 | 0.53 | 0.51066 | 0.5115 | 0.52 | 0.51 | 0.52 | 0.46 | 0.53213 | 0.53 | 0.5149 | 0.52 | 0.536 | 48
| 1 | 6 | 0.52 | 0.5 | 0.55 | 0.555 | 0.45 | 0.52 | 0.56 | 0.51 | 0.4988 | 0.493 | 0.52 | 0.51843 | 0.5183 | 0.52 | 0.53 | 0.513 | 0.46 | 0.53292 | 0.53 | 0.5211 | 0.52 | 0.525
| Mean |  | 0.525 | 0.48667 | 0.53833 | 0.5445 | 0.45833 | 0.51333 | 0.55 | 0.50833 | 0.4953 | 0.50367 | 0.52 | 0.51618 | 0.5107 | 0.52333 | 0.52333 | 0.51633 | 0.46667 | 0.534 | 0.52833 | 0.51323 | 0.51333 | 0.52567
| Median |  | 0.525 | 0.49 | 0.535 | 0.543 | 0.46 | 0.515 | 0.55 | 0.51 | 0.49605 | 0.501 | 0.52 | 0.51655 | 0.51125 | 0.52 | 0.525 | 0.5165 | 0.465 | 0.53365 | 0.53 | 0.51355 | 0.51 | 0.526
| Std Dev. |  | 0.0054772 | 0.020656 | 0.01472 | 0.0053572 | 0.0075277 | 0.008165 | 0.010954 | 0.0040825 | 0.0064504 | 0.009331 | 0.010954 | 0.004348 | 0.0066882 | 0.018619 | 0.012111 | 0.0023381 | 0.008165 | 0.0016665 | 0.0075277 | 0.0053664 | 0.005164 | 0.0078401
| Rel.Std.Dev. |  | 1.043% | 4.244% | 2.734% | 0.9839% | 1.642% | 1.591% | 1.992% | 0.8031% | 1.302% | 1.853% | 2.107% | 0.8424% | 1.31% | 3.558% | 2.314% | 0.4528% | 1.75% | 0.3121% | 1.425% | 1.046% | 1.006% | 1.491%
| PDM3 |  | 1.301% | -6.096% | 3.874% | 5.063% | -11.56% | -0.9503% | 6.125% | -1.915% | -4.43% | -2.816% | 0.3361% | -0.4015% | -1.458% | 0.9792% | 0.9792% | -0.3714% | -9.955% | 3.037% | 1.944% | -0.9696% | -0.9503% | 1.429%
| Z-Score (Absolute) |  | 0.65 | 2.2 | 1.64 | 2.1 | 4.3 | 0.22 | 2.51 | 0.59 | 1.56 | 0.94 | 0.28 | 0.01 | 0.41 | 0.53 | 0.53 | 0.01 | 3.68 | 1.32 | 0.9 | 0.22 | 0.22 | 0.7
| Analytical results for Sb in OREAS 612 (Certified Value 14.1 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 14.39 | 14.65 | 12.98 | 14.61 | < 50 | 14.65 | < 50 | < 50 | 14.36 | 11 | 14.3 | 14.1221 | 13.27 | 15.35 | 12.91 | 14.2 | 13.6 | 14 | 15.02 | 13.84 | 17.2015 |  | 1
| 1 | 2 | 14.87 | 14.98 | 13.66 | 14.85 | < 50 | 14.15 | < 50 | < 50 | 14.33 | 9 | 13.7 | 14.1794 | 13.62 | 15.7 | 13.16 | 14.2 | 13.2 | 13.4 | 15.08 | 13.88 | 17.4303 |  | 33
| 1 | 3 | 14.67 | 15.34 | 13.44 | 14.87 | < 50 | 14.5 | < 50 | < 50 | 14.33 | 8 | 13.2 | 14.0759 | 13.64 | 14.3 | 12.76 | 14.3 | 12.2 | 13.7 | 14.83 | 13.65 | 18.2157 |  | 16
| 1 | 4 | 14.64 | 15.71 | 13.79 | 14.79 | < 50 | 14.6 | < 50 | < 50 | 14.16 | 9 | 13.8 | 13.9109 | 13.71 | 14.65 | 12.69 | 14.4 | 14 | 13.7 | 15.19 | 13.64 | 16.866 |  | 14.1426
| 1 | 5 | 14.42 | 15.69 | 14.21 | 14.84 | < 50 | 14 | < 50 | < 50 | 14.3 | 11 | 13.8 | 13.9215 | 13.71 | 15.45 | 13.54 | 14.3 | 12.5 | 13.9 | 14.71 | 13.07 | 16.3205 |  | 49
| 1 | 6 | 14.84 | 15.45 | 12.58 | 15.05 | < 50 | 14.05 | < 50 | < 50 | 14.48 | 7 | 13.6 | 13.9396 | 14.11 | 14.55 | 13.31 | 14.4 | 13.8 | 14.1 | 14.82 | 13.85 | 17.38
| Mean |  | 14.6383 | 15.3033 | 13.4433 | 14.835 |  | 14.325 |  |  | 14.3267 | 9.1667 | 13.7333 | 14.0249 | 13.6767 | 15 | 13.0617 | 14.3 | 13.2167 | 13.8 | 14.9417 | 13.655 | 17.2357
| Median |  | 14.655 | 15.395 | 13.55 | 14.845 |  | 14.325 |  |  | 14.33 | 9 | 13.75 | 14.0077 | 13.675 | 15 | 13.035 | 14.3 | 13.4 | 13.8 | 14.925 | 13.745 | 17.2908
| Std Dev. |  | 0.20233 | 0.41693 | 0.58565 | 0.14167 |  | 0.29112 |  |  | 0.10309 | 1.6021 | 0.3559 | 0.11566 | 0.2683 | 0.57096 | 0.33235 | 0.089443 | 0.72778 | 0.25298 | 0.18324 | 0.30507 | 0.63158
| Rel.Std.Dev. |  | 1.382% | 2.724% | 4.356% | 0.955% |  | 2.032% |  |  | 0.7195% | 17.48% | 2.592% | 0.8247% | 1.962% | 3.806% | 2.544% | 0.6255% | 5.507% | 1.833% | 1.226% | 2.234% | 3.664%
| PDM3 |  | 3.505% | 8.207% | -4.944% | 4.896% |  | 1.29% |  |  | 1.302% | -35.18% | -2.894% | -0.8324% | -3.295% | 6.063% | -7.643% | 1.113% | -6.547% | -2.422% | 5.65% | -3.448% | 21.87%
| Z-Score (Absolute) |  | 0.32 | 1 | 0.9 | 0.52 | 10.91 | 0 | 10.91 | 10.91 | 0 | NA | 0.61 | 0.31 | 0.66 | 0.69 | 1.29 | 0.03 | 1.13 | 0.54 | 0.63 | 0.69 | 2.97
Results from laboratory 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 612 (Certified Value 5.28 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 5.7 | 2.7 | 5.26 | 5.3 | < 10 | 5.7 | < 10 | < 10 | 5.3 | 5.3 | 6 | 5.3205 | 4.88 | 5.8 | 5 | 5 | 4.6 | 5.7186 | 4.7 | 5.5 | 5.1 | 4.4887 | 1
| 1 | 2 | 5.7 | 4 | 5.5 | 5.4 | < 10 | 5.5 | < 10 | < 10 | 5.2 | 5.1 | 6 | 5.4784 | 5 | 6 | 5 | 5 | 5 | 5.7333 | 4.9 | 5.5 | 5.05 | 4.475 | 34
| 1 | 3 | 5.7 | 3.5 | 5.33 | 5.4 | < 10 | 5.7 | < 10 | < 10 | 5.4 | 5.2 | 6 | 5.3342 | 4.9 | 5.4 | 5 | 5 | 4.8 | 5.7434 | 4.8 | 5.5 | 5.1 | 4.4563 | 16
| 1 | 4 | 5.4 | 3.2 | 5.47 | 5.3 | < 10 | 5.6 | < 10 | < 10 | 5.2 | 5.1 | 6 | 5.3974 | 4.96 | 5.3 | 5 | 5 | 5.1 | 5.7497 | 5 | 5.6 | 5.05 | 4.5367 | 5.2761
| 1 | 5 | 5.3 | 4.6 | 5.5 | 5.4 | < 10 | 5.6 | < 10 | < 10 | 5.2 | 5.1 | 6 | 5.2782 | 4.93 | 5.6 | 5 | 5 | 4.9 | 5.7194 | 5.1 | 5.7 | 5.15 | 4.5755 | 50
| 1 | 6 | 5.7 | 4.5 | 5.52 | 5.4 | < 10 | 5.6 | < 10 | < 10 | 5.3 | 5 | 6 | 5.3477 | 4.95 | 5.3 | 5 | 5 | 4.8 | 5.7698 | 4.9 | 5.7 | 5.15 | 4.5953
| Mean |  | 5.5833 | 3.75 | 5.43 | 5.3667 |  | 5.6167 |  |  | 5.2667 | 5.1333 | 6 | 5.3594 | 4.9367 | 5.5667 | 5 | 5 | 4.8667 | 5.7391 | 4.9 | 5.5833 | 5.1 | 4.5213
| Median |  | 5.7 | 3.75 | 5.485 | 5.4 |  | 5.6 |  |  | 5.25 | 5.1 | 6 | 5.341 | 4.94 | 5.5 | 5 | 5 | 4.85 | 5.7384 | 4.9 | 5.55 | 5.1 | 4.5127
| Std Dev. |  | 0.18348 | 0.75033 | 0.10807 | 0.05164 |  | 0.075277 |  |  | 0.08165 | 0.10328 | 0 | 0.069975 | 0.043205 | 0.28752 | 0 | 0 | 0.17512 | 0.019571 | 0.14142 | 0.098319 | 0.044721 | 0.056717
| Rel.Std.Dev. |  | 3.286% | 20.01% | 1.99% | 0.9622% |  | 1.34% |  |  | 1.55% | 2.012% | 0% | 1.306% | 0.8752% | 5.165% | 0% | 0% | 3.598% | 0.341% | 2.886% | 1.761% | 0.8769% | 1.254%
| PDM3 |  | 5.823% | -28.92% | 2.917% | 1.717% |  | 6.455% |  |  | -0.1787% | -2.706% | 13.72% | 1.579% | -6.433% | 5.507% | -5.233% | -5.233% | -7.76% | 8.775% | -7.128% | 5.823% | -3.338% | -14.31%
| Z-Score (Absolute) |  | 1.28 | 3.74 | 0.86 | 0.68 | 0.32 | 1.37 | 0.32 | 0.32 | 0.41 | 0.05 | 2.42 | 0.66 | 0.49 | 1.23 | 0.32 | 0.32 | 0.68 | 1.7 | 0.59 | 1.28 | 0.05 | 1.63
| Analytical results for Se in OREAS 612 (Indicative Value 0.72 ppm) | [23 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | < 1 | < 2 | < 3 | < 1 | 1 | < 0.5 | 0.3 | < 1.5 | 0.73 | 1 | 2 | < 1 | < 2 | < 0.5 | 0.71 | [7 empty] | 1
| 1 | 2 | < 1 | < 2 | < 3 | < 1 | 1 | < 0.5 | 0.6 | < 1.5 | 0.74 | < 1 | 2 | 1 | < 2 | < 0.5 | 0.68 | [7 empty] | 3
| 1 | 3 | < 1 | < 2 | < 3 | < 1 | 1 | < 0.5 | 0.5 | < 1.5 | 0.78 | 1 | 2 | 1 | < 2 | < 0.5 | 0.66 | [7 empty] | 16
| 1 | 4 | < 1 | < 2 | < 3 | < 1 | 1 | < 0.5 | 0.6 | < 1.5 | 0.77 | 1 | 2 | 1 | < 2 | 0.5 | 0.62 | [7 empty] | 0.72381
| 1 | 5 | < 1 | < 2 | < 3 | < 1 | 1 | < 0.5 | 0.6 | < 1.5 | 0.7 | 1 | 2 | 1 | < 2 | < 0.5 | 0.56 | [7 empty] | 9
| 1 | 6 | < 1 | < 2 | < 3 | < 1 | 1 | < 0.5 | 0.6 | < 1.5 | 0.82 | 1 | 2 | 1 | < 2 | < 0.5 | 0.65
| Mean | [5 empty] | 1 |  | 0.53333 |  | 0.75667 | 1 | 2 | 1 |  | 0.5 | 0.64667
| Median | [5 empty] | 1 |  | 0.6 |  | 0.755 | 1 | 2 | 1 |  | 0.5 | 0.655
| Std Dev. | [5 empty] | 0 |  | 0.12111 |  | 0.042269 | 0 | 0 | 0 |  |  | 0.052026
| Rel.Std.Dev. | [5 empty] | 0% |  | 22.71% |  | 5.586% | 0% | 0% | 0% |  |  | 8.045%
| PDM3 | [5 empty] | 38.16% |  | -26.32% |  | 4.539% | 38.16% | 176.3% | 38.16% |  | -30.92% | -10.66%
| Z-Score (Absolute) |  | 0.71 | 0.67 | 2.06 | 0.71 | 0.67 | 1.4 | 0.62 | 0.02 | 0 | 0.44 | 3.45 | 0.44 | 0.67 | 1.29 | 0.3
| Analytical results for Sm in OREAS 612 (Certified Value 6.75 ppm) | [23 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 7.11 | 6.95 | 6.9 | 6.8108 | 6.54 | 6.68 | 6.551 | 7.64 | [14 empty] | 1
| 1 | 2 | 7.03 | 6.47 | 7.1 | 6.5863 | 6.77 | 6.77 | 6.6043 | 7.39 | [14 empty] | 35
| 1 | 3 | 7.25 | 6.89 | 7.2 | 6.4806 | 6.69 | 6.97 | 6.573 | 7.53 | [14 empty] | 16
| 1 | 4 | 7.29 | 6.72 | 7 | 6.4999 | 6.7 | 6.96 | 6.5174 | 7.56 | [14 empty] | 6.7511
| 1 | 5 | 6.94 | 6.57 | 7 | 6.4943 | 6.79 | 6.64 | 6.4258 | 7.4 | [14 empty] | 51
| 1 | 6 | 7.17 | 6.78 | 7 | 6.6553 | 6.79 | 6.86 | 6.5107 | 7.54
| Mean |  | 7.1317 | 6.73 | 7.0333 | 6.5879 | 6.7133 | 6.8133 | 6.5304 | 7.51
| Median |  | 7.14 | 6.75 | 7 | 6.5431 | 6.735 | 6.815 | 6.5342 | 7.535
| Std Dev. |  | 0.13273 | 0.18428 | 0.10328 | 0.12824 | 0.095638 | 0.13995 | 0.061971 | 0.09716
| Rel.Std.Dev. |  | 1.861% | 2.738% | 1.468% | 1.947% | 1.425% | 2.054% | 0.949% | 1.294%
| PDM3 |  | 5.637% | -0.3125% | 4.181% | -2.418% | -0.5593% | 0.9219% | -3.269% | 11.24%
| Z-Score (Absolute) |  | 1.91 | 0 | NA | 0.67 | 0.08 | 0.4 | 0.95 | 3.7
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sn in OREAS 612 (Certified Value 3.62 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  |  |  | 2
3
| 1 | 1 | 3.4 | 3.9 | 4.47 | 3.6 | 3.8 | 3.6 | < 10 | 4 | 3.5377 | 3.43 | 3.8 | 3.5 | 4.3 | 3.6 | 3.6 | 3.8 | 3.5 | 3.2287 |  |  |  |  | 1
| 1 | 2 | 3.7 | 4.1 | 4.64 | 3.8 | 3.7 | 3.6 | < 10 | 4 | 3.6399 | 3.52 | 3.8 | 3.6 | 3.5 | 3.5 | 3.1 | 3.9 | 3.45 | 3.1573 |  |  |  |  | 19
| 1 | 3 | 3.5 | 4.1 | 4.08 | 3.7 | 3.7 | 3.8 | < 10 | 4 | 3.6476 | 3.59 | 3.5 | 3.4 | 3.5 | 3.2 | 3.6 | 3.8 | 3.55 | 2.984 |  |  |  |  | 16
| 1 | 4 | 3.7 | 4.2 | 4.28 | 3.7 | 3.8 | 3.6 | < 10 | 4 | 3.6279 | 3.46 | 3.6 | 3.9 | 3.6 | 3.6 | 3.8 | 3.9 | 3.5 | 3.0363 |  |  |  |  | 3.6198
| 1 | 5 | 3.5 | 4.2 | 3.64 | 3.7 | 3.5 | 3.7 | < 10 | 4 | 3.6884 | 3.56 | 3.7 | 3.6 | 3.4 | 3.3 | 3.6 | 3.8 | 3.4 | 2.836 |  |  |  |  | 52
| 1 | 6 | 3.6 | 4.1 | 3.29 | 3.7 | 3.6 | 3.7 | < 10 | 4 | 3.6847 | 3.71 | 3.6 | 3.5 | 3.6 | 3.6 | 3.7 | 3.7 | 3.55 | 2.938
| Mean |  | 3.5667 | 4.1 | 4.0667 | 3.7 | 3.6833 | 3.6667 |  | 4 | 3.6377 | 3.545 | 3.6667 | 3.5833 | 3.65 | 3.4667 | 3.5667 | 3.8167 | 3.4917 | 3.0301
| Median |  | 3.55 | 4.1 | 4.18 | 3.7 | 3.7 | 3.65 |  | 4 | 3.6438 | 3.54 | 3.65 | 3.55 | 3.55 | 3.55 | 3.6 | 3.8 | 3.5 | 3.0102
| Std Dev. |  | 0.12111 | 0.10954 | 0.51426 | 0.063246 | 0.1169 | 0.08165 |  | 0 | 0.054732 | 0.10055 | 0.12111 | 0.17224 | 0.32711 | 0.17512 | 0.24221 | 0.075277 | 0.058452 | 0.14418
| Rel.Std.Dev. |  | 3.395% | 2.672% | 12.65% | 1.709% | 3.174% | 2.227% |  | 0% | 1.505% | 2.836% | 3.303% | 4.807% | 8.962% | 5.052% | 6.791% | 1.972% | 1.674% | 4.758%
| PDM3 |  | -1.469% | 13.27% | 12.34% | 2.215% | 1.754% | 1.294% |  | 10.5% | 0.4941% | -2.067% | 1.294% | -1.008% | 0.8336% | -4.231% | -1.469% | 5.438% | -3.54% | -16.29%
| Z-Score (Absolute) |  | 0.67 | 3.64 | 3.37 | 0.4 | 0.27 | 0.13 | 10.92 | NA | 0.1 | 0.85 | 0.13 | 0.54 | 0 | 1.48 | 0.67 | 1.35 | 1.28 | 5.02
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 612 (Certified Value 228 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 223.3 | 99.4 | 219.09 | 233 | 170 | 226 | 210 | 190 | 236.52 | 240 | 227 | 224.2612 | 258 | 232 | 227 | 229 | 223 | 223.3134 | 225.2 | 231.83 | 230 | 230.391 | 1
| 1 | 2 | 216.1 | 168 | 234.36 | 231 | 170 | 224 | 210 | 190 | 230.8 | 231 | 233 | 225.5155 | 262 | 237 | 231 | 227 | 234 | 222.1397 | 228.6 | 230.48 | 230 | 231.6085 | 20
| 1 | 3 | 225 | 133.9 | 223.29 | 233 | 180 | 229 | 210 | 200 | 236.77 | 234 | 230 | 227.5694 | 260 | 233 | 230 | 228 | 224 | 221.981 | 221.8 | 230.53 | 228 | 232.3253 | 16
| 1 | 4 | 225.9 | 130.4 | 230.72 | 231 | 170 | 229 | 210 | 190 | 231.77 | 229 | 239 | 225.4081 | 262 | 228 | 235 | 232 | 237 | 221.3715 | 224.9 | 232.16 | 229 | 251.802 | 228.1358
| 1 | 5 | 213.9 | 168.5 | 223.48 | 235 | 180 | 228 | 220 | 200 | 229.1 | 229 | 234 | 223.9515 | 262 | 237 | 225 | 231 | 223 | 222.9537 | 233.8 | 230.79 | 224 | 264.141 | 53
| 1 | 6 | 216.8 | 166.7 | 230.13 | 239 | 170 | 229 | 220 | 190 | 235.81 | 227 | 236 | 226.7863 | 264 | 240 | 243 | 227 | 222 | 221.9096 | 234.8 | 231.96 | 231 | 275.1
| Mean |  | 220.1667 | 144.4833 | 226.845 | 233.6667 | 173.3333 | 227.5 | 213.3333 | 193.3333 | 233.4617 | 231.6667 | 233.1667 | 225.582 | 261.3333 | 234.5 | 231.8333 | 229 | 227.1667 | 222.2782 | 228.1833 | 231.2917 | 228.6667 | 247.5613
| Median |  | 220.05 | 150.3 | 226.805 | 233 | 170 | 228.5 | 210 | 190 | 233.79 | 230 | 233.5 | 225.4618 | 262 | 235 | 230.5 | 228.5 | 223.5 | 222.0604 | 226.9 | 231.31 | 229.5 | 242.0637
| Std Dev. |  | 5.1613 | 28.1648 | 5.7686 | 3.0111 | 5.164 | 2.0736 | 5.164 | 5.164 | 3.3101 | 4.7188 | 4.2622 | 1.4027 | 2.0656 | 4.3243 | 6.4627 | 2.0976 | 6.5549 | 0.72021 | 5.214 | 0.77215 | 2.5033 | 19.1452
| Rel.Std.Dev. |  | 2.344% | 19.49% | 2.543% | 1.289% | 2.979% | 0.9115% | 2.421% | 2.671% | 1.418% | 2.037% | 1.828% | 0.6218% | 0.7904% | 1.844% | 2.788% | 0.916% | 2.886% | 0.324% | 2.285% | 0.3338% | 1.095% | 7.734%
| PDM3 |  | -3.493% | -36.67% | -0.5658% | 2.424% | -24.02% | -0.2787% | -6.488% | -15.26% | 2.335% | 1.548% | 2.205% | -1.119% | 14.55% | 2.79% | 1.621% | 0.3788% | -0.4248% | -2.568% | 0.02084% | 1.383% | 0.2327% | 8.515%
| Z-Score (Absolute) |  | 1.14 | 11.58 | 0.22 | 0.72 | 7.6 | 0.13 | 2.08 | 4.84 | 0.69 | 0.45 | 0.65 | 0.39 | 4.54 | 0.84 | 0.47 | 0.08 | 0.17 | 0.85 | 0.03 | 0.4 | 0.03 | 2.64
| Analytical results for Ta in OREAS 612 (Certified Value 1.08 ppm) | [23 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
3
| 1 | 1 | 0.98 | 1.15 | 1.1 | 1.1 | 1.1 | 1.1 | 1.1282 | 0.83 | 1.16 | 0.5 | 0.95 | 1.2 | 1.37 | 1.1 | 1.35 | [7 empty] | 1
| 1 | 2 | 1.02 | 1.14 | 1.08 | 1.03 | 1.07 | 1 | 1.091 | 0.84 | 1.17 | 0.49 | 0.99 | 1.2 | 1.36 | 1.08 | 1.35 | [7 empty] | 21
| 1 | 3 | 0.97 | 1.16 | 1.07 | 1.07 | 1.05 | 1.1 | 1.0593 | 0.78 | 1.05 | 0.5 | 0.97 | 1 | 1.32 | 1.07 | 1.3 | [7 empty] | 16
| 1 | 4 | 1.03 | 1.22 | 1.1 | 1.08 | 1.06 | 1.1 | 1.1449 | 0.81 | 1.04 | 0.58 | 0.98 | 1.2 | 1.38 | 1.08 | 1.25 | [7 empty] | 1.0809
| 1 | 5 | 0.95 | 1.18 | 1.1 | 1.07 | 1.08 | 1.2 | 1.0767 | 0.85 | 1.08 | 0.48 | 0.95 | 1 | 1.26 | 1.1 | 1.25 | [7 empty] | 54
| 1 | 6 | 0.98 | 1.18 | 1.12 | 1.06 | 1.15 | 1.1 | 1.0609 | 0.83 | 1.09 | 0.58 | 0.97 | 1.2 | 1.33 | 1.1 | 1.25
| Mean |  | 0.98833 | 1.1717 | 1.095 | 1.0683 | 1.085 | 1.1 | 1.0935 | 0.82333 | 1.0983 | 0.52167 | 0.96833 | 1.1333 | 1.3367 | 1.0883 | 1.2917
| Median |  | 0.98 | 1.17 | 1.1 | 1.07 | 1.075 | 1.1 | 1.0839 | 0.83 | 1.085 | 0.5 | 0.97 | 1.2 | 1.345 | 1.09 | 1.275
| Std Dev. |  | 0.030605 | 0.028577 | 0.017607 | 0.023166 | 0.036194 | 0.063246 | 0.035681 | 0.025033 | 0.054924 | 0.045789 | 0.016021 | 0.10328 | 0.044121 | 0.013292 | 0.04916
| Rel.Std.Dev. |  | 3.097% | 2.439% | 1.608% | 2.168% | 3.336% | 5.75% | 3.263% | 3.04% | 5.001% | 8.778% | 1.654% | 9.113% | 3.301% | 1.221% | 3.806%
| PDM3 |  | -8.566% | 8.395% | 1.302% | -1.165% | 0.3771% | 1.765% | 1.163% | -23.83% | 1.611% | -51.74% | -10.42% | 4.849% | 23.66% | 0.6855% | 19.5%
| Z-Score (Absolute) |  | 1.78 | 1.32 | 0.03 | 0.43 | 0.14 | 0.11 | 0 | 4.57 | 0.08 | 9.68 | 2.12 | 0.67 | 4.12 | 0.09 | 3.35
| Analytical results for Tb in OREAS 612 (Certified Value 0.7 ppm) | [23 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.20 | 1.22 | 1.23 | [10 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 0.76 | 0.41 | 0.69 | 0.71 | 0.7 | 0.67066 | < 0.2 | 0.74 | 0.69162 | 0.63 | 0.74 | 0.65 | [10 empty] | 1
| 1 | 2 | 0.74 | 0.52 | 0.67 | 0.69 | 0.7 | 0.69055 | < 0.2 | 0.74 | 0.65653 | 0.63 | 0.73 | 0.7 | [10 empty] | 22
| 1 | 3 | 0.75 | 0.48 | 0.68 | 0.68 | 0.8 | 0.67957 | < 0.2 | 0.77 | 0.68008 | 0.64 | 0.74 | 0.7 | [10 empty] | 16
| 1 | 4 | 0.77 | 0.45 | 0.66 | 0.68 | 0.7 | 0.69892 | < 0.2 | 0.76 | 0.68879 | 0.65 | 0.75 | 0.65 | [10 empty] | 0.69827
| 1 | 5 | 0.72 | 0.58 | 0.68 | 0.69 | 0.7 | 0.6614 | < 0.2 | 0.74 | 0.69652 | 0.66 | 0.74 | 0.65 | [10 empty] | 55
| 1 | 6 | 0.74 | 0.56 | 0.68 | 0.68 | 0.7 | 0.67788 | < 0.2 | 0.78 | 0.67391 | 0.63 | 0.75 | 0.7
| Mean |  | 0.74667 | 0.5 | 0.67667 | 0.68833 | 0.71667 | 0.67983 |  | 0.755 | 0.68124 | 0.64 | 0.74167 | 0.675
| Median |  | 0.745 | 0.5 | 0.68 | 0.685 | 0.7 | 0.67873 |  | 0.75 | 0.68443 | 0.635 | 0.74 | 0.675
| Std Dev. |  | 0.017512 | 0.065422 | 0.010328 | 0.01169 | 0.040825 | 0.01347 |  | 0.017607 | 0.014587 | 0.012649 | 0.0075277 | 0.027386
| Rel.Std.Dev. |  | 2.345% | 13.08% | 1.526% | 1.698% | 5.696% | 1.981% |  | 2.332% | 2.141% | 1.976% | 1.015% | 4.057%
| PDM3 |  | 6.931% | -28.39% | -3.093% | -1.423% | 2.635% | -2.64% |  | 8.125% | -2.438% | -8.345% | 6.215% | -3.332%
| Z-Score (Absolute) |  | 1.13 | 3.04 | 0.05 | 0.14 | NA | 0 | 9.82 | 1.27 | 0.02 | 0.67 | 1.05 | 0.08
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 612 (Certified Value < 0.05 ppm) | [23 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | < 0.05 | 0.09 | < 1 | < 0.05 | < 0.05 | < 0.2 | < 0.1 | < 0.2 | 0.05 | < 0.05 | 1.22 | < 0.1 | < 0.1 | < 0.05 | < 0.2 | [7 empty] | 1
| 1 | 2 | < 0.05 | 0.13 | < 1 | < 0.05 | < 0.05 | < 0.2 | < 0.1 | < 0.2 | 0.04 | < 0.05 | 1.42 | < 0.1 | < 0.1 | < 0.05 | < 0.2 | [7 empty] | 23
| 1 | 3 | < 0.05 | 0.08 | < 1 | < 0.05 | < 0.05 | < 0.2 | < 0.1 | < 0.2 | < 0.03 | < 0.05 | 1.26 | < 0.1 | < 0.1 | < 0.05 | < 0.2 | [7 empty] | 16
| 1 | 4 | < 0.05 | 0.1 | < 1 | < 0.05 | < 0.05 | < 0.2 | < 0.1 | < 0.2 | < 0.03 | < 0.05 | 1.21 | < 0.1 | < 0.1 | < 0.05 | < 0.2 | [7 empty] | < 0.05
| 1 | 5 | < 0.05 | 0.08 | < 1 | < 0.05 | < 0.05 | < 0.2 | < 0.1 | < 0.2 | < 0.03 | < 0.05 | 1.26 | < 0.1 | < 0.1 | < 0.05 | < 0.2 | [7 empty] | 56
| 1 | 6 | < 0.05 | 0.09 | < 1 | < 0.05 | < 0.05 | < 0.2 | < 0.1 | < 0.2 | < 0.03 | < 0.05 | 1.38 | < 0.1 | < 0.1 | < 0.05 | < 0.2
| Mean |  |  | 0.095 | [6 empty] | 0.045 |  | 1.2917
| Median |  |  | 0.09 | [6 empty] | 0.045 |  | 1.26
| Std Dev. |  |  | 0.018708 | [6 empty] | 0.0070711 |  | 0.087274
| Rel.Std.Dev. |  |  | 19.69% | [6 empty] | 15.71% |  | 6.757%
PDM3
| Z-Score (Absolute) |  | 0.67 | 1.21 | 12.14 | 0.67 | 0.67 | 1.35 | 0 | 1.35 | 0.67 | 0.67 | 33.49 | 0 | 0 | 0.67 | 1.35
| Analytical results for Th in OREAS 612 (Certified Value 12.9 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 13.68 | 8.2 | 13.31 | < 50 | 14.8 | < 50 | < 50 | 13.21 | 13.1 | 13.587 | 12.93 | 14.25 | 12.87 | 12.2 | 13.2 | 12.6149 | 11.7 | 12.89 | 13.18 |  |  |  | 1
| 1 | 2 | 13.06 | 9 | 13.06 | < 50 | 14.35 | < 50 | < 50 | 12.97 | 13.4 | 13.6924 | 12.97 | 14.45 | 13.06 | 12.1 | 13.5 | 12.485 | 11.4 | 12.53 | 13.09 |  |  |  | 24
| 1 | 3 | 13.22 | 9.5 | 12.82 | < 50 | 14.7 | < 50 | < 50 | 13.26 | 13.2 | 13.4193 | 12.87 | 12.05 | 12.81 | 12.2 | 12.1 | 12.5891 | 11.7 | 12.62 | 12.83 |  |  |  | 16
| 1 | 4 | 13.8 | 9.1 | 12.9 | < 50 | 14.9 | < 50 | < 50 | 13.53 | 13.8 | 13.1522 | 12.7 | 12.1 | 12.48 | 12.4 | 15 | 12.508 | 11.7 | 12.82 | 13.04 |  |  |  | 12.8998
| 1 | 5 | 14.04 | 10.7 | 13.06 | < 50 | 14.6 | < 50 | < 50 | 12.87 | 13.5 | 13.307 | 12.81 | 12.75 | 12.61 | 12.2 | 13 | 12.4568 | 12.3 | 12.47 | 12.68 |  |  |  | 57
| 1 | 6 | 13.03 | 10 | 13.24 | < 50 | 14.8 | < 50 | < 50 | 13.28 | 13.6 | 13.3606 | 13 | 12.3 | 13.41 | 12.1 | 14.1 | 12.5274 | 11.5 | 12.8 | 12.99
| Mean |  | 13.4717 | 9.4167 | 13.065 |  | 14.6917 |  |  | 13.1867 | 13.4333 | 13.4198 | 12.88 | 12.9833 | 12.8733 | 12.2 | 13.4833 | 12.5302 | 11.7167 | 12.6883 | 12.9683
| Median |  | 13.45 | 9.3 | 13.06 |  | 14.75 |  |  | 13.235 | 13.45 | 13.39 | 12.9 | 12.525 | 12.84 | 12.2 | 13.35 | 12.5177 | 11.7 | 12.71 | 13.015
| Std Dev. |  | 0.42475 | 0.86583 | 0.18865 |  | 0.196 |  |  | 0.23653 | 0.2582 | 0.19499 | 0.11171 | 1.0889 | 0.33194 | 0.10954 | 0.99079 | 0.060929 | 0.31252 | 0.17198 | 0.18302
| Rel.Std.Dev. |  | 3.153% | 9.195% | 1.444% |  | 1.334% |  |  | 1.794% | 1.922% | 1.453% | 0.8673% | 8.387% | 2.579% | 0.8979% | 7.348% | 0.4863% | 2.667% | 1.355% | 1.411%
| PDM3 |  | 4.433% | -27% | 1.281% |  | 13.89% |  |  | 2.224% | 4.136% | 4.031% | -0.1532% | 0.6479% | -0.2049% | -5.425% | 4.524% | -2.865% | -9.171% | -1.639% | 0.5316%
| Z-Score (Absolute) |  | 0.67 | 6.05 | 0 | 19.79 | 2.7 | 19.79 | 19.79 | 0.2 | 0.61 | 0.59 | 0.31 | 0.14 | 0.32 | 1.43 | 0.69 | 0.89 | 2.24 | 0.62 | 0.16
| Analytical results for Ti in OREAS 612 (Certified Value 0.201 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 |  | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.203 | 0.2 | 0.19 | 0.206 | 0.18 | 0.2 | 0.19 | 0.19 | 0.199 | 0.183 | 0.218 | 0.2019 | 0.1282 | 0.203 | 0.22 | 0.199 | 0.2 | 0.19798 | 0.19 | 0.2 | 0.2 |  | 1
| 1 | 2 | 0.202 | 0.2 | 0.2 | 0.204 | 0.18 | 0.201 | 0.19 | 0.19 | 0.194 | 0.174 | 0.214 | 0.20152 | 0.133 | 0.208 | 0.2 | 0.198 | 0.2 | 0.19742 | 0.2 | 0.2001 | 0.2 |  | 25
| 1 | 3 | 0.204 | 0.2 | 0.2 | 0.206 | 0.19 | 0.204 | 0.19 | 0.19 | 0.1986 | 0.178 | 0.22 | 0.19549 | 0.1264 | 0.198 | 0.2 | 0.192 | 0.2 | 0.1978 | 0.19 | 0.198 | 0.21 |  | 16
| 1 | 4 | 0.205 | 0.21 | 0.2 | 0.204 | 0.18 | 0.204 | 0.19 | 0.19 | 0.1945 | 0.174 | 0.2 | 0.19633 | 0.1311 | 0.2 | 0.2 | 0.201 | 0.201 | 0.19762 | 0.19 | 0.2016 | 0.21 |  | 0.20057
| 1 | 5 | 0.201 | 0.21 | 0.2 | 0.206 | 0.18 | 0.202 | 0.19 | 0.19 | 0.1949 | 0.175 | 0.22 | 0.2018 | 0.1289 | 0.203 | 0.2 | 0.204 | 0.2 | 0.19851 | 0.19 | 0.2009 | 0.21 |  | 58
| 1 | 6 | 0.201 | 0.21 | 0.2 | 0.211 | 0.17 | 0.203 | 0.19 | 0.19 | 0.1988 | 0.173 | 0.212 | 0.20288 | 0.1237 | 0.208 | 0.21 | 0.195 | 0.201 | 0.19755 | 0.2 | 0.2028 | 0.21
| Mean |  | 0.20267 | 0.205 | 0.19833 | 0.20617 | 0.18 | 0.20233 | 0.19 | 0.19 | 0.19663 | 0.17617 | 0.214 | 0.19999 | 0.12855 | 0.20333 | 0.205 | 0.19817 | 0.20033 | 0.19781 | 0.19333 | 0.20057 | 0.20667
| Median |  | 0.2025 | 0.205 | 0.2 | 0.206 | 0.18 | 0.2025 | 0.19 | 0.19 | 0.19675 | 0.1745 | 0.216 | 0.20166 | 0.12855 | 0.203 | 0.2 | 0.1985 | 0.2 | 0.19771 | 0.19 | 0.2005 | 0.21
| Std Dev. |  | 0.001633 | 0.0054772 | 0.0040825 | 0.0025626 | 0.0063246 | 0.001633 | 3.0405e-17 | 3.0405e-17 | 0.0023939 | 0.0037639 | 0.0075895 | 0.0032003 | 0.0033074 | 0.0040825 | 0.0083666 | 0.0042622 | 0.0005164 | 0.0003957 | 0.005164 | 0.0016305 | 0.005164
| Rel.Std.Dev. |  | 0.8058% | 2.672% | 2.058% | 1.243% | 3.514% | 0.8071% | 1.6e-14% | 1.6e-14% | 1.217% | 2.137% | 3.546% | 1.6% | 2.573% | 2.008% | 4.081% | 2.151% | 0.2578% | 0.2% | 2.671% | 0.813% | 2.499%
| PDM3 |  | 1.043% | 2.207% | -1.117% | 2.788% | -10.26% | 0.8771% | -5.272% | -5.272% | -1.965% | -12.17% | 6.694% | -0.2928% | -35.91% | 1.376% | 2.207% | -1.2% | -0.12% | -1.377% | -3.61% | -0.003679% | 3.038%
| Z-Score (Absolute) |  | 0.36 | 0.67 | 0.22 | 0.83 | 2.69 | 0.32 | 1.34 | 1.34 | 0.45 | 3.2 | 1.88 | 0 | 9.61 | 0.45 | 0.67 | 0.24 | 0.05 | 0.29 | 0.89 | 0.08 | 0.9
| Analytical results for Tl in OREAS 612 (Certified Value 3.77 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  | 2
3
| 1 | 1 | 3.9 | 3.82 | < 3 | 3.74 | < 50 | 4 | < 50 | < 50 | 3.9 | 3.7 | 3.7637 | 3.686 | 3.68 | 3.79 | 3.55 | 3.5 | 3.777 | 3.86 | 3.9 | 3.7 |  |  | 1
| 1 | 2 | 4.01 | 3.81 | < 3 | 3.77 | < 50 | 3.97 | < 50 | < 50 | 3.82 | 3.82 | 3.6014 | 3.783 | 3.73 | 3.88 | 3.59 | 3.3 | 3.7606 | 3.72 | 3.78 | 3.68 |  |  | 26
| 1 | 3 | 3.9 | 3.98 | < 3 | 3.58 | < 50 | 4 | < 50 | < 50 | 3.86 | 3.79 | 3.5637 | 3.67 | 3.22 | 3.8 | 3.59 | 3 | 3.7936 | 3.77 | 3.83 | 3.73 |  |  | 16
| 1 | 4 | 4.01 | 4.07 | < 3 | 3.68 | < 50 | 4.03 | < 50 | < 50 | 3.83 | 3.95 | 3.673 | 3.772 | 3.28 | 3.72 | 3.62 | 3.6 | 3.8358 | 3.78 | 3.89 | 3.64 |  |  | 3.7683
| 1 | 5 | 3.9 | 4.1 | < 3 | 3.63 | < 50 | 3.89 | < 50 | < 50 | 3.66 | 3.82 | 3.506 | 3.655 | 3.45 | 3.74 | 3.56 | 3.1 | 3.6905 | 3.92 | 3.79 | 3.55 |  |  | 59
| 1 | 6 | 4.07 | 3.97 | < 3 | 3.79 | < 50 | 3.9 | < 50 | < 50 | 3.91 | 3.82 | 3.5213 | 3.742 | 3.35 | 3.86 | 3.51 | 3.4 | 3.7579 | 3.99 | 3.81 | 3.66
| Mean |  | 3.965 | 3.9583 |  | 3.6983 |  | 3.965 |  |  | 3.83 | 3.8167 | 3.6048 | 3.718 | 3.4517 | 3.7983 | 3.57 | 3.3167 | 3.7693 | 3.84 | 3.8333 | 3.66
| Median |  | 3.955 | 3.975 |  | 3.71 |  | 3.985 |  |  | 3.845 | 3.82 | 3.5826 | 3.714 | 3.4 | 3.795 | 3.575 | 3.35 | 3.7688 | 3.82 | 3.82 | 3.67
| Std Dev. |  | 0.074498 | 0.12189 |  | 0.082805 |  | 0.057533 |  |  | 0.090774 | 0.080166 | 0.098323 | 0.054798 | 0.21123 | 0.063377 | 0.038471 | 0.23166 | 0.04793 | 0.10218 | 0.050859 | 0.06229
| Rel.Std.Dev. |  | 1.879% | 3.079% |  | 2.239% |  | 1.451% |  |  | 2.37% | 2.1% | 2.728% | 1.474% | 6.12% | 1.669% | 1.078% | 6.985% | 1.272% | 2.661% | 1.327% | 1.702%
| PDM3 |  | 5.219% | 5.042% |  | -1.858% |  | 5.219% |  |  | 1.636% | 1.282% | -4.339% | -1.336% | -8.404% | 0.7959% | -5.263% | -11.99% | 0.02466% | 1.902% | 1.725% | -2.875%
| Z-Score (Absolute) |  | 0.69 | 0.66 | 10.09 | 0.48 | 92.69 | 0.69 | 92.69 | 92.69 | 0.1 | 0.04 | 0.89 | 0.39 | 1.56 | 0.04 | 1.04 | 2.15 | 0.17 | 0.14 | 0.11 | 0.65
| Analytical results for Tm in OREAS 612 (Certified Value 0.17 ppm) | [23 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 0.17 | 0.18 | 0.2 | 0.15959 | 0.17 | 0.19 | 0.13534 | 0.17 | [14 empty] | 1
| 1 | 2 | 0.16 | 0.16 | 0.2 | 0.16265 | 0.17 | 0.19 | 0.14079 | 0.16 | [14 empty] | 27
| 1 | 3 | 0.16 | 0.16 | 0.2 | 0.16217 | 0.17 | 0.18 | 0.12576 | 0.17 | [14 empty] | 16
| 1 | 4 | 0.17 | 0.16 | 0.2 | 0.16038 | 0.17 | 0.19 | 0.14738 | 0.17 | [14 empty] | 0.16548
| 1 | 5 | 0.16 | 0.16 | 0.2 | 0.1599 | 0.17 | 0.2 | 0.1507 | 0.16 | [14 empty] | 60
| 1 | 6 | 0.17 | 0.17 | 0.2 | 0.15976 | 0.17 | 0.19 | 0.14401 | 0.17
| Mean |  | 0.165 | 0.165 | 0.2 | 0.16074 | 0.17 | 0.19 | 0.14066 | 0.16667
| Median |  | 0.165 | 0.16 | 0.2 | 0.16014 | 0.17 | 0.19 | 0.1424 | 0.17
| Std Dev. |  | 0.0054772 | 0.0083666 | 3.0405e-17 | 0.0013274 | 0 | 0.0063246 | 0.0090295 | 0.005164
| Rel.Std.Dev. |  | 3.32% | 5.071% | 1.52e-14% | 0.8258% | 0% | 3.329% | 6.419% | 3.098%
| PDM3 |  | -0.2911% | -0.2911% | 20.86% | -2.864% | 2.73% | 14.82% | -15% | 0.7161%
| Z-Score (Absolute) |  | 0 | 0 | NA | 0.67 | 0.79 | 3.96 | 3.85 | 0.26
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 612 (Certified Value 4.36 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 4.3 | 4 | 4.42 | < 50 | 5 | < 50 | < 50 | 4.42 | 4.4 | 4.5364 | 4.33 | 4.7 | 4.3 | 3.9 | 4.5 | 4.2292 | 4.3 | 4.55 | 4.4 | 1.694 |  |  | 1
| 1 | 2 | 4.4 | 3.5 | 4.43 | < 50 | 4.8 | < 50 | < 50 | 4.42 | 4.5 | 4.6952 | 4.43 | 4.6 | 4.5 | 3.96 | 4.5 | 4.1936 | 4.1 | 4.55 | 4.41 | 1.5927 |  |  | 28
| 1 | 3 | 4.5 | 4.3 | 4.43 | < 50 | 4.8 | < 50 | < 50 | 4.4 | 4.5 | 4.5648 | 4.4 | 3.9 | 4.3 | 3.99 | 4.1 | 4.2621 | 4.2 | 4.6 | 4.35 | 1.6512 |  |  | 16
| 1 | 4 | 4.6 | 4.3 | 4.43 | < 50 | 4.9 | < 50 | < 50 | 4.43 | 4.7 | 4.5973 | 4.33 | 3.8 | 4.3 | 3.99 | 4.8 | 4.321 | 4.2 | 4.66 | 4.36 | 1.5217 |  |  | 4.3589
| 1 | 5 | 4.3 | 4.5 | 4.52 | < 50 | 4.8 | < 50 | < 50 | 4.32 | 4.5 | 4.4994 | 4.34 | 4.4 | 4.3 | 3.86 | 4.3 | 4.2411 | 4.3 | 4.52 | 4.24 | 1.6565 |  |  | 61
| 1 | 6 | 4.4 | 4.1 | 4.52 | < 50 | 4.8 | < 50 | < 50 | 4.54 | 4.5 | 4.6083 | 4.35 | 4 | 4.4 | 3.88 | 4.6 | 4.2371 | 4.3 | 4.6 | 4.3 | 1.6103
| Mean |  | 4.4167 | 4.1167 | 4.4583 |  | 4.85 |  |  | 4.4217 | 4.5167 | 4.5836 | 4.3633 | 4.2333 | 4.35 | 3.93 | 4.4667 | 4.2473 | 4.2333 | 4.58 | 4.3433 | 1.6211
| Median |  | 4.4 | 4.2 | 4.43 |  | 4.8 |  |  | 4.42 | 4.5 | 4.581 | 4.345 | 4.2 | 4.3 | 3.93 | 4.5 | 4.2391 | 4.25 | 4.575 | 4.355 | 1.6308
| Std Dev. |  | 0.1169 | 0.34881 | 0.047924 |  | 0.083666 |  |  | 0.070545 | 0.098319 | 0.067704 | 0.041793 | 0.38297 | 0.083666 | 0.057271 | 0.24221 | 0.04242 | 0.08165 | 0.0502 | 0.064083 | 0.060473
| Rel.Std.Dev. |  | 2.647% | 8.473% | 1.075% |  | 1.725% |  |  | 1.595% | 2.177% | 1.477% | 0.9578% | 9.047% | 1.923% | 1.457% | 5.423% | 0.9987% | 1.929% | 1.096% | 1.475% | 3.73%
| PDM3 |  | 1.324% | -5.558% | 2.28% |  | 11.27% |  |  | 1.439% | 3.618% | 5.153% | 0.1006% | -2.882% | -0.2053% | -9.841% | 2.471% | -2.561% | -2.882% | 5.071% | -0.3582% | -62.81%
| Z-Score (Absolute) |  | 0.01 | 1.21 | 0.16 | 82.55 | 1.73 | 82.55 | 82.55 | 0.01 | 0.39 | 0.66 | 0.22 | 0.75 | 0.28 | 1.96 | 0.19 | 0.69 | 0.75 | 0.65 | 0.3 | 11.22
| Analytical results for V in OREAS 612 (Certified Value 28.9 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 29 | 28 | < 1 | 30 | 30 | 29 | 30 | 30 | 27 | 29 | 29 | 29.6888 | 28 | 29 | 29 | 28 | 28 | 27.2944 | 29 | 29 | 29 | 25.1687 | 1
| 1 | 2 | 29 | 27 | < 1 | 30 | 30 | 28 | 30 | 30 | 28 | 28 | 30 | 29.1039 | 28 | 30 | 29 | 29 | 27 | 27.0408 | 29 | 29 | 29 | 24.9903 | 29
| 1 | 3 | 29 | 28 | < 1 | 30 | 30 | 29 | 30 | 30 | 29 | 29 | 30 | 29.0967 | 29 | 29 | 29 | 30 | 25 | 27.1824 | 29 | 29 | 29 | 24.963 | 16
| 1 | 4 | 30 | 28 | < 1 | 30 | 30 | 29 | 30 | 30 | 29 | 28 | 30 | 29.3026 | 28 | 28 | 29 | 29 | 28 | 27.2048 | 29 | 29 | 29 | 25.1703 | 28.9217
| 1 | 5 | 29 | 28 | < 1 | 30 | 30 | 28 | 30 | 30 | 28 | 28 | 30 | 28.9472 | 28 | 29 | 29 | 29 | 26 | 27.502 | 30 | 30 | 29 | 25.519 | 62
| 1 | 6 | 29 | 26 | < 1 | 31 | 30 | 28 | 30 | 30 | 29 | 27 | 29 | 29.5014 | 28 | 30 | 29 | 29 | 28 | 27.3436 | 30 | 30 | 29 | 25.2763
| Mean |  | 29.1667 | 27.5 |  | 30.1667 | 30 | 28.5 | 30 | 30 | 28.3333 | 28.1667 | 29.6667 | 29.2734 | 28.1667 | 29.1667 | 29 | 29 | 27 | 27.2613 | 29.3333 | 29.3333 | 29 | 25.1813
| Median |  | 29 | 28 |  | 30 | 30 | 28.5 | 30 | 30 | 28.5 | 28 | 30 | 29.2032 | 28 | 29 | 29 | 29 | 27.5 | 27.2496 | 29 | 29 | 29 | 25.1695
| Std Dev. |  | 0.40825 | 0.83666 |  | 0.40825 | 0 | 0.54772 | 0 | 0 | 0.8165 | 0.75277 | 0.5164 | 0.2799 | 0.40825 | 0.75277 | 0 | 0.63246 | 1.2649 | 0.15738 | 0.5164 | 0.5164 | 0 | 0.2037
| Rel.Std.Dev. |  | 1.4% | 3.042% |  | 1.353% | 0% | 1.922% | 0% | 0% | 2.882% | 2.673% | 1.741% | 0.9561% | 1.449% | 2.581% | 0% | 2.181% | 4.685% | 0.5773% | 1.76% | 1.76% | 0% | 0.8089%
| PDM3 |  | 0.8469% | -4.916% |  | 4.304% | 3.728% | -1.458% | 3.728% | 3.728% | -2.034% | -2.611% | 2.576% | 1.216% | -2.611% | 0.8469% | 0.2706% | 0.2706% | -6.645% | -5.741% | 1.423% | 1.423% | 0.2706% | -12.93%
| Z-Score (Absolute) |  | 0.15 | 1.35 | 25.62 | 1.05 | 0.9 | 0.45 | 0.9 | 0.9 | 0.6 | 0.75 | 0.6 | 0.25 | 0.75 | 0.15 | 0 | 0 | 1.8 | 1.56 | 0.3 | 0.3 | 0 | 3.43
| Analytical results for W in OREAS 612 (Certified Value 3.07 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 2.8 | 3.1 | 8.68 | 3.2 | < 50 | 3 | < 50 | < 50 | 3.4 | 16 | 3 | 3.1446 | 2.66 | 3.3 | 4.6 | 3.1 | 3 | 3.5 | 3.3 | 3.4 | 2.906 |  | 1
| 1 | 2 | 3.2 | 3 | 8.46 | 3.3 | < 50 | 2.9 | < 50 | < 50 | 3.3 | 18 | 3 | 3.1383 | 2.79 | 3.3 | 4.9 | 3.3 | 2.7 | 3.2 | 3.3 | 3.3 | 2.6377 |  | 30
| 1 | 3 | 3 | 3.3 | 8.22 | 3.2 | < 50 | 3 | < 50 | < 50 | 3.3 | 16 | 3.1 | 3.1402 | 2.68 | 3.1 | 4.9 | 3.1 | 2.5 | 3.2 | 3.4 | 3.2 | 2.8687 |  | 16
| 1 | 4 | 3 | 3.2 | 9.02 | 3.2 | < 50 | 3 | < 50 | < 50 | 2.9 | 13 | 3 | 3.1457 | 2.75 | 3.2 | 5 | 3 | 2.4 | 3.2 | 3.3 | 3.2 | 2.6357 |  | 3.0657
| 1 | 5 | 2.9 | 3.4 | 9.28 | 3.2 | < 50 | 3 | < 50 | < 50 | 3.2 | 18 | 3.1 | 3.1385 | 2.66 | 3.1 | 4.9 | 3 | 2.6 | 3.3 | 3.2 | 3.1 | 2.659 |  | 63
| 1 | 6 | 2.9 | 3.1 | 9.7 | 3.3 | < 50 | 2.9 | < 50 | < 50 | 3.1 | 15 | 3 | 3.1399 | 2.7 | 3.1 | 4.8 | 3.3 | 2.8 | 3.4 | 3.4 | 3.1 | 2.7228
| Mean |  | 2.9667 | 3.1833 | 8.8933 | 3.2333 |  | 2.9667 |  |  | 3.2 | 16 | 3.0333 | 3.1412 | 2.7067 | 3.1833 | 4.85 | 3.1333 | 2.6667 | 3.3 | 3.3167 | 3.2167 | 2.7383
| Median |  | 2.95 | 3.15 | 8.85 | 3.2 |  | 3 |  |  | 3.25 | 16 | 3 | 3.14 | 2.69 | 3.15 | 4.9 | 3.1 | 2.65 | 3.25 | 3.3 | 3.2 | 2.6909
| Std Dev. |  | 0.13663 | 0.1472 | 0.54826 | 0.05164 |  | 0.05164 |  |  | 0.17889 | 1.8974 | 0.05164 | 0.0031706 | 0.052789 | 0.098319 | 0.13784 | 0.13663 | 0.21602 | 0.12649 | 0.075277 | 0.1169 | 0.12026
| Rel.Std.Dev. |  | 4.605% | 4.624% | 6.165% | 1.597% |  | 1.741% |  |  | 5.59% | 11.86% | 1.702% | 0.1009% | 1.95% | 3.089% | 2.842% | 4.36% | 8.101% | 3.833% | 2.27% | 3.634% | 4.392%
| PDM3 |  | -3.232% | 3.836% | 190.1% | 5.467% |  | -3.232% |  |  | 4.379% | 421.9% | -1.057% | 2.461% | -11.71% | 3.836% | 58.2% | 2.205% | -13.02% | 7.641% | 8.185% | 4.923% | -10.68%
| Z-Score (Absolute) |  | 0.67 | 0.05 | 16.45 | 0.1 | 63 | 0.67 | 63 | 63 | 0 | 36.99 | 0.48 | 0.17 | 1.43 | 0.05 | 4.77 | 0.19 | 1.54 | 0.29 | 0.34 | 0.05 | 1.33
| Analytical results for Y in OREAS 612 (Certified Value 14.3 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 1
2
| 1 | 1 | 15.2 | 8.8 | 13.14 | 14.1 | 15.7 | 14.29 | 13.2 | 14.1 | 13.9347 | 14.08 | 15.1 | 14.4 | 14 | 13.1 | 13.7642 | 14.7 | 15.04 | 14 |  |  |  |  | 1
| 1 | 2 | 15.5 | 10.7 | 14.13 | 14.5 | 15 | 14.29 | 12.6 | 15.1 | 13.9534 | 14.47 | 15.5 | 14.6 | 13.9 | 13.1 | 13.813 | 12.6 | 15.11 | 14.1 |  |  |  |  | 31
| 1 | 3 | 15.4 | 10.6 | 13.65 | 14.3 | 15.4 | 14.42 | 12.9 | 14.8 | 14.0238 | 14.22 | 14.3 | 14.5 | 14 | 12.6 | 13.4362 | 13.9 | 15.17 | 14.1 |  |  |  |  | 16
| 1 | 4 | 15.1 | 9.5 | 13.81 | 14 | 15.6 | 14.14 | 12.7 | 15.2 | 13.9339 | 14.31 | 14.4 | 14.5 | 13.9 | 14.7 | 13.6501 | 14.1 | 15.19 | 13.9 |  |  |  |  | 14.2579
| 1 | 5 | 14.6 | 12.7 | 14.09 | 14.2 | 15.2 | 14.29 | 12.6 | 15.3 | 13.9505 | 14.35 | 15.2 | 14.3 | 14.1 | 12.5 | 13.6504 | 14.4 | 15.07 | 14.2 |  |  |  |  | 64
| 1 | 6 | 14.7 | 12.2 | 14.12 | 14.5 | 15.4 | 14.64 | 12.5 | 14.9 | 13.9827 | 14.38 | 14.4 | 15.2 | 14 | 13.4 | 13.6741 | 14.6 | 15.34 | 14.2
| Mean |  | 15.0833 | 10.75 | 13.8233 | 14.2667 | 15.3833 | 14.345 | 12.75 | 14.9 | 13.9632 | 14.3017 | 14.8167 | 14.5833 | 13.9833 | 13.2333 | 13.6647 | 14.05 | 15.1533 | 14.0833
| Median |  | 15.15 | 10.65 | 13.95 | 14.25 | 15.4 | 14.29 | 12.65 | 15 | 13.9519 | 14.33 | 14.75 | 14.5 | 14 | 13.1 | 13.6622 | 14.25 | 15.14 | 14.1
| Std Dev. |  | 0.3656 | 1.503 | 0.38738 | 0.20656 | 0.25626 | 0.16956 | 0.25884 | 0.43359 | 0.034576 | 0.13615 | 0.51153 | 0.31885 | 0.075277 | 0.79415 | 0.13007 | 0.77136 | 0.10783 | 0.1169
| Rel.Std.Dev. |  | 2.424% | 13.98% | 2.802% | 1.448% | 1.666% | 1.182% | 2.03% | 2.91% | 0.2476% | 0.952% | 3.452% | 2.186% | 0.5383% | 6.001% | 0.9519% | 5.49% | 0.7116% | 0.8301%
| PDM3 |  | 5.789% | -24.6% | -3.048% | 0.06141% | 7.893% | 0.6108% | -10.58% | 4.503% | -2.067% | 0.3069% | 3.919% | 2.282% | -1.926% | -7.186% | -4.161% | -1.458% | 6.28% | -1.224%
| Z-Score (Absolute) |  | 1.33 | 5.03 | 0.52 | 0.13 | 1.77 | 0.25 | 2.09 | 1.06 | 0.31 | 0.19 | 0.94 | 0.6 | 0.28 | 1.38 | 0.75 | 0.18 | 1.44 | 0.13
| Analytical results for Yb in OREAS 612 (Certified Value 1.02 ppm) | [23 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.18 | 1.20 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.03 | 0.6 | 1.1 | 1.02 | 1 | 1.0123 | 0.94 | 1.09 | 0.96743 | 1 | 1.1 | [11 empty] | 1
| 1 | 2 | 1.07 | 0.8 | 1 | 1.01 | 1 | 0.98918 | 0.98 | 1.05 | 0.89655 | 0.9 | 1.09 | [11 empty] | 32
| 1 | 3 | 1.07 | 0.7 | 1 | 1.03 | 1.1 | 0.97009 | 0.98 | 1.02 | 0.91549 | 0.9 | 1.07 | [11 empty] | 16
| 1 | 4 | 1.11 | 0.7 | 1 | 0.96 | 1.1 | 0.991 | 0.97 | 1.09 | 1.118 | 1 | 1.11 | [11 empty] | 1.0154
| 1 | 5 | 1.04 | 0.8 | 1.1 | 0.97 | 1.1 | 0.96933 | 0.96 | 1.01 | 0.92841 | 1 | 1.07 | [11 empty] | 65
| 1 | 6 | 1.07 | 0.8 | 1.1 | 1 | 1.1 | 0.95961 | 0.98 | 1 | 0.82663 | 0.9 | 1.27
| Mean |  | 1.065 | 0.73333 | 1.05 | 0.99833 | 1.0667 | 0.98192 | 0.96833 | 1.0433 | 0.94208 | 0.95 | 1.1183
| Median |  | 1.07 | 0.75 | 1.05 | 1.005 | 1.1 | 0.97964 | 0.975 | 1.035 | 0.92195 | 0.95 | 1.095
| Std Dev. |  | 0.028107 | 0.08165 | 0.054772 | 0.027869 | 0.05164 | 0.019255 | 0.016021 | 0.039833 | 0.097846 | 0.054772 | 0.076004
| Rel.Std.Dev. |  | 2.639% | 11.13% | 5.216% | 2.792% | 4.841% | 1.961% | 1.654% | 3.818% | 10.39% | 5.766% | 6.796%
| PDM3 |  | 4.888% | -27.78% | 3.411% | -1.678% | 5.052% | -3.294% | -4.632% | 2.754% | -7.218% | -6.438% | 10.14%
| Z-Score (Absolute) |  | 0.87 | 3.46 | 0.67 | 0 | 0.89 | 0.21 | 0.39 | 0.59 | 0.73 | 0.63 | 1.57
| Analytical results for Zn in OREAS 612 (Certified Value 0.367 wt.%) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.3664 | 0.3571 | 0.34509 | 0.352 | 0.349 | 0.361 | 0.363 | 0.369 | 0.3591 | 0.3692 | 0.367 | 0.35566 | 0.377 | 0.358 | 0.3702 | 0.357 | 0.3818 | 0.36807 | 0.38 | 0.3695 | 0.3638 | 0.48899 | 1
| 1 | 2 | 0.3684 | 0.3526 | 0.36637 | 0.3497 | 0.351 | 0.353 | 0.369 | 0.37 | 0.349 | 0.3563 | 0.387 | 0.36039 | 0.3814 | 0.371 | 0.3787 | 0.361 | 0.3818 | 0.36726 | 0.3757 | 0.3701 | 0.3636 | 0.4853 | 33
| 1 | 3 | 0.3712 | 0.3611 | 0.35453 | 0.3526 | 0.353 | 0.365 | 0.37 | 0.375 | 0.3631 | 0.3601 | 0.379 | 0.36026 | 0.379 | 0.371 | 0.3703 | 0.359 | 0.3823 | 0.36912 | 0.3766 | 0.3705 | 0.3691 | 0.48662 | 16
| 1 | 4 | 0.3704 | 0.3741 | 0.3583 | 0.3478 | 0.355 | 0.36 | 0.359 | 0.37 | 0.3531 | 0.3556 | 0.389 | 0.35689 | 0.3809 | 0.366 | 0.3718 | 0.358 | 0.3834 | 0.366 | 0.3701 | 0.3734 | 0.3646 | 0.48638 | 0.36662
| 1 | 5 | 0.3669 | 0.382 | 0.35493 | 0.3529 | 0.351 | 0.358 | 0.369 | 0.366 | 0.3517 | 0.3556 | 0.38 | 0.35599 | 0.3784 | 0.38 | 0.3713 | 0.361 | 0.3816 | 0.36865 | 0.3762 | 0.3741 | 0.3678 | 0.48044 | 66
| 1 | 6 | 0.3678 | 0.3689 | 0.36298 | 0.3607 | 0.358 | 0.364 | 0.373 | 0.372 | 0.3606 | 0.3537 | 0.388 | 0.35898 | 0.3871 | 0.385 | 0.3707 | 0.359 | 0.3849 | 0.36986 | 0.3814 | 0.3776 | 0.3698 | 0.49115
| Mean |  | 0.36852 | 0.36597 | 0.35703 | 0.35262 | 0.35283 | 0.36017 | 0.36717 | 0.37033 | 0.3561 | 0.35842 | 0.38167 | 0.35803 | 0.38063 | 0.37183 | 0.37217 | 0.35917 | 0.38263 | 0.36816 | 0.37667 | 0.37253 | 0.36645 | 0.48648
| Median |  | 0.3681 | 0.365 | 0.35661 | 0.3523 | 0.352 | 0.3605 | 0.369 | 0.37 | 0.3561 | 0.35595 | 0.3835 | 0.35793 | 0.37995 | 0.371 | 0.371 | 0.359 | 0.38205 | 0.36836 | 0.3764 | 0.37195 | 0.3662 | 0.4865
| Std Dev. |  | 0.0019167 | 0.011071 | 0.0074477 | 0.0044161 | 0.0032506 | 0.0043551 | 0.0051543 | 0.0030111 | 0.0056039 | 0.005687 | 0.0083347 | 0.0021236 | 0.003559 | 0.0096626 | 0.003258 | 0.0016021 | 0.0012879 | 0.001381 | 0.0039444 | 0.0031027 | 0.0027797 | 0.0036339
| Rel.Std.Dev. |  | 0.5201% | 3.025% | 2.086% | 1.252% | 0.9213% | 1.209% | 1.404% | 0.8131% | 1.574% | 1.587% | 2.184% | 0.5931% | 0.935% | 2.599% | 0.8754% | 0.4461% | 0.3366% | 0.3751% | 1.047% | 0.8329% | 0.7586% | 0.747%
| PDM3 |  | 0.5164% | -0.1791% | -2.616% | -3.82% | -3.761% | -1.761% | 0.1482% | 1.012% | -2.87% | -2.238% | 4.103% | -2.344% | 3.821% | 1.421% | 1.512% | -2.034% | 4.367% | 0.4192% | 2.739% | 1.612% | -0.04727% | 32.69%
| Z-Score (Absolute) |  | 0.07 | 0.13 | 0.82 | 1.16 | 1.14 | 0.58 | 0.04 | 0.21 | 0.89 | 0.71 | 1.08 | 0.74 | 1 | 0.32 | 0.35 | 0.65 | 1.15 | 0.04 | 0.69 | 0.38 | 0.09 | 9.16
| Analytical results for Zr in OREAS 612 (Certified Value 149 ppm) | [23 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 0
0
| 1 | 1 | 147.6 | 138 | 155.53 | 150 | 153.5 | 154.6 | 136 | 149 | 144.2392 | 138.8 | 153 | 150 | 144 | 137.7 | 154.2984 | 152.3 | 161 | 144 |  |  |  |  | 1
| 1 | 2 | 146.5 | 137.1 | 166.83 | 148 | 148.5 | 152.3 | 131 | 152 | 146.0057 | 144.8 | 156.5 | 151 | 145 | 139 | 155.4374 | 152.4 | 158.8 | 144 |  |  |  |  | 34
| 1 | 3 | 148.9 | 145.6 | 158.38 | 149 | 151.5 | 148.6 | 136 | 150 | 147.2299 | 141.7 | 143.5 | 150 | 143 | 139.3 | 155.4609 | 149.7 | 160.1 | 145 |  |  |  |  | 16
| 1 | 4 | 150 | 151.5 | 164.18 | 149 | 152.5 | 150.8 | 131 | 158 | 145.9349 | 143.1 | 145.5 | 150 | 144 | 136.6 | 153.4492 | 148.3 | 155.8 | 143 |  |  |  |  | 149.4674
| 1 | 5 | 147 | 152.5 | 164.47 | 151 | 149 | 158.6 | 130 | 152 | 144.5426 | 143.2 | 151.5 | 150 | 143 | 139.7 | 154.8673 | 152.6 | 156 | 144 |  |  |  |  | 67
| 1 | 6 | 149.7 | 150 | 161.64 | 152 | 150.5 | 147.6 | 129 | 161 | 146.511 | 142.6 | 145.5 | 150 | 145 | 138.3 | 153.8638 | 151.6 | 160.5 | 145
| Mean |  | 148.2833 | 145.7833 | 161.8383 | 149.8333 | 150.9167 | 152.0833 | 132.1667 | 153.6667 | 145.7439 | 142.3667 | 149.25 | 150.1667 | 144 | 138.4333 | 154.5628 | 151.15 | 158.7 | 144.1667
| Median |  | 148.25 | 147.8 | 162.91 | 149.5 | 151 | 151.55 | 131 | 152 | 145.9703 | 142.85 | 148.5 | 150 | 144 | 138.65 | 154.5828 | 151.95 | 159.45 | 144
| Std Dev. |  | 1.458 | 6.8057 | 4.2219 | 1.472 | 1.96 | 4.0676 | 3.0605 | 4.761 | 1.1497 | 2.0186 | 5.1551 | 0.40825 | 0.89443 | 1.1483 | 0.83223 | 1.7558 | 2.2891 | 0.75277
| Rel.Std.Dev. |  | 0.9832% | 4.668% | 2.609% | 0.9824% | 1.299% | 2.675% | 2.316% | 3.098% | 0.7889% | 1.418% | 3.454% | 0.2719% | 0.6211% | 0.8295% | 0.5384% | 1.162% | 1.442% | 0.5222%
| PDM3 |  | -0.7922% | -2.465% | 8.277% | 0.2449% | 0.9697% | 1.75% | -11.57% | 2.81% | -2.491% | -4.751% | -0.1454% | 0.4679% | -3.658% | -7.382% | 3.409% | 1.126% | 6.177% | -3.546%
| Z-Score (Absolute) |  | 0.21 | 0.64 | 2.09 | 0.05 | 0.23 | 0.43 | 2.96 | 0.7 | 0.65 | 1.22 | 0.05 | 0.11 | 0.94 | 1.89 | 0.85 | 0.27 | 1.56 | 0.91
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 612 (Certified Value 47.2 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*AAS | AR*OES | AR*OES/MS | AR*OES | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 2
| 1 | 1 | 48.24 | > 10 | 48.55 | 47 | 50 | 45 | 48 | 46.79 | 45 | 46.3 | 49.158 | 47.2 | 52 | 47.5 | 47.6 | 48.2 | 43.3336 | 50 | 46.2 | 47.75 | 48.7293 | 1
| 1 | 2 | 48.56 | > 10 | 49.56 | 47 | 49.7 | 46 | 48 | 45.92 | 44 | 44.7 | 48.8527 | 46.3 | 51.6 | 47.5 | 47.9 | 48.7 | 43.5377 | 48 | 47.3 | 47.4 | 44.6363 | 1
| 1 | 3 | 48.57 | > 10 | 49.21 | 48 | 50 | 47 | 50 | 45.66 | 45 | 46.8 | 48.4629 | 46.8 | 53.7 | 47.2 | 47.6 | 48.4 | 43.6055 | 49 | 46.1 | 47.7 | 44.7195 | 16
| 1 | 4 | 48.23 | > 10 | 49.43 | 48 | 49.9 | 47 | 47 | 47.81 | 46 | 45.5 | 48.3251 | 46.9 | 52.4 | 47 | 47.4 | 49 | 43.4585 | 49 | 44.8 | 48.7 | 47.261 | 47.2395
| 1 | 5 | 47.99 | > 10 | 47.78 | 47 | 49.3 | 47 | 48 | 47.09 | 44 | 43.9 | 48.1095 | 46.3 | 51.3 | 47.2 | 47.6 | 48.9 | 43.4657 | 48 | 46.3 | 46.7 | 48.9693 | 69
| 1 | 6 | 48.45 | > 10 | 48.91 | 48 | 49.8 | 45 | 48 | 45.45 | 44 | 46.5 | 48.1796 | 47.1 | 53.5 | 47.4 | 48.4 | 49.1 | 43.7105 | 48 | 46 | 47.3 | 47.7433
| Mean |  | 48.34 |  | 48.9067 | 47.5 | 49.7833 | 46.1667 | 48.1667 | 46.4533 | 44.6667 | 45.6167 | 48.5146 | 46.7667 | 52.4167 | 47.3 | 47.75 | 48.7167 | 43.5186 | 48.6667 | 46.1167 | 47.5917 | 47.0098
| Median |  | 48.345 |  | 49.06 | 47.5 | 49.85 | 46.5 | 48 | 46.355 | 44.5 | 45.9 | 48.394 | 46.85 | 52.2 | 47.3 | 47.6 | 48.8 | 43.5017 | 48.5 | 46.15 | 47.55 | 47.5022
| Std Dev. |  | 0.22716 |  | 0.66171 | 0.54772 | 0.26394 | 0.98319 | 0.98319 | 0.92517 | 0.8165 | 1.1356 | 0.41069 | 0.38816 | 0.99079 | 0.2 | 0.35637 | 0.35449 | 0.13067 | 0.8165 | 0.79854 | 0.66062 | 1.9119
| Rel.Std.Dev. |  | 0.4699% |  | 1.353% | 1.153% | 0.5302% | 2.13% | 2.041% | 1.992% | 1.828% | 2.49% | 0.8465% | 0.83% | 1.89% | 0.4228% | 0.7463% | 0.7277% | 0.3003% | 1.678% | 1.732% | 1.388% | 4.067%
| PDM3 |  | 2.33% |  | 3.529% | 0.5513% | 5.385% | -2.271% | 1.963% | -1.664% | -5.446% | -3.435% | 2.699% | -1.001% | 10.96% | 0.128% | 1.081% | 3.127% | -7.877% | 3.021% | -2.377% | 0.7454% | -0.4863%
| Z-Score (Absolute) |  | 0.48 | NA | 0.83 | 0.03 | 1.36 | 0.84 | 0.38 | 0.67 | 1.75 | 1.18 | 0.59 | 0.47 | 2.97 | 0.15 | 0.12 | 0.71 | 2.45 | 0.68 | 0.87 | 0.03 | 0.33
| Analytical results for Al in OREAS 612 (Certified Value 0.874 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 0.85 | 0.8 | 0.865 | 1.18 | 0.89 | 1 | 1.14 | 0.8666 | 0.822 | 1.01 | 0.91446 | 0.89 | 0.92 | 1.05 | 0.87 | 0.93466 | 0.91 | 1.1561 | 0.82 | 0.84639 |  | 1
| 1 | 2 | 0.89 | 0.79 | 0.873 | 1.17 | 0.88 | 1.03 | 1.15 | 0.877 | 0.816 | 0.94 | 0.90895 | 0.89 | 0.92 | 1.02 | 0.88 | 0.93437 | 0.86 | 1.2454 | 0.81 | 0.82318 |  | #N/A
| 1 | 3 | 0.9 | 0.8 | 0.855 | 1.19 | 0.88 | 1.06 | 1.17 | 0.8587 | 0.847 | 0.88 | 0.90119 | 0.89 | 0.89 | 1.03 | 0.88 | 0.93433 | 0.85 | 1.1376 | 0.82 | 0.81891 |  | 16
| 1 | 4 | 0.87 | 0.81 | 0.853 | 1.17 | 0.9 | 1.08 | 1.16 | 0.8886 | 0.886 | 0.87 | 0.89764 | 0.89 | 0.91 | 1.08 | 0.89 | 0.93525 | 0.93 | 1.1213 | 0.84 | 0.81916 |  | 0.87352
| 1 | 5 | 0.89 | 0.81 | 0.85 | 1.17 | 0.88 | 1.06 | 1.16 | 0.8574 | 0.826 | 0.88 | 0.89171 | 0.87 | 0.9 | 1.04 | 0.9 | 0.93351 | 0.95 | 1.1491 | 0.79 | 0.84405 |  | 70
| 1 | 6 | 0.91 | 0.81 | 0.884 | 1.18 | 0.89 | 1.03 | 1.16 | 0.847 | 0.828 | 0.91 | 0.9145 | 0.88 | 0.9 | 1.06 | 0.9 | 0.93506 | 0.94 | 1.1513 | 0.81 | 0.80302
| Mean |  | 0.885 | 0.80333 | 0.86333 | 1.1767 | 0.88667 | 1.0433 | 1.1567 | 0.86588 | 0.8375 | 0.915 | 0.90474 | 0.885 | 0.90667 | 1.0467 | 0.88667 | 0.93453 | 0.90667 | 1.1601 | 0.815 | 0.82579
| Median |  | 0.89 | 0.805 | 0.86 | 1.175 | 0.885 | 1.045 | 1.16 | 0.86265 | 0.827 | 0.895 | 0.90507 | 0.89 | 0.905 | 1.045 | 0.885 | 0.93452 | 0.92 | 1.1502 | 0.815 | 0.82117
| Std Dev. |  | 0.021679 | 0.008165 | 0.013246 | 0.008165 | 0.008165 | 0.028752 | 0.010328 | 0.014969 | 0.025952 | 0.053198 | 0.0093847 | 0.0083666 | 0.012111 | 0.021602 | 0.012111 | 0.00062194 | 0.042269 | 0.043594 | 0.016432 | 0.016581
| Rel.Std.Dev. |  | 2.45% | 1.016% | 1.534% | 0.6939% | 0.9209% | 2.756% | 0.8929% | 1.729% | 3.099% | 5.814% | 1.037% | 0.9454% | 1.336% | 2.064% | 1.366% | 0.06655% | 4.662% | 3.758% | 2.016% | 2.008%
| PDM3 |  | 1.314% | -8.035% | -1.166% | 34.7% | 1.505% | 19.44% | 32.41% | -0.874% | -4.123% | 4.749% | 3.574% | 1.314% | 3.795% | 19.82% | 1.505% | 6.985% | 3.795% | 32.81% | -6.699% | -5.464%
| Z-Score (Absolute) |  | 0.2 | 1.75 | 0.61 | 5.32 | 0.17 | 2.8 | 4.94 | 0.56 | 1.1 | 0.37 | 0.17 | 0.2 | 0.21 | 2.86 | 0.17 | 0.74 | 0.21 | 5.01 | 1.53 | 1.32
| Analytical results for As in OREAS 612 (Certified Value 322 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | 0
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 0
| 1 | 1 | 319.8 | 311 | 315 | 330 | 341 | 310 | 330 | 331 | 312 | 327 | 323.7557 | 313.4 | 332 | 323 | 298 | 286.5 | 328.9882 | 319 | 317 | 325 | 349.188 | 1
| 1 | 2 | 320.6 | 312 | 328 | 330 | 346 | 320 | 340 | 327 | 314 | 320 | 322.7228 | 308.1 | 328 | 324 | 305 | 290.8 | 322.2173 | 315 | 325 | 321 | 348.3283 | 20
| 1 | 3 | 324.8 | 315 | 302 | 330 | 338 | 330 | 340 | 326 | 312 | 328 | 322.585 | 312.1 | 338 | 320 | 301 | 290.5 | 322.8746 | 317 | 316 | 323 | 349.5457 | 16
| 1 | 4 | 321.4 | 317 | 310 | 330 | 348 | 320 | 330 | 338 | 309 | 326 | 322.9934 | 311.9 | 336 | 321 | 298 | 294.3 | 323.5597 | 314 | 308 | 328 | 349.3243 | 321.6841
| 1 | 5 | 318.8 | 322 | 321 | 330 | 340 | 320 | 330 | 338 | 300 | 309 | 322.6566 | 307 | 329 | 323 | 301 | 293.2 | 321.5584 | 316 | 318 | 314 | 348.5847 | 71
| 1 | 6 | 329.3 | 319 | 326 | 320 | 337 | 310 | 330 | 324 | 322 | 331 | 321.7858 | 311.7 | 331 | 324 | 307 | 294.3 | 323.3926 | 320 | 317 | 320 | 350.287
| Mean |  | 322.45 | 316 | 317 | 328.3333 | 341.6667 | 318.3333 | 333.3333 | 330.6667 | 311.5 | 323.5 | 322.7499 | 310.7 | 332.3333 | 322.5 | 301.6667 | 291.6 | 323.7651 | 316.8333 | 316.8333 | 321.8333 | 349.2097
| Median |  | 321 | 316 | 318 | 330 | 340.5 | 320 | 330 | 329 | 312 | 326.5 | 322.6897 | 311.8 | 331.5 | 323 | 301 | 292 | 323.1336 | 316.5 | 317 | 322 | 349.2562
| Std Dev. |  | 3.9323 | 4.1952 | 9.9599 | 4.0825 | 4.4121 | 7.5277 | 5.164 | 6.121 | 7.1484 | 7.9687 | 0.6383 | 2.5354 | 3.9328 | 1.6432 | 3.6697 | 3 | 2.6653 | 2.3166 | 5.4191 | 4.7924 | 0.70065
| Rel.Std.Dev. |  | 1.22% | 1.328% | 3.142% | 1.243% | 1.291% | 2.365% | 1.549% | 1.851% | 2.295% | 2.463% | 0.1978% | 0.816% | 1.183% | 0.5095% | 1.216% | 1.029% | 0.8232% | 0.7312% | 1.71% | 1.489% | 0.2006%
| PDM3 |  | 0.2381% | -1.767% | -1.456% | 2.067% | 6.212% | -1.042% | 3.621% | 2.792% | -3.166% | 0.5645% | 0.3313% | -3.415% | 3.31% | 0.2536% | -6.223% | -9.352% | 0.6469% | -1.508% | -1.508% | 0.04638% | 8.557%
| Z-Score (Absolute) |  | 0 | 0.74 | 0.62 | 0.67 | 2.2 | 0.47 | 1.25 | 0.94 | 1.26 | 0.12 | 0.03 | 1.35 | 1.13 | 0.01 | 2.38 | 3.54 | 0.15 | 0.64 | 0.64 | 0.07 | 3.07
| Analytical results for B in OREAS 612 (Indicative Value < 10 ppm) | [22 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.20 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [13 empty] | 1
|  |  | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | [13 empty] | 1
| 1 | 1 | < 10 | < 10 | < 10 | 4 | < 10 | < 10 | < 10 | < 10 | [13 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | 4 | < 10 | < 10 | < 10 | < 10 | [13 empty] | 3
| 1 | 3 | < 10 | < 10 | < 10 | 4 | < 10 | < 10 | < 10 | < 10 | [13 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | 4 | < 10 | < 10 | < 10 | < 10 | [13 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 10 | 4 | < 10 | < 10 | < 10 | < 10 | [13 empty] | 12
| 1 | 6 | < 10 | < 10 | < 10 | 4 | < 10 | < 10 | < 10 | < 10
| Mean |  |  |  |  | 4
| Median |  |  |  |  | 4
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 612 (Certified Value 324 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS |  |  |  | 0
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g |  |  |  | 0
| 1 | 1 | 318 | 299 | 320 | 380 | 320 | 330 | 390 | 349 | 300 | 59.1 | 326.0655 | 330 | 239 | 322 | 303 | 306 | 379 | 265 |  |  |  | 1
| 1 | 2 | 329 | 297 | 322 | 380 | 320 | 340 | 390 | 340 | 297 | 62.4 | 321.358 | 320 | 228 | 330 | 311 | 307 | 383 | 263 |  |  |  | 21
| 1 | 3 | 335 | 300 | 314 | 380 | 310 | 350 | 400 | 325 | 307 | 54.8 | 321.96 | 330 | 219 | 332 | 307 | 311 | 372 | 269 |  |  |  | 16
| 1 | 4 | 325 | 301 | 315 | 380 | 320 | 350 | 390 | 342 | 315 | 63.5 | 318.2881 | 320 | 230 | 347 | 301 | 303 | 366 | 269 |  |  |  | 323.9871
| 1 | 5 | 327 | 304 | 312 | 380 | 320 | 350 | 380 | 342 | 298 | 56.1 | 318.4112 | 320 | 227 | 344 | 303 | 311 | 371 | 261 |  |  |  | 72
| 1 | 6 | 336 | 306 | 324 | 380 | 310 | 340 | 390 | 326 | 301 | 72.9 | 318.7543 | 320 | 224 | 340 | 305 | 313 | 375 | 259
| Mean |  | 328.3333 | 301.1667 | 317.8333 | 380 | 316.6667 | 343.3333 | 390 | 337.3333 | 303 | 61.4667 | 320.8062 | 323.3333 | 227.8333 | 335.8333 | 305 | 308.5 | 374.3333 | 264.3333
| Median |  | 328 | 300.5 | 317.5 | 380 | 320 | 345 | 390 | 341 | 300.5 | 60.75 | 320.0561 | 320 | 227.5 | 336 | 304 | 309 | 373.5 | 264
| Std Dev. |  | 6.6833 | 3.3116 | 4.8339 | 0 | 5.164 | 8.165 | 6.3246 | 9.6678 | 6.8411 | 6.5509 | 3.0194 | 5.164 | 6.6758 | 9.4745 | 3.5777 | 3.7815 | 6.0553 | 4.1312
| Rel.Std.Dev. |  | 2.036% | 1.1% | 1.521% | 0% | 1.631% | 2.378% | 1.622% | 2.866% | 2.258% | 10.66% | 0.9412% | 1.597% | 2.93% | 2.821% | 1.173% | 1.226% | 1.618% | 1.563%
| PDM3 |  | 1.341% | -7.044% | -1.899% | 17.29% | -2.259% | 5.971% | 20.38% | 4.119% | -6.478% | -81.03% | -0.9818% | -0.2018% | -29.68% | 3.656% | -5.86% | -4.78% | 15.54% | -18.41%
| Z-Score (Absolute) |  | 0.35 | 0.71 | 0.06 | 2.37 | 0.1 | 0.94 | 2.76 | 0.7 | 0.64 | 10.07 | 0.06 | 0.16 | 3.57 | 0.65 | 0.56 | 0.42 | 2.15 | 2.15
| Analytical results for Be in OREAS 612 (Certified Value 0.63 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 0.63 | 0.7 | 0.62 | < 5 | 0.64 | < 5 | < 5 | 0.61 | 0.6 | 0.6 | 0.67939 | 0.63 | 0.58 | 0.78 | 1 | 0.613 | 0.7 | 0.82 | 0.65 | 0.644 |  | 1
| 1 | 2 | 0.64 | 0.6 | 0.77 | < 5 | 0.65 | < 5 | < 5 | 0.64 | 0.6 | 0.6 | 0.67028 | 0.62 | 0.59 | 0.78 | 1 | 0.68453 | 0.7 | 0.85 | 0.64 | 0.64433 |  | 22
| 1 | 3 | 0.59 | 0.6 | 0.67 | < 5 | 0.63 | < 5 | < 5 | 0.62 | 0.6 | 0.7 | 0.66555 | 0.64 | 0.59 | 0.75 | 1 | 0.63103 | 0.7 | 0.84 | 0.65 | 0.64333 |  | 16
| 1 | 4 | 0.64 | 0.6 | 0.66 | < 5 | 0.65 | < 5 | < 5 | 0.66 | 0.6 | 0.7 | 0.65948 | 0.64 | 0.6 | 0.76 | < 1 | 0.61902 | 0.7 | 0.81 | 0.64 | 0.65033 |  | 0.63379
| 1 | 5 | 0.62 | 0.6 | 0.66 | < 5 | 0.64 | < 5 | < 5 | 0.66 | 0.6 | 0.5 | 0.66913 | 0.62 | 0.6 | 0.78 | 1 | 0.63027 | 0.8 | 0.81 | 0.62 | 0.65367 |  | 73
| 1 | 6 | 0.59 | 0.6 | 0.65 | < 5 | 0.64 | < 5 | < 5 | 0.63 | 0.6 | 0.8 | 0.65815 | 0.62 | 0.58 | 0.75 | 1 | 0.59668 | 0.8 | 0.85 | 0.63 | 0.64967
| Mean |  | 0.61833 | 0.61667 | 0.67167 |  | 0.64167 |  |  | 0.63667 | 0.6 | 0.65 | 0.667 | 0.62833 | 0.59 | 0.76667 | 1 | 0.62909 | 0.73333 | 0.83 | 0.63833 | 0.64756
| Median |  | 0.625 | 0.6 | 0.66 |  | 0.64 |  |  | 0.635 | 0.6 | 0.65 | 0.66734 | 0.625 | 0.59 | 0.77 | 1 | 0.62464 | 0.7 | 0.83 | 0.64 | 0.647
| Std Dev. |  | 0.023166 | 0.040825 | 0.051153 |  | 0.0075277 |  |  | 0.020656 | 0 | 0.10488 | 0.0078193 | 0.0098319 | 0.0089443 | 0.015055 | 0 | 0.029967 | 0.05164 | 0.018974 | 0.01169 | 0.0042514
| Rel.Std.Dev. |  | 3.747% | 6.62% | 7.616% |  | 1.173% |  |  | 3.244% | 0% | 16.14% | 1.172% | 1.565% | 1.516% | 1.964% | 0% | 4.764% | 7.042% | 2.286% | 1.831% | 0.6565%
| PDM3 |  | -2.438% | -2.701% | 5.977% |  | 1.243% |  |  | 0.4542% | -5.331% | 2.558% | 5.24% | -0.8607% | -6.909% | 20.97% | 57.78% | -0.7418% | 15.71% | 30.96% | 0.7171% | 2.172%
| Z-Score (Absolute) |  | 0.82 | NA | 0.67 | 51.81 | 0.16 | 51.81 | 51.81 | 0.3 | NA | NA | 0.54 | 0.54 | 1.61 | 3.33 | NA | 0.52 | NA | 5.1 | 0.26 | 0
Results from laboratories 1.02, 1.10, 1.11 and 1.20 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 612 (Certified Value 0.85 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 0.8 | 0.98 | 0.82 | < 10 | 0.8 | < 10 | < 10 | 0.89 | < 5 | 0.93 | 0.81737 | 0.81 | 0.84 | < 2 | 0.91 | 0.7 | 0.82 | 0.86 | 0.91 | < 1 |  | 1
| 1 | 2 | 0.77 | 0.85 | 0.91 | < 10 | 0.83 | < 10 | < 10 | 0.88 | < 5 | 0.85 | 0.79403 | 0.77 | 1 | < 2 | 0.88 | 0.8 | 0.89 | 0.88 | 0.89 | < 1 |  | 23
| 1 | 3 | 0.76 | 0.83 | 0.77 | < 10 | 0.81 | < 10 | < 10 | 0.84 | < 5 | 0.87 | 0.81209 | 0.8 | 0.92 | < 2 | 0.89 | 0.8 | 0.84 | 0.87 | 0.88 | < 1 |  | 16
| 1 | 4 | 0.79 | 0.84 | 0.83 | 10 | 0.82 | < 10 | < 10 | 0.92 | < 5 | 0.84 | 0.82259 | 0.77 | 0.9 | < 2 | 0.87 | 0.7 | 0.84 | 0.86 | 0.93 | < 1 |  | 0.84803
| 1 | 5 | 0.83 | 0.85 | 0.82 | < 10 | 0.8 | < 10 | < 10 | 0.89 | < 5 | 0.82 | 0.78986 | 0.94 | 0.88 | < 2 | 0.88 | 0.7 | 0.86 | 0.88 | 0.79 | < 1 |  | 74
| 1 | 6 | 0.83 | 0.83 | 0.87 | < 10 | 0.84 | < 10 | < 10 | 0.85 | < 5 | 0.88 | 0.81815 | 0.81 | 0.96 | < 2 | 0.88 | 0.7 | 0.86 | 0.89 | 0.88 | < 1
| Mean |  | 0.79667 | 0.86333 | 0.83667 | 10 | 0.81667 |  |  | 0.87833 |  | 0.865 | 0.80902 | 0.81667 | 0.91667 |  | 0.885 | 0.73333 | 0.85167 | 0.87333 | 0.88
| Median |  | 0.795 | 0.845 | 0.825 | 10 | 0.815 |  |  | 0.885 |  | 0.86 | 0.81473 | 0.805 | 0.91 |  | 0.88 | 0.7 | 0.85 | 0.875 | 0.885
| Std Dev. |  | 0.029439 | 0.05785 | 0.048028 |  | 0.01633 |  |  | 0.029269 |  | 0.038341 | 0.013701 | 0.06314 | 0.057155 |  | 0.013784 | 0.05164 | 0.024014 | 0.012111 | 0.048166
| Rel.Std.Dev. |  | 3.695% | 6.701% | 5.74% |  | 2% |  |  | 3.332% |  | 4.432% | 1.694% | 7.731% | 6.235% |  | 1.558% | 7.042% | 2.82% | 1.387% | 5.473%
| PDM3 |  | -6.056% | 1.805% | -1.34% | 1079% | -3.698% |  |  | 3.574% |  | 2.001% | -4.6% | -3.698% | 8.094% |  | 4.36% | -13.52% | 0.4292% | 2.984% | 3.77%
| Z-Score (Absolute) |  | 0.98 | 0.08 | 0.44 | NA | 0.71 | 55.71 | 55.71 | 0.12 | 21.99 | 0.06 | 0.81 | 0.71 | 0.64 | 1.76 | 0.21 | NA | 0.24 | 0.06 | 0.15 | 4.98
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.05 were removed due to their 10 ppm reading resolution.
| Analytical results for Ca in OREAS 612 (Certified Value 0.501 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 3
| 1 | 1 | 0.49 | 0.49 | 0.509 | 0.5 | 0.5 | 0.46 | 0.49 | 0.51 | 0.493 | 0.52 | 0.51399 | 0.4887 | 0.5 | 0.51 | 0.51 | 0.47 | 0.49971 | 0.52 | 0.51 | 0.52 | 0.59887 | 1
| 1 | 2 | 0.5 | 0.49 | 0.512 | 0.5 | 0.51 | 0.47 | 0.5 | 0.51 | 0.478 | 0.47 | 0.50681 | 0.4808 | 0.5 | 0.51 | 0.51 | 0.46 | 0.49534 | 0.5 | 0.54 | 0.51 | 0.57071 | #N/A
| 1 | 3 | 0.51 | 0.49 | 0.501 | 0.5 | 0.49 | 0.48 | 0.51 | 0.5 | 0.495 | 0.49 | 0.50176 | 0.4844 | 0.51 | 0.49 | 0.51 | 0.46 | 0.49462 | 0.54 | 0.5 | 0.52 | 0.57848 | 16
| 1 | 4 | 0.5 | 0.49 | 0.496 | 0.51 | 0.5 | 0.48 | 0.5 | 0.52 | 0.5 | 0.51 | 0.49468 | 0.4856 | 0.5 | 0.5 | 0.52 | 0.47 | 0.49917 | 0.53 | 0.49 | 0.53 | 0.59083 | 0.50119
| 1 | 5 | 0.5 | 0.5 | 0.496 | 0.51 | 0.5 | 0.48 | 0.5 | 0.52 | 0.479 | 0.47 | 0.50539 | 0.4812 | 0.5 | 0.5 | 0.52 | 0.47 | 0.49274 | 0.54 | 0.51 | 0.5 | 0.544 | 75
| 1 | 6 | 0.51 | 0.5 | 0.512 | 0.51 | 0.5 | 0.47 | 0.5 | 0.5 | 0.486 | 0.49 | 0.50273 | 0.4858 | 0.5 | 0.5 | 0.51 | 0.48 | 0.49552 | 0.54 | 0.51 | 0.51 | 0.53744
| Mean |  | 0.50167 | 0.49333 | 0.50433 | 0.505 | 0.5 | 0.47333 | 0.5 | 0.51 | 0.4885 | 0.49167 | 0.50423 | 0.48442 | 0.50167 | 0.50167 | 0.51333 | 0.46833 | 0.49618 | 0.52833 | 0.51 | 0.515 | 0.57006
| Median |  | 0.5 | 0.49 | 0.505 | 0.505 | 0.5 | 0.475 | 0.5 | 0.51 | 0.4895 | 0.49 | 0.50406 | 0.485 | 0.5 | 0.5 | 0.51 | 0.47 | 0.49543 | 0.535 | 0.51 | 0.515 | 0.57459
| Std Dev. |  | 0.0075277 | 0.005164 | 0.007607 | 0.0054772 | 0.0063246 | 0.008165 | 0.0063246 | 0.0089443 | 0.008961 | 0.020412 | 0.0063696 | 0.0030029 | 0.0040825 | 0.0075277 | 0.005164 | 0.0075277 | 0.0027147 | 0.016021 | 0.016733 | 0.010488 | 0.024802
| Rel.Std.Dev. |  | 1.501% | 1.047% | 1.508% | 1.085% | 1.265% | 1.725% | 1.265% | 1.754% | 1.834% | 4.152% | 1.263% | 0.6199% | 0.8138% | 1.501% | 1.006% | 1.607% | 0.5471% | 3.032% | 3.281% | 2.037% | 4.351%
| PDM3 |  | 0.09457% | -1.568% | 0.6266% | 0.7596% | -0.238% | -5.559% | -0.238% | 1.757% | -2.533% | -1.901% | 0.6053% | -3.347% | 0.09457% | 0.09457% | 2.422% | -6.556% | -0.9992% | 5.415% | 1.757% | 2.755% | 13.74%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.22 | 0.27 | 0.13 | 2.29 | 0.13 | 0.67 | 1.07 | 0.81 | 0.21 | 1.4 | 0 | 0 | 0.94 | 2.7 | 0.44 | 2.16 | 0.67 | 1.08 | 5.53
| Analytical results for Cd in OREAS 612 (Certified Value 13.9 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 2
| 1 | 1 | 13.84 | 13.55 | 14.33 | 14 | 14.1 | 13 | 14 | 14.61 | 14 | 14.9 | 13.8447 | 14.35 | 13.48 | 13.6 | 12.7 | 13.91 | 13.98 | 13.78 | 15.2607 |  |  | 1
| 1 | 2 | 14.11 | 13.75 | 14.4 | 15 | 14.25 | 14 | 14 | 14.14 | 14 | 14.4 | 13.7945 | 14.65 | 13.39 | 13.6 | 12.8 | 14.01 | 14.09 | 13.39 | 15.3483 |  |  | 24
| 1 | 3 | 14.18 | 13.7 | 14.22 | 14 | 14.25 | 14 | 15 | 13.97 | 14 | 14.3 | 13.762 | 15.6 | 13.26 | 13.5 | 12.7 | 14.12 | 13.87 | 14.1 | 15.3163 |  |  | 16
| 1 | 4 | 14.03 | 13.59 | 14.24 | 15 | 14.1 | 14 | 14 | 14.02 | 14 | 13.4 | 13.483 | 15.2 | 13.3 | 14 | 12.8 | 13.98 | 13.58 | 14.03 | 15.423 |  |  | 13.9156
| 1 | 5 | 13.89 | 14.19 | 14.15 | 15 | 13.7 | 14 | 14 | 13.9 | 14 | 14 | 13.5769 | 14.65 | 13.43 | 14 | 12.7 | 14.21 | 13.96 | 13.38 | 15.1407 |  |  | 76
| 1 | 6 | 14.14 | 13.89 | 14.36 | 14 | 14.3 | 13 | 15 | 13.63 | 14 | 14.1 | 13.7527 | 15.65 | 13.03 | 13.8 | 12.8 | 14.54 | 13.72 | 14.04 | 15.243
| Mean |  | 14.0317 | 13.7783 | 14.2833 | 14.5 | 14.1167 | 13.6667 | 14.3333 | 14.045 | 14 | 14.1833 | 13.7023 | 15.0167 | 13.315 | 13.75 | 12.75 | 14.1283 | 13.8667 | 13.7867 | 15.2887
| Median |  | 14.07 | 13.725 | 14.285 | 14.5 | 14.175 | 14 | 14 | 13.995 | 14 | 14.2 | 13.7573 | 14.925 | 13.345 | 13.7 | 12.75 | 14.065 | 13.915 | 13.905 | 15.2885
| Std Dev. |  | 0.13906 | 0.23515 | 0.095219 | 0.54772 | 0.22061 | 0.5164 | 0.5164 | 0.32488 | 0 | 0.49565 | 0.14049 | 0.54559 | 0.16159 | 0.21679 | 0.054772 | 0.22798 | 0.18715 | 0.32995 | 0.097102
| Rel.Std.Dev. |  | 0.991% | 1.707% | 0.6666% | 3.777% | 1.563% | 3.779% | 3.603% | 2.313% | 0% | 3.495% | 1.025% | 3.633% | 1.214% | 1.577% | 0.4296% | 1.614% | 1.35% | 2.393% | 0.6351%
| PDM3 |  | 0.834% | -0.9865% | 2.643% | 4.2% | 1.445% | -1.789% | 3.002% | 0.9298% | 0.6065% | 1.924% | -1.533% | 7.912% | -4.316% | -1.19% | -8.376% | 1.529% | -0.3517% | -0.9266% | 9.867%
| Z-Score (Absolute) |  | 0 | 0.68 | 0.67 | NA | 0.23 | NA | NA | 0.04 | NA | 0.41 | 0.88 | 2.64 | 1.92 | 0.75 | 3.43 | 0.26 | 0.44 | 0.66 | 3.37
Results from laboratories 1.05, 1.07, 1.08 and 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 612 (Certified Value 48.6 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | [7 empty] | 1
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | [7 empty] | 2
| 1 | 1 | 48.76 | 45.21 | 45.79 | 47.8 | 54.992 | 52.7 | 48.2855 | 49 | 46.6 | 52.3 | 45 | 46.41 | 58 | 47.4 | [7 empty] | 1
| 1 | 2 | 49.39 | 43.72 | 48.27 | 46.4 | 54.459 | 50.8 | 47.8888 | 49.1 | 47.1 | 54 | 46 | 45.99 | 65 | 47.3 | [7 empty] | 25
| 1 | 3 | 48.87 | 43.59 | 45.42 | 50 | 52.793 | 52.9 | 47.8612 | 51.5 | 46.5 | 53.3 | 45 | 46.15 | 59 | 47.6 | [7 empty] | 16
| 1 | 4 | 48.1 | 44.58 | 46.19 | 47.4 | 55.463 | 52 | 48.3824 | 49.9 | 46.3 | 55.5 | 46 | 46.27 | 57 | 48.4 | [7 empty] | 48.6151
| 1 | 5 | 48.65 | 44.78 | 47.6 | 47.4 | 53.316 | 49 | 48.2231 | 47.6 | 46.4 | 55.2 | 45 | 46.64 | 61 | 46.5 | [7 empty] | 77
| 1 | 6 | 50.87 | 45.67 | 48.11 | 48.9 | 52.27 | 53.4 | 47.2809 | 51.7 | 46.1 | 54.7 | 47 | 47.33 | 61 | 45.7
| Mean |  | 49.1067 | 44.5917 | 46.8967 | 47.9833 | 53.8822 | 51.8 | 47.987 | 49.8 | 46.5 | 54.1667 | 45.6667 | 46.465 | 60.1667 | 47.15
| Median |  | 48.815 | 44.68 | 46.895 | 47.6 | 53.8875 | 52.35 | 48.0559 | 49.5 | 46.45 | 54.35 | 45.5 | 46.34 | 60 | 47.35
| Std Dev. |  | 0.95789 | 0.81779 | 1.2456 | 1.275 | 1.2782 | 1.6407 | 0.40626 | 1.5799 | 0.34059 | 1.216 | 0.8165 | 0.47849 | 2.8577 | 0.93541
| Rel.Std.Dev. |  | 1.951% | 1.834% | 2.656% | 2.657% | 2.372% | 3.167% | 0.8466% | 3.172% | 0.7324% | 2.245% | 1.788% | 1.03% | 4.75% | 1.984%
| PDM3 |  | 1.011% | -8.276% | -3.535% | -1.299% | 10.83% | 6.551% | -1.292% | 2.437% | -4.351% | 11.42% | -6.065% | -4.423% | 23.76% | -3.014%
| Z-Score (Absolute) |  | 0.45 | 1.37 | 0.44 | 0 | 2.38 | 1.54 | 0 | 0.73 | 0.6 | 2.5 | 0.94 | 0.61 | 4.93 | 0.34
| Analytical results for Co in OREAS 612 (Certified Value 3.69 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 3.7 | 3.7 | 3.7 | 5 | 3.8 | < 5 | < 5 | 3.7 | 3 | 4.1 | 3.5524 | 3.8 | 3.3 | 3.7 | 3.5 | 3.3215 | 3.8 | 3.7 | 3.63 | 3.8413 |  | 1
| 1 | 2 | 3.6 | 3.6 | 3.9 | < 5 | 3.8 | < 5 | 6 | 3.7 | 3 | 3.9 | 3.4935 | 4 | 3.4 | 3.7 | 3.5 | 3.2882 | 3.7 | 3.8 | 3.59 | 3.9877 |  | 26
| 1 | 3 | 3.6 | 3.7 | 3.6 | 5 | 3.8 | < 5 | 5 | 3.7 | 3 | 4.1 | 3.5136 | 4.1 | 3.4 | 3.7 | 3.5 | 3.3232 | 3.8 | 3.7 | 3.73 | 3.8263 |  | 16
| 1 | 4 | 3.6 | 3.7 | 3.7 | < 5 | 3.6 | < 5 | 5 | 3.7 | 3 | 4.1 | 3.6382 | 4.1 | 3.4 | 3.6 | 3.6 | 3.4144 | 3.7 | 3.7 | 3.7 | 3.906 |  | 3.6881
| 1 | 5 | 3.6 | 3.8 | 3.8 | < 5 | 3.7 | < 5 | < 5 | 3.7 | 3 | 3.7 | 3.5553 | 3.9 | 3.4 | 3.7 | 3.6 | 3.4273 | 3.8 | 3.7 | 3.57 | 3.702 |  | 78
| 1 | 6 | 3.7 | 3.6 | 3.7 | < 5 | 3.8 | < 5 | 5 | 3.6 | 3 | 3.9 | 3.6045 | 4.1 | 3.4 | 3.7 | 3.6 | 3.2798 | 3.9 | 3.8 | 3.56 | 3.9
| Mean |  | 3.6333 | 3.6833 | 3.7333 | 5 | 3.75 |  | 5.25 | 3.6833 | 3 | 3.9667 | 3.5596 | 4 | 3.3833 | 3.6833 | 3.55 | 3.3424 | 3.7833 | 3.7333 | 3.63 | 3.8606
| Median |  | 3.6 | 3.7 | 3.7 | 5 | 3.8 |  | 5 | 3.7 | 3 | 4 | 3.5538 | 4.05 | 3.4 | 3.7 | 3.55 | 3.3223 | 3.8 | 3.7 | 3.61 | 3.8707
| Std Dev. |  | 0.05164 | 0.075277 | 0.10328 | 0 | 0.083666 |  | 0.5 | 0.040825 | 0 | 0.1633 | 0.054303 | 0.12649 | 0.040825 | 0.040825 | 0.054772 | 0.063345 | 0.075277 | 0.05164 | 0.070711 | 0.096394
| Rel.Std.Dev. |  | 1.421% | 2.044% | 2.766% | 0% | 2.231% |  | 9.524% | 1.108% | 0% | 4.117% | 1.526% | 3.162% | 1.207% | 1.108% | 1.543% | 1.895% | 1.99% | 1.383% | 1.948% | 2.497%
| PDM3 |  | -1.485% | -0.1294% | 1.226% | 35.57% | 1.678% |  | 42.35% | -0.1294% | -18.66% | 7.553% | -3.485% | 8.457% | -8.264% | -0.1294% | -3.745% | -9.374% | 2.582% | 1.226% | -1.575% | 4.676%
| Z-Score (Absolute) |  | 0.34 | 0 | 0.34 | NA | 0.45 | 7.98 | NA | 0 | NA | 1.91 | 0.83 | 2.14 | 2.02 | 0 | 0.9 | 2.3 | 0.67 | 0.34 | 0.36 | 1.2
Results from laboratories 1.05, 1.08 and 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 612 (Certified Value 19.5 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 2
| 1 | 1 | 20 | 19 | 19 | 20 | 18 | 19 | 22 | 20 | 20 | 21 | 18.4933 | 19 | 18 | 21 | 19 | 18.496 | 22 | 19 | 18 | 15.559 |  | 1
| 1 | 2 | 19 | 18 | 19 | 23 | 19 | 20 | 20 | 21 | 18 | 19 | 18.7146 | 23 | 18 | 21 | 19 | 18.592 | 21 | 21 | 18 | 13.8525 |  | 27
| 1 | 3 | 19 | 18 | 19 | 19 | 18 | 22 | 20 | 20 | 18 | 20 | 18.9966 | 20 | 20 | 20 | 18 | 18.765 | 19 | 21 | 20 | 12.91 |  | 16
| 1 | 4 | 18 | 18 | 17 | 20 | 19 | 21 | 23 | 21 | 20 | 21 | 18.8474 | 19 | 18 | 21 | 20 | 18.53 | 21 | 20 | 18 | 15.829 |  | 19.5023
| 1 | 5 | 18 | 19 | 20 | 20 | 18 | 19 | 22 | 21 | 18 | 18 | 19.5079 | 19 | 20 | 21 | 20 | 18.347 | 21 | 20 | 18 | 15.3023 |  | 79
| 1 | 6 | 18 | 18 | 21 | 23 | 18 | 18 | 21 | 21 | 19 | 21 | 19.3854 | 19 | 18 | 21 | 18 | 19.388 | 21 | 19 | 18 | 15.4123
| Mean |  | 18.6667 | 18.3333 | 19.1667 | 20.8333 | 18.3333 | 19.8333 | 21.3333 | 20.6667 | 18.8333 | 20 | 18.9909 | 19.8333 | 18.6667 | 20.8333 | 19 | 18.6863 | 20.8333 | 20 | 18.3333 | 14.8109
| Median |  | 18.5 | 18 | 19 | 20 | 18 | 19.5 | 21.5 | 21 | 18.5 | 20.5 | 18.922 | 19 | 18 | 21 | 19 | 18.561 | 21 | 20 | 18 | 15.3573
| Std Dev. |  | 0.8165 | 0.5164 | 1.3292 | 1.7224 | 0.5164 | 1.472 | 1.2111 | 0.5164 | 0.98319 | 1.2649 | 0.39178 | 1.6021 | 1.0328 | 0.40825 | 0.89443 | 0.36966 | 0.98319 | 0.89443 | 0.8165 | 1.1603
| Rel.Std.Dev. |  | 4.374% | 2.817% | 6.935% | 8.268% | 2.817% | 7.422% | 5.677% | 2.499% | 5.22% | 6.325% | 2.063% | 8.078% | 5.533% | 1.96% | 4.708% | 1.978% | 4.719% | 4.472% | 4.454% | 7.834%
| PDM3 |  | -4.285% | -5.994% | -1.721% | 6.825% | -5.994% | 1.697% | 9.389% | 5.97% | -3.43% | 2.552% | -2.622% | 1.697% | -4.285% | 6.825% | -2.576% | -4.184% | 6.825% | 2.552% | -5.994% | -24.06%
| Z-Score (Absolute) |  | 0.37 | 0.67 | 0.07 | 1.57 | 0.67 | 0.67 | 2.02 | 1.42 | 0.22 | 0.82 | 0.08 | 0.67 | 0.37 | 1.57 | 0.07 | 0.36 | 1.57 | 0.82 | 0.67 | 3.84
| Analytical results for Cs in OREAS 612 (Certified Value 2.25 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | [6 empty] | 3
| 1 | 1 | 2.19 | 2.16 | 2.26 | 2.09 | 2.28 | 2.37 | 2.3842 | 2.19 | 2.27 | 2.6 | 2.2 | 2.2011 | 2.26 | 3.21 | 2.23 | [6 empty] | 1
| 1 | 2 | 2.09 | 2.07 | 2.35 | 2.03 | 2.27 | 2.2 | 2.3484 | 2.16 | 2.28 | 2.65 | 2.3 | 2.4145 | 2.32 | 3.3 | 2.23 | [6 empty] | 28
| 1 | 3 | 2.31 | 2.08 | 2.32 | 2.1 | 2.24 | 2.17 | 2.3568 | 2.34 | 2.25 | 2.65 | 2.3 | 2.2783 | 2.42 | 3.16 | 2.25 | [6 empty] | 16
| 1 | 4 | 2.18 | 2.13 | 2.27 | 2.03 | 2.28 | 2.2 | 2.3227 | 2.2 | 2.21 | 2.76 | 2.3 | 2.4353 | 2.32 | 3.13 | 2.23 | [6 empty] | 2.2466
| 1 | 5 | 2.27 | 2.12 | 2.32 | 2.08 | 2.2 | 2.24 | 2.2568 | 2.23 | 2.25 | 2.74 | 2.3 | 2.4169 | 2.4 | 3.13 | 2.2 | [6 empty] | 80
| 1 | 6 | 2.28 | 2.14 | 2.31 | 2.09 | 2.23 | 2.25 | 2.3491 | 2.31 | 2.23 | 2.7 | 2.3 | 2.1781 | 2.4 | 3.16 | 2.21
| Mean |  | 2.22 | 2.1167 | 2.305 | 2.07 | 2.25 | 2.2383 | 2.3363 | 2.2383 | 2.2483 | 2.6833 | 2.2833 | 2.3207 | 2.3533 | 3.1817 | 2.225
| Median |  | 2.23 | 2.125 | 2.315 | 2.085 | 2.255 | 2.22 | 2.3487 | 2.215 | 2.25 | 2.675 | 2.3 | 2.3464 | 2.36 | 3.16 | 2.23
| Std Dev. |  | 0.081976 | 0.035024 | 0.033912 | 0.031623 | 0.032249 | 0.070828 | 0.043642 | 0.071391 | 0.025626 | 0.060882 | 0.040825 | 0.11629 | 0.062823 | 0.064936 | 0.017607
| Rel.Std.Dev. |  | 3.693% | 1.655% | 1.471% | 1.528% | 1.433% | 3.164% | 1.868% | 3.189% | 1.14% | 2.269% | 1.788% | 5.011% | 2.67% | 2.041% | 0.7913%
| PDM3 |  | -1.183% | -5.782% | 2.601% | -7.859% | 0.1528% | -0.3665% | 3.996% | -0.3665% | 0.0786% | 19.44% | 1.637% | 3.3% | 4.752% | 41.62% | -0.96%
| Z-Score (Absolute) |  | 0.37 | 1.63 | 0.67 | 2.21 | 0 | 0.14 | 1.06 | 0.14 | 0.02 | 5.31 | 0.41 | 0.87 | 1.27 | 11.42 | 0.31
| Analytical results for Cu in OREAS 612 (Certified Value 92 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | 0
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 1
| 1 | 1 | 94.2 | 92.5 | 93.1 | 94 | 92.3 | 86 | 97 | 92.1 | 90 | 102 | 88.9651 | 99 | 95.7 | 93 | 82.9 | 87.5 | 90.7323 | 103.2 | 92 | 91.2 | 91.833 | 1
| 1 | 2 | 96.8 | 92.3 | 95 | 93 | 91.7 | 87 | 96 | 92.7 | 89 | 101 | 88.7535 | 93 | 95.5 | 94 | 82.4 | 83.6 | 91.9742 | 104.9 | 98 | 90.1 | 92.8453 | 29
| 1 | 3 | 96.9 | 94.8 | 92.4 | 92 | 93.2 | 88 | 96 | 91.6 | 90 | 94 | 88.4153 | 92 | 97.8 | 90 | 81.8 | 84.9 | 92.4705 | 102.5 | 91 | 91.1 | 90.682 | 16
| 1 | 4 | 93.5 | 93.6 | 92.8 | 92 | 93.1 | 89 | 94 | 94.1 | 92 | 96.5 | 89.6388 | 93 | 100.5 | 92 | 82.9 | 86.8 | 91.9393 | 103.8 | 88 | 94.1 | 91.21 | 92.2795
| 1 | 5 | 93 | 97 | 91.8 | 92 | 92.8 | 86 | 95 | 91.7 | 87 | 96.2 | 88.4091 | 91 | 96.6 | 92 | 82.5 | 86.2 | 91.0266 | 103.7 | 91 | 89.1 | 91.0607 | 81
| 1 | 6 | 93.2 | 95.3 | 94.3 | 92 | 93.8 | 87 | 96 | 91.9 | 89 | 95.3 | 88.346 | 95 | 99.6 | 90 | 83.5 | 85.4 | 91.6376 | 105.7 | 92 | 89.7 | 91.0257
| Mean |  | 94.6 | 94.25 | 93.2333 | 92.5 | 92.8167 | 87.1667 | 95.6667 | 92.35 | 89.5 | 97.5 | 88.7546 | 93.8333 | 97.6167 | 91.8333 | 82.6667 | 85.7333 | 91.6301 | 103.9667 | 92 | 90.8833 | 91.4428
| Median |  | 93.85 | 94.2 | 92.95 | 92 | 92.95 | 87 | 96 | 92 | 89.5 | 96.35 | 88.5844 | 93 | 97.2 | 92 | 82.7 | 85.8 | 91.7885 | 103.75 | 91.5 | 90.6 | 91.1353
| Std Dev. |  | 1.79 | 1.8031 | 1.2011 | 0.83666 | 0.73598 | 1.169 | 1.0328 | 0.94181 | 1.6432 | 3.2336 | 0.49551 | 2.8577 | 2.0721 | 1.6021 | 0.57504 | 1.4024 | 0.64661 | 1.1587 | 3.2863 | 1.7713 | 0.78389
| Rel.Std.Dev. |  | 1.892% | 1.913% | 1.288% | 0.9045% | 0.7929% | 1.341% | 1.08% | 1.02% | 1.836% | 3.316% | 0.5583% | 3.046% | 2.123% | 1.745% | 0.6956% | 1.636% | 0.7057% | 1.115% | 3.572% | 1.949% | 0.8572%
| PDM3 |  | 2.515% | 2.135% | 1.034% | 0.2389% | 0.5821% | -5.541% | 3.671% | 0.07638% | -3.012% | 5.657% | -3.82% | 1.684% | 5.784% | -0.4835% | -10.42% | -7.094% | -0.7038% | 12.66% | -0.3029% | -1.513% | -0.9067%
| Z-Score (Absolute) |  | 0.8 | 0.67 | 0.31 | 0.05 | 0.17 | 1.84 | 1.18 | 0 | 1.01 | 1.83 | 1.28 | 0.53 | 1.87 | 0.18 | 3.44 | 2.35 | 0.26 | 4.12 | 0.12 | 0.52 | 0.32
| Analytical results for Dy in OREAS 612 (Certified Value 1.91 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 2
| 1 | 1 | 2.099 | 2 | 1.7766 | 1.65 | 2.05 | 1.7651 | 2.24 | [14 empty] | 1
| 1 | 2 | 1.993 | 1.9 | 1.7815 | 1.71 | 2 | 1.8005 | 2.32 | [14 empty] | 30
| 1 | 3 | 1.873 | 2 | 1.8255 | 1.65 | 2.09 | 1.7817 | 2.19 | [14 empty] | 16
| 1 | 4 | 1.978 | 2 | 1.8265 | 1.6 | 2.09 | 1.8184 | 2.27 | [14 empty] | 1.9117
| 1 | 5 | 1.913 | 1.8 | 1.7628 | 1.61 | 2.02 | 1.7256 | 2.2 | [14 empty] | 82
| 1 | 6 | 1.862 | 1.9 | 1.7565 | 1.61 | 2.04 | 1.7816 | 2.23
| Mean |  | 1.953 | 1.9333 | 1.7882 | 1.6383 | 2.0483 | 1.7788 | 2.2417
| Median |  | 1.9455 | 1.95 | 1.779 | 1.63 | 2.045 | 1.7817 | 2.235
| Std Dev. |  | 0.089243 | 0.08165 | 0.030606 | 0.041191 | 0.03656 | 0.03186 | 0.047924
| Rel.Std.Dev. |  | 4.57% | 4.223% | 1.712% | 2.514% | 1.785% | 1.791% | 2.138%
| PDM3 |  | 2.162% | 1.133% | -6.457% | -14.3% | 7.149% | -6.949% | 17.26%
| Z-Score (Absolute) |  | 0.09 | 0 | 0.67 | 1.37 | 0.53 | 0.72 | 1.43
| Analytical results for Er in OREAS 612 (Certified Value 0.59 ppm) | [22 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 3
| 1 | 1 | 0.663 | 0.6 | 0.56383 | 0.52 | 0.57 | 0.59846 | 0.67 | [14 empty] | 1
| 1 | 2 | 0.632 | 0.6 | 0.53862 | 0.51 | 0.57 | 0.63784 | 0.69 | [14 empty] | 31
| 1 | 3 | 0.608 | 0.6 | 0.56287 | 0.51 | 0.58 | 0.6023 | 0.68 | [14 empty] | 16
| 1 | 4 | 0.608 | 0.6 | 0.56781 | 0.5 | 0.57 | 0.6534 | 0.68 | [14 empty] | 0.59355
| 1 | 5 | 0.601 | 0.6 | 0.56598 | 0.51 | 0.57 | 0.62016 | 0.68 | [14 empty] | 83
| 1 | 6 | 0.581 | 0.7 | 0.55617 | 0.51 | 0.59 | 0.63719 | 0.66
| Mean |  | 0.6155 | 0.61667 | 0.55921 | 0.51 | 0.575 | 0.62489 | 0.67667
| Median |  | 0.608 | 0.6 | 0.56335 | 0.51 | 0.57 | 0.62868 | 0.68
| Std Dev. |  | 0.028431 | 0.040825 | 0.010842 | 0.0063246 | 0.0083666 | 0.02174 | 0.010328
| Rel.Std.Dev. |  | 4.619% | 6.62% | 1.939% | 1.24% | 1.455% | 3.479% | 1.526%
| PDM3 |  | 3.699% | 3.895% | -5.784% | -14.08% | -3.124% | 5.281% | 14%
| Z-Score (Absolute) |  | 0.42 | NA | 0.74 | 1.75 | 0.42 | 0.61 | 1.67
| Analytical results for Eu in OREAS 612 (Certified Value 0.84 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 2
| 1 | 1 | 0.892 | 0.8 | 0.73322 | 0.68 | 0.87 | 0.99896 | 0.92 | [14 empty] | 1
| 1 | 2 | 0.852 | 0.8 | 0.70782 | 0.68 | 0.87 | 1.0219 | 0.96 | [14 empty] | 32
| 1 | 3 | 0.785 | 0.8 | 0.74733 | 0.66 | 0.89 | 0.99843 | 0.93 | [14 empty] | 16
| 1 | 4 | 0.86 | 0.8 | 0.75374 | 0.69 | 0.89 | 0.99398 | 0.93 | [14 empty] | 0.84199
| 1 | 5 | 0.806 | 0.7 | 0.72393 | 0.68 | 0.88 | 0.98812 | 0.92 | [14 empty] | 84
| 1 | 6 | 0.786 | 0.8 | 0.72415 | 0.67 | 0.9 | 0.98887 | 0.93
| Mean |  | 0.83017 | 0.78333 | 0.7317 | 0.67667 | 0.88333 | 0.99838 | 0.93167
| Median |  | 0.829 | 0.8 | 0.72869 | 0.68 | 0.885 | 0.99621 | 0.93
| Std Dev. |  | 0.044192 | 0.040825 | 0.016853 | 0.010328 | 0.012111 | 0.01241 | 0.01472
| Rel.Std.Dev. |  | 5.323% | 5.212% | 2.303% | 1.526% | 1.371% | 1.243% | 1.58%
| PDM3 |  | -1.404% | -6.966% | -13.1% | -19.63% | 4.911% | 18.57% | 10.65%
| Z-Score (Absolute) |  | 0.18 | NA | 0.84 | 1.21 | 0.18 | 0.96 | 0.51
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 612 (Certified Value 1.84 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 3
| 1 | 1 | 1.74 | 1.78 | 1.88 | 1.84 | 1.76 | 1.71 | 1.87 | 1.82 | 1.79 | 2 | 1.8676 | 1.6175 | 1.79 | 1.88 | 2.14 | 1.8 | 1.876 | 1.99 | 1.85 | 1.79 | 1.9442 | 1
| 1 | 2 | 1.78 | 1.76 | 1.91 | 1.83 | 1.78 | 1.75 | 1.86 | 1.83 | 1.73 | 1.91 | 1.8466 | 1.5886 | 1.78 | 1.87 | 2.08 | 1.8 | 1.8997 | 1.98 | 1.99 | 1.78 | 1.9203 | #N/A
| 1 | 3 | 1.78 | 1.77 | 1.87 | 1.86 | 1.74 | 1.78 | 1.92 | 1.83 | 1.81 | 1.93 | 1.83 | 1.6036 | 1.81 | 1.82 | 2.07 | 1.8 | 1.8725 | 1.97 | 1.83 | 1.8 | 1.9032 | 16
| 1 | 4 | 1.77 | 1.78 | 1.86 | 1.84 | 1.77 | 1.79 | 1.86 | 1.84 | 1.84 | 1.91 | 1.8148 | 1.6126 | 1.8 | 1.85 | 2.14 | 1.82 | 1.8545 | 1.95 | 1.79 | 1.83 | 1.9512 | 1.8355
| 1 | 5 | 1.78 | 1.81 | 1.86 | 1.83 | 1.76 | 1.76 | 1.85 | 1.84 | 1.75 | 1.81 | 1.847 | 1.5875 | 1.78 | 1.86 | 2.13 | 1.84 | 1.8682 | 1.98 | 1.86 | 1.74 | 1.8723 | 85
| 1 | 6 | 1.8 | 1.79 | 1.89 | 1.83 | 1.74 | 1.72 | 1.88 | 1.8 | 1.79 | 1.9 | 1.8305 | 1.6069 | 1.78 | 1.83 | 2.1 | 1.83 | 1.8708 | 1.97 | 1.84 | 1.77 | 1.8585
| Mean |  | 1.775 | 1.7817 | 1.8783 | 1.8383 | 1.7583 | 1.7517 | 1.8733 | 1.8267 | 1.785 | 1.91 | 1.8394 | 1.6028 | 1.79 | 1.8517 | 2.11 | 1.815 | 1.8736 | 1.9733 | 1.86 | 1.785 | 1.9083
| Median |  | 1.78 | 1.78 | 1.875 | 1.835 | 1.76 | 1.755 | 1.865 | 1.83 | 1.79 | 1.91 | 1.8385 | 1.6053 | 1.785 | 1.855 | 2.115 | 1.81 | 1.8717 | 1.975 | 1.845 | 1.785 | 1.9117
| Std Dev. |  | 0.019748 | 0.017224 | 0.019408 | 0.01169 | 0.016021 | 0.031885 | 0.025033 | 0.015055 | 0.039875 | 0.060992 | 0.018317 | 0.012372 | 0.012649 | 0.023166 | 0.030984 | 0.017607 | 0.014755 | 0.013663 | 0.068118 | 0.030166 | 0.037606
| Rel.Std.Dev. |  | 1.113% | 0.9667% | 1.033% | 0.6359% | 0.9111% | 1.82% | 1.336% | 0.8242% | 2.234% | 3.193% | 0.9958% | 0.7719% | 0.7067% | 1.251% | 1.468% | 0.9701% | 0.7875% | 0.6924% | 3.662% | 1.69% | 1.971%
| PDM3 |  | -3.296% | -2.933% | 2.333% | 0.154% | -4.205% | -4.568% | 2.061% | -0.4817% | -2.752% | 4.058% | 0.2117% | -12.68% | -2.479% | 0.8804% | 14.95% | -1.117% | 2.077% | 7.509% | 1.334% | -2.752% | 3.965%
| Z-Score (Absolute) |  | 0.8 | 0.72 | 0.51 | 0 | 1.01 | 1.1 | 0.44 | 0.15 | 0.67 | 0.91 | 0.01 | 2.98 | 0.61 | 0.17 | 3.43 | 0.3 | 0.45 | 1.71 | 0.27 | 0.67 | 0.88
| Analytical results for Ga in OREAS 612 (Certified Value 4.33 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 3
| 1 | 1 | 4.5 | 2.9 | 3.8 | < 50 | 4.07 | < 50 | < 50 | 4.39 | 5.01 | 4.686 | 4.18 | 4.15 | 4.34 | 5 | 4.8563 | 4 | 6.02 | 3.9 | 4.6957 |  |  | 1
| 1 | 2 | 4.11 | 2.8 | 3.9 | < 50 | 4.01 | < 50 | < 50 | 4.29 | 4.77 | 4.5747 | 4.33 | 4.18 | 4.43 | 5 | 4.924 | 4 | 6.47 | 3.79 | 4.386 |  |  | 33
| 1 | 3 | 4.2 | 2.8 | 3.7 | < 50 | 4.18 | < 50 | < 50 | 4.13 | 4.58 | 4.5825 | 4.7 | 4.08 | 4.35 | 5 | 4.8646 | 4.1 | 5.9 | 3.85 | 4.2885 |  |  | 16
| 1 | 4 | 4.54 | 2.9 | 3.7 | < 50 | 4.12 | < 50 | < 50 | 4.38 | 4.62 | 4.762 | 4.49 | 4.22 | 4.59 | 5 | 4.938 | 4 | 5.96 | 3.84 | 4.709 |  |  | 4.3318
| 1 | 5 | 4.46 | 2.8 | 3.9 | < 50 | 4.09 | < 50 | < 50 | 4.25 | 4.79 | 4.5874 | 4.3 | 4.14 | 4.52 | 5 | 4.8682 | 4.1 | 5.97 | 3.75 | 4.4733 |  |  | 86
| 1 | 6 | 4.39 | 2.9 | 3.9 | < 50 | 4.32 | < 50 | < 50 | 4.16 | 4.77 | 4.5957 | 4.63 | 4.14 | 4.65 | 5 | 4.8319 | 4.2 | 6.02 | 3.85 | 4.427
| Mean |  | 4.3667 | 2.85 | 3.8167 |  | 4.1317 |  |  | 4.2667 | 4.7567 | 4.6314 | 4.4383 | 4.1517 | 4.48 | 5 | 4.8805 | 4.0667 | 6.0567 | 3.83 | 4.4966
| Median |  | 4.425 | 2.85 | 3.85 |  | 4.105 |  |  | 4.27 | 4.77 | 4.5915 | 4.41 | 4.145 | 4.475 | 5 | 4.8664 | 4.05 | 5.995 | 3.845 | 4.4502
| Std Dev. |  | 0.17363 | 0.054772 | 0.098319 |  | 0.10797 |  |  | 0.10857 | 0.15201 | 0.075956 | 0.20272 | 0.046655 | 0.12775 | 0 | 0.041355 | 0.08165 | 0.20733 | 0.052536 | 0.17068
| Rel.Std.Dev. |  | 3.976% | 1.922% | 2.576% |  | 2.613% |  |  | 2.545% | 3.196% | 1.64% | 4.568% | 1.124% | 2.852% | 0% | 0.8473% | 2.008% | 3.423% | 1.372% | 3.796%
| PDM3 |  | 0.8048% | -34.21% | -11.89% |  | -4.62% |  |  | -1.504% | 9.808% | 6.915% | 2.459% | -4.159% | 3.421% | 15.43% | 12.67% | -6.121% | 39.82% | -11.58% | 3.804%
| Z-Score (Absolute) |  | 0.17 | 3.01 | 1.2 | 38.47 | 0.61 | 38.47 | 38.47 | 0.36 | 0.56 | 0.32 | 0.04 | 0.58 | 0.04 | NA | 0.79 | 0.74 | 2.99 | 1.18 | 0.07
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 612 (Certified Value 3.5 ppm) | [22 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 3
| 1 | 1 | 3.775 | 3.7 | 3.2087 | 3.33 | 3.55 | 3.1068 | 4.11 | [14 empty] | 1
| 1 | 2 | 3.703 | 3.7 | 3.23 | 3.29 | 3.52 | 3.1379 | 4.25 | [14 empty] | 34
| 1 | 3 | 3.514 | 3.4 | 3.2928 | 3.26 | 3.5 | 3.1387 | 4.01 | [14 empty] | 16
| 1 | 4 | 3.719 | 3.3 | 3.2933 | 3.19 | 3.78 | 3.143 | 4.12 | [14 empty] | 3.5008
| 1 | 5 | 3.616 | 3.3 | 3.1614 | 3.13 | 3.73 | 3.1227 | 4.03 | [14 empty] | 87
| 1 | 6 | 3.455 | 3.8 | 3.19 | 3.13 | 3.87 | 3.1762 | 4.05
| Mean |  | 3.6303 | 3.5333 | 3.2294 | 3.2217 | 3.6583 | 3.1376 | 4.095
| Median |  | 3.6595 | 3.55 | 3.2193 | 3.225 | 3.64 | 3.1383 | 4.08
| Std Dev. |  | 0.12534 | 0.22509 | 0.054231 | 0.084479 | 0.15536 | 0.023212 | 0.087579
| Rel.Std.Dev. |  | 3.453% | 6.371% | 1.679% | 2.622% | 4.247% | 0.7398% | 2.139%
| PDM3 |  | 3.7% | 0.9294% | -7.754% | -7.973% | 4.5% | -10.38% | 16.97%
| Z-Score (Absolute) |  | 0.22 | 0 | 0.67 | 0.69 | 0.28 | 0.88 | 1.25
| Analytical results for Ge in OREAS 612 (Indicative Value 0.095 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.11 | 1.14 | 1.15 | 1.17 | 1.20 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [12 empty] | 3
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | [12 empty] | 3
| 1 | 1 | 0.05 | < 0.1 | < 0.1 | 0.08 | < 0.1 | 0.08 | 0.13 | 1.1 | < 0.1 | [12 empty] | 1
| 1 | 2 | 0.06 | < 0.1 | < 0.1 | 0.08 | < 0.1 | 0.12 | 0.14 | 1 | < 0.1 | [12 empty] | 4
| 1 | 3 | 0.05 | < 0.1 | < 0.1 | 0.09 | < 0.1 | 0.09 | 0.14 | 1 | < 0.1 | [12 empty] | 16
| 1 | 4 | 0.05 | < 0.1 | < 0.1 | 0.08 | < 0.1 | 0.1 | 0.14 | 1.1 | < 0.1 | [12 empty] | 0.094583
| 1 | 5 | 0.06 | < 0.1 | < 0.1 | 0.08 | < 0.1 | 0.1 | 0.14 | 1 | < 0.1 | [12 empty] | 13
| 1 | 6 | 0.06 | < 0.1 | < 0.1 | 0.09 | < 0.1 | 0.12 | 0.14 | 1 | < 0.1
| Mean |  | 0.055 |  |  | 0.083333 |  | 0.10167 | 0.13833 | 1.0333
| Median |  | 0.055 |  |  | 0.08 |  | 0.1 | 0.14 | 1
| Std Dev. |  | 0.0054772 |  |  | 0.005164 |  | 0.016021 | 0.0040825 | 0.05164
| Rel.Std.Dev. |  | 9.959% |  |  | 6.197% |  | 15.76% | 2.951% | 4.997%
| PDM3 |  | -41.85% |  |  | -11.89% |  | 7.489% | 46.26% | 992.5%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.67 | 3.82 | 0.67 | 6.29 | 11.24 | 131.94 | 0.67
| Analytical results for Hf in OREAS 612 (Certified Value 0.82 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | [7 empty] | 3
| 1 | 1 | 0.85 | 0.81 | 0.89 | 0.77 | 0.82 | < 0.1 | 0.82792 | 0.81 | 0.77 | 0.28 | 0.7 | 0.78 | 0.96 | 0.63 | [7 empty] | 1
| 1 | 2 | 0.9 | 0.76 | 0.86 | 0.8 | 0.82 | < 0.1 | 0.81132 | 0.82 | 0.76 | 0.29 | 0.7 | 0.81 | 0.99 | 0.69 | [7 empty] | 19
| 1 | 3 | 0.85 | 0.8 | 0.85 | 0.78 | 0.76 | < 0.1 | 0.83881 | 0.86 | 0.76 | 0.28 | 0.7 | 0.81 | 0.96 | 0.79 | [7 empty] | 16
| 1 | 4 | 0.9 | 0.81 | 0.89 | 0.82 | 0.87 | < 0.1 | 0.84832 | 0.84 | 0.73 | 0.28 | 0.7 | 0.8 | 0.96 | 0.79 | [7 empty] | 0.81543
| 1 | 5 | 0.9 | 0.79 | 0.9 | 0.8 | 0.83 | < 0.1 | 0.81564 | 0.8 | 0.72 | 0.28 | 0.7 | 0.82 | 0.93 | 0.77 | [7 empty] | 88
| 1 | 6 | 0.91 | 0.79 | 0.91 | 0.83 | 0.8 | < 0.1 | 0.81992 | 0.87 | 0.71 | 0.28 | 0.7 | 0.84 | 0.94 | 0.78
| Mean |  | 0.885 | 0.79333 | 0.88333 | 0.8 | 0.81667 |  | 0.82699 | 0.83333 | 0.74167 | 0.28167 | 0.7 | 0.81 | 0.95667 | 0.74167
| Median |  | 0.9 | 0.795 | 0.89 | 0.8 | 0.82 |  | 0.82392 | 0.83 | 0.745 | 0.28 | 0.7 | 0.81 | 0.96 | 0.775
| Std Dev. |  | 0.027386 | 0.018619 | 0.023381 | 0.022804 | 0.036148 |  | 0.014283 | 0.028048 | 0.024833 | 0.0040825 | 1.2162e-16 | 0.02 | 0.020656 | 0.066458
| Rel.Std.Dev. |  | 3.094% | 2.347% | 2.647% | 2.85% | 4.426% |  | 1.727% | 3.366% | 3.348% | 1.449% | 1.737e-14% | 2.469% | 2.159% | 8.961%
| PDM3 |  | 8.531% | -2.71% | 8.327% | -1.893% | 0.1514% |  | 1.417% | 2.195% | -9.046% | -65.46% | -14.16% | -0.6662% | 17.32% | -9.046%
| Z-Score (Absolute) |  | 0.74 | 0.16 | 0.72 | 0.1 | 0.07 | 7.5 | 0.17 | 0.23 | 0.67 | 5.21 | NA | 0 | 1.45 | 0.67
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 612 (Certified Value 0.19 ppm) | [22 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | [5 empty] | 3
| 1 | 1 | 0.2 | 0.18 | < 5 | 0.21 | < 5 | < 5 | 0.2 | < 1 | 0.21 | < 1 | 0.2 | 0.17 | 0.5 | 0.24 | 0.21 | 0.19 | [5 empty] | 1
| 1 | 2 | 0.18 | 0.19 | < 5 | 0.18 | < 5 | < 5 | 0.2 | < 1 | 0.22 | < 1 | 0.2 | 0.18 | 0.51 | 0.232 | 0.2 | 0.2 | [5 empty] | 3
| 1 | 3 | 0.18 | 0.19 | < 5 | 0.19 | < 5 | < 5 | 0.2 | < 1 | 0.17 | < 1 | 0.21 | 0.19 | 0.51 | 0.233 | 0.22 | 0.18 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.18 | < 5 | 0.19 | < 5 | < 5 | 0.2 | < 1 | 0.18 | < 1 | 0.2 | 0.18 | 0.49 | 0.238 | 0.22 | 0.2 | [5 empty] | 0.19458
| 1 | 5 | 0.19 | 0.19 | < 5 | 0.2 | < 5 | < 5 | 0.2 | < 1 | 0.19 | < 1 | 0.2 | 0.18 | 0.52 | 0.223 | 0.2 | 0.19 | [5 empty] | 89
| 1 | 6 | 0.19 | 0.2 | < 5 | 0.2 | < 5 | < 5 | 0.2 | < 1 | 0.21 | < 1 | 0.2 | 0.18 | 0.45 | 0.224 | 0.21 | 0.21
| Mean |  | 0.19 | 0.18833 |  | 0.195 |  |  | 0.2 |  | 0.19667 |  | 0.20167 | 0.18 | 0.49667 | 0.23167 | 0.21 | 0.195
| Median |  | 0.19 | 0.19 |  | 0.195 |  |  | 0.2 |  | 0.2 |  | 0.2 | 0.18 | 0.505 | 0.2325 | 0.21 | 0.195
| Std Dev. |  | 0.0089443 | 0.0075277 |  | 0.010488 |  |  | 3.0405e-17 |  | 0.019664 |  | 0.0040825 | 0.0063246 | 0.025033 | 0.0070048 | 0.0089443 | 0.010488
| Rel.Std.Dev. |  | 4.708% | 3.997% |  | 5.379% |  |  | 1.52e-14% |  | 9.999% |  | 2.024% | 3.514% | 5.04% | 3.024% | 4.259% | 5.379%
| PDM3 |  | -2.355% | -3.212% |  | 0.2141% |  |  | 2.784% |  | 1.071% |  | 3.64% | -7.495% | 155.2% | 19.06% | 7.923% | 0.2141%
| Z-Score (Absolute) |  | 0.62 | 0.67 | 71.27 | 0.47 | 71.27 | 71.27 | NA | 9.03 | 0.41 | 9.03 | 0.26 | 0.93 | 8.92 | 0.67 | 0 | 0.47
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ho in OREAS 612 (Certified Value 0.27 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 2
| 1 | 1 | 0.3 | 0.3 | 0.24498 | 0.22 | 0.94 | 0.28978 | 0.31 | [14 empty] | 1
| 1 | 2 | 0.3 | 0.3 | 0.24272 | 0.23 | 0.92 | 0.26944 | 0.32 | [14 empty] | 20
| 1 | 3 | 0.3 | 0.3 | 0.24761 | 0.23 | 0.95 | 0.24628 | 0.31 | [14 empty] | 16
| 1 | 4 | 0.3 | 0.3 | 0.25377 | 0.22 | 0.94 | 0.27728 | 0.31 | [14 empty] | 0.27442
| 1 | 5 | 0.3 | 0.3 | 0.24062 | 0.23 | 0.92 | 0.25966 | 0.3 | [14 empty] | 90
| 1 | 6 | 0.3 | 0.3 | 0.24342 | 0.22 | 0.96 | 0.2536 | 0.31
| Mean |  | 0.3 | 0.3 | 0.24552 | 0.225 | 0.93833 | 0.26601 | 0.31
| Median |  | 0.3 | 0.3 | 0.2442 | 0.225 | 0.94 | 0.26455 | 0.31
| Std Dev. |  | 0 | 0 | 0.004668 | 0.0054772 | 0.016021 | 0.016045 | 0.0063246
| Rel.Std.Dev. |  | 0% | 0% | 1.901% | 2.434% | 1.707% | 6.032% | 2.04%
| PDM3 |  | 9.321% | 9.321% | -10.53% | -18.01% | 241.9% | -3.066% | 12.97%
| Z-Score (Absolute) |  | 0 | 0 | 1.08 | 1.49 | 12.66 | 0.67 | 0.2
| Analytical results for In in OREAS 612 (Certified Value 0.064 ppm) | [22 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | [8 empty] | 3
| 1 | 1 | 0.118 | 0.07 | 0.055 | 0.063 | 0.06 | 0.07 | < 0.2 | 0.063 | 0.076 | 0.07 | 0.06 | 0.06 | 0.08 | [8 empty] | 1
| 1 | 2 | 0.108 | 0.06 | 0.057 | 0.056 | 0.06 | 0.07 | < 0.2 | 0.065 | 0.076 | 0.07 | 0.06 | 0.06 | 0.07 | [8 empty] | 21
| 1 | 3 | 0.115 | 0.06 | 0.06 | 0.055 | 0.06 | 0.08 | < 0.2 | 0.067 | 0.076 | 0.07 | 0.05 | 0.06 | 0.07 | [8 empty] | 16
| 1 | 4 | 0.108 | 0.06 | 0.054 | 0.057 | 0.06 | 0.08 | < 0.2 | 0.068 | 0.076 | 0.07 | 0.05 | 0.06 | 0.07 | [8 empty] | 0.064303
| 1 | 5 | 0.129 | 0.06 | 0.056 | 0.055 | 0.06 | 0.07 | < 0.2 | 0.069 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | [8 empty] | 91
| 1 | 6 | 0.117 | 0.06 | 0.058 | 0.056 | 0.06 | 0.06 | < 0.2 | 0.074 | 0.082 | 0.07 | 0.06 | 0.06 | 0.07
| Mean |  | 0.11583 | 0.061667 | 0.056667 | 0.057 | 0.06 | 0.071667 |  | 0.067667 | 0.076 | 0.07 | 0.056667 | 0.06 | 0.07
| Median |  | 0.116 | 0.06 | 0.0565 | 0.056 | 0.06 | 0.07 |  | 0.0675 | 0.076 | 0.07 | 0.06 | 0.06 | 0.07
| Std Dev. |  | 0.0077825 | 0.0040825 | 0.0021602 | 0.0030332 | 0 | 0.0075277 |  | 0.0037771 | 0.0037947 | 0 | 0.005164 | 0 | 0.0063246
| Rel.Std.Dev. |  | 6.719% | 6.62% | 3.812% | 5.321% | 0% | 10.5% |  | 5.582% | 4.993% | 0% | 9.113% | 0% | 9.035%
| PDM3 |  | 80.14% | -4.1% | -11.88% | -11.36% | -6.692% | 11.45% |  | 5.231% | 18.19% | 8.86% | -11.88% | -6.692% | 8.86%
| Z-Score (Absolute) |  | 4.24 | 0.53 | 0.97 | 0.94 | 0.67 | 0.35 | 2.84 | 0 | 0.73 | 0.21 | 0.97 | 0.67 | 0.21
| Analytical results for K in OREAS 612 (Certified Value 0.436 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 3
| 1 | 1 | 0.4 | 0.42 | 0.428 | 0.62 | 0.41 | 0.49 | 0.64 | 0.4507 | 0.45 | 0.49 | 0.44341 | 0.43 | 0.45 | 0.495 | 0.44 | 0.46 | 0.6049 | 0.41 | 0.42792 |  |  | 1
| 1 | 2 | 0.43 | 0.41 | 0.437 | 0.61 | 0.41 | 0.52 | 0.65 | 0.4532 | 0.42 | 0.46 | 0.43599 | 0.43 | 0.45 | 0.485 | 0.44 | 0.44 | 0.6507 | 0.41 | 0.40747 |  |  | #N/A
| 1 | 3 | 0.44 | 0.42 | 0.426 | 0.62 | 0.42 | 0.54 | 0.65 | 0.4433 | 0.42 | 0.45 | 0.43574 | 0.43 | 0.43 | 0.485 | 0.44 | 0.44 | 0.5947 | 0.41 | 0.40914 |  |  | 16
| 1 | 4 | 0.41 | 0.42 | 0.426 | 0.61 | 0.42 | 0.55 | 0.66 | 0.4503 | 0.45 | 0.44 | 0.43102 | 0.42 | 0.44 | 0.507 | 0.45 | 0.47 | 0.5884 | 0.42 | 0.39419 |  |  | 0.43563
| 1 | 5 | 0.42 | 0.43 | 0.424 | 0.61 | 0.42 | 0.54 | 0.65 | 0.4376 | 0.42 | 0.43 | 0.43985 | 0.42 | 0.44 | 0.481 | 0.45 | 0.47 | 0.603 | 0.4 | 0.43192 |  |  | 92
| 1 | 6 | 0.43 | 0.43 | 0.438 | 0.62 | 0.42 | 0.53 | 0.65 | 0.443 | 0.41 | 0.44 | 0.43152 | 0.42 | 0.44 | 0.495 | 0.46 | 0.47 | 0.601 | 0.41 | 0.40221
| Mean |  | 0.42167 | 0.42167 | 0.42983 | 0.615 | 0.41667 | 0.52833 | 0.65 | 0.44635 | 0.42833 | 0.45167 | 0.43625 | 0.425 | 0.44167 | 0.49133 | 0.44667 | 0.45833 | 0.60712 | 0.41 | 0.41214
| Median |  | 0.425 | 0.42 | 0.427 | 0.615 | 0.42 | 0.535 | 0.65 | 0.4468 | 0.42 | 0.445 | 0.43586 | 0.425 | 0.44 | 0.49 | 0.445 | 0.465 | 0.602 | 0.41 | 0.40831
| Std Dev. |  | 0.01472 | 0.0075277 | 0.00608 | 0.0054772 | 0.005164 | 0.02137 | 0.0063246 | 0.0059755 | 0.017224 | 0.02137 | 0.0047809 | 0.0054772 | 0.0075277 | 0.0095847 | 0.008165 | 0.01472 | 0.022196 | 0.0063246 | 0.014778
| Rel.Std.Dev. |  | 3.491% | 1.785% | 1.415% | 0.8906% | 1.239% | 4.045% | 0.973% | 1.339% | 4.021% | 4.731% | 1.096% | 1.289% | 1.704% | 1.951% | 1.828% | 3.212% | 3.656% | 1.543% | 3.586%
| PDM3 |  | -3.205% | -3.205% | -1.331% | 41.17% | -4.353% | 21.28% | 49.21% | 2.461% | -1.675% | 3.681% | 0.1432% | -2.44% | 1.386% | 12.79% | 2.534% | 5.212% | 39.37% | -5.883% | -5.392%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.4 | 5.84 | 0.84 | 2.92 | 7.02 | 0.16 | 0.45 | 0.34 | 0.18 | 0.56 | 0 | 1.67 | 0.17 | 0.56 | 5.58 | 1.07 | 1
| Analytical results for La in OREAS 612 (Certified Value 23.6 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 2
| 1 | 1 | 23.3 | 21.8 | 23.3 | < 50 | 23.4 | < 50 | < 50 | 26.461 | 21.9 | 26.5 | 23.073 | 23.8 | 24 | 25.3 | 24 | 22.2558 | 23.6 | 30 | 20.4 | 23.4357 |  | 1
| 1 | 2 | 24.4 | 21.4 | 23 | < 50 | 22.2 | < 50 | < 50 | 26.653 | 21.7 | 25.6 | 22.6642 | 24.4 | 24 | 26 | 24 | 22.098 | 23.5 | 31 | 20.1 | 23.6297 |  | 22
| 1 | 3 | 25 | 21.3 | 23.5 | < 50 | 23.9 | < 50 | < 50 | 25.395 | 22.5 | 26.6 | 22.8294 | 26.1 | 23 | 25.1 | 24 | 22.0795 | 23.6 | 29 | 20.6 | 23.3037 |  | 16
| 1 | 4 | 23.7 | 21.6 | 23.6 | < 50 | 23.2 | < 50 | < 50 | 26.763 | 22.6 | 26 | 22.7784 | 24.9 | 23 | 26.3 | 24 | 22.1689 | 23.5 | 29 | 20.9 | 23.5447 |  | 23.5825
| 1 | 5 | 23.9 | 22 | 23.5 | < 50 | 22.8 | < 50 | < 50 | 25.853 | 21.9 | 24.3 | 22.7249 | 24.1 | 23 | 26.4 | 24 | 22.7428 | 23.9 | 29 | 19.7 | 23.639 |  | 93
| 1 | 6 | 24.7 | 21.8 | 24.5 | < 50 | 24.2 | < 50 | < 50 | 25.353 | 22 | 26.4 | 22.7048 | 26.2 | 23 | 26.3 | 24 | 22.2603 | 24 | 29 | 19.6 | 23.206
| Mean |  | 24.1667 | 21.65 | 23.5667 |  | 23.2833 |  |  | 26.0797 | 22.1 | 25.9 | 22.7958 | 24.9167 | 23.3333 | 25.9 | 24 | 22.2675 | 23.6833 | 29.5 | 20.2167 | 23.4598
| Median |  | 24.15 | 21.7 | 23.5 |  | 23.3 |  |  | 26.157 | 21.95 | 26.2 | 22.7516 | 24.65 | 23 | 26.15 | 24 | 22.2124 | 23.6 | 29 | 20.25 | 23.4902
| Std Dev. |  | 0.64395 | 0.26646 | 0.50465 |  | 0.72778 |  |  | 0.63076 | 0.36332 | 0.86718 | 0.14756 | 1.0226 | 0.5164 | 0.56214 | 0 | 0.24489 | 0.2137 | 0.83666 | 0.51153 | 0.17751
| Rel.Std.Dev. |  | 2.665% | 1.231% | 2.141% |  | 3.126% |  |  | 2.419% | 1.644% | 3.348% | 0.6473% | 4.104% | 2.213% | 2.17% | 0% | 1.1% | 0.9023% | 2.836% | 2.53% | 0.7566%
| PDM3 |  | 2.477% | -8.194% | -0.06699% |  | -1.268% |  |  | 10.59% | -6.286% | 9.827% | -3.336% | 5.658% | -1.056% | 9.827% | 1.771% | -5.576% | 0.4277% | 25.09% | -14.27% | -0.5202%
| Z-Score (Absolute) |  | 0.19 | 1.28 | 0.16 | 0.67 | 0.33 | 0.67 | 0.67 | 1.3 | 1.01 | 1.2 | 0.61 | 0.63 | 0.3 | 1.2 | 0.09 | 0.92 | 0.09 | 3.29 | 2.11 | 0.22
| Analytical results for Li in OREAS 612 (Certified Value 10.4 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  |  |  | 2
| 1 | 1 | 10.1 | 10 | 11 | 10 | 9.9 | 9 | 10.6 | 10.681 | 9.8 | 10.1 | 10.9 | 10.9 | 11.761 | 12 | 12.2 | 9.7 | 9.7097 |  |  |  |  | 1
| 1 | 2 | 10.5 | 9 | 10 | 10.2 | 9.5 | 9 | 10.4 | 10.5863 | 9.8 | 10.1 | 10.8 | 10.8 | 11.752 | 11 | 12.6 | 9.4 | 10.1117 |  |  |  |  | #N/A
| 1 | 3 | 10.5 | 10 | 10 | 10 | 9.5 | 10 | 10.5 | 10.4209 | 9.9 | 10.1 | 10.7 | 11.6 | 11.789 | 11 | 12.1 | 9.6 | 10.0503 |  |  |  |  | 16
| 1 | 4 | 10.3 | 10 | 10 | 10.2 | 10 | 8 | 10 | 10.3391 | 9.9 | 10.2 | 11.4 | 12.6 | 11.499 | 12 | 12.3 | 9.5 | 9.9803 |  |  |  |  | 10.3974
| 1 | 5 | 10.3 | 10 | 10 | 10 | 9.7 | 10 | 9.8 | 10.3739 | 9.7 | 10.2 | 11.5 | 12.5 | 11.585 | 12 | 12.3 | 9.3 | 10.1155 |  |  |  |  | 94
| 1 | 6 | 10.4 | 10 | 10 | 10 | 9.3 | 10 | 10.3 | 10.1019 | 9.7 | 10.2 | 11.3 | 11.9 | 11.466 | 12 | 12.3 | 9.6 | 10.2073
| Mean |  | 10.35 | 9.8333 | 10.1667 | 10.0667 | 9.65 | 9.3333 | 10.2667 | 10.4172 | 9.8 | 10.15 | 11.1 | 11.7167 | 11.642 | 11.6667 | 12.3 | 9.5167 | 10.0291
| Median |  | 10.35 | 10 | 10 | 10 | 9.6 | 9.5 | 10.35 | 10.3974 | 9.8 | 10.15 | 11.1 | 11.75 | 11.6685 | 12 | 12.3 | 9.55 | 10.081
| Std Dev. |  | 0.15166 | 0.40825 | 0.40825 | 0.10328 | 0.26646 | 0.8165 | 0.30768 | 0.20277 | 0.089443 | 0.054772 | 0.34059 | 0.7679 | 0.14321 | 0.5164 | 0.16733 | 0.1472 | 0.17373
| Rel.Std.Dev. |  | 1.465% | 4.152% | 4.016% | 1.026% | 2.761% | 8.748% | 2.997% | 1.947% | 0.9127% | 0.5396% | 3.068% | 6.554% | 1.23% | 4.426% | 1.36% | 1.547% | 1.732%
| PDM3 |  | -0.4559% | -5.425% | -2.219% | -3.181% | -7.188% | -10.23% | -1.257% | 0.1902% | -5.746% | -2.38% | 6.757% | 12.69% | 11.97% | 12.21% | 18.3% | -8.471% | -3.542%
| Z-Score (Absolute) |  | 0.12 | NA | NA | 0.29 | 0.89 | NA | 0 | 0.22 | 0.67 | 0.17 | 1.2 | 2.1 | 1.99 | NA | 2.94 | 1.08 | 0.34
Results from laboratories 1.02, 1.04, 1.10 and 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 612 (Certified Value 0.057 ppm) | [22 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.04 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.20 | 1.22 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 3
|  |  | 0.25g | 0.25g | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [12 empty] | 3
| 1 | 1 | 0.1 | 0.05 | 0.061 | < 0.1 | 0.050734 | 0.05 | 0.06 | 0.06 | 0.07 | [12 empty] | 1
| 1 | 2 | 0.05 | 0.06 | 0.062 | < 0.1 | 0.051378 | 0.05 | 0.06 | 0.06 | 0.07 | [12 empty] | 23
| 1 | 3 | 0.05 | 0.05 | 0.059 | < 0.1 | 0.053034 | 0.05 | 0.06 | 0.06 | 0.07 | [12 empty] | 16
| 1 | 4 | 0.06 | 0.05 | 0.059 | < 0.1 | 0.05622 | 0.05 | 0.06 | 0.06 | 0.07 | [12 empty] | 0.057411
| 1 | 5 | 0.05 | 0.06 | 0.057 | < 0.1 | 0.053931 | 0.05 | 0.06 | 0.06 | 0.07 | [12 empty] | 95
| 1 | 6 | 0.05 | 0.06 | 0.058 | < 0.1 | 0.052442 | 0.05 | 0.06 | 0.06 | 0.07
| Mean |  | 0.06 | 0.055 | 0.059333 |  | 0.052956 | 0.05 | 0.06 | 0.06 | 0.07
| Median |  | 0.05 | 0.055 | 0.059 |  | 0.052738 | 0.05 | 0.06 | 0.06 | 0.07
| Std Dev. |  | 0.02 | 0.0054772 | 0.0018619 |  | 0.0019648 | 7.6012e-18 | 0 | 0 | 0
| Rel.Std.Dev. |  | 33.33% | 9.959% | 3.138% |  | 3.71% | 1.52e-14% | 0% | 0% | 0%
| PDM3 |  | 4.509% | -4.2% | 3.348% |  | -7.759% | -12.91% | 4.509% | 4.509% | 21.93%
| Z-Score (Absolute) |  | 0.1 | 0.67 | 0 | 1.45 | 0.99 | 1.45 | 0.1 | 0.1 | 1.66
| Analytical results for Mg in OREAS 612 (Certified Value 0.165 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 0.16 | 0.16 | 0.167 | 0.17 | 0.15 | 0.15 | 0.17 | 0.16 | 0.16 | 0.17 | 0.16717 | 0.15 | 0.17 | 0.17 | 0.16 | 0.13623 | 0.18 | 0.18 | 0.17 | 0.1695 |  | 1
| 1 | 2 | 0.16 | 0.16 | 0.167 | 0.17 | 0.16 | 0.15 | 0.17 | 0.16 | 0.158 | 0.18 | 0.16539 | 0.15 | 0.16 | 0.17 | 0.16 | 0.13654 | 0.16 | 0.19 | 0.16 | 0.17104 |  | #N/A
| 1 | 3 | 0.17 | 0.16 | 0.164 | 0.17 | 0.15 | 0.16 | 0.17 | 0.16 | 0.163 | 0.18 | 0.16323 | 0.15 | 0.16 | 0.16 | 0.16 | 0.13714 | 0.17 | 0.18 | 0.17 | 0.16945 |  | 16
| 1 | 4 | 0.16 | 0.16 | 0.162 | 0.17 | 0.16 | 0.16 | 0.17 | 0.16 | 0.165 | 0.17 | 0.16456 | 0.15 | 0.17 | 0.17 | 0.16 | 0.13643 | 0.18 | 0.18 | 0.17 | 0.17094 |  | 0.16474
| 1 | 5 | 0.16 | 0.16 | 0.163 | 0.17 | 0.15 | 0.16 | 0.17 | 0.16 | 0.16 | 0.16 | 0.16472 | 0.15 | 0.16 | 0.17 | 0.16 | 0.13865 | 0.18 | 0.18 | 0.16 | 0.17127 |  | 96
| 1 | 6 | 0.17 | 0.16 | 0.168 | 0.17 | 0.15 | 0.15 | 0.17 | 0.15 | 0.162 | 0.19 | 0.16309 | 0.15 | 0.15 | 0.17 | 0.17 | 0.13737 | 0.18 | 0.18 | 0.16 | 0.17104
| Mean |  | 0.16333 | 0.16 | 0.16517 | 0.17 | 0.15333 | 0.155 | 0.17 | 0.15833 | 0.16133 | 0.175 | 0.16469 | 0.15 | 0.16167 | 0.16833 | 0.16167 | 0.13706 | 0.175 | 0.18167 | 0.165 | 0.17054
| Median |  | 0.16 | 0.16 | 0.1655 | 0.17 | 0.15 | 0.155 | 0.17 | 0.16 | 0.161 | 0.175 | 0.16464 | 0.15 | 0.16 | 0.17 | 0.16 | 0.13684 | 0.18 | 0.18 | 0.165 | 0.17099
| Std Dev. |  | 0.005164 | 0 | 0.0024833 | 0 | 0.005164 | 0.0054772 | 0 | 0.0040825 | 0.0025033 | 0.010488 | 0.0015068 | 0 | 0.0075277 | 0.0040825 | 0.0040825 | 0.00089304 | 0.0083666 | 0.0040825 | 0.0054772 | 0.00083403
| Rel.Std.Dev. |  | 3.162% | 0% | 1.503% | 0% | 3.368% | 3.534% | 0% | 2.578% | 1.552% | 5.993% | 0.9149% | 0% | 4.656% | 2.425% | 2.525% | 0.6516% | 4.781% | 2.247% | 3.32% | 0.4891%
| PDM3 |  | -0.854% | -2.877% | 0.2588% | 3.193% | -6.924% | -5.912% | 3.193% | -3.889% | -2.068% | 6.228% | -0.02912% | -8.948% | -1.866% | 2.181% | -1.866% | -16.8% | 6.228% | 10.27% | 0.1577% | 3.52%
| Z-Score (Absolute) |  | 0.08 | 0.46 | 0.13 | 0.69 | 1.23 | 1.04 | 0.69 | 0.66 | 0.31 | 1.27 | 0.08 | 1.62 | 0.27 | 0.5 | 0.27 | 3.12 | 1.27 | 2.04 | 0.11 | 0.75
| Analytical results for Mn in OREAS 612 (Certified Value 0.041 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 0.0403 | 0.0406 | 0.0416 | 0.041 | 0.0402 | 0.038 | 0.041 | 0.0409 | 0.0392 | 0.0425 | 0.04089 | 0.041 | 0.0425 | 0.0416 | 0.0372 | 0.040084 | 0.0412 | 0.0415 | 0.0416 | 0.041835 |  | 1
| 1 | 2 | 0.0408 | 0.0403 | 0.0418 | 0.04 | 0.0407 | 0.039 | 0.042 | 0.0417 | 0.0383 | 0.0411 | 0.040133 | 0.0408 | 0.0422 | 0.0422 | 0.0382 | 0.040292 | 0.0411 | 0.0443 | 0.0412 | 0.042075 |  | 24
| 1 | 3 | 0.0416 | 0.0404 | 0.0407 | 0.04 | 0.04 | 0.04 | 0.042 | 0.0413 | 0.0396 | 0.0396 | 0.040622 | 0.0411 | 0.0419 | 0.0417 | 0.0386 | 0.040068 | 0.0414 | 0.0411 | 0.0418 | 0.042043 |  | 16
| 1 | 4 | 0.0409 | 0.0406 | 0.0406 | 0.04 | 0.0409 | 0.039 | 0.041 | 0.042 | 0.0397 | 0.0394 | 0.040937 | 0.0411 | 0.042 | 0.0416 | 0.0397 | 0.040394 | 0.0408 | 0.0401 | 0.0423 | 0.041946 |  | 0.040858
| 1 | 5 | 0.0405 | 0.0413 | 0.0404 | 0.04 | 0.0401 | 0.039 | 0.042 | 0.0408 | 0.0384 | 0.0396 | 0.040496 | 0.0403 | 0.0424 | 0.0422 | 0.0383 | 0.040319 | 0.0418 | 0.0415 | 0.0405 | 0.041805 |  | 97
| 1 | 6 | 0.0418 | 0.0413 | 0.0419 | 0.041 | 0.0397 | 0.039 | 0.042 | 0.0407 | 0.0389 | 0.0389 | 0.040093 | 0.0407 | 0.0423 | 0.0424 | 0.0394 | 0.040193 | 0.0426 | 0.0415 | 0.041 | 0.04189
| Mean |  | 0.040983 | 0.04075 | 0.041167 | 0.040333 | 0.040267 | 0.039 | 0.041667 | 0.041233 | 0.039017 | 0.040183 | 0.040529 | 0.040833 | 0.042217 | 0.04195 | 0.038567 | 0.040225 | 0.041483 | 0.041667 | 0.0414 | 0.041932
| Median |  | 0.04085 | 0.0406 | 0.04115 | 0.04 | 0.04015 | 0.039 | 0.042 | 0.0411 | 0.03905 | 0.0396 | 0.040559 | 0.0409 | 0.04225 | 0.04195 | 0.03845 | 0.040242 | 0.0413 | 0.0415 | 0.0414 | 0.041918
| Std Dev. |  | 0.00059805 | 0.00044159 | 0.00067132 | 0.0005164 | 0.00045019 | 0.00063246 | 0.0005164 | 0.00052789 | 0.00059133 | 0.0013527 | 0.00036145 | 0.00030768 | 0.00023166 | 0.00035637 | 0.00090037 | 0.00013234 | 0.00064005 | 0.0013995 | 0.00063561 | 0.00010997
| Rel.Std.Dev. |  | 1.459% | 1.084% | 1.631% | 1.28% | 1.118% | 1.622% | 1.239% | 1.28% | 1.516% | 3.366% | 0.8918% | 0.7535% | 0.5487% | 0.8495% | 2.335% | 0.329% | 1.543% | 3.359% | 1.535% | 0.2623%
| PDM3 |  | 0.3058% | -0.2653% | 0.7545% | -1.285% | -1.448% | -4.548% | 1.978% | 0.9176% | -4.508% | -1.652% | -0.8071% | -0.06136% | 3.324% | 2.672% | -5.609% | -1.55% | 1.53% | 1.978% | 1.326% | 2.629%
| Z-Score (Absolute) |  | 0.08 | 0.16 | 0.26 | 0.59 | 0.65 | 1.94 | 0.77 | 0.33 | 1.93 | 0.74 | 0.39 | 0.08 | 1.33 | 1.06 | 2.38 | 0.7 | 0.59 | 0.77 | 0.5 | 1.04
| Analytical results for Mo in OREAS 612 (Certified Value 6.49 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 3
| 1 | 1 | 6.7 | 6.73 | 6.49 | 8 | 6.55 | 6 | 8 | 6.3 | 5 | 6.71 | 6.545 | 6 | 7.44 | 6.74 | 5.9 | 6.4 | 6.87 | 6.5 | 6.41 | 6.291 |  | 1
| 1 | 2 | 6.84 | 6.52 | 6.92 | 6 | 6.32 | 5 | 7 | 6.4 | 6 | 6.43 | 6.4648 | 5.9 | 7.58 | 6.72 | 6.3 | 6.4 | 6.84 | 6.6 | 6.29 | 5.52 |  | 25
| 1 | 3 | 6.76 | 6.51 | 6.41 | 7 | 6.58 | 5 | 8 | 6.4 | 6 | 6.61 | 6.437 | 6.1 | 7.99 | 6.77 | 6.2 | 6.3 | 6.89 | 6.5 | 6.39 | 5.776 |  | 16
| 1 | 4 | 6.92 | 6.78 | 6.59 | 8 | 6.54 | < 5 | 6 | 6.6 | 7 | 6.35 | 6.3717 | 5.9 | 7.93 | 6.65 | 6.2 | 6.6 | 6.79 | 6.4 | 6.49 | 6.528 |  | 6.488
| 1 | 5 | 6.84 | 7 | 6.59 | 8 | 6.41 | 7 | 7 | 6.4 | 6 | 6.04 | 6.6266 | 5.9 | 7.31 | 6.82 | 6.2 | 6.3 | 7 | 6.4 | 6.16 | 6.728 |  | 98
| 1 | 6 | 6.94 | 6.77 | 6.85 | 7 | 6.69 | 6 | 7 | 6.4 | 6 | 6.65 | 6.5305 | 6.1 | 8 | 6.76 | 6.1 | 6.2 | 6.89 | 6.6 | 6.1 | 6.1985
| Mean |  | 6.8333 | 6.7183 | 6.6417 | 7.3333 | 6.515 | 5.8 | 7.1667 | 6.4167 | 6 | 6.465 | 6.4959 | 5.9833 | 7.7083 | 6.7433 | 6.15 | 6.3667 | 6.88 | 6.5 | 6.3067 | 6.1736
| Median |  | 6.84 | 6.75 | 6.59 | 7.5 | 6.545 | 6 | 7 | 6.4 | 6 | 6.52 | 6.4977 | 5.95 | 7.755 | 6.75 | 6.2 | 6.35 | 6.88 | 6.5 | 6.34 | 6.2447
| Std Dev. |  | 0.091797 | 0.18368 | 0.20144 | 0.8165 | 0.13096 | 0.83666 | 0.75277 | 0.098319 | 0.63246 | 0.24914 | 0.090061 | 0.098319 | 0.30354 | 0.056804 | 0.13784 | 0.13663 | 0.069857 | 0.089443 | 0.15214 | 0.45445
| Rel.Std.Dev. |  | 1.343% | 2.734% | 3.033% | 11.13% | 2.01% | 14.43% | 10.5% | 1.532% | 10.54% | 3.854% | 1.386% | 1.643% | 3.938% | 0.8424% | 2.241% | 2.146% | 1.015% | 1.376% | 2.412% | 7.361%
| PDM3 |  | 5.322% | 3.55% | 2.368% | 13.03% | 0.4159% | -10.6% | 10.46% | -1.1% | -7.522% | -0.3548% | 0.1222% | -7.779% | 18.81% | 3.935% | -5.21% | -1.87% | 6.042% | 0.1847% | -2.795% | -4.846%
| Z-Score (Absolute) |  | 1.1 | 0.72 | 0.47 | NA | 0.06 | NA | NA | 0.27 | NA | 0.11 | 0.01 | 1.69 | 3.97 | 0.8 | 1.14 | 0.43 | 1.25 | 0.01 | 0.63 | 1.06
Results from laboratories 1.05, 1.07, 1.08 and 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 612 (Certified Value 0.06 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 3
| 1 | 1 | 0.06 | 0.06 | 0.062 | 0.08 | 0.06 | 0.05 | 0.05 | 0.06 | 0.056 | 0.06 | 0.06065 | 0.05 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 0.07 | 0.053931 |  |  | 1
| 1 | 2 | 0.06 | 0.06 | 0.065 | 0.08 | 0.06 | 0.06 | < 0.05 | 0.06 | 0.056 | 0.06 | 0.059448 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 0.06 | 0.053397 |  |  | #N/A
| 1 | 3 | 0.06 | 0.06 | 0.062 | 0.07 | 0.06 | 0.07 | 0.06 | 0.06 | 0.058 | 0.06 | 0.059532 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 0.052919 |  |  | 16
| 1 | 4 | 0.06 | 0.06 | 0.062 | 0.08 | 0.06 | 0.08 | < 0.05 | 0.06 | 0.052 | 0.07 | 0.05861 | 0.06 | 0.06 | 0.07 | 0.05 | 0.07 | 0.06 | 0.07 | 0.052004 |  |  | 0.060282
| 1 | 5 | 0.06 | 0.06 | 0.062 | 0.07 | 0.06 | 0.08 | 0.06 | 0.06 | 0.057 | 0.05 | 0.059937 | 0.05 | 0.06 | 0.07 | 0.05 | 0.07 | 0.07 | 0.07 | 0.054614 |  |  | 99
| 1 | 6 | 0.06 | 0.06 | 0.063 | 0.07 | 0.06 | 0.06 | 0.07 | 0.06 | 0.058 | 0.06 | 0.05887 | 0.05 | 0.06 | 0.07 | 0.06 | 0.06 | 0.07 | 0.06 | 0.053554
| Mean |  | 0.06 | 0.06 | 0.062667 | 0.075 | 0.06 | 0.066667 | 0.06 | 0.06 | 0.056167 | 0.06 | 0.059508 | 0.055 | 0.06 | 0.065 | 0.056667 | 0.063333 | 0.066667 | 0.066667 | 0.053403
| Median |  | 0.06 | 0.06 | 0.062 | 0.075 | 0.06 | 0.065 | 0.06 | 0.06 | 0.0565 | 0.06 | 0.05949 | 0.055 | 0.06 | 0.065 | 0.06 | 0.06 | 0.07 | 0.07 | 0.053475
| Std Dev. |  | 0 | 0 | 0.0012111 | 0.0054772 | 0 | 0.012111 | 0.008165 | 0 | 0.0022286 | 0.0063246 | 0.00073574 | 0.0054772 | 0 | 0.0054772 | 0.005164 | 0.005164 | 0.005164 | 0.005164 | 0.00089006
| Rel.Std.Dev. |  | 0% | 0% | 1.933% | 7.303% | 0% | 18.17% | 13.61% | 0% | 3.968% | 10.54% | 1.236% | 9.959% | 0% | 8.427% | 9.113% | 8.154% | 7.746% | 7.746% | 1.667%
| PDM3 |  | -0.468% | -0.468% | 3.956% | 24.42% | -0.468% | 10.59% | -0.468% | -0.468% | -6.827% | -0.468% | -1.284% | -8.762% | -0.468% | 7.826% | -5.998% | 5.062% | 10.59% | 10.59% | -11.41%
| Z-Score (Absolute) |  | 0 | 0 | 0.54 | 3.03 | 0 | 1.35 | 2.36 | 0 | 0.78 | 0 | 0.1 | 1.01 | 0 | 1.01 | 0.67 | 0.67 | 1.35 | 1.35 | 1.33
| Analytical results for Nb in OREAS 612 (Indicative Value 0.39 ppm) | [22 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | [6 empty] | 2
| 1 | 1 | 0.43 | 1.11 | 0.59 | 0.28 | 0.71 | 0.2 | 0.37371 | 0.32 | 0.89 | 0.39 | 0.2 | 0.23 | 0.42 | 0.3 | 12.7747 | [6 empty] | 1
| 1 | 2 | 0.48 | 0.53 | 0.51 | 0.27 | 0.71 | 0.2 | 0.37078 | 0.3 | 0.84 | 0.38 | 0.2 | 0.22 | 0.4 | 0.3 | 13.0125 | [6 empty] | 8
| 1 | 3 | 0.46 | 0.49 | 0.58 | 0.28 | 0.66 | 0.2 | 0.37227 | 0.29 | 0.85 | 0.39 | 0.2 | 0.22 | 0.38 | 0.3 | 11.5315 | [6 empty] | 16
| 1 | 4 | 0.42 | 0.61 | 0.5 | 0.28 | 0.65 | 0.2 | 0.39485 | 0.26 | 0.83 | 0.38 | 0.2 | 0.22 | 0.44 | 0.25 | 12.7935 | [6 empty] | 0.38863
| 1 | 5 | 0.45 | 0.49 | 0.53 | 0.29 | 0.59 | 0.2 | 0.38739 | 0.31 | 0.83 | 0.38 | 0.2 | 0.22 | 0.34 | 0.25 | 13.5147 | [6 empty] | 14
| 1 | 6 | 0.43 | 1.3 | 0.54 | 0.3 | 0.62 | 0.2 | 0.36871 | 0.31 | 1 | 0.39 | 0.2 | 0.22 | 0.38 | 0.3 | 12.6167
| Mean |  | 0.445 | 0.755 | 0.54167 | 0.28333 | 0.65667 | 0.2 | 0.37795 | 0.29833 | 0.87333 | 0.385 | 0.2 | 0.22167 | 0.39333 | 0.28333 | 12.7073
| Median |  | 0.44 | 0.57 | 0.535 | 0.28 | 0.655 | 0.2 | 0.37299 | 0.305 | 0.845 | 0.385 | 0.2 | 0.22 | 0.39 | 0.3 | 12.7841
| Std Dev. |  | 0.022583 | 0.35641 | 0.03656 | 0.010328 | 0.048028 | 3.0405e-17 | 0.010599 | 0.02137 | 0.065929 | 0.0054772 | 3.0405e-17 | 0.0040825 | 0.035024 | 0.02582 | 0.65539
| Rel.Std.Dev. |  | 5.075% | 47.21% | 6.75% | 3.645% | 7.314% | 1.52e-14% | 2.804% | 7.163% | 7.549% | 1.423% | 1.52e-14% | 1.842% | 8.904% | 9.113% | 5.158%
| PDM3 |  | 14.51% | 94.27% | 39.38% | -27.09% | 68.97% | -48.54% | -2.748% | -23.23% | 124.7% | -0.9336% | -48.54% | -42.96% | 1.211% | -27.09% | 3170%
| Z-Score (Absolute) |  | 0.32 | 2.22 | 0.91 | 0.67 | 1.61 | NA | 0.09 | 0.58 | 2.94 | 0.05 | NA | 1.05 | 0 | 0.67 | 75.49
Results from laboratories 1.11 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 612 (Certified Value 22.5 ppm) | [22 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 1
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 2
| 1 | 1 | 25.725 | 24.3 | 21.0271 | 20.8 | 25.49 | 20.0187 | 30.73 | [14 empty] | 1
| 1 | 2 | 25.018 | 23.2 | 21.2387 | 21.2 | 25.35 | 20.0803 | 31.77 | [14 empty] | 27
| 1 | 3 | 23.761 | 21.8 | 21.4056 | 21.1 | 25.28 | 20.0776 | 30.47 | [14 empty] | 16
| 1 | 4 | 25.424 | 22.2 | 21.6152 | 20.8 | 26.32 | 20.1051 | 30.46 | [14 empty] | 22.5337
| 1 | 5 | 24.409 | 22.4 | 20.5681 | 20.8 | 26.18 | 20.1776 | 30.29 | [14 empty] | 100
| 1 | 6 | 23.664 | 22.2 | 21.1221 | 20.7 | 25.62 | 20.0376 | 30.49
| Mean |  | 24.6668 | 22.6833 | 21.1628 | 20.9 | 25.7067 | 20.0828 | 30.7017
| Median |  | 24.7135 | 22.3 | 21.1804 | 20.8 | 25.555 | 20.0789 | 30.48
| Std Dev. |  | 0.86128 | 0.91742 | 0.35864 | 0.2 | 0.43907 | 0.055969 | 0.54194
| Rel.Std.Dev. |  | 3.492% | 4.044% | 1.695% | 0.9569% | 1.708% | 0.2787% | 1.765%
| PDM3 |  | 9.466% | 0.6639% | -6.084% | -7.25% | 14.08% | -10.88% | 36.25%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.52 | 0.61 | 1.03 | 0.88 | 2.73
| Analytical results for Ni in OREAS 612 (Certified Value 10.3 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 2
| 1 | 1 | 9.1 | 9.7 | 10.1 | 12 | 10.1 | 9 | 17 | 9.8 | 11 | 11.5 | 9.7544 | 11.2 | 9.2 | 10.8 | 10 | 11.5 | 10.5 | 10.4 | 12.8023 |  |  | 1
| 1 | 2 | 9.7 | 9.2 | 10.8 | 12 | 9.8 | 10 | 11 | 10 | 10 | 10.4 | 9.761 | 12.9 | 9.7 | 10.8 | 10.1 | 11 | 10.7 | 9.9 | 13.247 |  |  | 28
| 1 | 3 | 9 | 9.4 | 10.1 | 9 | 10.4 | 11 | 10 | 9.9 | 10 | 10.2 | 10.2692 | 12 | 9.7 | 10.6 | 9.8 | 10.4 | 11.8 | 10.8 | 13.0255 |  |  | 16
| 1 | 4 | 9.4 | 9.3 | 9.1 | 10 | 10 | 11 | 12 | 10.2 | 10 | 11.5 | 9.9168 | 11.2 | 9.3 | 10.9 | 10.7 | 10.4 | 11.1 | 10.1 | 12.4793 |  |  | 10.2688
| 1 | 5 | 9.6 | 9.7 | 10.6 | 11 | 10.1 | 10 | 10 | 9.8 | 10 | 9.8 | 10.142 | 10.9 | 9.9 | 10.8 | 10.6 | 10.3 | 10.8 | 9.8 | 13.0037 |  |  | 101
| 1 | 6 | 9.6 | 9.5 | 11 | 13 | 10.2 | 8 | 10 | 9.7 | 10 | 11.4 | 10.1469 | 11.4 | 9.2 | 10.7 | 9.6 | 10.7 | 10.3 | 10 | 13.213
| Mean |  | 9.4 | 9.4667 | 10.2833 | 11.1667 | 10.1 | 9.8333 | 11.6667 | 9.9 | 10.1667 | 10.8 | 9.9984 | 11.6 | 9.5 | 10.7667 | 10.1333 | 10.7167 | 10.8667 | 10.1667 | 12.9618
| Median |  | 9.5 | 9.45 | 10.35 | 11.5 | 10.1 | 10 | 10.5 | 9.85 | 10 | 10.9 | 10.0294 | 11.3 | 9.5 | 10.8 | 10.05 | 10.55 | 10.75 | 10.05 | 13.0146
| Std Dev. |  | 0.28983 | 0.20656 | 0.68532 | 1.472 | 0.2 | 1.169 | 2.7325 | 0.17889 | 0.40825 | 0.75631 | 0.21844 | 0.73485 | 0.30332 | 0.10328 | 0.43665 | 0.46224 | 0.53166 | 0.37238 | 0.28582
| Rel.Std.Dev. |  | 3.083% | 2.182% | 6.664% | 13.18% | 1.98% | 11.89% | 23.42% | 1.807% | 4.016% | 7.003% | 2.185% | 6.335% | 3.193% | 0.9593% | 4.309% | 4.313% | 4.893% | 3.663% | 2.205%
| PDM3 |  | -8.461% | -7.811% | 0.1415% | 8.744% | -1.644% | -4.241% | 13.61% | -3.591% | -0.9946% | 5.173% | -2.633% | 12.96% | -7.487% | 4.848% | -1.319% | 4.361% | 5.822% | -0.9946% | 26.23%
| Z-Score (Absolute) |  | 0.86 | 0.79 | 0.13 | 1.12 | 0.07 | 0.37 | 1.69 | 0.3 | 0 | 0.71 | 0.19 | 1.61 | 0.75 | 0.67 | 0.04 | 0.62 | 0.79 | 0 | 3.14
| Analytical results for P in OREAS 612 (Certified Value 0.037 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 3
| 1 | 1 | 0.0359 | 0.0331 | 0.035 | 0.04 | 0.036 | 0.036 | 0.039 | 0.0363 | 0.04 | 0.037 | 0.036889 | 0.037 | 0.0352 | 0.036 | 0.035 | 0.0364 | 0.0387 | 0.0348 | 0.046181 |  |  | 1
| 1 | 2 | 0.0362 | 0.0327 | 0.036 | 0.039 | 0.037 | 0.037 | 0.039 | 0.0361 | 0.039 | 0.036 | 0.036664 | 0.036 | 0.0355 | 0.036 | 0.035 | 0.037 | 0.04 | 0.0346 | 0.047372 |  |  | #N/A
| 1 | 3 | 0.0368 | 0.0337 | 0.035 | 0.04 | 0.036 | 0.037 | 0.041 | 0.0356 | 0.041 | 0.037 | 0.036086 | 0.037 | 0.0348 | 0.036 | 0.036 | 0.0366 | 0.0391 | 0.0348 | 0.045969 |  |  | 16
| 1 | 4 | 0.0361 | 0.0331 | 0.035 | 0.039 | 0.037 | 0.037 | 0.04 | 0.035 | 0.037 | 0.036 | 0.036307 | 0.037 | 0.0353 | 0.036 | 0.037 | 0.038 | 0.0378 | 0.0355 | 0.046714 |  |  | 0.036671
| 1 | 5 | 0.036 | 0.0342 | 0.035 | 0.039 | 0.036 | 0.036 | 0.039 | 0.0373 | 0.039 | 0.036 | 0.036423 | 0.035 | 0.0353 | 0.037 | 0.036 | 0.0387 | 0.0392 | 0.0336 | 0.046489 |  |  | 102
| 1 | 6 | 0.0367 | 0.0338 | 0.035 | 0.04 | 0.036 | 0.036 | 0.04 | 0.0363 | 0.039 | 0.037 | 0.035988 | 0.036 | 0.0349 | 0.037 | 0.038 | 0.0379 | 0.0392 | 0.0343 | 0.046411
| Mean |  | 0.036283 | 0.033433 | 0.035167 | 0.0395 | 0.036333 | 0.0365 | 0.039667 | 0.0361 | 0.039167 | 0.0365 | 0.036393 | 0.036333 | 0.035167 | 0.036333 | 0.036167 | 0.037433 | 0.039 | 0.0346 | 0.046523
| Median |  | 0.03615 | 0.0334 | 0.035 | 0.0395 | 0.036 | 0.0365 | 0.0395 | 0.0362 | 0.039 | 0.0365 | 0.036365 | 0.0365 | 0.03525 | 0.036 | 0.036 | 0.03745 | 0.03915 | 0.0347 | 0.04645
| Std Dev. |  | 0.00037639 | 0.00055737 | 0.00040825 | 0.00054772 | 0.0005164 | 0.00054772 | 0.0008165 | 0.00077201 | 0.0013292 | 0.00054772 | 0.00034225 | 0.0008165 | 0.00026583 | 0.0005164 | 0.001169 | 0.0009048 | 0.00072388 | 0.00062929 | 0.00048895
| Rel.Std.Dev. |  | 1.037% | 1.667% | 1.161% | 1.387% | 1.421% | 1.501% | 2.058% | 2.139% | 3.394% | 1.501% | 0.9404% | 2.247% | 0.7559% | 1.421% | 3.232% | 2.417% | 1.856% | 1.819% | 1.051%
| PDM3 |  | -1.057% | -8.829% | -4.102% | 7.715% | -0.9205% | -0.466% | 8.169% | -1.557% | 6.806% | -0.466% | -0.7582% | -0.9205% | -4.102% | -0.9205% | -1.375% | 2.079% | 6.351% | -5.647% | 26.87%
| Z-Score (Absolute) |  | 0.03 | 1.78 | 0.72 | 1.94 | 0 | 0.1 | 2.04 | 0.14 | 1.74 | 0.1 | 0.04 | 0 | 0.72 | 0 | 0.1 | 0.67 | 1.63 | 1.06 | 6.25
| Analytical results for Pb in OREAS 612 (Certified Value 0.248 wt.%) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 3
| 1 | 1 | 0.2508 | 0.23912 | 0.2392 | 0.248 | 0.253 | 0.233 | 0.253 | 0.25637 | 0.2483 | 0.245 | 0.2407 | 0.2396 | 0.255 | 0.2744 | 0.252 | 0.23203 | 0.22928 | 0.24658 | 0.25197 | 0.2557 | 0.32053 | 1
| 1 | 2 | 0.25154 | 0.23837 | 0.2469 | 0.247 | 0.255 | 0.237 | 0.254 | 0.25078 | 0.2425 | 0.234 | 0.24068 | 0.2348 | 0.252 | 0.2617 | 0.254 | 0.22968 | 0.2289 | 0.24637 | 0.25347 | 0.253 | 0.31791 | 29
| 1 | 3 | 0.25461 | 0.23914 | 0.2404 | 0.25 | 0.249 | 0.241 | 0.258 | 0.24288 | 0.2515 | 0.257 | 0.23906 | 0.2398 | 0.258 | 0.2536 | 0.246 | 0.22619 | 0.228 | 0.24641 | 0.24583 | 0.2559 | 0.31124 | 16
| 1 | 4 | 0.2517 | 0.24099 | 0.2326 | 0.251 | 0.255 | 0.241 | 0.253 | 0.25246 | 0.2477 | 0.249 | 0.24388 | 0.2383 | 0.257 | 0.255 | 0.254 | 0.22977 | 0.22933 | 0.24792 | 0.24603 | 0.2598 | 0.31416 | 0.24768
| 1 | 5 | 0.25044 | 0.24305 | 0.2449 | 0.248 | 0.254 | 0.24 | 0.25 | 0.25035 | 0.2415 | 0.228 | 0.23778 | 0.2365 | 0.253 | 0.2521 | 0.252 | 0.22706 | 0.22347 | 0.25128 | 0.24737 | 0.2461 | 0.30308 | 103
| 1 | 6 | 0.25758 | 0.24195 | 0.2439 | 0.248 | 0.249 | 0.236 | 0.256 | 0.24691 | 0.2461 | 0.25 | 0.23921 | 0.2382 | 0.253 | 0.2507 | 0.251 | 0.22782 | 0.22739 | 0.25199 | 0.25147 | 0.2521 | 0.30199
| Mean |  | 0.25278 | 0.24044 | 0.24132 | 0.24867 | 0.2525 | 0.238 | 0.254 | 0.24996 | 0.24627 | 0.24383 | 0.24022 | 0.23787 | 0.25467 | 0.25792 | 0.2515 | 0.22876 | 0.22773 | 0.24843 | 0.24936 | 0.25377 | 0.31148
| Median |  | 0.25162 | 0.24007 | 0.24215 | 0.248 | 0.2535 | 0.2385 | 0.2535 | 0.25057 | 0.2469 | 0.247 | 0.23994 | 0.23825 | 0.254 | 0.2543 | 0.252 | 0.22875 | 0.22845 | 0.24725 | 0.24942 | 0.25435 | 0.3127
| Std Dev. |  | 0.0027745 | 0.00185 | 0.0051379 | 0.0015055 | 0.0028107 | 0.0032249 | 0.0027568 | 0.0046388 | 0.0037554 | 0.010834 | 0.0021038 | 0.0019138 | 0.0024221 | 0.0089338 | 0.0029496 | 0.0021419 | 0.0022219 | 0.0025613 | 0.0033366 | 0.0046224 | 0.0076316
| Rel.Std.Dev. |  | 1.098% | 0.7694% | 2.129% | 0.6054% | 1.113% | 1.355% | 1.085% | 1.856% | 1.525% | 4.443% | 0.8758% | 0.8046% | 0.9511% | 3.464% | 1.173% | 0.9363% | 0.9757% | 1.031% | 1.338% | 1.822% | 2.45%
| PDM3 |  | 2.06% | -2.923% | -2.568% | 0.3998% | 1.947% | -3.907% | 2.553% | 0.9213% | -0.5692% | -1.552% | -3.011% | -3.961% | 2.822% | 4.134% | 1.544% | -7.638% | -8.054% | 0.3022% | 0.6784% | 2.459% | 25.76%
| Z-Score (Absolute) |  | 0.52 | 1.04 | 0.93 | 0 | 0.48 | 1.35 | 0.67 | 0.16 | 0.3 | 0.61 | 1.07 | 1.37 | 0.76 | 1.17 | 0.36 | 2.52 | 2.65 | 0.03 | 0.09 | 0.64 | 7.94
| Analytical results for Pd in OREAS 612 (Indicative Value < 10 ppb) | [22 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 1.09 | 1.22 | [19 empty] | ppb
|  |  | AR*MS | AR*MS | [19 empty] | 1
|  |  | 25g | 01g | [19 empty] | 1
| 1 | 1 | < 10 | < 10 | [19 empty] | 1
| 1 | 2 | < 10 | < 10 | [19 empty] | 6
| 1 | 3 | < 10 | < 10 | [19 empty] | 16
| 1 | 4 | < 10 | < 10 | [19 empty] | < 10
| 1 | 5 | 14 | < 10 | [19 empty] | 12
| 1 | 6 | 11 | < 10
| Mean |  | 12.5
| Median |  | 12.5
| Std Dev. |  | 2.1213
| Rel.Std.Dev. |  | 16.97%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 612 (Certified Value 5.78 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 3
| 1 | 1 | 6.82 | 6.5 | 5.6539 | 5.15 | 5.94 | 5.3129 | 8.17 | [14 empty] | 1
| 1 | 2 | 6.639 | 6.2 | 5.5443 | 5.12 | 5.97 | 5.2627 | 8.38 | [14 empty] | 30
| 1 | 3 | 6.404 | 6.1 | 5.6396 | 5.05 | 5.98 | 5.2356 | 8.04 | [14 empty] | 16
| 1 | 4 | 6.706 | 5.9 | 5.572 | 5 | 6.13 | 5.2375 | 8.14 | [14 empty] | 5.779
| 1 | 5 | 6.501 | 5.9 | 5.4581 | 5.02 | 6.07 | 5.3223 | 8.1 | [14 empty] | 104
| 1 | 6 | 6.376 | 6.1 | 5.62 | 5.08 | 6.26 | 5.2701 | 8.27
| Mean |  | 6.5743 | 6.1167 | 5.5813 | 5.07 | 6.0583 | 5.2735 | 8.1833
| Median |  | 6.57 | 6.1 | 5.596 | 5.065 | 6.025 | 5.2664 | 8.155
| Std Dev. |  | 0.17641 | 0.22286 | 0.07322 | 0.057966 | 0.12156 | 0.036864 | 0.12307
| Rel.Std.Dev. |  | 2.683% | 3.643% | 1.312% | 1.143% | 2.006% | 0.699% | 1.504%
| PDM3 |  | 13.76% | 5.843% | -3.421% | -12.27% | 4.833% | -8.748% | 41.6%
| Z-Score (Absolute) |  | 0.67 | 0.08 | 0.62 | 1.29 | 0 | 1.03 | 2.78
| Analytical results for Pt in OREAS 612 (Indicative Value < 5 ppb) | [22 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 1.09 | 1.22 | [19 empty] | ppb
|  |  | AR*MS | AR*MS | [19 empty] | 2
|  |  | 25g | 01g | [19 empty] | 2
| 1 | 1 | < 5 | < 5 | [19 empty] | 1
| 1 | 2 | < 5 | < 5 | [19 empty] | 7
| 1 | 3 | < 5 | < 5 | [19 empty] | 16
| 1 | 4 | < 5 | < 5 | [19 empty] | < 5
| 1 | 5 | < 5 | < 5 | [19 empty] | 13
| 1 | 6 | < 5 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Rb in OREAS 612 (Certified Value 23.5 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 1
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | [6 empty] | 2
| 1 | 1 | 22.3 | 39.9 | 24.05 | 22.7 | 24.59 | 25.4 | 24.6627 | 22 | 22.8 | 28.5 | 23.8 | 28.6387 | 23.5 | 39.85 | 21.4 | [6 empty] | 1
| 1 | 2 | 21.9 | 29.5 | 23.65 | 22.7 | 23.68 | 23.8 | 24.1098 | 21.9 | 23.4 | 28.9 | 24.3 | 28.7509 | 23.6 | 41.65 | 21.6 | [6 empty] | 31
| 1 | 3 | 22.8 | 26.5 | 23.48 | 23.7 | 22.95 | 23.2 | 24.5996 | 23.1 | 23.9 | 29.2 | 24.4 | 28.4537 | 23.5 | 39.42 | 22 | [6 empty] | 16
| 1 | 4 | 22.5 | 25.7 | 24.18 | 22.4 | 24.19 | 22.7 | 23.9719 | 22.8 | 23.5 | 29.8 | 24.7 | 28.378 | 23.4 | 39.44 | 22.5 | [6 empty] | 23.544
| 1 | 5 | 22.7 | 25 | 24.01 | 23 | 23.49 | 23.7 | 24.5906 | 21.7 | 23.6 | 30 | 24.9 | 28.6729 | 23.8 | 39.26 | 21.2 | [6 empty] | 105
| 1 | 6 | 21.6 | 24.9 | 24.61 | 23.7 | 22.91 | 23.9 | 23.9947 | 23.8 | 23.9 | 30 | 24.9 | 28.8561 | 24.3 | 39.98 | 21.4
| Mean |  | 22.3 | 28.5833 | 23.9967 | 23.0333 | 23.635 | 23.7833 | 24.3215 | 22.55 | 23.5167 | 29.4 | 24.5 | 28.625 | 23.6833 | 39.9333 | 21.6833
| Median |  | 22.4 | 26.1 | 24.03 | 22.85 | 23.585 | 23.75 | 24.3502 | 22.4 | 23.55 | 29.5 | 24.55 | 28.6558 | 23.55 | 39.645 | 21.5
| Std Dev. |  | 0.46904 | 5.796 | 0.39998 | 0.55015 | 0.66902 | 0.91086 | 0.32865 | 0.82158 | 0.40702 | 0.62929 | 0.42426 | 0.18006 | 0.33116 | 0.88502 | 0.48339
| Rel.Std.Dev. |  | 2.103% | 20.28% | 1.667% | 2.389% | 2.831% | 3.83% | 1.351% | 3.643% | 1.731% | 2.14% | 1.732% | 0.629% | 1.398% | 2.216% | 2.229%
| PDM3 |  | -5.284% | 21.4% | 1.923% | -2.169% | 0.3864% | 1.016% | 3.302% | -4.222% | -0.1162% | 24.87% | 4.06% | 21.58% | 0.5917% | 69.61% | -7.903%
| Z-Score (Absolute) |  | 1.33 | 4.32 | 0.19 | 0.67 | 0.13 | 0 | 0.48 | 1.11 | 0.24 | 5.05 | 0.64 | 4.35 | 0.09 | 14.52 | 1.89
| Analytical results for Re in OREAS 612 (Indicative Value 0.001 ppm) | [22 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.2g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | [9 empty] | 3
| 1 | 1 | 0.005 | < 0.1 | 0.001 | < 0.001 | < 0.001 | < 0.1 | 0.001 | 0.001 | < 0.001 | < 0.005 | < 0.002 | < 0.001 | [9 empty] | 1
| 1 | 2 | 0.006 | < 0.1 | 0.001 | < 0.001 | < 0.001 | < 0.1 | 0.001 | 0.001 | < 0.001 | < 0.005 | < 0.002 | 0.001 | [9 empty] | 4
| 1 | 3 | 0.004 | < 0.1 | 0.001 | < 0.001 | < 0.001 | < 0.1 | < 0.001 | 0.001 | < 0.001 | < 0.005 | < 0.002 | 0.002 | [9 empty] | 16
| 1 | 4 | 0.006 | < 0.1 | 0.001 | < 0.001 | < 0.001 | < 0.1 | 0.001 | 0.001 | < 0.001 | < 0.005 | < 0.002 | < 0.001 | [9 empty] | 0.00087917
| 1 | 5 | 0.005 | < 0.1 | 0.001 | < 0.001 | < 0.001 | < 0.1 | 0.001 | 0.001 | < 0.001 | < 0.005 | < 0.002 | < 0.001 | [9 empty] | 14
| 1 | 6 | 0.004 | < 0.1 | 0.001 | < 0.001 | < 0.001 | < 0.1 | 0.001 | 0.001 | < 0.001 | < 0.005 | < 0.002 | < 0.001
| Mean |  | 0.005 |  | 0.001 |  |  |  | 0.001 | 0.001 |  |  |  | 0.0015
| Median |  | 0.005 |  | 0.001 |  |  |  | 0.001 | 0.001 |  |  |  | 0.0015
| Std Dev. |  | 0.00089443 |  | 0 |  |  |  | 0 | 0 |  |  |  | 0.00070711
| Rel.Std.Dev. |  | 17.89% |  | 0% |  |  |  | 0% | 0% |  |  |  | 47.14%
| PDM3 |  | 468.7% |  | 13.74% |  |  |  | 13.74% | 13.74% |  |  |  | 70.62%
| Z-Score (Absolute) |  | 5.39 | 66.08 | 0 | 0.67 | 0.67 | 66.08 | 0.11 | 0 | 0.67 | 2.02 | 0 | 0.22
| Analytical results for S in OREAS 612 (Certified Value 0.515 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 3
| 1 | 1 | 0.51 | 0.51 | 0.54 | 0.55 | 0.52 | 0.49 | 0.46 | 0.46 | 0.53 | 0.51 | 0.51912 | 0.5291 | 0.52 | 0.51 | 0.492 | 0.49 | 0.50445 | 0.53 | 0.5069 | 0.52 | 0.53 | 1
| 1 | 2 | 0.51 | 0.51 | 0.54 | 0.53 | 0.52 | 0.49 | 0.46 | 0.5 | 0.51 | 0.52 | 0.51478 | 0.5199 | 0.52 | 0.5 | 0.501 | 0.49 | 0.50697 | 0.51 | 0.5402 | 0.52 | 0.52 | 10
| 1 | 3 | 0.51 | 0.51 | 0.53 | 0.53 | 0.51 | 0.51 | 0.48 | 0.5 | 0.53 | 0.51 | 0.5109 | 0.5255 | 0.52 | 0.5 | 0.495 | 0.49 | 0.50374 | 0.53 | 0.504 | 0.52 | 0.52 | 16
| 1 | 4 | 0.51 | 0.52 | 0.53 | 0.53 | 0.52 | 0.51 | 0.45 | 0.53 | 0.53 | 0.51 | 0.50545 | 0.5258 | 0.53 | 0.5 | 0.502 | 0.49 | 0.50493 | 0.52 | 0.4907 | 0.53 | 0.52 | 0.5146
| 1 | 5 | 0.51 | 0.53 | 0.53 | 0.53 | 0.52 | 0.5 | 0.48 | 0.52 | 0.51 | 0.53 | 0.51404 | 0.518 | 0.51 | 0.51 | 0.494 | 0.5 | 0.50226 | 0.53 | 0.508 | 0.5 | 0.53 | 106
| 1 | 6 | 0.52 | 0.52 | 0.55 | 0.53 | 0.52 | 0.5 | 0.49 | 0.52 | 0.52 | 0.53 | 0.51256 | 0.526 | 0.51 | 0.51 | 0.505 | 0.5 | 0.50804 | 0.51 | 0.505 | 0.51 | 0.52
| Mean |  | 0.51167 | 0.51667 | 0.53667 | 0.53333 | 0.51833 | 0.5 | 0.47 | 0.505 | 0.52167 | 0.51833 | 0.51281 | 0.52405 | 0.51833 | 0.505 | 0.49817 | 0.49333 | 0.50506 | 0.52167 | 0.50913 | 0.51667 | 0.52333
| Median |  | 0.51 | 0.515 | 0.535 | 0.53 | 0.52 | 0.5 | 0.47 | 0.51 | 0.525 | 0.515 | 0.5133 | 0.52565 | 0.52 | 0.505 | 0.498 | 0.49 | 0.50469 | 0.525 | 0.50595 | 0.52 | 0.52
| Std Dev. |  | 0.0040825 | 0.008165 | 0.008165 | 0.008165 | 0.0040825 | 0.0089443 | 0.015492 | 0.0251 | 0.0098319 | 0.0098319 | 0.0045415 | 0.0042023 | 0.0075277 | 0.0054772 | 0.0051929 | 0.005164 | 0.0021228 | 0.0098319 | 0.01646 | 0.010328 | 0.005164
| Rel.Std.Dev. |  | 0.7979% | 1.58% | 1.521% | 1.531% | 0.7876% | 1.789% | 3.296% | 4.97% | 1.885% | 1.897% | 0.8856% | 0.8019% | 1.452% | 1.085% | 1.042% | 1.047% | 0.4203% | 1.885% | 3.233% | 1.999% | 0.9867%
| PDM3 |  | -0.5701% | 0.4015% | 4.288% | 3.64% | 0.7254% | -2.837% | -8.667% | -1.866% | 1.373% | 0.7254% | -0.348% | 1.836% | 0.7254% | -1.866% | -3.194% | -4.133% | -1.853% | 1.373% | -1.062% | 0.4015% | 1.697%
| Z-Score (Absolute) |  | 0.46 | 0 | 1.83 | 1.52 | 0.15 | 1.52 | 4.26 | 1.07 | 0.46 | 0.15 | 0.35 | 0.67 | 0.15 | 1.07 | 1.69 | 2.13 | 1.06 | 0.46 | 0.69 | 0 | 0.61
| Analytical results for Sb in OREAS 612 (Certified Value 10.5 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES |  |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g |  |  | 1
| 1 | 1 | 10.24 | 11.98 | 10.47 | 10 | 11.95 | 10 | 10 | 11.81 | 9 | 9.79 | 10.4105 | 9.14 | 10.4 | 11 | 9.61 | 8.83 | 9.84 | 9 | 14.151 |  |  | 1
| 1 | 2 | 10.69 | 11.78 | 9.99 | 20 | 11.9 | 10 | 10 | 11.87 | 9 | 10.3 | 10.3665 | 8.91 | 10.45 | 11 | 9.53 | 8.77 | 9.92 | 9.09 | 11.5597 |  |  | 33
| 1 | 3 | 10.17 | 11.94 | 10.59 | 20 | 12.3 | 10 | 10 | 11.78 | 9 | 9.49 | 10.2645 | 9.24 | 11.25 | 10 | 9.49 | 8.7 | 9.77 | 8.87 | 13.5463 |  |  | 16
| 1 | 4 | 10.3 | 12.03 | 10.49 | 10 | 12.05 | 10 | < 10 | 11.97 | 12 | 9.73 | 10.4308 | 8.98 | 10.25 | 10 | 9.49 | 8.85 | 9.82 | 8.67 | 11.489 |  |  | 10.4893
| 1 | 5 | 10.98 | 12.08 | 10.79 | 10 | 12 | 10 | 10 | 11.68 | 8 | 10.4 | 10.3573 | 9.18 | 10.3 | 11 | 9.29 | 8.64 | 9.99 | 8.79 | 17.714 |  |  | 107
| 1 | 6 | 10.72 | 12.03 | 11.11 | 10 | 12.2 | 10 | < 10 | 11.82 | 6 | 9.78 | 10.3344 | 9.36 | 11.15 | 11 | 9.59 | 8.75 | 9.76 | 8.88 | 15.441
| Mean |  | 10.5167 | 11.9733 | 10.5733 | 13.3333 | 12.0667 | 10 | 10 | 11.8217 | 8.8333 | 9.915 | 10.3607 | 9.135 | 10.6333 | 10.6667 | 9.5 | 8.7567 | 9.85 | 8.8833 | 13.9835
| Median |  | 10.495 | 12.005 | 10.54 | 10 | 12.025 | 10 | 10 | 11.815 | 9 | 9.785 | 10.3619 | 9.16 | 10.425 | 11 | 9.51 | 8.76 | 9.83 | 8.875 | 13.8487
| Std Dev. |  | 0.32549 | 0.10614 | 0.37233 | 5.164 | 0.15384 | 0 | 0 | 0.096212 | 1.9408 | 0.35557 | 0.058962 | 0.16634 | 0.44572 | 0.5164 | 0.11437 | 0.078909 | 0.089443 | 0.14882 | 2.3815
| Rel.Std.Dev. |  | 3.095% | 0.8865% | 3.521% | 38.73% | 1.275% | 0% | 0% | 0.8139% | 21.97% | 3.586% | 0.5691% | 1.821% | 4.192% | 4.841% | 1.204% | 0.9011% | 0.908% | 1.675% | 17.03%
| PDM3 |  | 0.2613% | 14.15% | 0.8016% | 27.11% | 15.04% | -4.664% | -4.664% | 12.7% | -15.79% | -5.475% | -1.226% | -12.91% | 1.374% | 1.691% | -9.431% | -16.52% | -6.094% | -15.31% | 33.31%
| Z-Score (Absolute) |  | 0 | 0.97 | 0.04 | NA | 1.03 | NA | NA | 0.87 | NA | 0.4 | 0.1 | 0.92 | 0.08 | 0.1 | 0.67 | 1.17 | 0.44 | 1.08 | 2.3
Results from laboratory 1.10 were removed due to their 1 ppm reading resolution. Results from laboratories 1.05, 1.07 and 1.08 were removed due to their 10 ppm reading resolution.
| Analytical results for Sc in OREAS 612 (Certified Value 1.61 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  | 2
| 1 | 1 | 1.6 | 1.4 | 1.5 | < 5 | 1.6 | < 5 | < 5 | 1.7 | 1.4 | 1.9 | 1.6035 | 1.8 | 1.6 | 2 | 1.7 | 1.7381 | 1.6 | 2.4 | 1.2 | 1.3547 |  | 1
| 1 | 2 | 1.7 | 1.3 | 1.6 | < 5 | 1.6 | < 5 | < 5 | 1.7 | 1.3 | 1.8 | 1.5653 | 1.9 | 1.7 | 1.9 | 1.7 | 1.7682 | 1.4 | 2.5 | 1.2 | 1.3437 |  | 34
| 1 | 3 | 1.8 | 1.3 | 1.5 | < 5 | 1.6 | < 5 | < 5 | 1.6 | 1.4 | 1.7 | 1.5667 | 2 | 1.6 | 2 | 1.7 | 1.679 | 1.5 | 2.4 | 1.25 | 1.3343 |  | 16
| 1 | 4 | 1.6 | 1.3 | 1.6 | < 5 | 1.6 | < 5 | < 5 | 1.7 | 1.4 | 1.8 | 1.5459 | 2 | 1.6 | 2 | 1.7 | 1.7567 | 1.5 | 2.3 | 1.25 | 1.347 |  | 1.6053
| 1 | 5 | 1.7 | 1.3 | 1.6 | < 5 | 1.6 | < 5 | < 5 | 1.7 | 1.4 | 1.7 | 1.5827 | 2 | 1.7 | 1.9 | 1.7 | 1.8199 | 1.5 | 2.4 | 1.2 | 1.3723 |  | 108
| 1 | 6 | 1.7 | 1.3 | 1.6 | < 5 | 1.7 | < 5 | < 5 | 1.6 | 1.4 | 1.9 | 1.5808 | 2.1 | 1.5 | 2 | 1.7 | 1.7011 | 1.6 | 2.4 | 1.2 | 1.346
| Mean |  | 1.6833 | 1.3167 | 1.5667 |  | 1.6167 |  |  | 1.6667 | 1.3833 | 1.8 | 1.5742 | 1.9667 | 1.6167 | 1.9667 | 1.7 | 1.7438 | 1.5167 | 2.4 | 1.2167 | 1.3497
| Median |  | 1.7 | 1.3 | 1.6 |  | 1.6 |  |  | 1.7 | 1.4 | 1.8 | 1.5738 | 2 | 1.6 | 2 | 1.7 | 1.7474 | 1.5 | 2.4 | 1.2 | 1.3465
| Std Dev. |  | 0.075277 | 0.040825 | 0.05164 |  | 0.040825 |  |  | 0.05164 | 0.040825 | 0.089443 | 0.019542 | 0.10328 | 0.075277 | 0.05164 | 0 | 0.050215 | 0.075277 | 0.063246 | 0.02582 | 0.012888
| Rel.Std.Dev. |  | 4.472% | 3.101% | 3.296% |  | 2.525% |  |  | 3.098% | 2.951% | 4.969% | 1.241% | 5.252% | 4.656% | 2.626% | 0% | 2.88% | 4.963% | 2.635% | 2.122% | 0.9549%
| PDM3 |  | 4.863% | -17.98% | -2.405% |  | 0.71% |  |  | 3.825% | -13.83% | 12.13% | -1.938% | 22.51% | 0.71% | 22.51% | 5.901% | 8.631% | -5.519% | 49.51% | -24.21% | -15.92%
| Z-Score (Absolute) |  | 0.02 | 1.07 | 0.32 | 2.47 | 0.17 | 2.47 | 2.47 | 0.02 | 0.87 | 0.37 | 0.3 | 0.87 | 0.17 | 0.87 | 0.07 | 0.21 | 0.47 | 2.17 | 1.37 | 0.98
| Analytical results for Se in OREAS 612 (Indicative Value 0.81 ppm) | [22 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | [7 empty] | 2
| 1 | 1 | 0.7 | < 1 | < 1 | 0.8 | < 1 | 0.6 | < 1.5 | 0.65 | 0.9 | 2 | < 1 | 1.4 | < 1 | < 1 | [7 empty] | 1
| 1 | 2 | 0.7 | < 1 | < 1 | 0.7 | < 1 | 0.6 | < 1.5 | 0.61 | 0.9 | 2 | < 1 | 1.2 | < 1 | < 1 | [7 empty] | 3
| 1 | 3 | 0.7 | < 1 | < 1 | 0.6 | < 1 | 0.9 | < 1.5 | 0.6 | 1 | 2 | < 1 | 0.9 | < 1 | < 1 | [7 empty] | 16
| 1 | 4 | 0.8 | < 1 | < 1 | 0.7 | < 1 | 0.9 | < 1.5 | 0.64 | 0.8 | 2 | < 1 | 1.1 | < 1 | < 1 | [7 empty] | 0.80583
| 1 | 5 | 0.8 | < 1 | < 1 | 0.7 | < 1 | 0.5 | < 1.5 | 0.61 | 0.7 | 2 | < 1 | 1.1 | < 1 | < 1 | [7 empty] | 12
| 1 | 6 | 0.7 | < 1 | < 1 | 0.7 | < 1 | 1.1 | < 1.5 | 0.6 | 1 | 2 | < 1 | 1.1 | < 1 | < 1
| Mean |  | 0.73333 |  |  | 0.7 |  | 0.76667 |  | 0.61833 | 0.88333 | 2 |  | 1.1333
| Median |  | 0.7 |  |  | 0.7 |  | 0.75 |  | 0.61 | 0.9 | 2 |  | 1.1
| Std Dev. |  | 0.05164 |  |  | 0.063246 |  | 0.23381 |  | 0.02137 | 0.1169 | 0 |  | 0.1633
| Rel.Std.Dev. |  | 7.042% |  |  | 9.035% |  | 30.5% |  | 3.456% | 13.23% | 0% |  | 14.41%
| PDM3 |  | -8.997% |  |  | -13.13% |  | -4.86% |  | -23.27% | 9.617% | 148.2% |  | 40.64%
| Z-Score (Absolute) |  | 0.31 | 0.67 | 0.67 | 0.17 | 0.67 | 0.46 | 0.38 | 0.17 | 0.95 | 5.68 | 0.67 | 2.01 | 0.67 | 0.67
| Analytical results for Sm in OREAS 612 (Certified Value 4.28 ppm) | [22 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | [14 empty] | 3
| 1 | 1 | 5.109 | 4.5 | 4.2058 | 3.88 | 4.57 | 3.7755 | 6.04 | [14 empty] | 1
| 1 | 2 | 4.828 | 4.5 | 4.0394 | 4.08 | 4.58 | 3.8462 | 6.28 | [14 empty] | 35
| 1 | 3 | 4.547 | 4.4 | 4.1298 | 4.07 | 4.47 | 3.8693 | 5.99 | [14 empty] | 16
| 1 | 4 | 4.799 | 4.2 | 4.16 | 4.09 | 4.56 | 3.7873 | 6.03 | [14 empty] | 4.2782
| 1 | 5 | 4.677 | 4.3 | 4.0195 | 3.97 | 4.63 | 3.8248 | 5.9 | [14 empty] | 109
| 1 | 6 | 4.539 | 4.4 | 4.0576 | 3.93 | 4.8 | 3.8722 | 5.94
| Mean |  | 4.7498 | 4.3833 | 4.102 | 4.0033 | 4.6017 | 3.8292 | 6.03
| Median |  | 4.738 | 4.4 | 4.0937 | 4.02 | 4.575 | 3.8355 | 6.01
| Std Dev. |  | 0.21374 | 0.1169 | 0.074311 | 0.088919 | 0.11017 | 0.040989 | 0.13357
| Rel.Std.Dev. |  | 4.5% | 2.667% | 1.812% | 2.221% | 2.394% | 1.07% | 2.215%
| PDM3 |  | 11.02% | 2.457% | -4.119% | -6.426% | 7.56% | -10.5% | 40.95%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.52 | 0.7 | 0.4 | 1.02 | 3.03
| Analytical results for Sn in OREAS 612 (Certified Value 0.98 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | [5 empty] | 2
| 1 | 1 | 1 | 0.9 | 0.8 | 0.8 | 0.98 | < 10 | 1.2 | 0.95027 | 1 | 1.1 | 0.96 | 0.9 | 0.9 | 1.24 | 0.65 | 3.4773 | [5 empty] | 1
| 1 | 2 | 1 | 0.9 | 0.9 | 0.8 | 1.03 | < 10 | 1.17 | 0.95031 | 1.1 | 1.2 | 1.01 | 0.9 | 1 | 1.26 | 0.65 | 3.283 | [5 empty] | 19
| 1 | 3 | 1.1 | 0.9 | 0.8 | 0.8 | 0.93 | < 10 | 1.12 | 0.9117 | 1.1 | 1.2 | 1.06 | 0.9 | 0.9 | 1.24 | 0.7 | 3.543 | [5 empty] | 16
| 1 | 4 | 1.1 | 0.9 | 0.8 | 0.8 | 1.06 | < 10 | 1.1 | 0.93465 | 1 | 1.2 | 1.11 | 0.9 | 0.9 | 1.24 | 0.7 | 4.493 | [5 empty] | 0.97504
| 1 | 5 | 1.2 | 0.9 | 0.8 | 0.8 | 0.95 | < 10 | 1.14 | 0.92326 | 1 | 1.1 | 1.06 | 0.9 | 0.9 | 1.22 | 0.6 | 3.511 | [5 empty] | 110
| 1 | 6 | 1.1 | 0.9 | 0.9 | 0.8 | 0.97 | < 10 | 1.14 | 0.95282 | 1.1 | 1.1 | 1.1 | 0.9 | 0.9 | 1.3 | 0.65 | 3.881
| Mean |  | 1.0833 | 0.9 | 0.83333 | 0.8 | 0.98667 |  | 1.145 | 0.93717 | 1.05 | 1.15 | 1.05 | 0.9 | 0.91667 | 1.25 | 0.65833 | 3.6981
| Median |  | 1.1 | 0.9 | 0.8 | 0.8 | 0.975 |  | 1.14 | 0.94246 | 1.05 | 1.15 | 1.06 | 0.9 | 0.9 | 1.24 | 0.65 | 3.527
| Std Dev. |  | 0.075277 | 0 | 0.05164 | 1.2162e-16 | 0.049261 |  | 0.035637 | 0.016954 | 0.054772 | 0.054772 | 0.056569 | 0 | 0.040825 | 0.027568 | 0.037639 | 0.43484
| Rel.Std.Dev. |  | 6.949% | 0% | 6.197% | 1.52e-14% | 4.993% |  | 3.112% | 1.809% | 5.216% | 4.763% | 5.387% | 0% | 4.454% | 2.205% | 5.717% | 11.76%
| PDM3 |  | 11.11% | -7.696% | -14.53% | -17.95% | 1.193% |  | 17.43% | -3.884% | 7.688% | 17.94% | 7.688% | -7.696% | -5.986% | 28.2% | -32.48% | 279.3%
| Z-Score (Absolute) |  | 0.36 | 0.65 | 1.02 | 1.2 | 0.17 | 21.92 | 0.7 | 0.45 | 0.17 | 0.72 | 0.17 | 0.65 | 0.56 | 1.28 | 1.98 | 14.75
| Analytical results for Sr in OREAS 612 (Certified Value 49.2 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 1
| 1 | 1 | 44.7 | 36.1 | 45.3 | 59 | 49.4 | 49 | 60 | 52.5 | 45 | 64.5 | 48.8527 | 46 | 49 | 53.1 | 45 | 43 | 60 | 36.8 | 44.7187 |  |  | 1
| 1 | 2 | 46.4 | 35.8 | 45.7 | 58 | 48.9 | 52 | 60 | 53.37 | 45 | 59.6 | 47.7034 | 46.6 | 49 | 51.5 | 46 | 44.4 | 64 | 36.5 | 44.8433 |  |  | 20
| 1 | 3 | 47.7 | 36.7 | 44.9 | 60 | 48.8 | 53 | 62 | 50.97 | 46.8 | 55.9 | 47.5803 | 49 | 49 | 51 | 46 | 43.6 | 59 | 37.36 | 44.0193 |  |  | 16
| 1 | 4 | 44.7 | 36.9 | 45.1 | 59 | 49.8 | 54 | 60 | 54.03 | 43.4 | 55.5 | 48.6847 | 48.2 | 48 | 54.2 | 46 | 44.1 | 57 | 37.65 | 44.443 |  |  | 49.2121
| 1 | 5 | 45.7 | 37.6 | 44.9 | 58 | 48.4 | 54 | 59 | 51.99 | 45.5 | 62.1 | 49.5615 | 46.5 | 49 | 54.6 | 47 | 44.9 | 59 | 35.7 | 45.3273 |  |  | 111
| 1 | 6 | 46.3 | 37.4 | 46.3 | 59 | 53.2 | 51 | 61 | 51.42 | 45.5 | 58.3 | 47.6598 | 49.4 | 49 | 52.9 | 47 | 45.1 | 59 | 35.65 | 44.9427
| Mean |  | 45.9167 | 36.75 | 45.3667 | 58.8333 | 49.75 | 52.1667 | 60.3333 | 52.38 | 45.2 | 59.3167 | 48.3404 | 47.6167 | 48.8333 | 52.8833 | 46.1667 | 44.1833 | 59.6667 | 36.61 | 44.7157
| Median |  | 46 | 36.8 | 45.2 | 59 | 49.15 | 52.5 | 60 | 52.245 | 45.25 | 58.95 | 48.1941 | 47.4 | 49 | 53 | 46 | 44.25 | 59 | 36.65 | 44.781
| Std Dev. |  | 1.1462 | 0.7064 | 0.5465 | 0.75277 | 1.7593 | 1.9408 | 1.0328 | 1.1644 | 1.1009 | 3.5205 | 0.8147 | 1.4372 | 0.40825 | 1.4275 | 0.75277 | 0.79352 | 2.3381 | 0.8297 | 0.44748
| Rel.Std.Dev. |  | 2.496% | 1.922% | 1.205% | 1.28% | 3.536% | 3.72% | 1.712% | 2.223% | 2.436% | 5.935% | 1.685% | 3.018% | 0.836% | 2.699% | 1.631% | 1.796% | 3.919% | 2.266% | 1.001%
| PDM3 |  | -6.696% | -25.32% | -7.814% | 19.55% | 1.093% | 6.004% | 22.6% | 6.437% | -8.153% | 20.53% | -1.771% | -3.242% | -0.7696% | 7.46% | -6.188% | -10.22% | 21.24% | -25.61% | -9.137%
| Z-Score (Absolute) |  | 0.43 | 2.04 | 0.52 | 1.85 | 0.25 | 0.67 | 2.11 | 0.71 | 0.55 | 1.93 | 0 | 0.13 | 0.09 | 0.8 | 0.38 | 0.73 | 2 | 2.07 | 0.64
| Analytical results for Ta in OREAS 612 (Certified Value < 0.01 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | [8 empty] | 3
| 1 | 1 | < 0.01 | 0.07 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | 0.03 | < 0.01 | < 0.1 | < 0.05 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | 0.03 | < 0.01 | < 0.1 | < 0.05 | < 0.01 | [8 empty] | 21
| 1 | 3 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | 0.03 | < 0.01 | < 0.1 | < 0.05 | < 0.01 | [8 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | 0.03 | < 0.01 | < 0.1 | < 0.05 | < 0.01 | [8 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | 0.03 | < 0.01 | < 0.1 | < 0.05 | < 0.01 | [8 empty] | 112
| 1 | 6 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | 0.03 | < 0.01 | < 0.1 | < 0.05 | < 0.01
| Mean |  |  | 0.07 | 0.01 | [5 empty] | 0.03
| Median |  |  | 0.07 | 0.01 | [5 empty] | 0.03
| Std Dev. | [9 empty] | 0
| Rel.Std.Dev. | [9 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 21.58 | 0 | 0.67 | 0.67 | 15.51 | 15.51 | 0.67 | 19.55 | 0.67 | 35.74 | 15.51 | 0.67
| Analytical results for Tb in OREAS 612 (Certified Value 0.41 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.04 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.20 | 1.22 | 1.23 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [11 empty] | 2
|  |  | 0.25g | 0.25g | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | [11 empty] | 3
| 1 | 1 | 0.45 | 0.41 | 0.458 | 0.4 | 0.40733 | 0.39 | 0.4 | 0.44 | 0.49 | 0.39 | [11 empty] | 1
| 1 | 2 | 0.38 | 0.43 | 0.438 | 0.4 | 0.41848 | 0.4 | 0.4 | 0.44 | 0.52 | 0.41 | [11 empty] | 22
| 1 | 3 | 0.37 | 0.41 | 0.426 | 0.5 | 0.40878 | 0.38 | 0.39 | 0.43 | 0.49 | 0.41 | [11 empty] | 16
| 1 | 4 | 0.38 | 0.4 | 0.45 | 0.4 | 0.41071 | 0.37 | 0.42 | 0.43 | 0.49 | 0.42 | [11 empty] | 0.411
| 1 | 5 | 0.39 | 0.42 | 0.436 | 0.4 | 0.40588 | 0.39 | 0.4 | 0.44 | 0.49 | 0.41 | [11 empty] | 113
| 1 | 6 | 0.38 | 0.42 | 0.423 | 0.4 | 0.40566 | 0.37 | 0.42 | 0.45 | 0.5 | 0.4
| Mean |  | 0.39167 | 0.415 | 0.4385 | 0.41667 | 0.40947 | 0.38333 | 0.405 | 0.43833 | 0.49667 | 0.40667
| Median |  | 0.38 | 0.415 | 0.437 | 0.4 | 0.40806 | 0.385 | 0.4 | 0.44 | 0.49 | 0.41
| Std Dev. |  | 0.029269 | 0.010488 | 0.013531 | 0.040825 | 0.0047987 | 0.012111 | 0.012247 | 0.0075277 | 0.012111 | 0.010328
| Rel.Std.Dev. |  | 7.473% | 2.527% | 3.086% | 9.798% | 1.172% | 3.159% | 3.024% | 1.717% | 2.438% | 2.54%
| PDM3 |  | -4.703% | 0.9741% | 6.692% | 1.38% | -0.3708% | -6.731% | -1.459% | 6.651% | 20.84% | -1.054%
| Z-Score (Absolute) |  | 0.67 | 0.21 | 1.1 | NA | 0 | 0.99 | 0.17 | 1.09 | 3.3 | 0.11
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 612 (Certified Value < 0.05 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | [7 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.05 | 0.01 | < 0.1 | < 0.02 | < 0.2 | < 0.03 | 0.01 | 0.22 | < 0.01 | < 0.1 | < 0.05 | < 0.05 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.05 | 0.01 | < 0.1 | < 0.02 | < 0.2 | < 0.03 | 0.01 | 0.23 | < 0.01 | < 0.1 | < 0.05 | < 0.05 | [7 empty] | 23
| 1 | 3 | < 0.01 | < 0.05 | < 0.05 | 0.02 | < 0.1 | < 0.02 | < 0.2 | < 0.03 | 0.02 | 0.22 | < 0.01 | < 0.1 | < 0.05 | < 0.05 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.05 | 0.01 | < 0.1 | < 0.02 | < 0.2 | < 0.03 | < 0.01 | 0.23 | < 0.01 | < 0.1 | < 0.05 | < 0.05 | [7 empty] | < 0.05
| 1 | 5 | < 0.01 | < 0.05 | < 0.05 | 0.01 | < 0.1 | 0.02 | < 0.2 | < 0.03 | 0.01 | 0.24 | < 0.01 | < 0.1 | < 0.05 | < 0.05 | [7 empty] | 114
| 1 | 6 | < 0.01 | < 0.05 | < 0.05 | 0.02 | < 0.1 | < 0.02 | < 0.2 | < 0.03 | 0.02 | 0.23 | < 0.01 | < 0.1 | < 0.05 | < 0.05
| Mean |  |  |  |  | 0.013333 |  | 0.02 |  |  | 0.014 | 0.22833
| Median |  |  |  |  | 0.01 |  | 0.02 |  |  | 0.01 | 0.23
| Std Dev. |  |  |  |  | 0.005164 |  |  |  |  | 0.0054772 | 0.0075277
| Rel.Std.Dev. |  |  |  |  | 38.73% |  |  |  |  | 39.12% | 3.297%
PDM3
| Z-Score (Absolute) |  | 1.04 | 0 | 0 | 0.61 | 1.31 | 0.7 | 3.92 | 0.52 | 0.65 | 10.61 | 1.04 | 1.31 | 0 | 0
| Analytical results for Th in OREAS 612 (Certified Value 7.79 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g |  |  |  |  | 2
| 1 | 1 | 8.2 | 7.5 | 7.13 | < 100 | 8.1 | < 100 | < 100 | 8.16 | 8.5 | 7.5261 | 7.2 | 7.79 | 8.88 | 6.7 | 7.6 | 9.61 | 6.39 |  |  |  |  | 1
| 1 | 2 | 8.6 | 6.5 | 7.63 | < 100 | 7.5 | < 100 | < 100 | 8.18 | 8.1 | 7.4419 | 7.6 | 7.48 | 8.82 | 6.7 | 7.6 | 9.79 | 6.56 |  |  |  |  | 24
| 1 | 3 | 8.4 | 6.6 | 7.38 | < 100 | 7.6 | < 100 | < 100 | 7.97 | 8.3 | 7.5842 | 7.9 | 7.48 | 8.75 | 6.8 | 7.5 | 9.27 | 6.98 |  |  |  |  | 16
| 1 | 4 | 8.2 | 6.5 | 6.97 | < 100 | 7.9 | < 100 | < 100 | 8.68 | 8.4 | 7.7352 | 7.9 | 7.47 | 9 | 6.7 | 7.7 | 9.42 | 7.33 |  |  |  |  | 7.787
| 1 | 5 | 8.5 | 6.7 | 7.33 | < 100 | 7.8 | < 100 | < 100 | 8.22 | 7.8 | 7.4678 | 7.4 | 7.4 | 8.95 | 6.8 | 7.8 | 9.24 | 6.98 |  |  |  |  | 115
| 1 | 6 | 8 | 6.7 | 7.48 | < 100 | 8.1 | < 100 | < 100 | 7.96 | 8.5 | 7.4 | 7.9 | 7.34 | 9.03 | 6.8 | 7.8 | 9.56 | 6.94
| Mean |  | 8.3167 | 6.75 | 7.32 |  | 7.8333 |  |  | 8.195 | 8.2667 | 7.5259 | 7.65 | 7.4933 | 8.905 | 6.75 | 7.6667 | 9.4817 | 6.8633
| Median |  | 8.3 | 6.65 | 7.355 |  | 7.85 |  |  | 8.17 | 8.35 | 7.497 | 7.75 | 7.475 | 8.915 | 6.75 | 7.65 | 9.49 | 6.96
| Std Dev. |  | 0.22286 | 0.37815 | 0.23833 |  | 0.25033 |  |  | 0.26197 | 0.27325 | 0.12121 | 0.30166 | 0.15565 | 0.10821 | 0.054772 | 0.12111 | 0.21198 | 0.33673
| Rel.Std.Dev. |  | 2.68% | 5.602% | 3.256% |  | 3.196% |  |  | 3.197% | 3.305% | 1.611% | 3.943% | 2.077% | 1.215% | 0.8114% | 1.58% | 2.236% | 4.906%
| PDM3 |  | 6.802% | -13.32% | -5.997% |  | 0.5954% |  |  | 5.24% | 6.16% | -3.353% | -1.759% | -3.771% | 14.36% | -13.32% | -1.545% | 21.76% | -11.86%
| Z-Score (Absolute) |  | 0.63 | 1.42 | 0.67 | 55.39 | 0 | 55.39 | 55.39 | 0.48 | 0.57 | 0.4 | 0.24 | 0.45 | 1.41 | 1.42 | 0.22 | 2.17 | 1.27
| Analytical results for Ti in OREAS 612 (Certified Value 0.016 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g |  |  |  | 3
| 1 | 1 | 0.014 | 0.01 | 0.016 | < 0.05 | 0.015 | < 0.05 | < 0.05 | 0.015 | 0.015 | 0.02 | 0.015901 | 0.015 | 0.014 | 0.0227 | 0.015 | 0.02 | 0.0219 | 0.01 |  |  |  | 1
| 1 | 2 | 0.015 | 0.01 | 0.016 | < 0.05 | 0.015 | < 0.05 | < 0.05 | 0.0163 | 0.015 | 0.02 | 0.015491 | 0.014 | 0.014 | 0.022 | 0.015 | 0.02 | 0.0228 | 0.01 |  |  |  | 25
| 1 | 3 | 0.016 | 0.01 | 0.016 | < 0.05 | 0.015 | < 0.05 | < 0.05 | 0.0157 | 0.015 | 0.02 | 0.015761 | 0.014 | 0.014 | 0.022 | 0.015 | 0.02 | 0.021 | 0.01 |  |  |  | 16
| 1 | 4 | 0.014 | 0.01 | 0.016 | < 0.05 | 0.015 | < 0.05 | < 0.05 | 0.0171 | 0.016 | 0.02 | 0.015846 | 0.014 | 0.014 | 0.0222 | 0.016 | 0.02 | 0.0207 | 0.01 |  |  |  | 0.016049
| 1 | 5 | 0.015 | 0.01 | 0.016 | < 0.05 | 0.015 | < 0.05 | < 0.05 | 0.0166 | 0.015 | 0.02 | 0.016049 | 0.014 | 0.014 | 0.0224 | 0.016 | 0.02 | 0.0206 | 0.01 |  |  |  | 116
| 1 | 6 | 0.015 | 0.01 | 0.016 | < 0.05 | 0.015 | < 0.05 | < 0.05 | 0.0161 | 0.015 | 0.02 | 0.015415 | 0.014 | 0.014 | 0.0222 | 0.016 | 0.02 | 0.0211 | 0.01
| Mean |  | 0.014833 | 0.01 | 0.016 |  | 0.015 |  |  | 0.016133 | 0.015167 | 0.02 | 0.015744 | 0.014167 | 0.014 | 0.02225 | 0.0155 | 0.02 | 0.02135 | 0.01
| Median |  | 0.015 | 0.01 | 0.016 |  | 0.015 |  |  | 0.0162 | 0.015 | 0.02 | 0.015804 | 0.014 | 0.014 | 0.0222 | 0.0155 | 0.02 | 0.02105 | 0.01
| Std Dev. |  | 0.00075277 | 0 | 0 |  | 0 |  |  | 0.00072847 | 0.00040825 | 0 | 0.00024504 | 0.00040825 | 1.9003e-18 | 0.00026646 | 0.00054772 | 0 | 0.00084558 | 0
| Rel.Std.Dev. |  | 5.075% | 0% | 0% |  | 0% |  |  | 4.515% | 2.692% | 0% | 1.556% | 2.882% | 1.357e-14% | 1.198% | 3.534% | 0% | 3.961% | 0%
| PDM3 |  | -7.577% | -37.69% | -0.3081% |  | -6.539% |  |  | 0.5227% | -5.5% | 24.61% | -1.904% | -11.73% | -12.77% | 38.63% | -3.423% | 24.61% | 33.03% | -37.69%
| Z-Score (Absolute) |  | 0.23 | 1.32 | 0.03 | 2.05 | 0.2 | 2.05 | 2.05 | 0.06 | 0.16 | 0.93 | 0.03 | 0.38 | 0.42 | 1.43 | 0.08 | 0.93 | 1.23 | 1.32
| Analytical results for Tl in OREAS 612 (Certified Value 0.86 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g |  |  | 3
| 1 | 1 | 1.38 | 1.04 | 0.87 | < 50 | 0.8 | < 50 | < 50 | 0.9 | 0.84 | 0.85268 | 0.87 | 0.78 | 0.83 | 0.91 | 0.73 | 0.82822 | 0.86 | 0.98 | 0.8 |  |  | 1
| 1 | 2 | 1.34 | 0.94 | 0.91 | < 50 | 0.82 | < 50 | < 50 | 0.92 | 0.84 | 0.85195 | 0.86 | 0.82 | 0.85 | 0.92 | 0.74 | 0.87413 | 0.87 | 1 | 0.8 |  |  | 26
| 1 | 3 | 1.39 | 0.92 | 0.88 | < 50 | 0.8 | < 50 | < 50 | 0.88 | 0.82 | 0.86775 | 0.86 | 0.87 | 0.84 | 0.87 | 0.74 | 0.82475 | 0.87 | 0.97 | 0.81 |  |  | 16
| 1 | 4 | 1.38 | 0.93 | 0.87 | < 50 | 0.85 | < 50 | < 50 | 0.88 | 0.81 | 0.88235 | 0.86 | 0.87 | 0.82 | 0.91 | 0.73 | 0.82 | 0.87 | 0.97 | 0.81 |  |  | 0.85888
| 1 | 5 | 1.42 | 0.94 | 0.9 | < 50 | 0.81 | < 50 | < 50 | 0.92 | 0.88 | 0.83679 | 0.85 | 0.81 | 0.82 | 0.93 | 0.73 | 0.84477 | 0.88 | 0.99 | 0.81 |  |  | 117
| 1 | 6 | 1.38 | 0.94 | 0.88 | < 50 | 0.86 | < 50 | < 50 | 0.87 | 0.81 | 0.86223 | 0.87 | 0.87 | 0.78 | 0.92 | 0.73 | 0.81974 | 0.88 | 0.97 | 0.78
| Mean |  | 1.3817 | 0.95167 | 0.885 |  | 0.82333 |  |  | 0.895 | 0.83333 | 0.85896 | 0.86167 | 0.83667 | 0.82333 | 0.91 | 0.73333 | 0.83527 | 0.87167 | 0.98 | 0.80167
| Median |  | 1.38 | 0.94 | 0.88 |  | 0.815 |  |  | 0.89 | 0.83 | 0.85746 | 0.86 | 0.845 | 0.825 | 0.915 | 0.73 | 0.82649 | 0.87 | 0.975 | 0.805
| Std Dev. |  | 0.025626 | 0.044008 | 0.016432 |  | 0.02582 |  |  | 0.021679 | 0.026583 | 0.01559 | 0.0075277 | 0.038816 | 0.024221 | 0.020976 | 0.005164 | 0.021143 | 0.0075277 | 0.012649 | 0.01169
| Rel.Std.Dev. |  | 1.855% | 4.624% | 1.857% |  | 3.136% |  |  | 2.422% | 3.19% | 1.815% | 0.8736% | 4.639% | 2.942% | 2.305% | 0.7042% | 2.531% | 0.8636% | 1.291% | 1.458%
| PDM3 |  | 60.87% | 10.8% | 3.041% |  | -4.139% |  |  | 4.205% | -2.975% | 0.008941% | 0.3242% | -2.587% | -4.139% | 5.952% | -14.62% | -2.749% | 1.489% | 14.1% | -6.662%
| Z-Score (Absolute) |  | 7.12 | 1.12 | 0.19 | 336.62 | 0.67 | 336.62 | 336.62 | 0.33 | 0.53 | 0.18 | 0.14 | 0.49 | 0.67 | 0.53 | 1.93 | 0.51 | 0 | 1.51 | 0.98
| Analytical results for Tm in OREAS 612 (Certified Value 0.073 ppm) | [22 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.22 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [15 empty] | 3
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 01g | [15 empty] | 3
| 1 | 1 | 0.085 | < 0.1 | 0.068273 | 0.06 | 0.08 | 0.08 | [15 empty] | 1
| 1 | 2 | 0.081 | < 0.1 | 0.066611 | 0.06 | 0.08 | 0.08 | [15 empty] | 27
| 1 | 3 | 0.074 | < 0.1 | 0.06714 | 0.06 | 0.08 | 0.08 | [15 empty] | 16
| 1 | 4 | 0.081 | < 0.1 | 0.068916 | 0.06 | 0.08 | 0.08 | [15 empty] | 0.072999
| 1 | 5 | 0.076 | < 0.1 | 0.063573 | 0.06 | 0.08 | 0.08 | [15 empty] | 118
| 1 | 6 | 0.073 | < 0.1 | 0.065444 | 0.06 | 0.08 | 0.08
| Mean |  | 0.078333 |  | 0.066659 | 0.06 | 0.08 | 0.08
| Median |  | 0.0785 |  | 0.066876 | 0.06 | 0.08 | 0.08
| Std Dev. |  | 0.0047188 |  | 0.001946 | 0 | 0 | 0
| Rel.Std.Dev. |  | 6.024% |  | 2.919% | 0% | 0% | 0%
| PDM3 |  | 7.308% |  | -8.684% | -17.81% | 9.591% | 9.591%
| Z-Score (Absolute) |  | 0.52 | 2.02 | 0.52 | 1.12 | 0.67 | 0.67
| Analytical results for U in OREAS 612 (Certified Value 2.22 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 3
| 1 | 1 | 2.36 | 1.93 | 2 | < 50 | 2 | < 50 | < 50 | 2.3 | 2.3 | 2.0481 | 2.45 | 2.15 | 2.4 | 2.14 | 2.3 | 2.18 | 2.4 | 2.19 | 1.4175 |  |  | 1
| 1 | 2 | 2.1 | 1.88 | 2.21 | < 50 | 2.02 | < 50 | < 50 | 2.31 | 2.2 | 2.044 | 2.39 | 2.28 | 2.29 | 2.15 | 2.3 | 2.19 | 2.45 | 2.31 | 1.5607 |  |  | 28
| 1 | 3 | 2.29 | 1.91 | 2.08 | < 50 | 2.08 | < 50 | < 50 | 2.3 | 2.2 | 2.0768 | 2.46 | 2.34 | 2.22 | 2.18 | 2.3 | 2.17 | 2.31 | 2.29 | 2.1415 |  |  | 16
| 1 | 4 | 2.13 | 1.95 | 2.13 | < 50 | 2.12 | < 50 | < 50 | 2.35 | 2.2 | 2.119 | 2.42 | 2.35 | 2.22 | 2.19 | 2.3 | 2.19 | 2.34 | 2.29 | 2.0583 |  |  | 2.2153
| 1 | 5 | 2.34 | 1.98 | 2.14 | < 50 | 2.09 | < 50 | < 50 | 2.35 | 2.1 | 2.0975 | 2.38 | 2.15 | 2.21 | 2.21 | 2.3 | 2.23 | 2.32 | 2.27 | 1.8923 |  |  | 119
| 1 | 6 | 2.3 | 1.99 | 2.15 | < 50 | 2.13 | < 50 | < 50 | 2.24 | 2.2 | 2.0901 | 2.44 | 2.34 | 2.19 | 2.24 | 2.3 | 2.23 | 2.39 | 2.2 | 1.858
| Mean |  | 2.2533 | 1.94 | 2.1183 |  | 2.0733 |  |  | 2.3083 | 2.2 | 2.0793 | 2.4233 | 2.2683 | 2.255 | 2.185 | 2.3 | 2.1983 | 2.3683 | 2.2583 | 1.8214
| Median |  | 2.295 | 1.94 | 2.135 |  | 2.085 |  |  | 2.305 | 2.2 | 2.0834 | 2.43 | 2.31 | 2.22 | 2.185 | 2.3 | 2.19 | 2.365 | 2.28 | 1.8752
| Std Dev. |  | 0.11057 | 0.041952 | 0.071391 |  | 0.052789 |  |  | 0.040702 | 0.063246 | 0.029147 | 0.03266 | 0.094956 | 0.078677 | 0.037283 | 4.8648e-16 | 0.025626 | 0.054191 | 0.050761 | 0.28146
| Rel.Std.Dev. |  | 4.907% | 2.162% | 3.37% |  | 2.546% |  |  | 1.763% | 2.875% | 1.402% | 1.348% | 4.186% | 3.489% | 1.706% | 2.115e-14% | 1.166% | 2.288% | 2.248% | 15.45%
| PDM3 |  | 1.718% | -12.43% | -4.376% |  | -6.408% |  |  | 4.2% | -0.6899% | -6.141% | 9.392% | 2.395% | 1.793% | -1.367% | 3.824% | -0.7651% | 6.909% | 1.943% | -17.78%
| Z-Score (Absolute) |  | 0.01 | 1.87 | 0.81 | 135.33 | 1.08 | 135.33 | 135.33 | 0.32 | 0.33 | 1.05 | 1 | 0.08 | 0 | 0.42 | 0.27 | 0.34 | 0.67 | 0.02 | 2.58
| Analytical results for V in OREAS 612 (Certified Value 9.41 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 2
| 1 | 1 | 9 | 8 | 9 | 11 | 9 | 10 | 12 | 9 | 7 | 12 | 8.9711 | 9 | 9 | 12 | 9 | 11 | 12 | 8 | 8.2003 |  |  | 1
| 1 | 2 | 9 | 8 | 9 | 11 | 9 | 9 | 13 | 9 | 7 | 11 | 8.7494 | 9 | 9 | 11 | 9 | 11 | 13 | 8 | 8.163 |  |  | 29
| 1 | 3 | 9 | 8 | 8 | 11 | 9 | 10 | 11 | 9 | 7 | 10 | 8.9751 | 9 | 9 | 12 | 9 | 11 | 12 | 8 | 8.076 |  |  | 16
| 1 | 4 | 9 | 8 | 8 | 10 | 9 | 10 | 11 | 10 | 9 | 10 | 8.9903 | 9 | 9 | 12 | 9 | 12 | 12 | 8 | 8.0977 |  |  | 9.4108
| 1 | 5 | 9 | 9 | 9 | 10 | 9 | 9 | 11 | 10 | 7 | 11 | 9.0568 | 8 | 9 | 12 | 9 | 12 | 12 | 8 | 8.2653 |  |  | 120
| 1 | 6 | 9 | 8 | 9 | 11 | 9 | 9 | 11 | 10 | 7 | 11 | 8.7133 | 9 | 9 | 12 | 9 | 12 | 12 | 8 | 8.1047
| Mean |  | 9 | 8.1667 | 8.6667 | 10.6667 | 9 | 9.5 | 11.5 | 9.5 | 7.3333 | 10.8333 | 8.9093 | 8.8333 | 9 | 11.8333 | 9 | 11.5 | 12.1667 | 8 | 8.1512
| Median |  | 9 | 8 | 9 | 11 | 9 | 9.5 | 11 | 9.5 | 7 | 11 | 8.9731 | 9 | 9 | 12 | 9 | 11.5 | 12 | 8 | 8.1338
| Std Dev. |  | 0 | 0.40825 | 0.5164 | 0.5164 | 0 | 0.54772 | 0.83666 | 0.54772 | 0.8165 | 0.75277 | 0.14175 | 0.40825 | 0 | 0.40825 | 0 | 0.54772 | 0.40825 | 0 | 0.072479
| Rel.Std.Dev. |  | 0% | 4.999% | 5.958% | 4.841% | 0% | 5.766% | 7.275% | 5.766% | 11.13% | 6.949% | 1.591% | 4.622% | 0% | 3.45% | 0% | 4.763% | 3.355% | 0% | 0.8892%
| PDM3 |  | -4.365% | -13.22% | -7.907% | 13.35% | -4.365% | 0.9482% | 22.2% | 0.9482% | -22.08% | 15.12% | -5.328% | -6.136% | -4.365% | 25.74% | -4.365% | 22.2% | 29.28% | -14.99% | -13.38%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.27 | 1.35 | 0 | 0.4 | 2.02 | 0.4 | 1.35 | 1.48 | 0.07 | 0.13 | 0 | 2.29 | 0 | 2.02 | 2.56 | 0.81 | 0.69
| Analytical results for W in OREAS 612 (Certified Value 1.03 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.24 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | 0.2g |  |  | 2
| 1 | 1 | 2 | 1.6 | 1.01 | < 50 | 1.14 | < 50 | < 50 | 1.32 | 25 | 0.5 | 0.98016 | 1.1 | 1.45 | 0.78 | 1 | 0.8 | < 2 | 0.8 | 2.3425 |  |  | 1
| 1 | 2 | 2.29 | 1.6 | 0.99 | < 50 | 1.12 | < 50 | < 50 | 1.3 | 22 | 0.5 | 1.0235 | 1.16 | 1.33 | 0.83 | 1 | 0.8 | < 2 | 0.8 | 2.588 |  |  | 30
| 1 | 3 | 1.98 | 1.4 | 1.03 | < 50 | 1.13 | < 50 | < 50 | 1.11 | 22 | 0.7 | 1.0119 | 1.17 | 1.37 | 0.85 | 1 | 0.8 | < 2 | 0.8 | 1.9243 |  |  | 16
| 1 | 4 | 2.34 | 1.4 | 1.02 | < 50 | 1.11 | < 50 | < 50 | 1.34 | < 10 | 0.8 | 1.0009 | 1.11 | 1.3 | 0.83 | 1 | 0.8 | < 2 | 0.8 | 2.381 |  |  | 1.0313
| 1 | 5 | 2.25 | 1.6 | 1.03 | < 50 | 1.11 | < 50 | < 50 | 1.17 | 21 | 0.4 | 0.95959 | 1.16 | 1.38 | 0.84 | 0.9 | 0.8 | < 2 | 0.7 | 3.448 |  |  | 121
| 1 | 6 | 2.23 | 1.5 | 1.06 | < 50 | 1.12 | < 50 | < 50 | 1.29 | 21 | 0.7 | 0.99944 | 1.17 | 1.36 | 0.79 | 0.9 | 0.8 | < 2 | 0.7 | 3.383
| Mean |  | 2.1817 | 1.5167 | 1.0233 |  | 1.1217 |  |  | 1.255 | 22.2 | 0.6 | 0.99592 | 1.145 | 1.365 | 0.82 | 0.96667 | 0.8 |  | 0.76667 | 2.6778
| Median |  | 2.24 | 1.55 | 1.025 |  | 1.12 |  |  | 1.295 | 22 | 0.6 | 1.0002 | 1.16 | 1.365 | 0.83 | 1 | 0.8 |  | 0.8 | 2.4845
| Std Dev. |  | 0.15329 | 0.098319 | 0.023381 |  | 0.01169 |  |  | 0.092682 | 1.6432 | 0.15492 | 0.022898 | 0.031464 | 0.050892 | 0.028284 | 0.05164 | 1.2162e-16 |  | 0.05164 | 0.61103
| Rel.Std.Dev. |  | 7.026% | 6.483% | 2.285% |  | 1.042% |  |  | 7.385% | 7.402% | 25.82% | 2.299% | 2.748% | 3.728% | 3.449% | 5.342% | 1.52e-14% |  | 6.736% | 22.82%
| PDM3 |  | 111.5% | 47.06% | -0.775% |  | 8.76% |  |  | 21.69% | 2053% | -41.82% | -3.433% | 11.02% | 32.35% | -20.49% | -6.27% | -22.43% |  | -25.66% | 159.6%
| Z-Score (Absolute) |  | 2.19 | 0.8 | 0.23 | 49.77 | 0.02 | 49.77 | 49.77 | 0.25 | NA | 1.11 | 0.29 | 0.02 | 0.48 | 0.65 | 0.35 | 0.7 | 0.28 | 0.76 | 3.22
Results from laboratory 1.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 612 (Certified Value 7.22 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | [5 empty] | 3
| 1 | 1 | 7.28 | 6.67 | 6.74 | 7.24 | 7.72 | 6.5 | 7.99 | 7.3601 | 7.12 | 6.5 | 8.05 | 6.7 | 6.9852 | 7.08 | 8.5 | 6.7 | [5 empty] | 1
| 1 | 2 | 7.39 | 6.47 | 7.38 | 7.2 | 7.58 | 6.4 | 7.26 | 7.2194 | 7.26 | 6.7 | 8.1 | 6.7 | 7.0256 | 7.18 | 8.73 | 6.6 | [5 empty] | 31
| 1 | 3 | 7.53 | 6.57 | 6.72 | 7.34 | 7.53 | 6.7 | 7.42 | 7.1851 | 7.71 | 6.8 | 8 | 6.8 | 6.993 | 7.11 | 8.37 | 6.7 | [5 empty] | 16
| 1 | 4 | 7.29 | 6.54 | 6.91 | 7.36 | 7.95 | 6.7 | 6.99 | 7.1557 | 7.59 | 7 | 8.05 | 6.8 | 7.063 | 7.05 | 8.23 | 6.8 | [5 empty] | 7.2186
| 1 | 5 | 7.31 | 6.62 | 6.86 | 7.28 | 7.66 | 6.5 | 7.39 | 7.3082 | 7.19 | 6.9 | 8.22 | 6.8 | 7.0347 | 7.17 | 8.33 | 6.5 | [5 empty] | 122
| 1 | 6 | 7.5 | 6.58 | 7.23 | 7.79 | 7.44 | 6.5 | 7.21 | 7.2362 | 7.84 | 6.8 | 8.2 | 6.9 | 6.9294 | 7.42 | 8.45 | 6.6
| Mean |  | 7.3833 | 6.575 | 6.9733 | 7.3683 | 7.6467 | 6.55 | 7.3767 | 7.2441 | 7.4517 | 6.7833 | 8.1033 | 6.7833 | 7.0051 | 7.1683 | 8.435 | 6.65
| Median |  | 7.35 | 6.575 | 6.885 | 7.31 | 7.62 | 6.5 | 7.325 | 7.2278 | 7.425 | 6.8 | 8.075 | 6.8 | 7.0093 | 7.14 | 8.41 | 6.65
| Std Dev. |  | 0.10948 | 0.068337 | 0.27083 | 0.21507 | 0.17795 | 0.12247 | 0.33738 | 0.076868 | 0.30063 | 0.17224 | 0.088694 | 0.075277 | 0.046727 | 0.13318 | 0.17248 | 0.10488
| Rel.Std.Dev. |  | 1.483% | 1.039% | 3.884% | 2.919% | 2.327% | 1.87% | 4.574% | 1.061% | 4.034% | 2.539% | 1.095% | 1.11% | 0.667% | 1.858% | 2.045% | 1.577%
| PDM3 |  | 2.282% | -8.916% | -3.398% | 2.074% | 5.93% | -9.262% | 2.19% | 0.3537% | 3.229% | -6.03% | 12.26% | -6.03% | -2.957% | -0.6964% | 16.85% | -7.877%
| Z-Score (Absolute) |  | 0.36 | 1.27 | 0.47 | 0.33 | 0.89 | 1.32 | 0.34 | 0.08 | 0.5 | 0.85 | 1.81 | 0.85 | 0.41 | 0.08 | 2.48 | 1.12
| Analytical results for Yb in OREAS 612 (Certified Value 0.4 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.04 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.18 | 1.20 | 1.22 | 1.23 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [10 empty] | 2
|  |  | 0.25g | 0.25g | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.2g | [10 empty] | 3
| 1 | 1 | 0.4 | 0.4 | 0.468 | 0.4 | 0.36875 | 0.4 | 0.46 | 0.4572 | 0.4 | 0.47 | 0.39 | [10 empty] | 1
| 1 | 2 | 0.4 | 0.4 | 0.443 | 0.4 | 0.36401 | 0.39 | 0.45 | 0.45391 | 0.4 | 0.48 | 0.38 | [10 empty] | 32
| 1 | 3 | 0.4 | 0.4 | 0.42 | 0.4 | 0.37065 | 0.37 | 0.46 | 0.45971 | 0.4 | 0.48 | 0.38 | [10 empty] | 16
| 1 | 4 | 0.4 | 0.4 | 0.457 | 0.4 | 0.3805 | 0.39 | 0.46 | 0.46527 | 0.4 | 0.47 | 0.41 | [10 empty] | 0.4046
| 1 | 5 | 0.4 | 0.4 | 0.45 | 0.4 | 0.36081 | 0.38 | 0.46 | 0.45981 | 0.4 | 0.47 | 0.42 | [10 empty] | 123
| 1 | 6 | 0.4 | 0.4 | 0.424 | 0.4 | 0.36194 | 0.38 | 0.44 | 0.45987 | 0.4 | 0.48 | 0.36
| Mean |  | 0.4 | 0.4 | 0.44367 | 0.4 | 0.36778 | 0.385 | 0.455 | 0.4593 | 0.4 | 0.475 | 0.39
| Median |  | 0.4 | 0.4 | 0.4465 | 0.4 | 0.36638 | 0.385 | 0.46 | 0.45976 | 0.4 | 0.475 | 0.385
| Std Dev. |  | 6.0809e-17 | 6.0809e-17 | 0.018747 | 6.0809e-17 | 0.0073201 | 0.010488 | 0.0083666 | 0.0037386 | 6.0809e-17 | 0.0054772 | 0.021909
| Rel.Std.Dev. |  | 1.52e-14% | 1.52e-14% | 4.226% | 1.52e-14% | 1.99% | 2.724% | 1.839% | 0.814% | 1.52e-14% | 1.153% | 5.618%
| PDM3 |  | -1.138% | -1.138% | 9.654% | -1.138% | -9.102% | -4.845% | 12.46% | 13.52% | -1.138% | 17.4% | -3.61%
| Z-Score (Absolute) |  | 0 | 0 | 1.96 | 0 | 1.45 | 0.67 | 2.47 | 2.67 | 0 | 3.37 | 0.45
| Analytical results for Zn in OREAS 612 (Certified Value 0.362 wt.%) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | 1.23 | 1.24 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | 3
|  |  | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 01g | 0.2g | 0.2g | 3
| 1 | 1 | 0.3562 | 0.3566 | 0.3668 | 0.367 | 0.363 | 0.337 | 0.363 | 0.3495 | 0.3583 | 0.364 | 0.35984 | 0.3717 | 0.362 | 0.3678 | 0.0028 | 0.3577 | 0.36423 | 0.3952 | 0.3637 | 0.3637 | 0.46561 | 1
| 1 | 2 | 0.3602 | 0.3489 | 0.3689 | 0.36 | 0.369 | 0.344 | 0.364 | 0.3537 | 0.3514 | 0.369 | 0.35469 | 0.3651 | 0.358 | 0.3645 | 0.003 | 0.3602 | 0.36857 | 0.3895 | 0.3717 | 0.3619 | 0.45491 | 33
| 1 | 3 | 0.3629 | 0.3494 | 0.3599 | 0.361 | 0.358 | 0.35 | 0.373 | 0.3481 | 0.3636 | 0.366 | 0.3535 | 0.3715 | 0.366 | 0.3623 | 0.0029 | 0.3641 | 0.36209 | 0.4003 | 0.3605 | 0.3647 | 0.45206 | 16
| 1 | 4 | 0.3617 | 0.3523 | 0.3619 | 0.362 | 0.368 | 0.346 | 0.36 | 0.3555 | 0.3496 | 0.367 | 0.35688 | 0.3716 | 0.366 | 0.3641 | 0.0031 | 0.3699 | 0.36465 | 0.3909 | 0.3598 | 0.3718 | 0.45583 | 0.36219
| 1 | 5 | 0.3565 | 0.3561 | 0.3576 | 0.357 | 0.362 | 0.346 | 0.36 | 0.348 | 0.3504 | 0.348 | 0.35829 | 0.3668 | 0.358 | 0.3679 | 0.003 | 0.3707 | 0.36377 | 0.4019 | 0.3631 | 0.353 | 0.44336 | 124
| 1 | 6 | 0.3589 | 0.3565 | 0.3689 | 0.361 | 0.358 | 0.339 | 0.368 | 0.3466 | 0.3578 | 0.374 | 0.35519 | 0.3701 | 0.359 | 0.3656 | 0.003 | 0.3741 | 0.36194 | 0.4038 | 0.3641 | 0.3604 | 0.43829
| Mean |  | 0.3594 | 0.3533 | 0.364 | 0.36133 | 0.363 | 0.34367 | 0.36467 | 0.35023 | 0.35518 | 0.36467 | 0.3564 | 0.36947 | 0.3615 | 0.36537 | 0.0029667 | 0.36612 | 0.36421 | 0.39693 | 0.36382 | 0.36258 | 0.45168
| Median |  | 0.35955 | 0.3542 | 0.36435 | 0.361 | 0.3625 | 0.345 | 0.3635 | 0.3488 | 0.3546 | 0.3665 | 0.35603 | 0.3708 | 0.3605 | 0.36505 | 0.003 | 0.367 | 0.364 | 0.39775 | 0.3634 | 0.3628 | 0.45348
| Std Dev. |  | 0.0027232 | 0.0035928 | 0.0048588 | 0.003266 | 0.0047329 | 0.0048442 | 0.0050465 | 0.0035506 | 0.0055815 | 0.0088468 | 0.0023868 | 0.0028374 | 0.0037815 | 0.0021979 | 0.00010328 | 0.0064648 | 0.0024083 | 0.0059634 | 0.00424 | 0.0061278 | 0.0096986
| Rel.Std.Dev. |  | 0.7577% | 1.017% | 1.335% | 0.9039% | 1.304% | 1.41% | 1.384% | 1.014% | 1.571% | 2.426% | 0.6697% | 0.768% | 1.046% | 0.6016% | 3.481% | 1.766% | 0.6612% | 1.502% | 1.165% | 1.69% | 2.147%
| PDM3 |  | -0.7697% | -2.454% | 0.5003% | -0.2359% | 0.2242% | -5.114% | 0.6844% | -3.301% | -1.934% | 0.6844% | -1.599% | 2.01% | -0.1899% | 0.8777% | -99.18% | 1.085% | 0.5578% | 9.593% | 0.4497% | 0.1092% | 24.71%
| Z-Score (Absolute) |  | 0.78 | 2.1 | 0.22 | 0.36 | 0 | 4.18 | 0.36 | 2.76 | 1.69 | 0.36 | 1.43 | 1.4 | 0.32 | 0.51 | 77.9 | 0.67 | 0.26 | 7.34 | 0.18 | 0.09 | 19.19
| Analytical results for Zr in OREAS 612 (Certified Value 28 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.06 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.2g | 0.25g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | [7 empty] | 2
| 1 | 1 | 29.5 | 24.7 | 26.7 | 29.9 | 28.6 | 29.2 | 2 | 29.967 | 26.8 | 26 | 3490 | 26 | 27.4 | 38.1 | [7 empty] | 1
| 1 | 2 | 29.5 | 25 | 27.6 | 28.4 | 29.7 | 28.9 | 2.3 | 30.2332 | 27.8 | 25 | 3520 | 27.2 | 28.2 | 38.6 | [7 empty] | 34
| 1 | 3 | 30.5 | 25.4 | 26.4 | 28.8 | 28.3 | 29.9 | 2.2 | 29.5041 | 29.2 | 24 | 3510 | 26.8 | 29 | 37.3 | [7 empty] | 16
| 1 | 4 | 29 | 25.5 | 26.3 | 29.3 | 29.6 | 29.2 | 1.7 | 29.6356 | 28.3 | 25 | 3570 | 27.2 | 28.6 | 37.2 | [7 empty] | 27.9699
| 1 | 5 | 28.8 | 25.5 | 26.4 | 28.8 | 29.4 | 29.7 | 1.6 | 29.678 | 27.8 | 25 | 3510 | 26.8 | 29.2 | 37.3 | [7 empty] | 125
| 1 | 6 | 29.1 | 25.9 | 27.3 | 30.6 | 28.6 | 29.9 | 3.3 | 29.7959 | 30 | 24 | 3570 | 26.6 | 29.4 | 37.1
| Mean |  | 29.4 | 25.3333 | 26.7833 | 29.3 | 29.0333 | 29.4667 | 2.1833 | 29.8023 | 28.3167 | 24.8333 | 3528.3333 | 26.7667 | 28.6333 | 37.6
| Median |  | 29.3 | 25.45 | 26.55 | 29.05 | 29 | 29.45 | 2.1 | 29.737 | 28.05 | 25 | 3515 | 26.8 | 28.8 | 37.3
| Std Dev. |  | 0.60663 | 0.42269 | 0.54191 | 0.81976 | 0.60222 | 0.42269 | 0.61128 | 0.26266 | 1.1356 | 0.75277 | 33.7145 | 0.44572 | 0.74207 | 0.60663
| Rel.Std.Dev. |  | 2.063% | 1.669% | 2.023% | 2.798% | 2.074% | 1.434% | 28% | 0.8814% | 4.01% | 3.031% | 0.9555% | 1.665% | 2.592% | 1.613%
| PDM3 |  | 5.113% | -9.426% | -4.242% | 4.755% | 3.802% | 5.351% | -92.19% | 6.551% | 1.24% | -11.21% | 12515% | -4.302% | 2.372% | 34.43%
| Z-Score (Absolute) |  | 0.25 | 1.56 | 0.92 | 0.21 | 0.09 | 0.28 | 11.9 | 0.43 | 0.23 | 1.79 | 1563.27 | 0.92 | 0.09 | 3.92
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1020: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1038: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 612 (Indicative Value 0.15 wt.%) | [9 empty] | Indicative
C | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 4.01 | [6 empty] | wt.%
|  |  | IRC | IRC | [6 empty] | 3
3
| 1 | 1 | N.A. | 0.15 | [6 empty] | 1
| 1 | 2 | N.A. | 0.15 | [6 empty] | 10
| 1 | 3 | N.A. | [7 empty] | 16
| 1 | 4 | N.A. | [7 empty] | 0.15
| 1 | 5 | N.A. | [7 empty] | 16
| 1 | 6 | N.A.
| Mean |  |  | 0.15
| Median |  |  | 0.15
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for S in OREAS 612 (Indicative Value 0.48 wt.%) | [9 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.16 | 1.18 | 1.19 | 1.20 | 4.01 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | N.A. | N.A. | N.A. | NA |  | --- | -- | 0.48 | 1
| 1 | 2 | N.A. | N.A. | N.A. | NA |  | --- | -- | 0.48 | 10
| 1 | 3 | N.A. | N.A. | N.A. | NA |  | --- | -- |  | 16
| 1 | 4 | N.A. | N.A. | N.A. | NA |  | --- | -- |  | 0.48
| 1 | 5 | N.A. | N.A. | N.A. | NA |  | --- | -- |  | 16
| 1 | 6 | N.A. | N.A. | N.A. | NA |  | --- | --
| Mean | [8 empty] | 0.48
| Median | [8 empty] | 0.48
| Std Dev. | [8 empty] | 0
| Rel.Std.Dev. | [8 empty] | 0%
| PDM3 | [8 empty] | 0%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 612 (Indicative Value 14.38 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.37 | 1
| 1 | 2 | 14.39 | 12
| Mean |  | 14.38 | 16
| Median |  | 14.38 | 14.38
| Std Dev. |  | 0.014142 | 18
| Rel.Std.Dev. |  | 0.09835%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 612 (Indicative Value 0.99 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.99 | 1
| 1 | 2 | 0.99 | 13
| Mean |  | 0.99 | 16
| Median |  | 0.99 | 0.99
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 612 (Indicative Value 3.3 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.29 | 1
| 1 | 2 | 3.3 | 14
| Mean |  | 3.295 | 16
| Median |  | 3.295 | 3.295
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.2146%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 612 (Indicative Value 5.36 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.35 | 1
| 1 | 2 | 5.37 | 15
| Mean |  | 5.36 | 16
| Median |  | 5.36 | 5.36
| Std Dev. |  | 0.014142 | 21
| Rel.Std.Dev. |  | 0.2638%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 612 (Indicative Value 0.46 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.46 | 1
| 1 | 2 | 0.46 | 12
| Mean |  | 0.46 | 16
| Median |  | 0.46 | 0.46
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 612 (Indicative Value 0.06 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.06 | 13
| Mean |  | 0.06 | 16
| Median |  | 0.06 | 0.06
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 612 (Indicative Value 2.28 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.28 | 1
| 1 | 2 | 2.27 | 14
| Mean |  | 2.275 | 16
| Median |  | 2.275 | 2.275
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.3108%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 612 (Indicative Value 0.108 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.107 | 1
| 1 | 2 | 0.108 | 15
| Mean |  | 0.1075 | 16
| Median |  | 0.1075 | 0.1075
| Std Dev. |  | 0.00070711 | 21
| Rel.Std.Dev. |  | 0.6578%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 612 (Indicative Value 0.519 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.5166 | 1
| 1 | 2 | 0.5206 | 10
| Mean |  | 0.5186 | 16
| Median |  | 0.5186 | 0.51858
| Std Dev. |  | 0.0028284 | 18
| Rel.Std.Dev. |  | 0.5454%
| PDM3 |  | 0.003326%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 612 (Indicative Value 69.76 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69.78 | 1
| 1 | 2 | 69.73 | 13
| Mean |  | 69.755 | 16
| Median |  | 69.755 | 69.755
| Std Dev. |  | 0.035355 | 19
| Rel.Std.Dev. |  | 0.05069%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 612 (Indicative Value 0.35 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.35 | 14
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 612 (Indicative Value 2.04 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.04 | 1
| 1 | 2 | 2.03 | 17
| Mean |  | 2.035 | 16
| Median |  | 2.035 | 2.035
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 0.3475%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 612 (Indicative Value 57.5 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 56.8 | 1
| 1 | 2 | 58.1 | 1
| Mean |  | 57.45 | 16
| Median |  | 57.45 | 57.45
| Std Dev. |  | 0.91924 | 25
| Rel.Std.Dev. |  | 1.6%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 612 (Indicative Value 329 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 327 | 1
| 1 | 2 | 330 | 20
| Mean |  | 328.5 | 16
| Median |  | 328.5 | 328.5
| Std Dev. |  | 2.1213 | 26
| Rel.Std.Dev. |  | 0.6458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 612 (Indicative Value 2165 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2190 | 1
| 1 | 2 | 2140 | 21
| Mean |  | 2165 | 16
| Median |  | 2165 | 2165
| Std Dev. |  | 35.3553 | 27
| Rel.Std.Dev. |  | 1.633%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 612 (Indicative Value 2.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 2.8 | 22
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.14142 | 28
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 612 (Indicative Value 0.97 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 0.94 | 23
| Mean |  | 0.97 | 16
| Median |  | 0.97 | 0.97
| Std Dev. |  | 0.042426 | 29
| Rel.Std.Dev. |  | 4.374%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 612 (Indicative Value 15.6 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.7 | 1
| 1 | 2 | 15.4 | 24
| Mean |  | 15.55 | 16
| Median |  | 15.55 | 15.55
| Std Dev. |  | 0.21213 | 30
| Rel.Std.Dev. |  | 1.364%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 612 (Indicative Value 75 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75.4 | 1
| 1 | 2 | 73.9 | 25
| Mean |  | 74.65 | 16
| Median |  | 74.65 | 74.65
| Std Dev. |  | 1.0607 | 31
| Rel.Std.Dev. |  | 1.421%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 612 (Indicative Value 4.3 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.2 | 26
| Mean |  | 4.3 | 16
| Median |  | 4.3 | 4.3
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 612 (Indicative Value 29 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28 | 1
| 1 | 2 | 30 | 27
| Mean |  | 29 | 16
| Median |  | 29 | 29
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 612 (Indicative Value 6.4 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.52 | 1
| 1 | 2 | 6.27 | 28
| Mean |  | 6.395 | 16
| Median |  | 6.395 | 6.395
| Std Dev. |  | 0.17678 | 34
| Rel.Std.Dev. |  | 2.764%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 612 (Indicative Value 104 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 104 | 1
| 1 | 2 | 104 | 29
| Mean |  | 104 | 16
| Median |  | 104 | 104
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 612 (Indicative Value 3.51 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.41 | 1
| 1 | 2 | 3.6 | 30
| Mean |  | 3.505 | 16
| Median |  | 3.505 | 3.505
| Std Dev. |  | 0.13435 | 36
| Rel.Std.Dev. |  | 3.833%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 612 (Indicative Value 1.33 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.39 | 1
| 1 | 2 | 1.27 | 31
| Mean |  | 1.33 | 16
| Median |  | 1.33 | 1.33
| Std Dev. |  | 0.084853 | 37
| Rel.Std.Dev. |  | 6.38%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 612 (Indicative Value 1.4 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.41 | 1
| 1 | 2 | 1.38 | 32
| Mean |  | 1.395 | 16
| Median |  | 1.395 | 1.395
| Std Dev. |  | 0.021213 | 38
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 612 (Indicative Value 19.4 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.8 | 1
| 1 | 2 | 19 | 33
| Mean |  | 19.4 | 16
| Median |  | 19.4 | 19.4
| Std Dev. |  | 0.56569 | 39
| Rel.Std.Dev. |  | 2.916%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 612 (Indicative Value 5.5 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.42 | 1
| 1 | 2 | 5.58 | 34
| Mean |  | 5.5 | 16
| Median |  | 5.5 | 5.5
| Std Dev. |  | 0.11314 | 40
| Rel.Std.Dev. |  | 2.057%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 612 (Indicative Value 1.33 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.3 | 4
| Mean |  | 1.325 | 16
| Median |  | 1.325 | 1.325
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 612 (Indicative Value 6.44 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.27 | 1
| 1 | 2 | 6.6 | 19
| Mean |  | 6.435 | 16
| Median |  | 6.435 | 6.435
| Std Dev. |  | 0.23335 | 25
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 612 (Indicative Value 0.58 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.57 | 1
| 1 | 2 | 0.58 | 20
| Mean |  | 0.575 | 16
| Median |  | 0.575 | 0.575
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 1.23%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 612 (Indicative Value 0.13 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.15 | 21
| Mean |  | 0.125 | 16
| Median |  | 0.125 | 0.125
| Std Dev. |  | 0.035355 | 27
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 612 (Indicative Value 39.6 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39.2 | 1
| 1 | 2 | 39.9 | 22
| Mean |  | 39.55 | 16
| Median |  | 39.55 | 39.55
| Std Dev. |  | 0.49497 | 28
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 612 (Indicative Value 0.15 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.15 | 23
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 612 (Indicative Value 0.047 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0462 | 1
| 1 | 2 | 0.047 | 24
| Mean |  | 0.0466 | 16
| Median |  | 0.0466 | 0.0466
| Std Dev. |  | 0.00056569 | 30
| Rel.Std.Dev. |  | 1.214%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 612 (Indicative Value 7 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7 | 1
| 1 | 2 | 7 | 25
| Mean |  | 7 | 16
| Median |  | 7 | 7
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 612 (Indicative Value 14.5 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.8 | 1
| 1 | 2 | 14.1 | 8
| Mean |  | 14.45 | 16
| Median |  | 14.45 | 14.45
| Std Dev. |  | 0.49497 | 32
| Rel.Std.Dev. |  | 3.425%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 612 (Indicative Value 36 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.2 | 1
| 1 | 2 | 35.7 | 27
| Mean |  | 35.95 | 16
| Median |  | 35.95 | 35.95
| Std Dev. |  | 0.35355 | 33
| Rel.Std.Dev. |  | 0.9835%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 612 (Indicative Value 11 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12 | 1
| 1 | 2 | 10 | 28
| Mean |  | 11 | 16
| Median |  | 11 | 11
| Std Dev. |  | 1.4142 | 34
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 612 (Indicative Value 0.278 wt.%) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.276 | 1
| 1 | 2 | 0.28 | 29
| Mean |  | 0.278 | 16
| Median |  | 0.278 | 0.278
| Std Dev. |  | 0.0028284 | 35
| Rel.Std.Dev. |  | 1.017%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 612 (Indicative Value 9.42 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.28 | 1
| 1 | 2 | 9.55 | 30
| Mean |  | 9.415 | 16
| Median |  | 9.415 | 9.415
| Std Dev. |  | 0.19092 | 36
| Rel.Std.Dev. |  | 2.028%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 612 (Indicative Value 210 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 208 | 1
| 1 | 2 | 212 | 31
| Mean |  | 210 | 16
| Median |  | 210 | 210
| Std Dev. |  | 2.8284 | 37
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 612 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 4
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 38
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 612 (Indicative Value 15.8 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.6 | 1
| 1 | 2 | 15.9 | 33
| Mean |  | 15.75 | 16
| Median |  | 15.75 | 15.75
| Std Dev. |  | 0.21213 | 39
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 612 (Indicative Value 5.4 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.3 | 1
| 1 | 2 | 5.5 | 34
| Mean |  | 5.4 | 16
| Median |  | 5.4 | 5.4
| Std Dev. |  | 0.14142 | 40
| Rel.Std.Dev. |  | 2.619%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 612 (Indicative Value 6.78 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.77 | 1
| 1 | 2 | 6.78 | 35
| Mean |  | 6.775 | 16
| Median |  | 6.775 | 6.775
| Std Dev. |  | 0.0070711 | 41
| Rel.Std.Dev. |  | 0.1044%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 612 (Indicative Value 3.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 3.8 | 19
| Mean |  | 3.8 | 16
| Median |  | 3.8 | 3.8
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 612 (Indicative Value 230 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 227 | 1
| 1 | 2 | 232 | 20
| Mean |  | 229.5 | 16
| Median |  | 229.5 | 229.5
| Std Dev. |  | 3.5355 | 26
| Rel.Std.Dev. |  | 1.541%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 612 (Indicative Value 1.13 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.12 | 1
| 1 | 2 | 1.14 | 21
| Mean |  | 1.13 | 16
| Median |  | 1.13 | 1.13
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 612 (Indicative Value 0.75 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.74 | 22
| Mean |  | 0.75 | 16
| Median |  | 0.75 | 0.75
| Std Dev. |  | 0.014142 | 28
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 612 (Indicative Value 0.15 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 29
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 612 (Indicative Value 13.5 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.3 | 1
| 1 | 2 | 13.7 | 24
| Mean |  | 13.5 | 16
| Median |  | 13.5 | 13.5
| Std Dev. |  | 0.28284 | 30
| Rel.Std.Dev. |  | 2.095%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 612 (Indicative Value 0.21 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.21 | 1
| 1 | 2 | 0.209 | 25
| Mean |  | 0.2095 | 16
| Median |  | 0.2095 | 0.2095
| Std Dev. |  | 0.00070711 | 31
| Rel.Std.Dev. |  | 0.3375%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 612 (Indicative Value 2.8 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 3 | 26
| Mean |  | 2.8 | 16
| Median |  | 2.8 | 2.8
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 10.1%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 612 (Indicative Value 0.18 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.18 | 1
| 1 | 2 | 0.17 | 27
| Mean |  | 0.175 | 16
| Median |  | 0.175 | 0.175
| Std Dev. |  | 0.0070711 | 33
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 612 (Indicative Value 4.4 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.47 | 1
| 1 | 2 | 4.33 | 28
| Mean |  | 4.4 | 16
| Median |  | 4.4 | 4.4
| Std Dev. |  | 0.098995 | 34
| Rel.Std.Dev. |  | 2.25%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 612 (Indicative Value 30.7 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.5 | 1
| 1 | 2 | 30.9 | 29
| Mean |  | 30.7 | 16
| Median |  | 30.7 | 30.7
| Std Dev. |  | 0.28284 | 35
| Rel.Std.Dev. |  | 0.9213%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 612 (Indicative Value 3.25 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3 | 30
| Mean |  | 3.25 | 16
| Median |  | 3.25 | 3.25
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 612 (Indicative Value 16.2 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.2 | 1
| 1 | 2 | 16.1 | 31
| Mean |  | 16.15 | 16
| Median |  | 16.15 | 16.15
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.4378%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 612 (Indicative Value 1.05 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.06 | 1
| 1 | 2 | 1.03 | 32
| Mean |  | 1.045 | 16
| Median |  | 1.045 | 1.045
| Std Dev. |  | 0.021213 | 38
| Rel.Std.Dev. |  | 2.03%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 612 (Indicative Value 0.371 wt.%) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.377 | 1
| 1 | 2 | 0.364 | 33
| Mean |  | 0.3705 | 16
| Median |  | 0.3705 | 0.3705
| Std Dev. |  | 0.0091924 | 39
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 612 (Indicative Value 220 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 221 | 1
| 1 | 2 | 218 | 34
| Mean |  | 219.5 | 16
| Median |  | 219.5 | 219.5
| Std Dev. |  | 2.1213 | 40
| Rel.Std.Dev. |  | 0.9664%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.