## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 265
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 1.0761 | 1.0634 | 1.0887 | 1.0679 | 1.0843
PhotonAssay
Au, Gold (ppm) | 1.092 | 1.0796 | 1.1045 | 1.0899 | 1.0941
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.0776 | 1.0494 | 1.1058 | 1.0686 | 1.0866
Cyanide Leach
Au, Gold (ppm) | 1.0388 | 1.0186 | 1.0591 | 1.0358 | 1.0419
4-Acid Digestion
Ag, Silver (ppm) | 0.68001 | 0.63479 | 0.72523 | 0.65225 | 0.70778
Al, Aluminium (wt.%) | 6.6232 | 6.4214 | 6.8251 | 6.4593 | 6.7871
As, Arsenic (ppm) | 220.7193 | 212.5832 | 228.8554 | 214.1902 | 227.2484
Ba, Barium (ppm) | 532.7349 | 514.961 | 550.5088 | 523.338 | 542.1318
Be, Beryllium (ppm) | 2.5493 | 2.3965 | 2.7021 | 2.4679 | 2.6307
Bi, Bismuth (ppm) | 2.1103 | 1.9524 | 2.2683 | 2.024 | 2.1967
Ca, Calcium (wt.%) | 0.85654 | 0.82425 | 0.88884 | 0.83684 | 0.87624
Ce, Cerium (ppm) | 92.6624 | 88.5632 | 96.7615 | 90.0318 | 95.2929
Co, Cobalt (ppm) | 10.9749 | 10.5774 | 11.3725 | 10.7368 | 11.213
Cr, Chromium (ppm) | 87.6835 | 79.4335 | 95.9334 | 83.9844 | 91.3826
Cs, Caesium (ppm) | 6.1963 | 6.0107 | 6.3819 | 6.0456 | 6.347
Cu, Copper (ppm) | 23.4881 | 22.2069 | 24.7693 | 22.7317 | 24.2445
Dy, Dysprosium (ppm) | 3.9631 | 3.7652 | 4.1609 | 3.7837 | 4.1424
Er, Erbium (ppm) | 1.7679 | 1.6684 | 1.8673 | 1.6916 | 1.8441
Eu, Europium (ppm) | 1.6106 | 1.5072 | 1.7139 | 1.5393 | 1.6818
Fe, Iron (wt.%) | 3.4435 | 3.3394 | 3.5477 | 3.3814 | 3.5056
Ga, Gallium (ppm) | 19.3499 | 18.7291 | 19.9706 | 18.8507 | 19.849
Gd, Gadolinium (ppm) | 5.7305 | 5.4165 | 6.0445 | 5.5439 | 5.9171
Hf, Hafnium (ppm) | 4.8943 | 4.6353 | 5.1532 | 4.6845 | 5.104
Ho, Holmium (ppm) | 0.69122 | 0.60921 | 0.77323 | 0.63603 | 0.74641
In, Indium (ppm) | 0.069946 | 0.064349 | 0.075543 | 0.065183 | 0.074709
K, Potassium (wt.%) | 1.9491 | 1.8936 | 2.0045 | 1.9083 | 1.9898
La, Lanthanum (ppm) | 49.801 | 47.5718 | 52.0301 | 48.6461 | 50.9559
Li, Lithium (ppm) | 35.6208 | 34.1607 | 37.0809 | 34.6252 | 36.6164
Lu, Lutetium (ppm) | 0.24036 | 0.21451 | 0.2662 | 0.22244 | 0.25828
Mg, Magnesium (wt.%) | 0.84101 | 0.8154 | 0.86663 | 0.8289 | 0.85313
Mn, Manganese (wt.%) | 0.030778 | 0.029688 | 0.031869 | 0.030151 | 0.031406
Mo, Molybdenum (ppm) | 4.219 | 4.0178 | 4.4202 | 4.0671 | 4.3709
Na, Sodium (wt.%) | 0.76041 | 0.73609 | 0.78473 | 0.74238 | 0.77844
Nb, Niobium (ppm) | 26.1889 | 24.9769 | 27.4009 | 25.2094 | 27.1685
Nd, Neodymium (ppm) | 41.1121 | 38.9048 | 43.3195 | 40.4282 | 41.7961
Ni, Nickel (ppm) | 53.0533 | 51.3525 | 54.7542 | 51.8032 | 54.3034
P, Phosphorus (wt.%) | 0.077551 | 0.074222 | 0.080881 | 0.075916 | 0.079186
Pb, Lead (ppm) | 16.1234 | 15.2393 | 17.0075 | 15.4714 | 16.7755
Pr, Praseodymium (ppm) | 10.9893 | 10.478 | 11.5006 | 10.5187 | 11.4599
Rb, Rubidium (ppm) | 100.1058 | 97.0576 | 103.1539 | 97.8386 | 102.373
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.0092278 | 0.0084145 | 0.010041 | 0.0088124 | 0.0096432
Sb, Antimony (ppm) | 2.9816 | 2.8214 | 3.1418 | 2.8424 | 3.1208
Sc, Scandium (ppm) | 11.7215 | 11.2242 | 12.2188 | 11.3754 | 12.0677
Sm, Samarium (ppm) | 7.4175 | 7.0292 | 7.8058 | 7.2021 | 7.6328
Sn, Tin (ppm) | 8.8218 | 8.3574 | 9.2861 | 8.5502 | 9.0933
Sr, Strontium (ppm) | 260.7875 | 251.9452 | 269.6298 | 256.3284 | 265.2467
Ta, Tantalum (ppm) | 1.65 | 1.5387 | 1.7612 | 1.5598 | 1.7402
Tb, Terbium (ppm) | 0.76366 | 0.72215 | 0.80516 | 0.73075 | 0.79656
Te, Tellurium (ppm) | 0.059722 | 0.043216 | 0.076229 | IND | IND
Th, Thorium (ppm) | 13.5889 | 12.9164 | 14.2613 | 13.1816 | 13.9961
Ti, Titanium (wt.%) | 0.4542 | 0.43029 | 0.47812 | 0.44269 | 0.46571
Tl, Thallium (ppm) | 0.49178 | 0.46474 | 0.51883 | 0.47222 | 0.51135
Tm, Thulium (ppm) | 0.24308 | 0.21199 | 0.27418 | 0.22404 | 0.26213
U, Uranium (ppm) | 2.525 | 2.4031 | 2.647 | 2.4216 | 2.6285
V, Vanadium (ppm) | 78.6616 | 75.4127 | 81.9105 | 77.0737 | 80.2495
W, Tungsten (ppm) | 11.2347 | 10.5638 | 11.9055 | 10.5084 | 11.9609
Y, Yttrium (ppm) | 17.4459 | 16.5486 | 18.3432 | 16.9375 | 17.9543
Yb, Ytterbium (ppm) | 1.5868 | 1.4683 | 1.7053 | 1.5097 | 1.6639
Zn, Zinc (ppm) | 61.0228 | 58.6589 | 63.3867 | 59.2804 | 62.7651
Zr, Zirconium (ppm) | 196.974 | 187.9304 | 206.0176 | 193.1698 | 200.7782
Aqua Regia Digestion
Ag, Silver (ppm) | 0.62905 | 0.59751 | 0.66059 | 0.60281 | 0.65529
Al, Aluminium (wt.%) | 1.1622 | 1.1139 | 1.2105 | 1.1406 | 1.1839
As, Arsenic (ppm) | 180.7574 | 176.0119 | 185.5029 | 176.7821 | 184.7327
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 70.0016 | 65.8114 | 74.1918 | 68.2736 | 71.7296
Be, Beryllium (ppm) | 1.0502 | 0.97789 | 1.1224 | 1.004 | 1.0963
Bi, Bismuth (ppm) | 1.6428 | 1.513 | 1.7726 | 1.5631 | 1.7226
Ca, Calcium (wt.%) | 0.25726 | 0.24745 | 0.26707 | 0.25067 | 0.26385
Cd, Cadmium (ppm) | 0.029489 | 0.019553 | 0.039425 | IND | IND
Ce, Cerium (ppm) | 53.0199 | 50.6509 | 55.3889 | 51.7281 | 54.3117
Co, Cobalt (ppm) | 8.1313 | 7.8043 | 8.4583 | 7.9508 | 8.3119
Cr, Chromium (ppm) | 35.9976 | 34.1408 | 37.8543 | 34.9583 | 37.0368
Cs, Caesium (ppm) | 1.2666 | 1.1794 | 1.3538 | 1.2247 | 1.3085
Cu, Copper (ppm) | 20.3415 | 19.5393 | 21.1438 | 19.8429 | 20.8402
Dy, Dysprosium (ppm) | 2.3257 | 2.1874 | 2.4641 | 2.2459 | 2.4056
Er, Erbium (ppm) | 0.87272 | 0.81694 | 0.9285 | 0.83948 | 0.90596
Eu, Europium (ppm) | 0.96964 | 0.8869 | 1.0524 | 0.93157 | 1.0077
Fe, Iron (wt.%) | 2.5259 | 2.4166 | 2.6353 | 2.4739 | 2.578
Ga, Gallium (ppm) | 4.7417 | 4.4136 | 5.0698 | 4.62 | 4.8634
Gd, Gadolinium (ppm) | 3.6635 | 3.3553 | 3.9717 | 3.5264 | 3.8006
Ge, Germanium (ppm) | < 0.05 | IND | IND | IND | IND
Hf, Hafnium (ppm) | 0.55511 | 0.45688 | 0.65335 | 0.53075 | 0.57948
Hg, Mercury (ppm) | 0.013244 | 0.010263 | 0.016224 | IND | IND
Ho, Holmium (ppm) | 0.36075 | 0.32112 | 0.40038 | 0.32838 | 0.39312
In, Indium (ppm) | 0.018897 | 0.01687 | 0.020925 | IND | IND
K, Potassium (wt.%) | 0.25383 | 0.24386 | 0.26381 | 0.24622 | 0.26144
La, Lanthanum (ppm) | 27.5074 | 26.246 | 28.7689 | 26.644 | 28.3709
Li, Lithium (ppm) | 6.191 | 5.8706 | 6.5114 | 5.9544 | 6.4275
Lu, Lutetium (ppm) | 0.076415 | 0.069316 | 0.083515 | IND | IND
Mg, Magnesium (wt.%) | 0.4164 | 0.39782 | 0.43498 | 0.40584 | 0.42697
Mn, Manganese (wt.%) | 0.019813 | 0.019062 | 0.020564 | 0.019407 | 0.020219
Mo, Molybdenum (ppm) | 3.5046 | 3.3586 | 3.6506 | 3.3736 | 3.6355
Na, Sodium (wt.%) | 0.23332 | 0.2229 | 0.24374 | 0.22596 | 0.24068
Nd, Neodymium (ppm) | 25.3886 | 24.2474 | 26.5298 | 24.9352 | 25.842
Ni, Nickel (ppm) | 42.4328 | 41.0082 | 43.8574 | 41.3051 | 43.5605
P, Phosphorus (wt.%) | 0.047391 | 0.045573 | 0.049208 | 0.046353 | 0.048429
Pb, Lead (ppm) | 8.4444 | 7.9383 | 8.9504 | 8.1396 | 8.7492
Pr, Praseodymium (ppm) | 6.7131 | 6.4128 | 7.0135 | 6.5709 | 6.8554
Rb, Rubidium (ppm) | 15.096 | 14.3123 | 15.8797 | 14.6935 | 15.4985
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | < 0.01 | IND | IND | IND | IND
Sb, Antimony (ppm) | 1.2433 | 1.1623 | 1.3242 | 1.183 | 1.3035
Sc, Scandium (ppm) | 2.8183 | 2.6655 | 2.9711 | 2.7124 | 2.9243
Sm, Samarium (ppm) | 4.4207 | 3.9699 | 4.8714 | 4.29 | 4.5513
Sn, Tin (ppm) | 1.4004 | 1.3059 | 1.4949 | 1.3173 | 1.4835
Sr, Strontium (ppm) | 39.8064 | 38.1505 | 41.4623 | 38.9638 | 40.6489
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.44915 | 0.42379 | 0.47451 | 0.43268 | 0.46562
Te, Tellurium (ppm) | 0.034015 | 0.021855 | 0.046175 | IND | IND
Th, Thorium (ppm) | 8.7598 | 8.3447 | 9.1749 | 8.5549 | 8.9646
Ti, Titanium (wt.%) | 0.072088 | 0.068165 | 0.076011 | 0.06985 | 0.074326
Tl, Thallium (ppm) | 0.080391 | 0.071014 | 0.089768 | IND | IND
Tm, Thulium (ppm) | 0.10151 | 0.089448 | 0.11356 | IND | IND
U, Uranium (ppm) | 1.2328 | 1.1793 | 1.2864 | 1.1955 | 1.2702
V, Vanadium (ppm) | 27.2565 | 26.1189 | 28.394 | 26.4629 | 28.0501
W, Tungsten (ppm) | 2.0302 | 1.8178 | 2.2426 | 1.8856 | 2.1748
Y, Yttrium (ppm) | 8.7041 | 8.3786 | 9.0296 | 8.5079 | 8.9002
Yb, Ytterbium (ppm) | 0.58151 | 0.53235 | 0.63068 | 0.55142 | 0.61161
Zn, Zinc (ppm) | 36.3123 | 34.8017 | 37.8229 | 35.2417 | 37.3829
Zr, Zirconium (ppm) | 31.0582 | 27.7549 | 34.3614 | 29.8495 | 32.2669
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 265
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Infrared Combustion
C | wt.% | 0.04 | S | wt.% | 0.02
4-Acid Digestion
B | ppm | 32 | Ge | ppm | 0.19505 | Se | ppm | 0.611
Cd | ppm | 0.04506 | Hg | ppm | 0.016667
Aqua Regia Digestion
Nb | ppm | 0.5541 | Pt | ppb | < 5
Pd | ppb | < 10 | Se | ppm | 0.31083
Borate Fusion XRF
Al2O3 | wt.% | 12.86 | Fe2O3 | wt.% | 5 | S | wt.% | 0.0084095
As | ppm | 200 | K2O | wt.% | 2.355 | SiO2 | wt.% | 71.925
BaO | ppm | 610 | MgO | wt.% | 1.47 | Sn | ppm | 20
CaO | wt.% | 1.205 | MnO | wt.% | 0.0395 | Sr | ppm | 211.3986
Cl | ppm | 175 | Na2O | wt.% | 1.025 | TiO2 | wt.% | 0.8495
Co | ppm | 20 | Ni | ppm | 60 | V2O5 | ppm | 155
Cr2O3 | ppm | 165 | P2O5 | wt.% | 0.175 | Zn | ppm | 65
Cu | ppm | 30 | Pb | ppm | 45 | Zr | ppm | 266.5159
Thermogravimetry
LOI1000 | wt.% | 2.795
Laser Ablation ICP-MS
Ag | ppm | 0.65 | Hf | ppm | 7.26 | Sn | ppm | 11
As | ppm | 199 | Ho | ppm | 0.995 | Sr | ppm | 254
Ba | ppm | 523.5 | In | ppm | 0.05 | Ta | ppm | 2.085
Be | ppm | 2.6 | La | ppm | 50.1 | Tb | ppm | 0.93
Bi | ppm | 2.16 | Lu | ppm | 0.345 | Te | ppm | < 0.2
Cd | ppm | 0.075 | Mn | wt.% | 0.03155 | Th | ppm | 13.7
Ce | ppm | 92.5 | Mo | ppm | 4.1 | Ti | wt.% | 0.5255
Co | ppm | 11.1 | Nb | ppm | 30.25 | Tl | ppm | 0.5
Cr | ppm | 113 | Nd | ppm | 41.45 | Tm | ppm | 0.38
Cs | ppm | 6.08 | Ni | ppm | 55 | U | ppm | 2.815
Cu | ppm | 24 | Pb | ppm | 15.5 | V | ppm | 83.95
Dy | ppm | 5.16 | Pr | ppm | 11.5 | W | ppm | 12.25
Er | ppm | 2.725 | Rb | ppm | 97.3 | Y | ppm | 25.65
Eu | ppm | 1.59 | Re | ppm | 0.0125 | Yb | ppm | 2.485
Ga | ppm | 19.05 | Sb | ppm | 2.95 | Zn | ppm | 55
Gd | ppm | 5.965 | Sc | ppm | 10.95 | Zr | ppm | 279
Ge | ppm | 1.4 | Sm | ppm | 7.755
Instrumental Neutron Activation Analysis
Ag | ppm | < 5 | Fe | wt.% | 3.6 | Sb | ppm | 3.1
As | ppm | 229 | Hg | ppm | < 10 | Sc | ppm | 12.7
Au | ppm | 1.19 | Ir | ppm | < 0.1 | Se | ppm | < 10
Ba | ppm | 600 | La | ppm | 55 | Sm | ppm | 7.9
Br | ppm | < 1 | Mn | wt.% | 0.0312 | Ta | ppm | 2
Cd | ppm | < 10 | Mo | ppm | < 10 | Th | ppm | 13.4
Co | ppm | < 10 | Na | wt.% | 0.75 | U | ppm | 3
Cr | ppm | 120 | Ni | ppm | < 100 | W | ppm | 16
Cu | ppm | < 10000 | Rb | ppm | 100 | Zn | ppm | < 200
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 265
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 1.0761 | 0.033148 | 1.0098 | 1.1424 | 0.97663 | 1.1755 | 3.08% | 6.161% | 9.241% | 1.0223 | 1.1299
PhotonAssay
Au, ppm | 1.092 | 0.046175 | 0.99966 | 1.1844 | 0.95349 | 1.2305 | 4.228% | 8.457% | 12.69% | 1.0374 | 1.1466
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.0776 | 0.05748 | 0.96266 | 1.1926 | 0.90518 | 1.2501 | 5.334% | 10.67% | 16% | 1.0237 | 1.1315
Cyanide Leach
Au, ppm | 1.0388 | 0.043112 | 0.95262 | 1.1251 | 0.90951 | 1.1682 | 4.15% | 8.3% | 12.45% | 0.98691 | 1.0908
4-Acid Digestion
Ag, ppm | 0.68001 | 0.043117 | 0.59378 | 0.76625 | 0.55066 | 0.80936 | 6.341% | 12.68% | 19.02% | 0.64601 | 0.71402
Al, wt.% | 6.6232 | 0.19563 | 6.232 | 7.0145 | 6.0363 | 7.2101 | 2.954% | 5.907% | 8.861% | 6.2921 | 6.9544
As, ppm | 220.7193 | 6.7696 | 207.1801 | 234.2585 | 200.4105 | 241.0281 | 3.067% | 6.134% | 9.201% | 209.6833 | 231.7552
Ba, ppm | 532.7349 | 15.2516 | 502.2316 | 563.2382 | 486.98 | 578.4898 | 2.863% | 5.726% | 8.589% | 506.0982 | 559.3716
Be, ppm | 2.5493 | 0.14947 | 2.2503 | 2.8482 | 2.1009 | 2.9977 | 5.863% | 11.73% | 17.59% | 2.4218 | 2.6767
Bi, ppm | 2.1103 | 0.1189 | 1.8725 | 2.3481 | 1.7536 | 2.467 | 5.634% | 11.27% | 16.9% | 2.0048 | 2.2159
Ca, wt.% | 0.85654 | 0.027227 | 0.80209 | 0.91099 | 0.77486 | 0.93822 | 3.179% | 6.357% | 9.536% | 0.81371 | 0.89937
Ce, ppm | 92.6624 | 5.5051 | 81.6522 | 103.6725 | 76.1471 | 109.1776 | 5.941% | 11.88% | 17.82% | 88.0292 | 97.2955
Co, ppm | 10.9749 | 0.31359 | 10.3477 | 11.6021 | 10.0342 | 11.9157 | 2.857% | 5.715% | 8.572% | 10.4262 | 11.5237
Cr, ppm | 87.6835 | 15.8396 | 56.0043 | 119.3627 | 40.1646 | 135.2023 | 18.06% | 36.13% | 54.19% | 83.2993 | 92.0676
Cs, ppm | 6.1963 | 0.23758 | 5.7211 | 6.6714 | 5.4835 | 6.909 | 3.834% | 7.669% | 11.5% | 5.8865 | 6.5061
Cu, ppm | 23.4881 | 1.6226 | 20.2428 | 26.7334 | 18.6202 | 28.356 | 6.908% | 13.82% | 20.72% | 22.3137 | 24.6625
Dy, ppm | 3.9631 | 0.12794 | 3.7072 | 4.219 | 3.5792 | 4.3469 | 3.228% | 6.457% | 9.685% | 3.7649 | 4.1612
Er, ppm | 1.7679 | 0.072885 | 1.6221 | 1.9136 | 1.5492 | 1.9865 | 4.123% | 8.246% | 12.37% | 1.6795 | 1.8562
Eu, ppm | 1.6106 | 0.09051 | 1.4296 | 1.7916 | 1.339 | 1.8821 | 5.62% | 11.24% | 16.86% | 1.53 | 1.6911
Fe, wt.% | 3.4435 | 0.090906 | 3.2617 | 3.6253 | 3.1708 | 3.7162 | 2.64% | 5.28% | 7.92% | 3.2713 | 3.6157
Ga, ppm | 19.3499 | 0.89836 | 17.5532 | 21.1466 | 16.6548 | 22.045 | 4.643% | 9.285% | 13.93% | 18.3824 | 20.3174
Gd, ppm | 5.7305 | 0.23297 | 5.2646 | 6.1965 | 5.0316 | 6.4294 | 4.066% | 8.131% | 12.2% | 5.444 | 6.017
Hf, ppm | 4.8943 | 0.27704 | 4.3402 | 5.4483 | 4.0631 | 5.7254 | 5.661% | 11.32% | 16.98% | 4.6495 | 5.139
Ho, ppm | 0.69122 | 0.056207 | 0.5788 | 0.80363 | 0.5226 | 0.85984 | 8.132% | 16.26% | 24.39% | 0.65666 | 0.72578
In, ppm | 0.069946 | 0.004499 | 0.060948 | 0.078944 | 0.056449 | 0.083443 | 6.432% | 12.86% | 19.3% | 0.066449 | 0.073443
K, wt.% | 1.9491 | 0.048587 | 1.8519 | 2.0462 | 1.8033 | 2.0948 | 2.493% | 4.986% | 7.479% | 1.8516 | 2.0465
La, ppm | 49.801 | 2.2722 | 45.2566 | 54.3454 | 42.9844 | 56.6175 | 4.563% | 9.125% | 13.69% | 47.3109 | 52.291
Li, ppm | 35.6208 | 1.4666 | 32.6876 | 38.554 | 31.221 | 40.0206 | 4.117% | 8.235% | 12.35% | 33.8398 | 37.4018
Lu, ppm | 0.24036 | 0.013967 | 0.21242 | 0.26829 | 0.19846 | 0.28226 | 5.811% | 11.62% | 17.43% | 0.22834 | 0.25238
Mg, wt.% | 0.84101 | 0.022156 | 0.7967 | 0.88533 | 0.77455 | 0.90748 | 2.634% | 5.269% | 7.903% | 0.79896 | 0.88307
Mn, wt.% | 0.030778 | 0.0010392 | 0.0287 | 0.032857 | 0.02766 | 0.033896 | 3.377% | 6.753% | 10.13% | 0.029239 | 0.032317
Mo, ppm | 4.219 | 0.19419 | 3.8306 | 4.6074 | 3.6364 | 4.8016 | 4.603% | 9.205% | 13.81% | 4.0081 | 4.43
Na, wt.% | 0.76041 | 0.026869 | 0.70667 | 0.81414 | 0.6798 | 0.84101 | 3.534% | 7.067% | 10.6% | 0.72239 | 0.79843
Nb, ppm | 26.1889 | 1.4903 | 23.2083 | 29.1695 | 21.718 | 30.6598 | 5.691% | 11.38% | 17.07% | 24.8795 | 27.4984
Nd, ppm | 41.1121 | 1.8329 | 37.4464 | 44.7779 | 35.6135 | 46.6108 | 4.458% | 8.916% | 13.37% | 39.0565 | 43.1677
Ni, ppm | 53.0533 | 1.5674 | 49.9185 | 56.1881 | 48.3511 | 57.7556 | 2.954% | 5.909% | 8.863% | 50.4007 | 55.706
P, wt.% | 0.077551 | 0.0024494 | 0.072653 | 0.08245 | 0.070203 | 0.0849 | 3.158% | 6.317% | 9.475% | 0.073674 | 0.081429
Pb, ppm | 16.1234 | 1.0002 | 14.123 | 18.1239 | 13.1227 | 19.1241 | 6.204% | 12.41% | 18.61% | 15.3173 | 16.9296
Pr, ppm | 10.9893 | 0.36785 | 10.2536 | 11.725 | 9.8857 | 12.0928 | 3.347% | 6.695% | 10.04% | 10.4398 | 11.5387
Rb, ppm | 100.1058 | 2.6143 | 94.8773 | 105.3343 | 92.263 | 107.9486 | 2.611% | 5.223% | 7.834% | 95.1005 | 105.1111
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.0092278 | 0.0010228 | 0.0071821 | 0.011274 | 0.0061593 | 0.012296 | 11.08% | 22.17% | 33.25% | 0.0087664 | 0.0096892
Sb, ppm | 2.9816 | 0.17476 | 2.6321 | 3.3311 | 2.4573 | 3.5059 | 5.861% | 11.72% | 17.58% | 2.8325 | 3.1307
Sc, ppm | 11.7215 | 0.60726 | 10.507 | 12.936 | 9.8997 | 13.5433 | 5.181% | 10.36% | 15.54% | 11.1354 | 12.3076
Sm, ppm | 7.4175 | 0.20858 | 7.0003 | 7.8346 | 6.7918 | 8.0432 | 2.812% | 5.624% | 8.436% | 7.0466 | 7.7884
Sn, ppm | 8.8218 | 0.47057 | 7.8806 | 9.7629 | 7.41 | 10.2335 | 5.334% | 10.67% | 16% | 8.3807 | 9.2628
Sr, ppm | 260.7875 | 7.2911 | 246.2054 | 275.3697 | 238.9143 | 282.6608 | 2.796% | 5.592% | 8.387% | 247.7481 | 273.8269
Ta, ppm | 1.65 | 0.15189 | 1.3462 | 1.9538 | 1.1943 | 2.1056 | 9.205% | 18.41% | 27.62% | 1.5675 | 1.7325
Tb, ppm | 0.76366 | 0.033266 | 0.69713 | 0.83019 | 0.66386 | 0.86346 | 4.356% | 8.712% | 13.07% | 0.72548 | 0.80184
Te, ppm | 0.059722 | 0.0089522 | 0.041818 | 0.077627 | 0.032866 | 0.086579 | 14.99% | 29.98% | 44.97% | 0.056736 | 0.062708
Th, ppm | 13.5889 | 0.76384 | 12.0612 | 15.1165 | 11.2973 | 15.8804 | 5.621% | 11.24% | 16.86% | 12.9094 | 14.2683
Ti, wt.% | 0.4542 | 0.018415 | 0.41737 | 0.49103 | 0.39896 | 0.50945 | 4.054% | 8.109% | 12.16% | 0.43149 | 0.47691
Tl, ppm | 0.49178 | 0.02649 | 0.4388 | 0.54476 | 0.41231 | 0.57126 | 5.387% | 10.77% | 16.16% | 0.46719 | 0.51637
Tm, ppm | 0.24308 | 0.026857 | 0.18937 | 0.2968 | 0.16251 | 0.32365 | 11.05% | 22.1% | 33.15% | 0.23093 | 0.25524
U, ppm | 2.525 | 0.10432 | 2.3164 | 2.7337 | 2.2121 | 2.838 | 4.131% | 8.262% | 12.39% | 2.3988 | 2.6513
V, ppm | 78.6616 | 2.5875 | 73.4867 | 83.8366 | 70.8992 | 86.424 | 3.289% | 6.579% | 9.868% | 74.7285 | 82.5947
W, ppm | 11.2347 | 0.50546 | 10.2237 | 12.2456 | 9.7183 | 12.751 | 4.499% | 8.998% | 13.5% | 10.6729 | 11.7964
Y, ppm | 17.4459 | 0.90819 | 15.6295 | 19.2623 | 14.7213 | 20.1705 | 5.206% | 10.41% | 15.62% | 16.5736 | 18.3182
Yb, ppm | 1.5868 | 0.081863 | 1.4231 | 1.7505 | 1.3412 | 1.8324 | 5.159% | 10.32% | 15.48% | 1.5074 | 1.6661
Zn, ppm | 61.0228 | 1.4702 | 58.0824 | 63.9631 | 56.6122 | 65.4333 | 2.409% | 4.818% | 7.228% | 57.9716 | 64.0739
Zr, ppm | 196.974 | 8.7141 | 179.5458 | 214.4021 | 170.8318 | 223.1162 | 4.424% | 8.848% | 13.27% | 187.1253 | 206.8227
Aqua Regia Digestion
Ag, ppm | 0.62905 | 0.026654 | 0.57574 | 0.68236 | 0.54909 | 0.70901 | 4.237% | 8.474% | 12.71% | 0.5976 | 0.6605
Al, wt.% | 1.1622 | 0.073744 | 1.0148 | 1.3097 | 0.94101 | 1.3835 | 6.345% | 12.69% | 19.03% | 1.1041 | 1.2204
As, ppm | 180.7574 | 5.9901 | 168.7771 | 192.7376 | 162.787 | 198.7278 | 3.314% | 6.628% | 9.942% | 171.7195 | 189.7953
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 70.0016 | 5.132 | 59.7377 | 80.2656 | 54.6057 | 85.3975 | 7.331% | 14.66% | 21.99% | 66.5015 | 73.5017
Be, ppm | 1.0502 | 0.069347 | 0.91146 | 1.1888 | 0.84211 | 1.2582 | 6.604% | 13.21% | 19.81% | 0.99765 | 1.1027
Bi, ppm | 1.6428 | 0.1299 | 1.383 | 1.9026 | 1.2532 | 2.0325 | 7.907% | 15.81% | 23.72% | 1.5607 | 1.725
Ca, wt.% | 0.25726 | 0.011432 | 0.2344 | 0.28012 | 0.22296 | 0.29156 | 4.444% | 8.888% | 13.33% | 0.2444 | 0.27012
Cd, ppm | 0.029489 | 0.0058544 | 0.01778 | 0.041198 | 0.011926 | 0.047052 | 19.85% | 39.71% | 59.56% | 0.028014 | 0.030963
Ce, ppm | 53.0199 | 2.717 | 47.586 | 58.4538 | 44.869 | 61.1708 | 5.124% | 10.25% | 15.37% | 50.3689 | 55.6709
Co, ppm | 8.1313 | 0.50239 | 7.1266 | 9.1361 | 6.6242 | 9.6385 | 6.178% | 12.36% | 18.54% | 7.7248 | 8.5379
Cr, ppm | 35.9976 | 3.2389 | 29.5197 | 42.4754 | 26.2808 | 45.7143 | 8.998% | 18% | 26.99% | 34.1977 | 37.7974
Cs, ppm | 1.2666 | 0.12503 | 1.0165 | 1.5167 | 0.89151 | 1.6417 | 9.872% | 19.74% | 29.61% | 1.2033 | 1.3299
Cu, ppm | 20.3415 | 0.79347 | 18.7546 | 21.9285 | 17.9611 | 22.722 | 3.901% | 7.801% | 11.7% | 19.3245 | 21.3586
Dy, ppm | 2.3257 | 0.094917 | 2.1359 | 2.5156 | 2.041 | 2.6105 | 4.081% | 8.162% | 12.24% | 2.2094 | 2.442
Er, ppm | 0.87272 | 0.037224 | 0.79827 | 0.94717 | 0.76105 | 0.98439 | 4.265% | 8.531% | 12.8% | 0.82908 | 0.91636
Eu, ppm | 0.96964 | 0.046291 | 0.87705 | 1.0622 | 0.83076 | 1.1085 | 4.774% | 9.548% | 14.32% | 0.92115 | 1.0181
Fe, wt.% | 2.5259 | 0.15917 | 2.2076 | 2.8443 | 2.0484 | 3.0034 | 6.301% | 12.6% | 18.9% | 2.3996 | 2.6522
Ga, ppm | 4.7417 | 0.51066 | 3.7204 | 5.763 | 3.2097 | 6.2737 | 10.77% | 21.54% | 32.31% | 4.5046 | 4.9788
Gd, ppm | 3.6635 | 0.20543 | 3.2527 | 4.0744 | 3.0472 | 4.2798 | 5.607% | 11.21% | 16.82% | 3.4803 | 3.8467
Ge, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Hf, ppm | 0.55511 | 0.15505 | 0.24502 | 0.86521 | 0.089973 | 1.0203 | 27.93% | 55.86% | 83.79% | 0.52736 | 0.58287
Hg, ppm | 0.013244 | 0.004759 | 0.0037257 | 0.022762 | 0 | 0.027521 | 35.93% | 71.87% | 107.8% | 0.012582 | 0.013906
Ho, ppm | 0.36075 | 0.03383 | 0.29309 | 0.42841 | 0.25926 | 0.46224 | 9.378% | 18.76% | 28.13% | 0.34271 | 0.37879
In, ppm | 0.018897 | 0.0016497 | 0.015598 | 0.022197 | 0.013948 | 0.023846 | 8.73% | 17.46% | 26.19% | 0.017953 | 0.019842
K, wt.% | 0.25383 | 0.015692 | 0.22245 | 0.28522 | 0.20675 | 0.30091 | 6.182% | 12.36% | 18.55% | 0.24114 | 0.26652
La, ppm | 27.5074 | 1.3172 | 24.873 | 30.1419 | 23.5557 | 31.4591 | 4.789% | 9.577% | 14.37% | 26.1321 | 28.8828
Li, ppm | 6.191 | 0.38476 | 5.4215 | 6.9605 | 5.0367 | 7.3452 | 6.215% | 12.43% | 18.64% | 5.8814 | 6.5005
Lu, ppm | 0.076415 | 0.0042637 | 0.067888 | 0.084943 | 0.063624 | 0.089206 | 5.58% | 11.16% | 16.74% | 0.072594 | 0.080236
Mg, wt.% | 0.4164 | 0.024491 | 0.36742 | 0.46538 | 0.34293 | 0.48988 | 5.882% | 11.76% | 17.64% | 0.39558 | 0.43722
Mn, wt.% | 0.019813 | 0.0011574 | 0.017498 | 0.022128 | 0.016341 | 0.023285 | 5.841% | 11.68% | 17.52% | 0.018822 | 0.020804
Mo, ppm | 3.5046 | 0.16728 | 3.17 | 3.8391 | 3.0027 | 4.0064 | 4.773% | 9.546% | 14.32% | 3.3293 | 3.6798
Na, wt.% | 0.23332 | 0.014078 | 0.20516 | 0.26148 | 0.19109 | 0.27555 | 6.034% | 12.07% | 18.1% | 0.22165 | 0.24499
Nd, ppm | 25.3886 | 0.80828 | 23.7721 | 27.0052 | 22.9638 | 27.8135 | 3.184% | 6.367% | 9.551% | 24.1192 | 26.6581
Ni, ppm | 42.4328 | 1.3742 | 39.6844 | 45.1812 | 38.3101 | 46.5554 | 3.239% | 6.477% | 9.716% | 40.3112 | 44.5544
P, wt.% | 0.047391 | 0.0019634 | 0.043464 | 0.051318 | 0.041501 | 0.053281 | 4.143% | 8.286% | 12.43% | 0.045021 | 0.04976
Pb, ppm | 8.4444 | 0.86212 | 6.7201 | 10.1686 | 5.858 | 11.0307 | 10.21% | 20.42% | 30.63% | 8.0221 | 8.8666
Pr, ppm | 6.7131 | 0.13086 | 6.4514 | 6.9749 | 6.3206 | 7.1057 | 1.949% | 3.899% | 5.848% | 6.3775 | 7.0488
Rb, ppm | 15.096 | 0.81054 | 13.4749 | 16.7171 | 12.6643 | 17.5276 | 5.369% | 10.74% | 16.11% | 14.3412 | 15.8508
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sb, ppm | 1.2433 | 0.12828 | 0.98669 | 1.4998 | 0.85841 | 1.6281 | 10.32% | 20.64% | 30.95% | 1.1811 | 1.3054
Sc, ppm | 2.8183 | 0.20235 | 2.4136 | 3.223 | 2.2113 | 3.4254 | 7.18% | 14.36% | 21.54% | 2.6774 | 2.9592
Sm, ppm | 4.4207 | 0.37058 | 3.6795 | 5.1618 | 3.3089 | 5.5324 | 8.383% | 16.77% | 25.15% | 4.1996 | 4.6417
Sn, ppm | 1.4004 | 0.083423 | 1.2336 | 1.5672 | 1.1501 | 1.6507 | 5.957% | 11.91% | 17.87% | 1.3304 | 1.4704
Sr, ppm | 39.8064 | 2.4048 | 34.9968 | 44.6159 | 32.5921 | 47.0206 | 6.041% | 12.08% | 18.12% | 37.816 | 41.7967
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.44915 | 0.020476 | 0.4082 | 0.4901 | 0.38772 | 0.51058 | 4.559% | 9.118% | 13.68% | 0.42669 | 0.47161
Te, ppm | 0.034015 | 0.0072434 | 0.019528 | 0.048502 | 0.012285 | 0.055745 | 21.29% | 42.59% | 63.88% | 0.032314 | 0.035716
Th, ppm | 8.7598 | 0.7466 | 7.2666 | 10.253 | 6.52 | 10.9996 | 8.523% | 17.05% | 25.57% | 8.3218 | 9.1978
Ti, wt.% | 0.072088 | 0.0047173 | 0.062653 | 0.081523 | 0.057936 | 0.08624 | 6.544% | 13.09% | 19.63% | 0.068484 | 0.075693
Tl, ppm | 0.080391 | 0.0081424 | 0.064106 | 0.096676 | 0.055963 | 0.10482 | 10.13% | 20.26% | 30.39% | 0.076371 | 0.08441
Tm, ppm | 0.10151 | 0.0063129 | 0.08888 | 0.11413 | 0.082567 | 0.12044 | 6.219% | 12.44% | 18.66% | 0.096431 | 0.10658
U, ppm | 1.2328 | 0.063801 | 1.1052 | 1.3604 | 1.0414 | 1.4242 | 5.175% | 10.35% | 15.53% | 1.1712 | 1.2945
V, ppm | 27.2565 | 1.6084 | 24.0398 | 30.4732 | 22.4314 | 32.0815 | 5.901% | 11.8% | 17.7% | 25.8936 | 28.6193
W, ppm | 2.0302 | 0.36122 | 1.3078 | 2.7527 | 0.94658 | 3.1139 | 17.79% | 35.58% | 53.38% | 1.9287 | 2.1317
Y, ppm | 8.7041 | 0.29946 | 8.1052 | 9.303 | 7.8057 | 9.6025 | 3.441% | 6.881% | 10.32% | 8.2689 | 9.1393
Yb, ppm | 0.58151 | 0.038969 | 0.50357 | 0.65945 | 0.46461 | 0.69842 | 6.701% | 13.4% | 20.1% | 0.55244 | 0.61059
Zn, ppm | 36.3123 | 1.846 | 32.6203 | 40.0044 | 30.7742 | 41.8504 | 5.084% | 10.17% | 15.25% | 34.4967 | 38.1279
Zr, ppm | 31.0582 | 4.5884 | 21.8813 | 40.2351 | 17.2928 | 44.8235 | 14.77% | 29.55% | 44.32% | 29.5053 | 32.6111
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 265
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PAAU02 | PhotonAssay (Au) two-cycle analysis
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 265
Lab Slot | Laboratory Name and Location
Accura Geomysore Labs Private Limited, Jonnagiri, Kurnool, India
Accura Gold Minerals Testing, Namakkal, Tamilnadu, India
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Thunder Bay, Ontario, Canada
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
BV Quito, Quito, Pichincha, Ecuador
Gekko Assay Labs, Ballarat, VIC, Australia
Gloriosa Laboratory, Otjiwarongo, Otjozondjupa, Namibia
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek Burkina Faso SAS, Ouagadougou, Zoundwéogo, Burkina Faso
Intertek Genalysis, Adelaide, SA, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Carlin, Carlin, Nevada, USA
MSALABS Elko, Elko, Nevada, USA
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Paragon Geochemical Laboratories, Surrey, BC, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Kalgoorlie, WA, Australia
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Mwanza, Mwanza, Mwanza, United Republic of Tanzania
SGS New Zealand Limited, Palmerston, Otago, New Zealand
SGS Orange, Orange, NSW, Australia
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 265 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 8 | 10 | 309654 | 0.087168 | 1.1863 | 1.1266
2 | 1 | ANSLu | 1 | 20 | 4 | 309655 | 0.084938 | 1.1485 | 1.1246
2 | 1 | ANSLu | 1 | 12 | 12 | 309656 | 0.084025 | 1.2423 | 1.1296
2 | 1 | ANSLu | 1 | 19 | 4 | 309657 | 0.084961 | 1.2386 | 1.1294
2 | 1 | ANSLu | 1 | 7 | 10 | 309658 | 0.086392 | 1.1428 | 1.1243
2 | 1 | ANSLu | 1 | 1 | 1 | 309659 | 0.08795 | 1.1136 | 1.1227
2 | 1 | ANSLu | 1 | 16 | 2 | 309660 | 0.084829 | 1.0893 | 1.1214
2 | 1 | ANSLu | 1 | 11 | 12 | 309661 | 0.086929 | 1.1092 | 1.1225
2 | 1 | ANSLu | 1 | 2 | 1 | 309662 | 0.084551 | 1.0614 | 1.1199
2 | 1 | ANSLu | 1 | 9 | 5 | 309663 | 0.087969 | 1.0864 | 1.1213
2 | 1 | ANSLu | 1 | 10 | 5 | 309664 | 0.087319 | 1.0835 | 1.1211
2 | 1 | ANSLu | 1 | 15 | 2 | 309665 | 0.085317 | 1.0865 | 1.1213
2 | 1 | ANSLu | 1 | 13 | 7 | 309666 | 0.085638 | 1.0679 | 1.1203
2 | 1 | ANSLu | 1 | 4 | 8 | 309667 | 0.088271 | 1.1212 | 1.1231
2 | 1 | ANSLu | 1 | 3 | 8 | 309668 | 0.088084 | 1.1398 | 1.1241
2 | 1 | ANSLu | 1 | 6 | 3 | 309669 | 0.083566 | 1.1368 | 1.124
2 | 1 | ANSLu | 1 | 5 | 3 | 309670 | 0.086389 | 1.104 | 1.1222
2 | 1 | ANSLu | 1 | 18 | 9 | 309671 | 0.084351 | 1.1073 | 1.1224
2 | 1 | ANSLu | 1 | 14 | 7 | 309672 | 0.085105 | 1.077 | 1.1208
2 | 1 | ANSLu | 1 | 17 | 9 | 309673 | 0.087862 | 1.1221 | 1.1232
Mean | [6 empty] | 0.086081 | 1.1232 | 1.1232
Median | [7 empty] | 1.1114 | 1.1226
SD | [7 empty] | 0.050433 | 0.0027015
RSD | [7 empty] | 4.49% | 0.2405%
RSD Upscale Factor | [7 empty] | 0.053566
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 265 (Certified Value 1.08 ppm) | [31 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 30g | 30g | 30g | 50g | 25g | 30g | 40g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 1.1136 | 1.052 | 1.08 | 1.12 | 1.0432 | 1.071 | 1.14 | 1.064 | 1.08 | 1.018 | 1.09 | 1.08 | 1.096 | 1.071 | 1.1 | 1.116 | 0.996 | 1.25 | 1.07 | 1.08 | 1.04 | 1.109 | 1.0681 | 1.024 | 1.0686 | 1.08 | 1.1088 | 1.112 | 1.05 | 1.024 | 1
| 1 | 2 | 1.0614 | 1.069 | 1.08 | 1.09 | 1.0507 | 1.084 | 1.13 | 1.065 | 1.09 | 1.05 | 1.08 | 1.07 | 1.077 | 1.052 | 1.12 | 1.107 | 1.09 | 1.2681 | 1.11 | 1.13 | 1.02 | 1.15 | 1.0686 | 1.036 | 1.0833 | 1.09 | 1.0762 | 1.0793 | 1.01 | 0.984 | 40
| 1 | 3 | 1.1398 | 1.057 | 1.09 | 1.08 | 1.041 | 1.081 | 1.13 | 1.056 | 1.08 | 1.06 | 1.08 | 1.06 | 1.087 | 1.059 | 1.13 | 1.083 | 1.11 | 1.2614 | 1.1 | 1.1 | 1.06 | 1.144 | 1.054 | 1.044 | 1.0875 | 1.07 | 1.0825 | 1.0856 | 1.02 | 1.045 | 16
| 1 | 4 | 1.1212 | 1.072 | 1.08 | 1.11 | 1.0346 | 1.084 | 1.07 | 1.063 | 1.09 | 1.048 | 1.08 | 1.06 | 1.089 | 1.068 | 1.11 | 1.083 | 1.07 | 1.2376 | 1.1 | 1.12 | 1.01 | 1.156 | 1.0405 | 1.027 | 1.0833 | 1.09 | 1.1119 | 1.1151 | 1.01 | 1.017 | 1.0761
| 1 | 5 | 1.104 | 1.074 | 1.11 | 1.09 | 1.0506 | 1.091 | 1.11 | 1.069 | 1.08 | 1.068 | 1.05 | 1.06 | 1.075 | 1.066 | 1.12 | 1.111 | 1.03 | 1.2681 | 1.07 | 1.11 | 1.06 | 1.139 | 1.053 | 1.032 | 1.0898 | 1.08 | 1.134 | 1.1372 | 1.05 | 1.042 | 7
| 1 | 6 | 1.1368 | 1.078 | 1.08 | 1.08 | 1.031 | 1.079 | 1.13 | 1.057 | 1.08 | 1.059 | 1.06 | 1.07 | 1.075 | 1.07 | 1.1 | 1.084 | 1.07 | 1.2681 | 1.11 | 1.1 | 0.984 | 1.107 | 1.0445 | 1.052 | 1.0716 | 1.06 | 1.1382 | 1.1415 | 1.01 | 1.043
|  | 7 | 1.1428
|  | 8 | 1.1863
|  | 9 | 1.0864
|  | 10 | 1.0835
|  | 11 | 1.1092
|  | 12 | 1.2423
|  | 13 | 1.0679
|  | 14 | 1.077
|  | 15 | 1.0865
|  | 16 | 1.0893
|  | 17 | 1.1221
|  | 18 | 1.1073
|  | 19 | 1.2386
|  | 20 | 1.1485
| Mean |  | 1.1232 | 1.067 | 1.0867 | 1.095 | 1.0419 | 1.0817 | 1.1183 | 1.0623 | 1.0833 | 1.0505 | 1.0733 | 1.0667 | 1.0832 | 1.0643 | 1.1133 | 1.0973 | 1.061 | 1.2589 | 1.0933 | 1.1067 | 1.029 | 1.1342 | 1.0548 | 1.0358 | 1.0807 | 1.0783 | 1.1086 | 1.1118 | 1.025 | 1.0258
| Median |  | 1.1114 | 1.0705 | 1.08 | 1.09 | 1.0421 | 1.0825 | 1.13 | 1.0635 | 1.08 | 1.0545 | 1.08 | 1.065 | 1.082 | 1.067 | 1.115 | 1.0955 | 1.07 | 1.2647 | 1.1 | 1.105 | 1.03 | 1.1415 | 1.0535 | 1.034 | 1.0833 | 1.08 | 1.1104 | 1.1135 | 1.015 | 1.033
| Std Dev. |  | 0.050433 | 0.010237 | 0.012111 | 0.016432 | 0.0081068 | 0.0066232 | 0.025626 | 0.0049666 | 0.005164 | 0.017501 | 0.015055 | 0.008165 | 0.0087731 | 0.0073937 | 0.012111 | 0.015603 | 0.041449 | 0.012604 | 0.018619 | 0.017512 | 0.030033 | 0.021066 | 0.011656 | 0.010591 | 0.0086119 | 0.01169 | 0.025525 | 0.025598 | 0.019748 | 0.023473
| Rel.Std.Dev. |  | 4.49% | 0.9594% | 1.114% | 1.501% | 0.7781% | 0.6123% | 2.291% | 0.4675% | 0.4767% | 1.666% | 1.403% | 0.7655% | 0.8099% | 0.6947% | 1.088% | 1.422% | 3.907% | 1.001% | 1.703% | 1.582% | 2.919% | 1.857% | 1.105% | 1.022% | 0.7969% | 1.084% | 2.302% | 2.302% | 1.927% | 2.288%
| PDM3 |  | 4.383% | -0.8427% | 0.985% | 1.759% | -3.179% | 0.5203% | 3.928% | -1.276% | 0.6752% | -2.376% | -0.2541% | -0.8737% | 0.6597% | -1.09% | 3.463% | 1.976% | -1.4% | 16.99% | 1.605% | 2.844% | -4.374% | 5.399% | -1.978% | -3.739% | 0.4272% | 0.2105% | 3.026% | 3.32% | -4.746% | -4.668%
| Z-Score (Absolute) |  | NA | 0.47 | 0.21 | 0.49 | 1.33 | 0.03 | 1.29 | 0.63 | 0.09 | 1.03 | 0.25 | 0.48 | 0.09 | 0.56 | 1.12 | 0.57 | 0.67 | 6.11 | 0.43 | 0.89 | 1.77 | 1.83 | 0.89 | 1.54 | 0 | 0.08 | 0.96 | 1.07 | 1.91 | 1.88
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 265 (Certified Value 1.09 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | 6.13 | 6.15 | 6.16 | 6.17 | 6.18 | 6.19 | 6.20 | 6.21 | 6.22 | 6.23 | 6.24 | ppm
|  |  | INAA | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | 2
|  |  | 0.085g | [23 empty] | 3
| 1 | 1 | 1.1136 | 1.12 | 1.06 | 1.105 | 1.08 | 1.07 | 1.07 | 1.186 | 1.112 | 1.134 | 1.131 | 1.117 | 1.053 | 1.09 | 1.1 | 1.094 | 1.191 | 1.107 | 1.092 | 1.073 | 1 | 1.1 | 1.06 | 1.035 | 1
| 1 | 2 | 1.0614 | 1.04 | 1 | 1.106 | 1.1 | 1.05 | 1.12 | 1.167 | 1.121 | 1.123 | 1.103 | 1.123 | 1.132 | 1.15 | 1.15 | 1.021 | 1.082 | 1.241 | 1.043 | 1.115 | 1.09 | 1.06 | 1.04 | 1.115 | 40
| 1 | 3 | 1.1398 | 1.12 | 0.98 | 1.137 | 1.08 | 1.12 | 1.04 | 1.103 | 1.065 | 1.162 | 1.115 | 1.089 | 1.128 | 1.14 | 1.08 | 1.006 | 1.132 | 1.117 | 1.084 | 1.132 | 1.13 | 1.04 | 1.12 | 1.03 | 16
| 1 | 4 | 1.1212 | 1.06 | 1.08 | 1.17 | 1.04 | 1.06 | 1.12 | 1.134 | 1.091 | 1.101 | 1.107 | 1.087 | 1.149 | 1.2 | 1.07 | 1.039 | 1.157 | 1.057 | 1.141 | 1.066 | 1.06 | 1 | 1.03 | 1.022 | 1.092
| 1 | 5 | 1.104 | 1.1 | 1.09 | 1.132 | 1.07 | 1.06 | 1.06 | 1.131 | 1.037 | 1.077 | 1.153 | 1.131 | 1.048 | 1.09 | 1.11 | 1.035 | 1.149 | 1.193 | 1.128 | 1 | 1.09 | 1.15 | 1.15 | 1.048 | 9
| 1 | 6 | 1.1368 | 1.11 | 1.02 | 1.082 | 1.03 | 1.09 | 1.06 | 1.108 | 1.01 | 1.158 | 1.115 | 1.07 | 1.147 | 1.08 | 1.07 | 1.076 | 1.039 | 1.168 | 1.066 | 1.079 | 1.01 | 1.07 | 1.14 | 1.117
|  | 7 | 1.1428
|  | 8 | 1.1863
|  | 9 | 1.0864
|  | 10 | 1.0835
|  | 11 | 1.1092
|  | 12 | 1.2423
|  | 13 | 1.0679
|  | 14 | 1.077
|  | 15 | 1.0865
|  | 16 | 1.0893
|  | 17 | 1.1221
|  | 18 | 1.1073
|  | 19 | 1.2386
|  | 20 | 1.1485
| Mean |  | 1.1232 | 1.0917 | 1.0383 | 1.122 | 1.0667 | 1.075 | 1.0783 | 1.1382 | 1.0727 | 1.1258 | 1.1207 | 1.1028 | 1.1095 | 1.125 | 1.0967 | 1.0452 | 1.125 | 1.1472 | 1.0923 | 1.0775 | 1.0633 | 1.07 | 1.09 | 1.0612
| Median |  | 1.1114 | 1.105 | 1.04 | 1.119 | 1.075 | 1.065 | 1.065 | 1.1325 | 1.078 | 1.1285 | 1.115 | 1.103 | 1.13 | 1.115 | 1.09 | 1.037 | 1.1405 | 1.1425 | 1.088 | 1.076 | 1.075 | 1.065 | 1.09 | 1.0415
| Std Dev. |  | 0.050433 | 0.033714 | 0.044907 | 0.0309 | 0.026583 | 0.025884 | 0.033714 | 0.032664 | 0.043537 | 0.032921 | 0.018522 | 0.02417 | 0.046453 | 0.046797 | 0.030768 | 0.033451 | 0.055227 | 0.066322 | 0.036979 | 0.045908 | 0.050465 | 0.051381 | 0.052915 | 0.04331
| Rel.Std.Dev. |  | 4.49% | 3.088% | 4.325% | 2.754% | 2.492% | 2.408% | 3.127% | 2.87% | 4.059% | 2.924% | 1.653% | 2.192% | 4.187% | 4.16% | 2.806% | 3.201% | 4.909% | 5.781% | 3.385% | 4.261% | 4.746% | 4.802% | 4.855% | 4.081%
| PDM3 |  | 2.859% | -0.03145% | -4.915% | 2.746% | -2.321% | -1.558% | -1.252% | 4.227% | -1.771% | 3.097% | 2.624% | 0.9911% | 1.602% | 3.021% | 0.4264% | -4.29% | 3.021% | 5.051% | 0.0296% | -1.329% | -2.626% | -2.016% | -0.1841% | -2.824%
| Z-Score (Absolute) |  | NA | 0 | 1.44 | 0.82 | 0.67 | 0.45 | 0.36 | 1.25 | 0.51 | 0.92 | 0.78 | 0.3 | 0.48 | 0.9 | 0.13 | 1.25 | 0.9 | 1.5 | 0.02 | 0.38 | 0.76 | 0.58 | 0.04 | 0.82
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 265 (Certified Value 1.08 ppm) | [20 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.29 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*OES/MS | 2
|  |  | 0.085g | 25g | 30g | 25g | 25g | 30g | 25g | 25g | 25g | 15g | 15g | 25g | 30g | 30g | 10g | 25g | 20g | 30g | 20g | 3
| 1 | 1 | 1.1136 | 1.03 | 1.08 | 1.0217 | 1.07 | 1.041 | 1 | 1.14 | 1.1 | 1.0723 | 1.0716 | 1.082 | 1.15 | 1.2192 | 0.2 | 1.08 | 1.016 | 1.1764 | 1.0332 | 1
| 1 | 2 | 1.0614 | 1.01 | 1.08 | 1.0506 | 1.09 | 1.034 | 1.06 | 1.16 | 1.06 | 1.0913 | 1.0154 | 1.098 | 1.13 | 1.1975 | < 0.02 | 1.08 | 1.018 | 1.0994 | 1.0248 | 40
| 1 | 3 | 1.1398 | 1.01 | 1.08 | 1.0392 | 1.08 | 1.028 | 1.01 | 1.16 | 1.08 | 1.1007 | 1.0992 | 1.083 | 1.19 | 1.1802 | 0.4 | 1.06 | 1.019 | 1.1259 | 1.0707 | 16
| 1 | 4 | 1.1212 | 0.99 | 1.08 | 1.066 | 1.1 | 1.037 | 1.06 | 1.12 | 1.08 | 1.0643 | 1.0164 | 1.081 | 1.15 | 1.2181 | 0.4 | 0.96 | 0.996 | 1.1236 | 1.0329 | 1.0776
| 1 | 5 | 1.104 | 1.01 | 1.07 | 1.0255 | 1.1 | 1.05 | 1.08 | 1.13 | 1.08 | 1.1068 | 1.0299 | 1.082 | 1.15 | 1.2242 | 0.2 | 1.09 | 1.001 | 1.1615 | 1.046 | 11
| 1 | 6 | 1.1368 | 1.03 | 1.08 | 1.0663 | 1.08 | 1.05 | 1.08 | 1.14 | 1.06 | 1.1036 | 1.0526 | 1.052 | 1.17 | 1.2146 | 0.2 | 0.95 | 1.011 | 1.0361 | 1.0307
|  | 7 | 1.1428
|  | 8 | 1.1863
|  | 9 | 1.0864
|  | 10 | 1.0835
|  | 11 | 1.1092
|  | 12 | 1.2423
|  | 13 | 1.0679
|  | 14 | 1.077
|  | 15 | 1.0865
|  | 16 | 1.0893
|  | 17 | 1.1221
|  | 18 | 1.1073
|  | 19 | 1.2386
|  | 20 | 1.1485
| Mean |  | 1.1232 | 1.0133 | 1.0783 | 1.0449 | 1.0867 | 1.04 | 1.0483 | 1.1417 | 1.0767 | 1.0898 | 1.0475 | 1.0797 | 1.1567 | 1.209 | 0.28 | 1.0367 | 1.0102 | 1.1205 | 1.0397
| Median |  | 1.1114 | 1.01 | 1.08 | 1.0449 | 1.085 | 1.039 | 1.06 | 1.14 | 1.08 | 1.096 | 1.0413 | 1.082 | 1.15 | 1.2164 | 0.2 | 1.07 | 1.0135 | 1.1247 | 1.033
| Std Dev. |  | 0.050433 | 0.015055 | 0.0040825 | 0.01939 | 0.012111 | 0.0088318 | 0.034881 | 0.016021 | 0.015055 | 0.017646 | 0.033428 | 0.015002 | 0.020656 | 0.01679 | 0.10954 | 0.064083 | 0.0095795 | 0.049819 | 0.016686
| Rel.Std.Dev. |  | 4.49% | 1.486% | 0.3786% | 1.856% | 1.114% | 0.8492% | 3.327% | 1.403% | 1.398% | 1.619% | 3.191% | 1.39% | 1.786% | 1.389% | 39.12% | 6.182% | 0.9483% | 4.446% | 1.605%
| PDM3 |  | 4.232% | -5.966% | 0.06596% | -3.038% | 0.8393% | -3.491% | -2.718% | 5.943% | -0.0887% | 1.133% | -2.794% | 0.1897% | 7.335% | 12.19% | -74.02% | -3.801% | -6.26% | 3.977% | -3.518%
| Z-Score (Absolute) |  | NA | 1.16 | 0.03 | 0.58 | 0.18 | 0.67 | 0.52 | 1.2 | 0 | 0.24 | 0.54 | 0.06 | 1.47 | 2.43 | NA | 0.74 | 1.22 | 0.81 | 0.68
Results from laboratory 1.18 were removed due to their 0.01 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 265 (Certified Value 1.04 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.22 | 1.28 | 1.30 | ppm
|  |  | INAA | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*OES | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 200g | 30g | 05g | 200g | 40g | 30g | 30g | 30g | 05g | 50g | 200g | 30g | 15g | 200g | 100g | 200g | 3
| 1 | 1 | 1.1136 | 1.085 | 0.96 | 1 | 1.07 | 1.06 | 1 | 1 | 1.04 | 1.039 | 0.999 | 0.86746 | 1.02 | 1.08 | 1.1058 | 1.046 | 1 | 1
| 1 | 2 | 1.0614 | 1.065 | 1.08 | 1 | 1.07 | 1.05 | 1.05 | 0.99 | 1.02 | 1.078 | 1.05 | 0.85932 | 1.02 | 0.93 | 1.0979 | 1.034 | 1 | 40
| 1 | 3 | 1.1398 | 1.162 | 1.1 | 1 | 1.11 | 1.05 | 1.02 | 1.01 | 1.03 | 1.075 | 1.05 | 0.85435 | 1.01 | 0.88 | 1.1362 | 1.024 | 1 | 16
| 1 | 4 | 1.1212 | 1.108 | 1.02 | 1 | 1.07 | 1.06 | 1.03 | 0.98 | 1.05 | 1.079 | 1 | 0.86666 | 1.01 | 0.99 | 1.1228 | 1.027 | 1 | 1.0388
| 1 | 5 | 1.104 | 1.084 | 1.14 | 1 | 1.08 | 1.06 | 1.03 | 1 | 1.04 | 1.044 | 1.05 | 0.87098 | 1.01 | 0.91 | 1.0897 | 1.049 | 1 | 13
| 1 | 6 | 1.1368 | 1.094 | 1.02 | 1 | 1.08 | 1.06 | 1.04 | 0.92 | 1.05 | 1.042 | 1.04 | 0.86422 | 1 | 1.05 | 1.1256 | 1.031 | 1
|  | 7 | 1.1428
|  | 8 | 1.1863
|  | 9 | 1.0864
|  | 10 | 1.0835
|  | 11 | 1.1092
|  | 12 | 1.2423
|  | 13 | 1.0679
|  | 14 | 1.077
|  | 15 | 1.0865
|  | 16 | 1.0893
|  | 17 | 1.1221
|  | 18 | 1.1073
|  | 19 | 1.2386
|  | 20 | 1.1485
| Mean |  | 1.1232 | 1.0997 | 1.0533 | 1 | 1.08 | 1.0567 | 1.0283 | 0.98333 | 1.0383 | 1.0595 | 1.0315 | 0.86383 | 1.0117 | 0.97333 | 1.113 | 1.0352 | 1
| Median |  | 1.1114 | 1.0895 | 1.05 | 1 | 1.075 | 1.06 | 1.03 | 0.995 | 1.04 | 1.0595 | 1.045 | 0.86544 | 1.01 | 0.96 | 1.1143 | 1.0325 | 1
| Std Dev. |  | 0.050433 | 0.033613 | 0.06532 | 0 | 0.015492 | 0.005164 | 0.017224 | 0.03266 | 0.01169 | 0.019644 | 0.02509 | 0.0060445 | 0.0075277 | 0.080166 | 0.017983 | 0.010187 | 0
| Rel.Std.Dev. |  | 4.49% | 3.057% | 6.201% | 0% | 1.434% | 0.4887% | 1.675% | 3.321% | 1.126% | 1.854% | 2.432% | 0.6997% | 0.7441% | 8.236% | 1.616% | 0.9841% | 0%
| PDM3 |  | 8.122% | 5.854% | 1.394% | -3.74% | 3.961% | 1.715% | -1.012% | -5.344% | -0.04954% | 1.988% | -0.7073% | -16.85% | -2.616% | -6.306% | 7.14% | -0.3544% | -3.74%
| Z-Score (Absolute) |  | NA | 1.5 | 0.45 | 0.76 | 1.06 | 0.53 | 0.11 | 1.13 | 0.11 | 0.59 | 0.04 | 3.84 | 0.49 | 1.36 | 1.81 | 0.04 | 0.76
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 265 (Indicative Value 0.04 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.04 | 1
| Mean |  | 0.04 | 16
| Median |  | 0.04 | 0.04
| Std Dev. |  | 0 | 7
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 265 (Indicative Value 0.02 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.02 | 1
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.02
| Std Dev. |  | 0 | 7
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 265 (Certified Value 0.68 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.67 | 0.74 | 0.7 | 0.66987 | 0.71 | 0.64 | 0.72 | 0.71 | 0.6 | 0.69 | 0.65 | 0.663 | 0.915 | 0.58 | 0.65 | 0.68 | 0.68 | 0.72 | 0.81 | 0.8 | 0.69965 | 1.248 | 0.67 |  | 1
| 1 | 2 | 0.61 | 0.74 | 0.6 | 0.63147 | 0.69 | 0.64 | 0.75 | 0.68 | 0.64 | 0.69 | 0.67 | 0.693 | 0.891 | 0.58 | 0.65 | 0.73 | 0.63 | 0.72 | 0.86 | 0.9 | 0.68403 | 1.256 | 0.72 |  | 20
| 1 | 3 | 0.71 | 0.72 | 0.7 | 0.62038 | 0.66 | 0.68 | 0.76 | 0.69 | 0.66 | 0.68 | 0.66 | 0.651 | 0.925 | 0.59 | 0.66 | 0.71 | 0.65 | 0.69 | 0.82 | 0.7 | 0.67426 | 1.249 | 0.69 |  | 16
| 1 | 4 | 0.72 | 0.75 | 0.6 | 0.68356 | 0.74 | 0.66 | 0.79 | 0.64 | 0.63 | 0.73 | 0.66 | 0.656 | 0.86 | 0.53 | 0.66 | 0.72 | 0.66 | 0.71 | 0.87 | 0.7 | 0.69898 | 1.2563 | 0.73 |  | 0.68001
| 1 | 5 | 0.7 | 0.7 | 0.6 | 0.66439 | 0.68 | 0.64 | 0.72 | 0.64 | 0.62 | 0.69 | 0.66 | 0.659 | 0.892 | 0.59 | 0.65 | 0.7 | 0.77 | 0.63 | 0.89 | 0.7 | 0.70595 | 1.2498 | 0.72 |  | 15
| 1 | 6 | 0.72 | 0.71 | 0.7 | 0.66339 | 0.7 | 0.65 | 0.72 | 0.67 | 0.63 | 0.72 | 0.64 | 0.681 | 0.797 | 0.59 | 0.63 | 0.72 | 0.79 | 0.68 | 0.87 | 0.7 | 0.72264 | 1.263 | 0.7
| Mean |  | 0.68833 | 0.72667 | 0.65 | 0.65551 | 0.69667 | 0.65167 | 0.74333 | 0.67167 | 0.63 | 0.7 | 0.65667 | 0.66717 | 0.88 | 0.57667 | 0.65 | 0.71 | 0.69667 | 0.69167 | 0.85333 | 0.75 | 0.69759 | 1.2537 | 0.705
| Median |  | 0.705 | 0.73 | 0.65 | 0.66389 | 0.695 | 0.645 | 0.735 | 0.675 | 0.63 | 0.69 | 0.66 | 0.661 | 0.8915 | 0.585 | 0.65 | 0.715 | 0.67 | 0.7 | 0.865 | 0.7 | 0.69932 | 1.2529 | 0.71
| Std Dev. |  | 0.042622 | 0.019664 | 0.054772 | 0.024273 | 0.027325 | 0.016021 | 0.028752 | 0.027869 | 0.02 | 0.02 | 0.010328 | 0.016302 | 0.046484 | 0.023381 | 0.010954 | 0.017889 | 0.066833 | 0.034303 | 0.031411 | 0.083666 | 0.016921 | 0.0058084 | 0.022583
| Rel.Std.Dev. |  | 6.192% | 2.706% | 8.427% | 3.703% | 3.922% | 2.458% | 3.868% | 4.149% | 3.175% | 2.857% | 1.573% | 2.444% | 5.282% | 4.054% | 1.685% | 2.52% | 9.593% | 4.959% | 3.681% | 11.16% | 2.426% | 0.4633% | 3.203%
| PDM3 |  | 1.223% | 6.86% | -4.414% | -3.604% | 2.449% | -4.169% | 9.311% | -1.228% | -7.355% | 2.939% | -3.433% | -1.889% | 29.41% | -15.2% | -4.414% | 4.41% | 2.449% | 1.714% | 25.49% | 10.29% | 2.584% | 84.36% | 3.674%
| Z-Score (Absolute) |  | 0.19 | 0.67 | NA | 0.93 | 0 | 1.01 | 1.05 | 0.56 | 1.5 | 0.07 | 0.9 | 0.66 | 4.12 | 2.7 | 1.05 | 0.3 | 0 | 0.11 | 3.52 | NA | 0.02 | 12.52 | 0.19
Results from laboratories 1.03 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 265 (Certified Value 6.62 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 5.7 | 6.9 | 6.52 | 6.8527 | 6.6263 | 6.65 | 7.11 | 6.77 | 6.5074 | 6.72 | 6.62 | 6.62 | 6.38 | 6.77 | 6.5 | 6.91 | 6.4958 | 6.8972 | 6.54 | 6.721 | 7.24 | 6.2854 | 6.3988 | 6.6 | 1
| 1 | 2 | 5.18 | 7.06 | 6.42 | 6.7233 | 6.6239 | 6.5 | 6.92 | 6.73 | 6.6835 | 6.51 | 6.52 | 6.8 | 6.29 | 6.43 | 6.73 | 6.83 | 6.4957 | 6.9634 | 6.58 | 6.986 | 6.91 | 6.2488 | 6.2573 | 6.49 | #N/A
| 1 | 3 | 5 | 6.96 | 6.54 | 6.5234 | 6.6303 | 6.46 | 6.54 | 6.8 | 6.5635 | 6.56 | 6.64 | 6.75 | 6.17 | 6.71 | 6.53 | 6.82 | 6.4721 | 7.0005 | 6.62 | 6.827 | 6.97 | 6.3776 | 6.2278 | 6.48 | 16
| 1 | 4 | 4.71 | 6.84 | 6.45 | 6.5901 | 6.5259 | 6.39 | 7.12 | 6.53 | 6.5018 | 6.74 | 6.78 | 6.66 | 6.67 | 6.69 | 6.75 | 7.01 | 6.4887 | 6.7926 | 6.62 | 6.774 | 7.13 | 6.3667 | 6.1848 | 6.79 | 6.6232
| 1 | 5 | 4.47 | 7.03 | 6.61 | 6.6921 | 6.6612 | 6.6 | 6.66 | 6.61 | 6.6092 | 6.58 | 6.66 | 6.65 | 6.34 | 6.63 | 6.56 | 6.81 | 6.4897 | 6.9545 | 6.55 | 6.827 | 6.69 | 6.5324 | 6.3745 | 6.65 | 16
| 1 | 6 | 4.4 | 6.97 | 6.63 | 6.746 | 6.6349 | 6.57 | 6.73 | 6.66 | 6.5183 | 6.71 | 6.51 | 6.77 | 6.48 | 6.39 | 6.59 | 7.05 | 6.4729 | 7.0061 | 6.56 | 7.092 | 6.92 | 6.4946 | 6.2673 | 6.7
| Mean |  | 4.91 | 6.96 | 6.5283 | 6.6879 | 6.6171 | 6.5283 | 6.8467 | 6.6833 | 6.564 | 6.6367 | 6.6217 | 6.7083 | 6.3883 | 6.6033 | 6.61 | 6.905 | 6.4858 | 6.9357 | 6.5783 | 6.8712 | 6.9767 | 6.3842 | 6.2851 | 6.6183
| Median |  | 4.855 | 6.965 | 6.53 | 6.7077 | 6.6283 | 6.535 | 6.825 | 6.695 | 6.5409 | 6.645 | 6.63 | 6.705 | 6.36 | 6.66 | 6.575 | 6.87 | 6.4892 | 6.9589 | 6.57 | 6.827 | 6.945 | 6.3721 | 6.2623 | 6.625
| Std Dev. |  | 0.48965 | 0.08124 | 0.083766 | 0.11699 | 0.046657 | 0.096212 | 0.24163 | 0.1027 | 0.071392 | 0.098116 | 0.099683 | 0.07414 | 0.17175 | 0.1568 | 0.10526 | 0.10387 | 0.010718 | 0.080294 | 0.034881 | 0.13989 | 0.19117 | 0.11187 | 0.084087 | 0.1209
| Rel.Std.Dev. |  | 9.973% | 1.167% | 1.283% | 1.749% | 0.7051% | 1.474% | 3.529% | 1.537% | 1.088% | 1.478% | 1.505% | 1.105% | 2.688% | 2.375% | 1.592% | 1.504% | 0.1653% | 1.158% | 0.5302% | 2.036% | 2.74% | 1.752% | 1.338% | 1.827%
| PDM3 |  | -25.87% | 5.085% | -1.433% | 0.977% | -0.09278% | -1.433% | 3.374% | 0.9075% | -0.895% | 0.2029% | -0.02357% | 1.285% | -3.547% | -0.3004% | -0.1997% | 4.254% | -2.075% | 4.718% | -0.6778% | 3.743% | 5.336% | -3.608% | -5.105% | -0.0739%
| Z-Score (Absolute) |  | 12.79 | 2.56 | 0.67 | 0.53 | 0 | 0.67 | 1.72 | 0.49 | 0.4 | 0.14 | 0.03 | 0.68 | 1.72 | 0.11 | 0.06 | 2.15 | 0.99 | 2.38 | 0.3 | 1.9 | 2.69 | 1.75 | 2.49 | 0
| Analytical results for As in OREAS 265 (Certified Value 221 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 286 | 229 | 222 | 216.2583 | 224.4 | 218 | 235 | 220 | 218.5 | 235 | 228 | 217.2 | 216.3 | 226.7 | 216 | 218 | 218.4 | 213 | 194.6 | 214.266 | 129 | 211.1 |  |  | 1
| 1 | 2 | 268 | 235 | 223 | 211.6638 | 221.4 | 223 | 229 | 221 | 221.6 | 226 | 225 | 223.1 | 211.6 | 213.3 | 217 | 213 | 221.5 | 219 | 190.4 | 213.9949 | 123 | 206.7 |  |  | 10
| 1 | 3 | 252 | 239 | 221 | 208.8202 | 223.1 | 215 | 231 | 223 | 221.8 | 231 | 230 | 220.4 | 209.2 | 228.7 | 216 | 215 | 222.4 | 218 | 191.8 | 217.6302 | 121 | 206.1 |  |  | 16
| 1 | 4 | 271 | 228 | 225 | 208.8476 | 219.4 | 217 | 235 | 213 | 215.3 | 235 | 229 | 224 | 219.1 | 223 | 221 | 217 | 217.7 | 217 | 195.3 | 217.4255 | 122 | 217 |  |  | 220.7193
| 1 | 5 | 233 | 233 | 222 | 213.8446 | 222.3 | 218 | 227 | 216 | 212 | 228 | 228 | 221.2 | 215.8 | 229.5 | 220 | 216 | 222.3 | 216 | 190.4 | 222.5078 | 125.45 | 216.2 |  |  | 17
| 1 | 6 | 237 | 233 | 224 | 214.5807 | 223 | 222 | 219 | 217 | 217.6 | 233 | 214 | 224 | 220.1 | 216.5 | 212 | 220 | 225.9 | 216 | 191.3 | 225.5583 | 119.62 | 216.2
| Mean |  | 257.8333 | 232.8333 | 222.8333 | 212.3359 | 222.2667 | 218.8333 | 229.3333 | 218.3333 | 217.8 | 231.3333 | 225.6667 | 221.65 | 215.35 | 222.95 | 217 | 216.5 | 221.3667 | 216.5 | 192.3 | 218.5638 | 123.345 | 212.2167
| Median |  | 260 | 233 | 222.5 | 212.7542 | 222.65 | 218 | 230 | 218.5 | 218.05 | 232 | 228 | 222.15 | 216.05 | 224.85 | 216.5 | 216.5 | 221.9 | 216.5 | 191.55 | 217.5279 | 122.5 | 213.65
| Std Dev. |  | 20.7597 | 4.0208 | 1.472 | 3.0881 | 1.7178 | 3.0605 | 5.9889 | 3.6697 | 3.7646 | 3.7238 | 5.9554 | 2.635 | 4.2336 | 6.7046 | 3.2249 | 2.429 | 2.9918 | 2.0736 | 2.1335 | 4.6042 | 3.3974 | 4.9741
| Rel.Std.Dev. |  | 8.052% | 1.727% | 0.6606% | 1.454% | 0.7728% | 1.399% | 2.611% | 1.681% | 1.728% | 1.61% | 2.639% | 1.189% | 1.966% | 3.007% | 1.486% | 1.122% | 1.351% | 0.9578% | 1.109% | 2.107% | 2.754% | 2.344%
| PDM3 |  | 16.82% | 5.488% | 0.9578% | -3.798% | 0.7011% | -0.8545% | 3.903% | -1.081% | -1.323% | 4.809% | 2.241% | 0.4217% | -2.433% | 1.011% | -1.685% | -1.912% | 0.2933% | -1.912% | -12.88% | -0.9766% | -44.12% | -3.852%
| Z-Score (Absolute) |  | 6.85 | 2.47 | 0.72 | 1.11 | 0.62 | 0.02 | 1.86 | 0.06 | 0.16 | 2.21 | 1.22 | 0.52 | 0.59 | 0.74 | 0.3 | 0.38 | 0.47 | 0.38 | 4.62 | 0.02 | 16.69 | 1.13
| Analytical results for B in OREAS 265 (Indicative Value 32 ppm) | [25 empty] | Indicative
B | Round | Replicate | Lab | [23 empty] | 1
| No | No | 1.07 | [23 empty] | ppm
|  |  | 4A*OES | [23 empty] | 1
1
| 1 | 1 | 34 | [23 empty] | 1
| 1 | 2 | 32 | [23 empty] | 3
| 1 | 3 | 30 | [23 empty] | 16
| 1 | 4 | 30 | [23 empty] | 32
| 1 | 5 | 31 | [23 empty] | 9
| 1 | 6 | 35
| Mean |  | 32
| Median |  | 31.5
| Std Dev. |  | 2.0976
| Rel.Std.Dev. |  | 6.555%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 265 (Certified Value 533 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 536 | 550 | 541 | 522.2004 | 540.9 | 534 | 551 | 540 | 512.8 | 550 | 550 | 542 | 537 | 547 | 533 | 578 | 530.1644 | 581 | 500 | 626 | 530 | 532.9556 | 486 | 536 | 1
| 1 | 2 | 517 | 570 | 546 | 548.0591 | 528.6 | 523 | 545 | 540 | 520.7 | 530 | 550 | 538 | 523 | 517 | 529 | 569 | 529.8793 | 586 | 520 | 640 | 535 | 521.6139 | 501 | 531 | 22
| 1 | 3 | 503 | 560 | 542 | 522.6579 | 540.5 | 519 | 545 | 530 | 517.6 | 550 | 550 | 519 | 516 | 542 | 526 | 568 | 528.9048 | 588 | 510 | 616 | 529 | 533.4337 | 506 | 527 | 16
| 1 | 4 | 516 | 550 | 541 | 525.4541 | 524.4 | 515 | 556 | 520 | 503.1 | 550 | 560 | 545 | 535 | 540 | 537 | 585 | 528.7206 | 573 | 520 | 616 | 537 | 530.5428 | 475 | 552 | 532.7349
| 1 | 5 | 501 | 560 | 542 | 530.9872 | 537.1 | 503 | 531 | 520 | 505.1 | 540 | 560 | 559 | 530 | 520 | 527 | 568 | 528.9069 | 588 | 510 | 619 | 531 | 542.3619 | 478 | 541 | 18
| 1 | 6 | 492 | 550 | 547 | 546.0836 | 532.5 | 501 | 559 | 530 | 511.8 | 550 | 540 | 541 | 534 | 516 | 526 | 586 | 530.4131 | 589 | 510 | 647 | 527 | 543.7483 | 490 | 545
| Mean |  | 510.8333 | 556.6667 | 543.1667 | 532.5737 | 534 | 515.8333 | 547.8333 | 530 | 511.85 | 545 | 551.6667 | 540.6667 | 529.1667 | 530.3333 | 529.6667 | 575.6667 | 529.4982 | 584.1667 | 511.6667 | 627.3333 | 531.5 | 534.1094 | 489.3333 | 538.6667
| Median |  | 509.5 | 555 | 542 | 528.2206 | 534.8 | 517 | 548 | 530 | 512.3 | 550 | 550 | 541.5 | 532 | 530 | 528 | 573.5 | 529.3931 | 587 | 510 | 622.5 | 530.5 | 533.1946 | 488 | 538.5
| Std Dev. |  | 15.5617 | 8.165 | 2.6394 | 11.6743 | 6.6789 | 12.4646 | 10.0083 | 8.9443 | 6.8471 | 8.3666 | 7.5277 | 12.9099 | 8.1343 | 14.1233 | 4.4572 | 8.501 | 0.73927 | 6.1779 | 7.5277 | 13.2313 | 3.7815 | 8.1444 | 12.3234 | 9.2232
| Rel.Std.Dev. |  | 3.046% | 1.467% | 0.4859% | 2.192% | 1.251% | 2.416% | 1.827% | 1.688% | 1.338% | 1.535% | 1.365% | 2.388% | 1.537% | 2.663% | 0.8415% | 1.477% | 0.1396% | 1.058% | 1.471% | 2.109% | 0.7115% | 1.525% | 2.518% | 1.712%
| PDM3 |  | -4.111% | 4.492% | 1.958% | -0.03025% | 0.2375% | -3.173% | 2.834% | -0.5134% | -3.92% | 2.302% | 3.554% | 1.489% | -0.6698% | -0.4508% | -0.5759% | 8.059% | -0.6076% | 9.654% | -3.955% | 17.76% | -0.2318% | 0.258% | -8.147% | 1.113%
| Z-Score (Absolute) |  | 1.4 | 1.46 | 0.62 | 0.04 | 0.04 | 1.09 | 0.91 | 0.21 | 1.34 | 0.73 | 1.15 | 0.46 | 0.26 | 0.18 | 0.23 | 2.65 | 0.24 | 3.18 | 1.35 | 5.87 | 0.11 | 0.05 | 2.75 | 0.34
| Analytical results for Be in OREAS 265 (Certified Value 2.55 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.4 | 2.59 | 2.6 | 2.6625 | 2.68 | 2.79 | < 5 | 2.52 | 2.48 | 2.58 | 2.57 | 3 | 3 | 2.61 | 2.7 | 2.89 | 2.4058 | 2.14 | 2.44 | 2.9 | 3 | 2.4091 | 2 | 2.83 | 1
| 1 | 2 | 2.2 | 2.64 | 2.5 | 2.4245 | 2.46 | 2.63 | < 5 | 2.5 | 2.38 | 2.5 | 2.57 | 3 | 3 | 2.43 | 2.6 | 2.65 | 2.4499 | 2.15 | 2.45 | 2.8 | 3 | 2.399 | 2 | 2.75 | 23
| 1 | 3 | 2.3 | 2.68 | 2.6 | 2.6115 | 2.58 | 2.63 | < 5 | 2.49 | 2.48 | 2.55 | 2.57 | 3 | 3 | 2.5 | 2.7 | 2.4 | 2.4068 | 2.16 | 2.43 | 2.7 | 3 | 2.4451 | 2 | 2.84 | 16
| 1 | 4 | 2.4 | 2.58 | 2.5 | 2.3991 | 2.36 | 2.69 | < 5 | 2.39 | 2.57 | 2.56 | 2.61 | 3 | 4 | 2.61 | 2.8 | 2.77 | 2.4364 | 2.11 | 2.45 | 2.9 | 3 | 2.439 | 2 | 3.03 | 2.5493
| 1 | 5 | 2.3 | 2.62 | 2.5 | 2.6006 | 2.37 | 2.74 | < 5 | 2.4 | 2.58 | 2.52 | 2.62 | 3 | 3 | 2.49 | 2.7 | 2.86 | 2.378 | 2.17 | 2.44 | 2.9 | 3 | 2.4993 | 2 | 2.92 | 19
| 1 | 6 | 2.2 | 2.61 | 2.6 | 2.6292 | 2.4 | 2.81 | < 5 | 2.45 | 2.47 | 2.57 | 2.48 | 3 | 4 | 2.46 | 2.7 | 2.68 | 2.366 | 2.17 | 2.44 | 2.8 | 3 | 2.4953 | 2 | 3.02
| Mean |  | 2.3 | 2.62 | 2.55 | 2.5546 | 2.475 | 2.715 |  | 2.4583 | 2.4933 | 2.5467 | 2.57 | 3 | 3.3333 | 2.5167 | 2.7 | 2.7083 | 2.4071 | 2.15 | 2.4417 | 2.8333 | 3 | 2.4478 | 2 | 2.8983
| Median |  | 2.3 | 2.615 | 2.55 | 2.606 | 2.43 | 2.715 |  | 2.47 | 2.48 | 2.555 | 2.57 | 3 | 3 | 2.495 | 2.7 | 2.725 | 2.4063 | 2.155 | 2.44 | 2.85 | 3 | 2.4421 | 2 | 2.88
| Std Dev. |  | 0.089443 | 0.036332 | 0.054772 | 0.11284 | 0.12896 | 0.07791 |  | 0.054191 | 0.073666 | 0.030768 | 0.049396 | 0 | 0.5164 | 0.076333 | 0.063246 | 0.17837 | 0.032342 | 0.022804 | 0.0075277 | 0.08165 | 0 | 0.042128 | 0 | 0.11197
| Rel.Std.Dev. |  | 3.889% | 1.387% | 2.148% | 4.417% | 5.21% | 2.87% |  | 2.204% | 2.955% | 1.208% | 1.922% | 0% | 15.49% | 3.033% | 2.342% | 6.586% | 1.344% | 1.061% | 0.3083% | 2.882% | 0% | 1.721% | 0% | 3.863%
| PDM3 |  | -9.779% | 2.774% | 0.02811% | 0.2067% | -2.914% | 6.501% |  | -3.568% | -2.195% | -0.1026% | 0.8126% | 17.68% | 30.76% | -1.279% | 5.912% | 6.239% | -5.576% | -15.66% | -4.221% | 11.14% | 17.68% | -3.981% | -21.55% | 13.69%
| Z-Score (Absolute) |  | 1.95 | 0.93 | 0.3 | 0.34 | 0.37 | 1.78 | 0.15 | 0.52 | 0.21 | 0.27 | 0.48 | NA | NA | 0 | 1.65 | 1.72 | 0.98 | 3.3 | 0.67 | 2.85 | NA | 0.62 | 4.65 | 3.43
Results from laboratories 1.12, 1.13 and 1.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 265 (Certified Value 2.11 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.18 | 2.31 | 2.21 | 2.2417 | 2.14 | 2.23 | 1.8 | 2.1 | 2.25 | 1.98 | 1.97 | 2.22 | 2.06 | 2.1 | 2.19 | 2.08 | 2.03 | 1.93 | 2.7 | 2.1 | 2.0716 | < 5 | 2.05 |  | 1
| 1 | 2 | 2.04 | 2.19 | 2.19 | 2.143 | 2.1 | 2.24 | 2 | 2.11 | 2.13 | 1.94 | 2.25 | 2.23 | 2.14 | 2.2 | 2.12 | 2.09 | 2.28 | 1.98 | 2.7 | 2.1 | 2.0791 | < 5 | 2.07 |  | 24
| 1 | 3 | 2.03 | 2.22 | 2.21 | 2.1104 | 2.12 | 2.16 | 1.91 | 2.1 | 2.21 | 1.94 | 2.05 | 2.25 | 2.07 | 2.22 | 2.14 | 2.01 | 2.12 | 2.1 | 2.4 | 2.1 | 2.1112 | < 5 | 2.15 |  | 16
| 1 | 4 | 2.03 | 2.44 | 2.2 | 2.0609 | 2.14 | 2.14 | 2.03 | 1.99 | 2.21 | 2.05 | 2.02 | 2.3 | 1.98 | 1.98 | 2.23 | 2 | 2.16 | 1.87 | 2.4 | 2.3 | 2.0413 | < 5 | 2.13 |  | 2.1103
| 1 | 5 | 2.09 | 2.2 | 2.2 | 2.2167 | 2.13 | 2.22 | 1.81 | 2.01 | 2.17 | 1.98 | 2.02 | 2.33 | 2.04 | 2.12 | 2.13 | 2.04 | 2.37 | 1.92 | 2.4 | 2 | 2.1736 | < 5 | 2.07 |  | 20
| 1 | 6 | 2.13 | 2.11 | 2.21 | 2.2687 | 2.14 | 2.19 | 1.76 | 2.15 | 2.14 | 2.01 | 1.91 | 2.28 | 2.75 | 1.91 | 2.1 | 2.28 | 2.56 | 1.88 | 2.5 | 2.1 | 2.2446 | < 5 | 2.1
| Mean |  | 2.0833 | 2.245 | 2.2033 | 2.1736 | 2.1283 | 2.1967 | 1.885 | 2.0767 | 2.185 | 1.9833 | 2.0367 | 2.2683 | 2.1733 | 2.0883 | 2.1517 | 2.0833 | 2.2533 | 1.9467 | 2.5167 | 2.1167 | 2.1202 |  | 2.095
| Median |  | 2.065 | 2.21 | 2.205 | 2.1799 | 2.135 | 2.205 | 1.86 | 2.1 | 2.19 | 1.98 | 2.02 | 2.265 | 2.065 | 2.11 | 2.135 | 2.06 | 2.22 | 1.925 | 2.45 | 2.1 | 2.0952 |  | 2.085
| Std Dev. |  | 0.06186 | 0.11502 | 0.008165 | 0.081456 | 0.016021 | 0.040332 | 0.11256 | 0.062503 | 0.046368 | 0.042269 | 0.11553 | 0.042622 | 0.28717 | 0.12205 | 0.048751 | 0.10289 | 0.19222 | 0.084774 | 0.1472 | 0.098319 | 0.07573 |  | 0.038859
| Rel.Std.Dev. |  | 2.969% | 5.123% | 0.3706% | 3.748% | 0.7527% | 1.836% | 5.971% | 3.01% | 2.122% | 2.131% | 5.672% | 1.879% | 13.21% | 5.844% | 2.266% | 4.939% | 8.53% | 4.355% | 5.849% | 4.645% | 3.572% |  | 1.855%
| PDM3 |  | -1.28% | 6.381% | 4.406% | 2.998% | 0.8525% | 4.091% | -10.68% | -1.596% | 3.538% | -6.018% | -3.491% | 7.487% | 2.985% | -1.043% | 1.958% | -1.28% | 6.776% | -7.756% | 19.25% | 0.2997% | 0.4691% |  | -0.727%
| Z-Score (Absolute) |  | 0.54 | 1.39 | 0.89 | 0.54 | 0 | 0.81 | 2.9 | 0.61 | 0.67 | 1.73 | 1.09 | 1.67 | 0.54 | 0.48 | 0.28 | 0.54 | 1.49 | 2.16 | 4.62 | 0.14 | 0.1 | 4.42 | 0.4
| Analytical results for Ca in OREAS 265 (Certified Value 0.857 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.83 | 0.88 | 0.83 | 0.86298 | 0.8522 | 0.895 | 0.91 | 0.9 | 0.8213 | 0.87 | 0.83 | 0.86 | 0.86 | 0.86 | 0.82 | 0.89 | 0.86436 | 0.9919 | 0.85 | 0.822 | 0.82 | 0.82498 | 0.8526 | 0.87 | 1
| 1 | 2 | 0.76 | 0.91 | 0.844 | 0.86223 | 0.8479 | 0.877 | 0.9 | 0.88 | 0.8242 | 0.84 | 0.83 | 0.89 | 0.88 | 0.84 | 0.85 | 0.87 | 0.86349 | 1.0003 | 0.87 | 0.85 | 0.78 | 0.82196 | 0.8463 | 0.85 | #N/A
| 1 | 3 | 0.74 | 0.9 | 0.848 | 0.84514 | 0.8543 | 0.866 | 0.9 | 0.88 | 0.8223 | 0.86 | 0.84 | 0.89 | 0.83 | 0.89 | 0.83 | 0.88 | 0.86516 | 0.9969 | 0.86 | 0.829 | 0.78 | 0.83913 | 0.853 | 0.85 | 16
| 1 | 4 | 0.7 | 0.88 | 0.841 | 0.85388 | 0.8398 | 0.865 | 0.92 | 0.85 | 0.8011 | 0.86 | 0.86 | 0.88 | 0.9 | 0.89 | 0.86 | 0.91 | 0.86466 | 0.9721 | 0.87 | 0.836 | 0.81 | 0.83578 | 0.8261 | 0.9 | 0.85654
| 1 | 5 | 0.65 | 0.9 | 0.839 | 0.84689 | 0.8427 | 0.872 | 0.85 | 0.85 | 0.8165 | 0.85 | 0.86 | 0.88 | 0.86 | 0.85 | 0.81 | 0.88 | 0.86378 | 0.9881 | 0.87 | 0.829 | 0.76 | 0.86496 | 0.8585 | 0.87 | 21
| 1 | 6 | 0.66 | 0.88 | 0.845 | 0.86248 | 0.8441 | 0.867 | 0.86 | 0.89 | 0.8288 | 0.87 | 0.81 | 0.89 | 0.87 | 0.84 | 0.82 | 0.91 | 0.86406 | 0.9854 | 0.87 | 0.865 | 0.8 | 0.86274 | 0.83453 | 0.88
| Mean |  | 0.72333 | 0.89167 | 0.84117 | 0.8556 | 0.84683 | 0.87367 | 0.89 | 0.875 | 0.81903 | 0.85833 | 0.83833 | 0.88167 | 0.86667 | 0.86167 | 0.83167 | 0.89 | 0.86425 | 0.98912 | 0.865 | 0.8385 | 0.79167 | 0.84159 | 0.84517 | 0.87
| Median |  | 0.72 | 0.89 | 0.8425 | 0.85805 | 0.846 | 0.8695 | 0.9 | 0.88 | 0.8218 | 0.86 | 0.835 | 0.885 | 0.865 | 0.855 | 0.825 | 0.885 | 0.86421 | 0.99 | 0.87 | 0.8325 | 0.79 | 0.83746 | 0.84945 | 0.87
| Std Dev. |  | 0.067725 | 0.013292 | 0.0063061 | 0.0081737 | 0.0056518 | 0.011378 | 0.028284 | 0.020736 | 0.0096531 | 0.01169 | 0.019408 | 0.01169 | 0.023381 | 0.023166 | 0.019408 | 0.016733 | 0.00060612 | 0.0099781 | 0.0083666 | 0.016084 | 0.022286 | 0.018408 | 0.012429 | 0.018974
| Rel.Std.Dev. |  | 9.363% | 1.491% | 0.7497% | 0.9553% | 0.6674% | 1.302% | 3.178% | 2.37% | 1.179% | 1.362% | 2.315% | 1.326% | 2.698% | 2.689% | 2.334% | 1.88% | 0.07013% | 1.009% | 0.9672% | 1.918% | 2.815% | 2.187% | 1.471% | 2.181%
| PDM3 |  | -15.55% | 4.101% | -1.795% | -0.1097% | -1.133% | 1.999% | 3.906% | 2.155% | -4.379% | 0.2093% | -2.126% | 2.933% | 1.182% | 0.5985% | -2.904% | 3.906% | 0.9003% | 15.48% | 0.9876% | -2.106% | -7.574% | -1.745% | -1.327% | 1.571%
| Z-Score (Absolute) |  | 4.95 | 1.15 | 0.68 | 0.16 | 0.48 | 0.49 | 1.09 | 0.54 | 1.48 | 0.06 | 0.78 | 0.78 | 0.24 | 0.06 | 1.03 | 1.09 | 0.15 | 4.68 | 0.18 | 0.78 | 2.47 | 0.67 | 0.54 | 0.36
| Analytical results for Cd in OREAS 265 (Indicative Value 0.045 ppm) | [25 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.06 | 0.04 | < 0.05 | < 0.05 | 0.05 | 0.04 | 0.13 | 0.02 | 0.05 | 0.05 | < 0.02 | < 0.02 | 0.03 | 0.34 | 0.03 | 0.08 | 0.06 | 0.04 | 0.2 | < 0.1 | < 0.2 | 0.4 | 0.06 |  | 1
| 1 | 2 | 0.05 | 0.06 | < 0.05 | < 0.05 | 0.07 | 0.04 | 0.13 | < 0.02 | 0.04 | 0.03 | < 0.02 | 0.03 | 0.03 | 0.34 | 0.04 | 0.06 | 0.06 | 0.02 | 0.2 | < 0.1 | < 0.2 | 0.5 | 0.05 |  | 4
| 1 | 3 | 0.05 | 0.04 | < 0.05 | < 0.05 | 0.08 | 0.04 | 0.1 | 0.03 | 0.06 | 0.05 | < 0.02 | < 0.02 | < 0.02 | 0.34 | 0.03 | 0.08 | 0.06 | 0.05 | 0.1 | < 0.1 | < 0.2 | 0.4 | 0.05 |  | 16
| 1 | 4 | 0.07 | 0.05 | < 0.05 | < 0.05 | 0.06 | 0.04 | 0.14 | < 0.02 | 0.03 | 0.04 | < 0.02 | 0.04 | 0.04 | 0.34 | 0.04 | 0.06 | 0.07 | 0.02 | 0.2 | < 0.1 | < 0.2 | 0.4 | 0.05 |  | 0.04506
| 1 | 5 | 0.05 | 0.03 | < 0.05 | < 0.05 | 0.05 | 0.04 | 0.15 | 0.02 | 0.04 | 0.03 | < 0.02 | 0.05 | < 0.02 | 0.37 | 0.03 | 0.04 | 0.07 | 0.03 | 0.2 | < 0.1 | < 0.2 | 0.5 | 0.06 |  | 10
| 1 | 6 | 0.04 | 0.04 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.12 | 0.03 | 0.05 | 0.04 | < 0.02 | 0.05 | < 0.02 | 0.35 | 0.04 | 0.06 | 0.07 | 0.02 | 0.2 | < 0.1 | < 0.2 | 0.482 | 0.06
| Mean |  | 0.053333 | 0.043333 |  |  | 0.058333 | 0.041667 | 0.12833 | 0.025 | 0.045 | 0.04 |  | 0.0425 | 0.033333 | 0.34667 | 0.035 | 0.063333 | 0.065 | 0.03 | 0.18333 |  |  | 0.447 | 0.055
| Median |  | 0.05 | 0.04 |  |  | 0.055 | 0.04 | 0.13 | 0.025 | 0.045 | 0.04 |  | 0.045 | 0.03 | 0.34 | 0.035 | 0.06 | 0.065 | 0.025 | 0.2 |  |  | 0.441 | 0.055
| Std Dev. |  | 0.010328 | 0.010328 |  |  | 0.01472 | 0.0040825 | 0.017224 | 0.0057735 | 0.010488 | 0.0089443 |  | 0.0095743 | 0.0057735 | 0.012111 | 0.0054772 | 0.015055 | 0.0054772 | 0.012649 | 0.040825 |  |  | 0.051904 | 0.0054772
| Rel.Std.Dev. |  | 19.36% | 23.83% |  |  | 25.23% | 9.798% | 13.42% | 23.09% | 23.31% | 22.36% |  | 22.53% | 17.32% | 3.493% | 15.65% | 23.77% | 8.427% | 42.16% | 22.27% |  |  | 11.61% | 9.959%
| PDM3 |  | 18.36% | -3.831% |  |  | 29.46% | -7.53% | 184.8% | -44.52% | -0.1321% | -11.23% |  | -5.68% | -26.02% | 669.4% | -22.32% | 40.55% | 44.25% | -33.42% | 306.9% |  |  | 892% | 22.06%
| Z-Score (Absolute) |  | 0.37 | 0.03 | 0.78 | 0.78 | 0.57 | 0.1 | 3.41 | 0.98 | 0.03 | 0.17 | 1.38 | 0.51 | 0.91 | 12.24 | 0.37 | 0.78 | 0.84 | 0.57 | NA | 0.24 | 2.26 | 16.3 | 0.44
Results from laboratory 1.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 265 (Certified Value 93 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 0
1
| 1 | 1 | 100.26 | 97.5 | 98.6 | 89.72 | 97 | 101 | 92.41 | 99.3 | 93.3 | 95.39 | 107.71 | 89.12 | 89.61 | 91.9 | 92.4097 | 79.7 | 97 | 90 | 88 | 92.5522 | 84.31 |  |  |  | 1
| 1 | 2 | 93.47 | 98.9 | 97.6 | 91.28 | 96 | 97.9 | 95.93 | 98 | 91 | 95.54 | 106.93 | 85.58 | 87.95 | 91.9 | 92.4667 | 80.4 | 97.3 | 93 | 84 | 89.7048 | 85.04 |  |  |  | 26
| 1 | 3 | 90.32 | 100 | 96.23 | 86.31 | 92 | 96.8 | 94.79 | 97.4 | 94.5 | 92.92 | 103.67 | 86.43 | 87.91 | 93.3 | 92.2966 | 81.7 | 98 | 89 | 90 | 92.7591 | 84.36 |  |  |  | 16
| 1 | 4 | 89.5 | 99.1 | 96.33 | 86.5 | 95 | 95.5 | 92.01 | 101.5 | 95.7 | 95.51 | 91.56 | 89.32 | 89.52 | 92.4 | 92.3164 | 78.9 | 97.5 | 88 | 90 | 91.9954 | 88 |  |  |  | 92.6624
| 1 | 5 | 90.16 | 100 | 96.5 | 88.64 | 96 | 96.2 | 93.72 | 99.9 | 95.3 | 99.52 | 105.38 | 89.22 | 87.86 | 91.7 | 92.2633 | 81.1 | 95.6 | 90 | 80 | 94.7805 | 89.49 |  |  |  | 22
| 1 | 6 | 86.95 | 95.2 | 97.56 | 89.58 | 92 | 99.6 | 94.79 | 97 | 90.6 | 98.96 | 90.57 | 85.72 | 86.47 | 93.9 | 92.4574 | 81.9 | 97.5 | 93 | 91 | 95.3748 | 88.63
| Mean |  | 91.7767 | 98.45 | 97.1367 | 88.6717 | 94.6667 | 97.8333 | 93.9417 | 98.85 | 93.4 | 96.3067 | 100.97 | 87.565 | 88.22 | 92.5167 | 92.3684 | 80.6167 | 97.15 | 90.5 | 87.1667 | 92.8611 | 86.6383
| Median |  | 90.24 | 99 | 97.03 | 89.11 | 95.5 | 97.35 | 94.255 | 98.65 | 93.9 | 95.525 | 104.525 | 87.775 | 87.93 | 92.15 | 92.3631 | 80.75 | 97.4 | 90 | 89 | 92.6557 | 86.52
| Std Dev. |  | 4.6486 | 1.8382 | 0.9395 | 1.9505 | 2.1602 | 2.1134 | 1.5178 | 1.7073 | 2.1781 | 2.4859 | 7.8025 | 1.8368 | 1.1818 | 0.89088 | 0.087397 | 1.1737 | 0.82644 | 2.0736 | 4.3089 | 2.0406 | 2.3289
| Rel.Std.Dev. |  | 5.065% | 1.867% | 0.9672% | 2.2% | 2.282% | 2.16% | 1.616% | 1.727% | 2.332% | 2.581% | 7.728% | 2.098% | 1.34% | 0.9629% | 0.09462% | 1.456% | 0.8507% | 2.291% | 4.943% | 2.197% | 2.688%
| PDM3 |  | -0.9558% | 6.246% | 4.829% | -4.307% | 2.163% | 5.58% | 1.381% | 6.678% | 0.7961% | 3.933% | 8.965% | -5.501% | -4.794% | -0.1572% | -0.3173% | -13% | 4.843% | -2.334% | -5.931% | 0.2145% | -6.501%
| Z-Score (Absolute) |  | 0.17 | 0.88 | 0.67 | 0.66 | 0.28 | 0.78 | 0.17 | 0.94 | 0.08 | 0.54 | 1.28 | 0.84 | 0.73 | 0.05 | 0.08 | 1.93 | 0.68 | 0.37 | 0.9 | 0 | 0.98
| Analytical results for Co in OREAS 265 (Certified Value 11 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 11.4 | 11.3 | 11 | 11.1898 | 11.3 | 11 | 11.1 | 10.6 | 10.8 | 10.9 | 10.8 | 10.8 | 10.1 | 10.7 | 11 | 10.4 | 11.1458 | 12.4 | 11.1 | 11.4 | 10.7 | 10.8493 | 10.46 | 11 | 1
| 1 | 2 | 11.2 | 12.1 | 11 | 11.0208 | 10.9 | 11 | 10.8 | 10.2 | 11 | 10.6 | 10.8 | 11.3 | 9.7 | 10.7 | 10.9 | 10.6 | 11.1743 | 12.4 | 11 | 11.4 | 10.9 | 10.7719 | 10.52 | 11.4 | 14
| 1 | 3 | 11 | 11.4 | 11 | 10.9672 | 11 | 11 | 10.3 | 10.2 | 10.8 | 10.5 | 11.1 | 11.3 | 10.4 | 10.6 | 10.9 | 10.7 | 11.1558 | 12.6 | 11.2 | 11.2 | 11.1 | 11.0112 | 10.63 | 11.2 | 16
| 1 | 4 | 11.3 | 11.6 | 11 | 11.1538 | 11.2 | 11.1 | 11.3 | 10.6 | 11 | 10.9 | 11.2 | 11.3 | 10.1 | 10.7 | 11.1 | 10.7 | 11.1865 | 12 | 11 | 11.3 | 10.8 | 10.7677 | 10.724 | 11.6 | 10.9749
| 1 | 5 | 10.9 | 11.8 | 11 | 11.2698 | 10.9 | 11.3 | 11.1 | 10 | 10.8 | 10.6 | 10.9 | 11.4 | 9.9 | 10.8 | 11 | 10.7 | 11.1184 | 12.6 | 11 | 11.4 | 10.7 | 11.2096 | 10.56 | 11.5 | 23
| 1 | 6 | 10.5 | 11.2 | 11 | 10.9067 | 11.1 | 11.2 | 10.7 | 11 | 10.8 | 10.7 | 10.8 | 11.1 | 10.6 | 10.8 | 10.7 | 10.9 | 11.1557 | 12.7 | 11.1 | 11.5 | 10.7 | 11.3713 | 10.76 | 11.5
| Mean |  | 11.05 | 11.5667 | 11 | 11.0847 | 11.0667 | 11.1 | 10.8833 | 10.4333 | 10.8667 | 10.7 | 10.9333 | 11.2 | 10.1333 | 10.7167 | 10.9333 | 10.6667 | 11.1561 | 12.45 | 11.0667 | 11.3667 | 10.8167 | 10.9968 | 10.609 | 11.3667
| Median |  | 11.1 | 11.5 | 11 | 11.0873 | 11.05 | 11.05 | 10.95 | 10.4 | 10.8 | 10.65 | 10.85 | 11.3 | 10.1 | 10.7 | 10.95 | 10.7 | 11.1557 | 12.5 | 11.05 | 11.4 | 10.75 | 10.9303 | 10.595 | 11.45
| Std Dev. |  | 0.32711 | 0.33862 | 0 | 0.14119 | 0.1633 | 0.12649 | 0.36009 | 0.36697 | 0.10328 | 0.16733 | 0.17512 | 0.21909 | 0.3266 | 0.075277 | 0.13663 | 0.1633 | 0.023597 | 0.251 | 0.08165 | 0.10328 | 0.16021 | 0.24925 | 0.11746 | 0.22509
| Rel.Std.Dev. |  | 2.96% | 2.928% | 0% | 1.274% | 1.476% | 1.14% | 3.309% | 3.517% | 0.9504% | 1.564% | 1.602% | 1.956% | 3.223% | 0.7024% | 1.25% | 1.531% | 0.2115% | 2.016% | 0.7378% | 0.9086% | 1.481% | 2.267% | 1.107% | 1.98%
| PDM3 |  | 0.6841% | 5.392% | 0.2285% | 1% | 0.836% | 1.14% | -0.8345% | -4.935% | -0.9864% | -2.505% | -0.3789% | 2.051% | -7.668% | -2.353% | -0.3789% | -2.809% | 1.651% | 13.44% | 0.836% | 3.569% | -1.442% | 0.1998% | -3.334% | 3.569%
| Z-Score (Absolute) |  | 0.2 | 2.13 | NA | 0.33 | 0.26 | 0.39 | 0.42 | 2.11 | 0.49 | 1.11 | 0.24 | 0.76 | 3.23 | 1.05 | 0.24 | 1.24 | 0.6 | 5.44 | 0.26 | 1.38 | 0.67 | 0 | 1.45 | 1.38
Results from laboratory 1.03 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 265 (Certified Value 88 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 75 | 93 | 82 | 88.6739 | 103 | 78 | 81 | 99 | 107 | 99 | 95 | 65 | 72 | 84 | 94 | 71 | 98.6 | 91 | 124 | 90 | 96.7014 | 54.25 | 98 |  | 1
| 1 | 2 | 71 | 98 | 80 | 90.0878 | 102 | 78 | 78 | 100 | 99 | 91 | 94 | 66 | 65 | 78 | 97 | 69 | 99.1 | 95 | 124 | 87 | 96.7886 | 54.86 | 98 |  | 28
| 1 | 3 | 65 | 100 | 80 | 83.8074 | 101 | 72 | 76 | 97 | 100 | 93 | 97 | 63 | 67 | 84 | 96 | 74 | 99.5 | 96 | 122 | 90 | 98.0515 | 55.22 | 93 |  | 16
| 1 | 4 | 63 | 98 | 80 | 89.9818 | 105 | 70 | 75 | 94 | 106 | 92 | 97 | 65 | 69 | 79 | 99 | 73 | 97 | 95 | 110 | 90 | 98.1527 | 55.12 | 97 |  | 87.6835
| 1 | 5 | 55 | 99 | 87 | 83.2519 | 111 | 74 | 78 | 96 | 113 | 91 | 97 | 70 | 67 | 81 | 93 | 72 | 99.7 | 95 | 124 | 89 | 100.2843 | 54.86 | 96 |  | 24
| 1 | 6 | 54 | 99 | 86 | 88.6038 | 105 | 79 | 79 | 96 | 112 | 93 | 93 | 68 | 67 | 78 | 96 | 75 | 99.4 | 95 | 121 | 88 | 102.3329 | 53.99 | 96
| Mean |  | 63.8333 | 97.8333 | 82.5 | 87.4011 | 104.5 | 75.1667 | 77.8333 | 97 | 106.1667 | 93.1667 | 95.5 | 66.1667 | 67.8333 | 80.6667 | 95.8333 | 72.3333 | 98.8833 | 94.5 | 120.8333 | 89 | 98.7186 | 54.7167 | 96.3333
| Median |  | 64 | 98.5 | 81 | 88.6389 | 104 | 76 | 78 | 96.5 | 106.5 | 92.5 | 96 | 65.5 | 67 | 80 | 96 | 72.5 | 99.25 | 95 | 123 | 89.5 | 98.1021 | 54.86 | 96.5
| Std Dev. |  | 8.4004 | 2.4833 | 3.2094 | 3.0684 | 3.5637 | 3.7103 | 2.137 | 2.1909 | 5.8452 | 2.9944 | 1.7607 | 2.4833 | 2.4014 | 2.8048 | 2.137 | 2.1602 | 0.99883 | 1.7607 | 5.4559 | 1.2649 | 2.1943 | 0.4905 | 1.8619
| Rel.Std.Dev. |  | 13.16% | 2.538% | 3.89% | 3.511% | 3.41% | 4.936% | 2.746% | 2.259% | 5.506% | 3.214% | 1.844% | 3.753% | 3.54% | 3.477% | 2.23% | 2.987% | 1.01% | 1.863% | 4.515% | 1.421% | 2.223% | 0.8964% | 1.933%
| PDM3 |  | -27.2% | 11.58% | -5.912% | -0.322% | 19.18% | -14.27% | -11.23% | 10.63% | 21.08% | 6.253% | 8.914% | -24.54% | -22.64% | -8.002% | 9.295% | -17.51% | 12.77% | 7.774% | 37.81% | 1.501% | 12.59% | -37.6% | 9.865%
| Z-Score (Absolute) |  | 1.85 | 0.3 | 0.67 | 0.36 | 0.72 | 1.14 | 0.97 | 0.24 | 0.82 | 0 | 0.15 | 1.71 | 1.6 | 0.79 | 0.17 | 1.32 | 0.36 | 0.08 | 1.75 | 0.26 | 0.35 | 2.43 | 0.2
| Analytical results for Cs in OREAS 265 (Certified Value 6.2 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 7 | 6.56 | 6.4 | 6.4501 | 6.24 | 6.08 | 6.29 | 6.01 | 6.2 | 6.04 | 6.21 | 6.3 | 5.8 | 5.88 | 6.16 | 6.21 | 5.9599 | 6 | 5.88 | 6.3 | 7 | 6.0474 | 5.66 |  | 1
| 1 | 2 | 6 | 6.67 | 6.2 | 6.3952 | 6.14 | 6.05 | 6.45 | 5.85 | 6.27 | 6.03 | 6.06 | 6.4 | 5.7 | 5.8 | 6.12 | 6.3 | 5.9423 | 6.2 | 5.96 | 6.5 | 6.4 | 5.938 | 5.75 |  | 29
| 1 | 3 | 6 | 6.7 | 6.2 | 6.2602 | 6.3 | 5.94 | 6.09 | 6.24 | 6.23 | 6.06 | 6.23 | 6.2 | 5.8 | 5.97 | 6.1 | 6.31 | 5.9709 | 6.3 | 5.96 | 6.2 | 6.6 | 6.0942 | 5.67 |  | 16
| 1 | 4 | 6 | 6.68 | 6.3 | 6.2944 | 6.21 | 6.06 | 6.72 | 6.37 | 6.07 | 6.14 | 6.37 | 6.5 | 5.3 | 5.95 | 6.12 | 6.65 | 5.916 | 6.3 | 5.98 | 6.3 | 6.8 | 5.9832 | 5.93 |  | 6.1963
| 1 | 5 | 6 | 6.86 | 6.3 | 6.3301 | 6.19 | 6.08 | 6.2 | 6.26 | 6.15 | 6.07 | 6.3 | 6.6 | 5.8 | 5.98 | 6.08 | 6.13 | 5.9004 | 6.6 | 5.9 | 6.3 | 6.2 | 6.1595 | 6.09 |  | 25
| 1 | 6 | 6 | 6.6 | 6.2 | 6.3685 | 6.14 | 6.04 | 6.25 | 6.5 | 6.11 | 6.04 | 6.19 | 6.5 | 5.5 | 5.81 | 5.99 | 6.16 | 5.9291 | 6.4 | 5.85 | 6.4 | 6.7 | 6.2085 | 5.98
| Mean |  | 6.1667 | 6.6783 | 6.2667 | 6.3497 | 6.2033 | 6.0417 | 6.3333 | 6.205 | 6.1717 | 6.0633 | 6.2267 | 6.4167 | 5.65 | 5.8983 | 6.095 | 6.2933 | 5.9364 | 6.3 | 5.9217 | 6.3333 | 6.6167 | 6.0718 | 5.8467
| Median |  | 6 | 6.675 | 6.25 | 6.3493 | 6.2 | 6.055 | 6.27 | 6.25 | 6.175 | 6.05 | 6.22 | 6.45 | 5.75 | 5.915 | 6.11 | 6.255 | 5.9357 | 6.3 | 5.93 | 6.3 | 6.65 | 6.0708 | 5.84
| Std Dev. |  | 0.40825 | 0.10362 | 0.08165 | 0.069207 | 0.061536 | 0.052313 | 0.22313 | 0.23772 | 0.075476 | 0.040332 | 0.1052 | 0.1472 | 0.20736 | 0.080353 | 0.057879 | 0.18917 | 0.026611 | 0.2 | 0.052313 | 0.10328 | 0.28577 | 0.10317 | 0.17851
| Rel.Std.Dev. |  | 6.62% | 1.552% | 1.303% | 1.09% | 0.992% | 0.8659% | 3.523% | 3.831% | 1.223% | 0.6652% | 1.689% | 2.294% | 3.67% | 1.362% | 0.9496% | 3.006% | 0.4483% | 3.175% | 0.8834% | 1.631% | 4.319% | 1.699% | 3.053%
| PDM3 |  | -0.4777% | 7.78% | 1.136% | 2.477% | 0.114% | -2.495% | 2.212% | 0.1409% | -0.397% | -2.145% | 0.4906% | 3.557% | -8.816% | -4.808% | -1.634% | 1.567% | -4.193% | 1.674% | -4.432% | 2.212% | 6.785% | -2.009% | -5.642%
| Z-Score (Absolute) |  | NA | 2.34 | 0.31 | 0.72 | 0 | 0.8 | 0.64 | 0 | 0.16 | 0.7 | 0.11 | 1.05 | 2.74 | 1.51 | 0.54 | 0.44 | 1.32 | 0.47 | 1.39 | 0.64 | 2.04 | 0.65 | 1.76
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 265 (Certified Value 23.5 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 24.6 | 25.6 | 24 | 22.4559 | 23 | 26.4 | 24.7 | 23.3 | 23.2 | 24.3 | 21 | 23.3 | 24.3 | 28.7 | 25.4 | 23.4 | 22.414 | 12.4 | 20.5 | 25.1 | 21.3 | 24.1028 | 19.968 | 23.3 | 1
| 1 | 2 | 24.5 | 25.6 | 24 | 23.0391 | 23 | 26.1 | 26.6 | 22.9 | 23.3 | 23.9 | 21 | 24 | 23.8 | 29.1 | 25.4 | 23.1 | 22.3834 | 12.8 | 20.7 | 25.1 | 21.5 | 22.5192 | 20.35 | 24 | 16
| 1 | 3 | 23.5 | 27.6 | 23 | 22.0592 | 23 | 26.1 | 24.2 | 22.2 | 22.9 | 24 | 23.8 | 23.5 | 23.3 | 27.8 | 25 | 23.6 | 22.2617 | 15.3 | 20.8 | 23.9 | 21.4 | 22.6258 | 20.85 | 23.6 | 16
| 1 | 4 | 23.5 | 25.4 | 23 | 23.2199 | 22 | 25 | 27.6 | 22 | 25.6 | 24.4 | 24.6 | 23.6 | 24.7 | 27.1 | 25.6 | 23.6 | 22.2304 | 15.7 | 20.8 | 24.2 | 21.4 | 23.2996 | 20.89 | 24.2 | 23.4881
| 1 | 5 | 22.3 | 27 | 24 | 22.3176 | 23 | 23.6 | 23.9 | 22 | 23.4 | 23.9 | 21.3 | 24.3 | 24.2 | 29.1 | 25.2 | 23.2 | 22.1791 | 17.7 | 20.6 | 26.1 | 21 | 23.6636 | 20.856 | 24.2 | 26
| 1 | 6 | 23.6 | 30.9 | 24 | 23.1439 | 23 | 24.2 | 28.2 | 22.8 | 24.1 | 24.4 | 21.6 | 23.8 | 24.7 | 27.4 | 25.2 | 24.4 | 22.4786 | 17.6 | 20.7 | 26.5 | 21.5 | 23.6668 | 20.76 | 24.2
| Mean |  | 23.6667 | 27.0167 | 23.6667 | 22.7059 | 22.8333 | 25.2333 | 25.8667 | 22.5333 | 23.75 | 24.15 | 22.2167 | 23.75 | 24.1667 | 28.2 | 25.3 | 23.55 | 22.3245 | 15.25 | 20.6833 | 25.15 | 21.35 | 23.313 | 20.6123 | 23.9167
| Median |  | 23.55 | 26.3 | 24 | 22.7475 | 23 | 25.55 | 25.65 | 22.5 | 23.35 | 24.15 | 21.45 | 23.7 | 24.25 | 28.25 | 25.3 | 23.5 | 22.3226 | 15.5 | 20.7 | 25.1 | 21.4 | 23.4816 | 20.805 | 24.1
| Std Dev. |  | 0.83586 | 2.0999 | 0.5164 | 0.4896 | 0.40825 | 1.1535 | 1.8435 | 0.54283 | 0.98944 | 0.2429 | 1.5728 | 0.36194 | 0.54283 | 0.88091 | 0.20976 | 0.46368 | 0.1176 | 2.274 | 0.1169 | 1.0193 | 0.18708 | 0.62841 | 0.37383 | 0.38166
| Rel.Std.Dev. |  | 3.532% | 7.773% | 2.182% | 2.156% | 1.788% | 4.572% | 7.127% | 2.409% | 4.166% | 1.006% | 7.079% | 1.524% | 2.246% | 3.124% | 0.8291% | 1.969% | 0.5268% | 14.91% | 0.5652% | 4.053% | 0.8763% | 2.696% | 1.814% | 1.596%
| PDM3 |  | 0.7603% | 15.02% | 0.7603% | -3.33% | -2.788% | 7.43% | 10.13% | -4.065% | 1.115% | 2.818% | -5.413% | 1.115% | 2.889% | 20.06% | 7.714% | 0.2636% | -4.954% | -35.07% | -11.94% | 7.076% | -9.103% | -0.7456% | -12.24% | 1.825%
| Z-Score (Absolute) |  | 0 | 1.83 | 0 | 0.52 | 0.45 | 0.85 | 1.2 | 0.62 | 0.05 | 0.26 | 0.79 | 0.05 | 0.27 | 2.47 | 0.89 | 0.06 | 0.73 | 4.59 | 1.63 | 0.81 | 1.26 | 0.19 | 1.66 | 0.14
| Analytical results for Dy in OREAS 265 (Certified Value 3.96 ppm) | [25 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 3.88 | 4.07 | 4 | 3.4 | 3.83 | 3.72 | 3.8723 | 3.92 | 4.31 | 3.9537 | [14 empty] | 1
| 1 | 2 | 3.95 | 4 | 4 | 3.6 | 3.7 | 3.85 | 3.9158 | 3.83 | 4.43 | 3.8651 | [14 empty] | 31
| 1 | 3 | 3.99 | 3.65 | 4 | 3.6 | 3.73 | 4.23 | 3.9003 | 3.94 | 4.32 | 4.0067 | [14 empty] | 16
| 1 | 4 | 3.94 | 4.02 | 4 | 3.6 | 3.91 | 3.91 | 3.952 | 3.87 | 4.55 | 3.9681 | [14 empty] | 3.9631
| 1 | 5 | 4.08 | 4.09 | 4.2 | 3.6 | 3.65 | 3.85 | 3.9429 | 4.07 | 4.43 | 4.0813 | [14 empty] | 27
| 1 | 6 | 4.69 | 3.63 | 4.2 | 3.8 | 3.69 | 3.99 | 3.973 | 4.04 | 4.57 | 4.1297
| Mean |  | 4.0883 | 3.91 | 4.0667 | 3.6 | 3.7517 | 3.925 | 3.9261 | 3.945 | 4.435 | 4.0008
| Median |  | 3.97 | 4.01 | 4 | 3.6 | 3.715 | 3.88 | 3.9293 | 3.93 | 4.43 | 3.9874
| Std Dev. |  | 0.30209 | 0.21175 | 0.10328 | 0.12649 | 0.098472 | 0.17364 | 0.036929 | 0.093968 | 0.10986 | 0.094688
| Rel.Std.Dev. |  | 7.389% | 5.416% | 2.54% | 3.514% | 2.625% | 4.424% | 0.9406% | 2.382% | 2.477% | 2.367%
| PDM3 |  | 3.161% | -1.339% | 2.614% | -9.161% | -5.334% | -0.9606% | -0.9337% | -0.456% | 11.91% | 0.951%
| Z-Score (Absolute) |  | 1.05 | 0.18 | 0.9 | 2.3 | 1.26 | 0.07 | 0.07 | 0.07 | 3.43 | 0.45
| Analytical results for Er in OREAS 265 (Certified Value 1.77 ppm) | [25 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 1.89 | 1.71 | 1.9 | 1.5 | 1.7 | 1.67 | 1.7575 | 1.7 | 2.11 | 1.749 | [14 empty] | 1
| 1 | 2 | 1.89 | 1.7 | 1.7 | 1.5 | 1.71 | 1.64 | 1.81 | 1.75 | 2.09 | 1.7292 | [14 empty] | 32
| 1 | 3 | 1.86 | 1.72 | 1.7 | 1.5 | 1.73 | 1.67 | 1.7903 | 1.78 | 2.06 | 1.7996 | [14 empty] | 16
| 1 | 4 | 1.84 | 1.66 | 1.8 | 1.5 | 1.77 | 1.95 | 1.8057 | 1.74 | 2.25 | 1.7649 | [14 empty] | 1.7679
| 1 | 5 | 1.93 | 1.68 | 1.9 | 1.6 | 1.77 | 1.8 | 1.8125 | 1.83 | 2.17 | 1.7976 | [14 empty] | 28
| 1 | 6 | 1.87 | 1.71 | 1.8 | 1.5 | 1.7 | 1.69 | 1.7726 | 1.81 | 2.19 | 1.8539
| Mean |  | 1.88 | 1.6967 | 1.8 | 1.5167 | 1.73 | 1.7367 | 1.7914 | 1.7683 | 2.145 | 1.7824
| Median |  | 1.88 | 1.705 | 1.8 | 1.5 | 1.72 | 1.68 | 1.798 | 1.765 | 2.14 | 1.7813
| Std Dev. |  | 0.030984 | 0.022509 | 0.089443 | 0.040825 | 0.032863 | 0.11827 | 0.022366 | 0.047924 | 0.070922 | 0.044493
| Rel.Std.Dev. |  | 1.648% | 1.327% | 4.969% | 2.692% | 1.9% | 6.81% | 1.249% | 2.71% | 3.306% | 2.496%
| PDM3 |  | 6.344% | -4.027% | 1.819% | -14.21% | -2.141% | -1.764% | 1.334% | 0.02726% | 21.33% | 0.8217%
| Z-Score (Absolute) |  | 1.68 | 1.26 | 0.4 | 4.15 | 0.73 | 0.62 | 0.26 | 0.11 | 5.93 | 0.11
| Analytical results for Eu in OREAS 265 (Certified Value 1.61 ppm) | [25 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 1.64 | 1.71 | 1.6 | 1.5 | 1.54 | 1.56 | 1.7044 | 1.54 | 1.75 | 1.5985 | [14 empty] | 1
| 1 | 2 | 1.64 | 1.71 | 1.5 | 1.4 | 1.48 | 1.61 | 1.6479 | 1.6 | 1.79 | 1.5567 | [14 empty] | 33
| 1 | 3 | 1.6 | 1.69 | 1.5 | 1.5 | 1.5 | 1.62 | 1.6777 | 1.64 | 1.71 | 1.6097 | [14 empty] | 16
| 1 | 4 | 1.6 | 1.91 | 1.6 | 1.5 | 1.52 | 1.65 | 1.6872 | 1.62 | 1.74 | 1.5641 | [14 empty] | 1.6106
| 1 | 5 | 1.61 | 1.62 | 1.6 | 1.4 | 1.54 | 1.58 | 1.6658 | 1.69 | 1.75 | 1.6434 | [14 empty] | 29
| 1 | 6 | 1.7 | 1.67 | 1.6 | 1.4 | 1.49 | 1.67 | 1.6774 | 1.67 | 1.76 | 1.6115
| Mean |  | 1.6317 | 1.7183 | 1.5667 | 1.45 | 1.5117 | 1.615 | 1.6767 | 1.6267 | 1.75 | 1.5973
| Median |  | 1.625 | 1.7 | 1.6 | 1.45 | 1.51 | 1.615 | 1.6776 | 1.63 | 1.75 | 1.6041
| Std Dev. |  | 0.038166 | 0.099683 | 0.05164 | 0.054772 | 0.025626 | 0.041352 | 0.019116 | 0.053541 | 0.026077 | 0.032369
| Rel.Std.Dev. |  | 2.339% | 5.801% | 3.296% | 3.777% | 1.695% | 2.561% | 1.14% | 3.291% | 1.49% | 2.026%
| PDM3 |  | 1.31% | 6.691% | -2.726% | -9.97% | -6.141% | 0.275% | 4.108% | 0.9994% | 8.657% | -0.8236%
| Z-Score (Absolute) |  | 0.13 | 1.19 | 0.66 | 2.09 | 1.34 | 0.07 | 0.68 | 0.07 | 1.58 | 0.29
| Analytical results for Fe in OREAS 265 (Certified Value 3.44 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 3.52 | 3.49 | 3.41 | 3.4727 | 3.39 | 3.64 | 3.52 | 3.5 | 3.37 | 3.45 | 3.49 | 3.38 | 3.59 | 3.33 | 3.39 | 3.56 | 3.353 | 3.5756 | 3.32 | 3.399 | 3.44 | 3.3007 | 3.2061 | 3.45 | 1
| 1 | 2 | 3.38 | 3.6 | 3.42 | 3.5051 | 3.38 | 3.56 | 3.53 | 3.44 | 3.49 | 3.32 | 3.49 | 3.52 | 3.52 | 3.35 | 3.51 | 3.52 | 3.3628 | 3.5841 | 3.39 | 3.504 | 3.47 | 3.2898 | 3.2982 | 3.4 | 38
| 1 | 3 | 3.3 | 3.57 | 3.49 | 3.4909 | 3.39 | 3.53 | 3.49 | 3.43 | 3.43 | 3.41 | 3.53 | 3.47 | 3.45 | 3.44 | 3.43 | 3.51 | 3.3578 | 3.5832 | 3.35 | 3.413 | 3.43 | 3.367 | 3.3485 | 3.39 | 16
| 1 | 4 | 3.35 | 3.5 | 3.44 | 3.4791 | 3.33 | 3.51 | 3.59 | 3.33 | 3.39 | 3.42 | 3.56 | 3.52 | 3.55 | 3.43 | 3.58 | 3.63 | 3.3567 | 3.4839 | 3.4 | 3.413 | 3.46 | 3.3452 | 3.2971 | 3.54 | 3.4435
| 1 | 5 | 3.29 | 3.58 | 3.46 | 3.4844 | 3.35 | 3.56 | 3.4 | 3.34 | 3.43 | 3.37 | 3.57 | 3.5 | 3.57 | 3.37 | 3.35 | 3.53 | 3.3691 | 3.5545 | 3.36 | 3.406 | 3.39 | 3.4319 | 3.1095 | 3.49 | 30
| 1 | 6 | 3.21 | 3.53 | 3.53 | 3.5121 | 3.37 | 3.53 | 3.39 | 3.44 | 3.41 | 3.42 | 3.4 | 3.44 | 3.52 | 3.4 | 3.37 | 3.64 | 3.3619 | 3.5342 | 3.36 | 3.56 | 3.49 | 3.4266 | 3.2071 | 3.5
| Mean |  | 3.3417 | 3.545 | 3.4583 | 3.4907 | 3.3683 | 3.555 | 3.4867 | 3.4133 | 3.42 | 3.3983 | 3.5067 | 3.4717 | 3.5333 | 3.3867 | 3.4383 | 3.565 | 3.3602 | 3.5526 | 3.3633 | 3.4492 | 3.4467 | 3.3602 | 3.2444 | 3.4617
| Median |  | 3.325 | 3.55 | 3.45 | 3.4877 | 3.375 | 3.545 | 3.505 | 3.435 | 3.42 | 3.415 | 3.51 | 3.485 | 3.535 | 3.385 | 3.41 | 3.545 | 3.3599 | 3.5651 | 3.36 | 3.413 | 3.45 | 3.3561 | 3.2521 | 3.47
| Std Dev. |  | 0.10496 | 0.045056 | 0.04535 | 0.015246 | 0.024014 | 0.045935 | 0.078145 | 0.065625 | 0.041473 | 0.046224 | 0.062183 | 0.054559 | 0.049261 | 0.044121 | 0.089536 | 0.056833 | 0.0056132 | 0.03878 | 0.028752 | 0.066764 | 0.035024 | 0.060512 | 0.086677 | 0.059133
| Rel.Std.Dev. |  | 3.141% | 1.271% | 1.311% | 0.4368% | 0.7129% | 1.292% | 2.241% | 1.923% | 1.213% | 1.36% | 1.773% | 1.572% | 1.394% | 1.303% | 2.604% | 1.594% | 0.167% | 1.092% | 0.8549% | 1.936% | 1.016% | 1.801% | 2.672% | 1.708%
| PDM3 |  | -2.958% | 2.947% | 0.4301% | 1.371% | -2.184% | 3.237% | 1.253% | -0.8767% | -0.6831% | -1.312% | 1.834% | 0.8173% | 2.608% | -1.651% | -0.1507% | 3.528% | -2.419% | 3.167% | -2.329% | 0.1639% | 0.09126% | -2.42% | -5.782% | 0.5269%
| Z-Score (Absolute) |  | 1.2 | 1.09 | 0.12 | 0.48 | 0.9 | 1.2 | 0.43 | 0.39 | 0.32 | 0.56 | 0.66 | 0.27 | 0.96 | 0.69 | 0.11 | 1.31 | 0.99 | 1.17 | 0.95 | 0.02 | 0.02 | 0.99 | 2.29 | 0.15
| Analytical results for Ga in OREAS 265 (Certified Value 19.3 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 19 | 20.7 | 19.5 | 19.4222 | 19.44 | 19.35 | 20 | 18.15 | 18.99 | 18.1 | 18.65 | 18.54 | 17.6 | 20.94 | 19.3 | 19.3 | 19.76 | 20.06 | 20 | 19.5 | 21 | 20.073 | 19.16 |  | 1
| 1 | 2 | 18.3 | 20.8 | 19.3 | 19.9392 | 18.79 | 19.45 | 20 | 17.75 | 19.47 | 17.9 | 18.4 | 18.98 | 16.74 | 19.59 | 19.1 | 19.1 | 19.71 | 20.14 | 20 | 19.9 | 21 | 19.8005 | 19.35 |  | 34
| 1 | 3 | 18 | 20.7 | 19.7 | 19.5066 | 19.05 | 19.23 | 19 | 17.55 | 19.26 | 17.7 | 18.65 | 18.47 | 17.86 | 19.71 | 19.2 | 19.2 | 19.93 | 20.12 | 20.2 | 19.7 | 21 | 19.943 | 19.14 |  | 16
| 1 | 4 | 18.6 | 20.5 | 19.3 | 19.8682 | 18.91 | 19.76 | 20 | 17.4 | 18.59 | 18.25 | 19.25 | 18.81 | 18.77 | 20.5 | 19.5 | 19.2 | 19.64 | 19.63 | 20.1 | 19.8 | 21 | 19.6383 | 19.92 |  | 19.3499
| 1 | 5 | 18.1 | 20.4 | 19.1 | 19.3768 | 19.01 | 19.49 | 19 | 17.5 | 18.8 | 17.65 | 18.65 | 18.95 | 17.51 | 21.23 | 19.2 | 19.8 | 19.57 | 19.84 | 19.55 | 19.7 | 21 | 20.6088 | 20.4 |  | 31
| 1 | 6 | 17.3 | 20.2 | 18.9 | 19.7163 | 19.17 | 19.46 | 19 | 19.1 | 19.07 | 17.9 | 18.4 | 18.74 | 19.4 | 20.38 | 18.6 | 19.8 | 19.85 | 20.18 | 19.9 | 19.8 | 21 | 20.241 | 19.94
| Mean |  | 18.2167 | 20.55 | 19.3 | 19.6382 | 19.0617 | 19.4567 | 19.5 | 17.9083 | 19.03 | 17.9167 | 18.6667 | 18.7483 | 17.98 | 20.3917 | 19.15 | 19.4 | 19.7433 | 19.995 | 19.9583 | 19.7333 | 21 | 20.0507 | 19.6517
| Median |  | 18.2 | 20.6 | 19.3 | 19.6115 | 19.03 | 19.455 | 19.5 | 17.65 | 19.03 | 17.9 | 18.65 | 18.775 | 17.73 | 20.44 | 19.2 | 19.25 | 19.735 | 20.09 | 20 | 19.75 | 21 | 20.008 | 19.635
| Std Dev. |  | 0.57764 | 0.22583 | 0.28284 | 0.23745 | 0.2256 | 0.1766 | 0.54772 | 0.64142 | 0.31502 | 0.22949 | 0.31091 | 0.20932 | 0.95476 | 0.65156 | 0.30166 | 0.31623 | 0.13292 | 0.21557 | 0.22454 | 0.13663 | 0 | 0.34423 | 0.5118
| Rel.Std.Dev. |  | 3.171% | 1.099% | 1.466% | 1.209% | 1.184% | 0.9076% | 2.809% | 3.582% | 1.655% | 1.281% | 1.666% | 1.116% | 5.31% | 3.195% | 1.575% | 1.63% | 0.6732% | 1.078% | 1.125% | 0.6924% | 0% | 1.717% | 2.604%
| PDM3 |  | -5.856% | 6.202% | -0.2578% | 1.49% | -1.489% | 0.5519% | 0.7758% | -7.45% | -1.653% | -7.407% | -3.531% | -3.109% | -7.08% | 5.384% | -1.033% | 0.259% | 2.033% | 3.334% | 3.144% | 1.982% | 8.528% | 3.622% | 1.56%
| Z-Score (Absolute) |  | 1.67 | 1.47 | 0.21 | 0.24 | 0.53 | 0 | 0.06 | 2.08 | 0.57 | 2.07 | 1.06 | 0.95 | 1.98 | 1.26 | 0.41 | 0.08 | 0.39 | 0.72 | 0.67 | 0.37 | 2.07 | 0.8 | 0.26
| Analytical results for Gd in OREAS 265 (Certified Value 5.73 ppm) | [25 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 5.76 | 6 | 5.7 | 5.6 | 5.5 | 5.61 | 6.9858 | 5.58 | 6 | 5.558 | [14 empty] | 1
| 1 | 2 | 6.01 | 6.13 | 5.7 | 5.4 | 5.26 | 5.76 | 6.9128 | 5.68 | 6.1 | 5.4023 | [14 empty] | 35
| 1 | 3 | 5.78 | 5.81 | 5.6 | 5.5 | 5.4 | 5.57 | 6.9473 | 5.87 | 5.9 | 5.438 | [14 empty] | 16
| 1 | 4 | 5.63 | 6.01 | 5.8 | 5.7 | 5.48 | 5.84 | 6.9553 | 5.76 | 6.1 | 5.5372 | [14 empty] | 5.7305
| 1 | 5 | 5.83 | 6.05 | 5.9 | 5.6 | 5.67 | 5.47 | 6.9423 | 6.04 | 6 | 5.7156 | [14 empty] | 32
| 1 | 6 | 5.81 | 5.35 | 5.9 | 5.5 | 5.34 | 6.15 | 6.9743 | 6.04 | 5.9 | 5.7065
| Mean |  | 5.8033 | 5.8917 | 5.7667 | 5.55 | 5.4417 | 5.7333 | 6.9529 | 5.8283 | 6 | 5.5596
| Median |  | 5.795 | 6.005 | 5.75 | 5.55 | 5.44 | 5.685 | 6.9513 | 5.815 | 6 | 5.5476
| Std Dev. |  | 0.12323 | 0.28555 | 0.12111 | 0.10488 | 0.14289 | 0.24369 | 0.025682 | 0.18957 | 0.089443 | 0.13115
| Rel.Std.Dev. |  | 2.124% | 4.847% | 2.1% | 1.89% | 2.626% | 4.25% | 0.3694% | 3.253% | 1.491% | 2.359%
| PDM3 |  | 1.271% | 2.812% | 0.6309% | -3.15% | -5.04% | 0.04924% | 21.33% | 1.707% | 4.703% | -2.982%
| Z-Score (Absolute) |  | 0.08 | 0.45 | 0.08 | 0.99 | 1.44 | 0.22 | 4.9 | 0.18 | 0.9 | 0.94
| Analytical results for Ge in OREAS 265 (Indicative Value 0.2 ppm) | [25 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.29 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | < 0.1 | 0.07 | 1.3 | 0.42725 | 1.5 | 0.28 | 0.05 | 1.3 | 0.07 | 0.08 | 0.12 | 0.14 | 0.3 | 0.68 | 0.39 | 0.07 | 4.1 | < 0.2 | < 0.05 | [5 empty] | 1
| 1 | 2 | < 0.1 | 0.07 | 1.3 | 0.46798 | 1.5 | 0.36 | 0.05 | 1.3 | 0.08 | 0.09 | < 0.05 | 0.14 | 0.4 | 0.37 | 0.5 | 0.07 | 3.9 | < 0.2 | < 0.05 | [5 empty] | 36
| 1 | 3 | < 0.1 | 0.06 | 1.3 | 0.45201 | 1.4 | 0.33 | < 0.05 | 1.3 | 0.07 | 0.07 | 0.06 | 0.14 | 0.3 | 0.48 | 0.15 | 0.08 | 4 | < 0.2 | < 0.05 | [5 empty] | 16
| 1 | 4 | < 0.1 | 0.06 | 1.3 | 0.41601 | 1.6 | 0.34 | 0.05 | 1.3 | 0.07 | 0.08 | 0.08 | 0.15 | 0.3 | 0.74 | 0.35 | 0.07 | 4 | < 0.2 | < 0.05 | [5 empty] | 0.19505
| 1 | 5 | < 0.1 | 0.07 | 1.3 | 0.39736 | 1.5 | 0.31 | 0.05 | 1.2 | 0.07 | 0.09 | 0.06 | 0.14 | 0.3 | 0.34 | 0.14 | 0.08 | 3.8 | < 0.2 | < 0.05 | [5 empty] | 9
| 1 | 6 | < 0.1 | 0.05 | 1.3 | 0.4478 | 1.6 | 0.31 | 0.05 | 1.3 | 0.07 | 0.09 | 0.05 | 0.13 | 0.3 | 0.38 | 0.12 | 0.06 | 4.2 | < 0.2 | < 0.05
| Mean |  |  | 0.063333 | 1.3 | 0.43473 | 1.5167 | 0.32167 | 0.05 | 1.2833 | 0.071667 | 0.083333 | 0.074 | 0.14 | 0.31667 | 0.49833 | 0.275 | 0.071667 | 4
| Median |  |  | 0.065 | 1.3 | 0.43753 | 1.5 | 0.32 | 0.05 | 1.3 | 0.07 | 0.085 | 0.06 | 0.14 | 0.3 | 0.43 | 0.25 | 0.07 | 4
| Std Dev. |  |  | 0.008165 | 0 | 0.025991 | 0.075277 | 0.027869 | 0 | 0.040825 | 0.0040825 | 0.008165 | 0.027928 | 0.0063246 | 0.040825 | 0.17163 | 0.15959 | 0.0075277 | 0.14142
| Rel.Std.Dev. |  |  | 12.89% | 0% | 5.979% | 4.963% | 8.664% | 0% | 3.181% | 5.696% | 9.798% | 37.74% | 4.518% | 12.89% | 34.44% | 58.03% | 10.5% | 3.536%
| PDM3 |  |  | -67.53% | 566.5% | 122.9% | 677.6% | 64.92% | -74.37% | 558% | -63.26% | -57.28% | -62.06% | -28.22% | 62.35% | 155.5% | 40.99% | -63.26% | 1951%
| Z-Score (Absolute) |  | 0.53 | 0.45 | 6.8 | 1.73 | 8.07 | 1.07 | 0.55 | 6.7 | 0.4 | 0.33 | 0.43 | 0 | 1.04 | 2.1 | 0.79 | 0.4 | 22.63 | 0.23 | 0.67
| Analytical results for Hf in OREAS 265 (Certified Value 4.89 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.37 | 5.2 | 4.8 | 5.0173 | 4.96 | 5.12 | 4.4 | 4.7 | 4.9 | 4.6 | 4.9 | 4.8 | 4.81 | 4.6 | 5.38 | 4.7 | 4.52 | 5.2 | 5.3 | 4 | 4.8289 | 4.7 |  |  | 1
| 1 | 2 | 5.2 | 5.3 | 4.8 | 4.7302 | 5.11 | 4.87 | 4.4 | 4.5 | 4.84 | 4.5 | 4.7 | 4.88 | 4.78 | 4.6 | 5.19 | 4.9 | 4.6 | 5.3 | 5.4 | 3.8 | 4.6853 | 5 |  |  | 20
| 1 | 3 | 5.04 | 5.3 | 4.8 | 4.9469 | 5.02 | 4.74 | 4.9 | 4.5 | 5.17 | 4.5 | 4.9 | 4.83 | 4.53 | 4.6 | 5.2 | 4.9 | 4.76 | 5.2 | 5.5 | 3.9 | 4.8353 | 4.9 |  |  | 16
| 1 | 4 | 5.23 | 5.2 | 4.9 | 4.662 | 5.02 | 4.67 | 4.6 | 4.5 | 4.94 | 4.7 | 4.9 | 4.88 | 4.79 | 4.6 | 5.45 | 4.9 | 4.6 | 5 | 5.4 | 3.9 | 4.8122 | 5.1 |  |  | 4.8943
| 1 | 5 | 4.97 | 5.4 | 4.8 | 4.6576 | 5.39 | 4.95 | 4.3 | 4.5 | 5.13 | 4.5 | 4.9 | 5.14 | 4.73 | 4.7 | 5.19 | 5 | 4.9 | 5 | 5.3 | 3.8 | 4.9308 | 5.1 |  |  | 33
| 1 | 6 | 4.74 | 5 | 4.9 | 4.7858 | 5.16 | 4.87 | 4.1 | 4.6 | 5.05 | 4.6 | 4.7 | 4.98 | 4.7 | 4.4 | 5.36 | 5 | 4.8 | 5 | 5.3 | 3.8 | 5.0243 | 5.1
| Mean |  | 5.0917 | 5.2333 | 4.8333 | 4.8 | 5.11 | 4.87 | 4.45 | 4.55 | 5.005 | 4.5667 | 4.8333 | 4.9183 | 4.7233 | 4.5833 | 5.295 | 4.9 | 4.6967 | 5.1167 | 5.3667 | 3.8667 | 4.8528 | 4.9833
| Median |  | 5.12 | 5.25 | 4.8 | 4.758 | 5.065 | 4.87 | 4.4 | 4.5 | 4.995 | 4.55 | 4.9 | 4.88 | 4.755 | 4.6 | 5.28 | 4.9 | 4.68 | 5.1 | 5.35 | 3.85 | 4.8321 | 5.05
| Std Dev. |  | 0.22338 | 0.13663 | 0.05164 | 0.15045 | 0.15466 | 0.15862 | 0.27386 | 0.083666 | 0.13217 | 0.08165 | 0.10328 | 0.12465 | 0.10309 | 0.098319 | 0.11537 | 0.10954 | 0.14556 | 0.13292 | 0.08165 | 0.08165 | 0.11498 | 0.16021
| Rel.Std.Dev. |  | 4.387% | 2.611% | 1.068% | 3.134% | 3.027% | 3.257% | 6.154% | 1.839% | 2.641% | 1.788% | 2.137% | 2.534% | 2.182% | 2.145% | 2.179% | 2.236% | 3.099% | 2.598% | 1.521% | 2.112% | 2.369% | 3.215%
| PDM3 |  | 4.033% | 6.928% | -1.245% | -1.927% | 4.408% | -0.4957% | -9.077% | -7.034% | 2.263% | -6.693% | -1.245% | 0.4919% | -3.492% | -6.353% | 8.188% | 0.1173% | -4.037% | 4.544% | 9.652% | -21% | -0.8469% | 1.82%
| Z-Score (Absolute) |  | 0.79 | 1.27 | 0.1 | 0.21 | 0.85 | 0.03 | 1.4 | 1.06 | 0.49 | 1.01 | 0.1 | 0.19 | 0.47 | 0.95 | 1.48 | 0.13 | 0.56 | 0.87 | 1.73 | 3.4 | 0.03 | 0.42
| Analytical results for Hg in OREAS 265 (Indicative Value 0.017 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [22 empty] | 1
| No | No | 1.18 | 1.22 | [22 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [22 empty] | 3
3
| 1 | 1 | 0.02 | < 1 | [22 empty] | 1
| 1 | 2 | 0.02 | < 1 | [22 empty] | 4
| 1 | 3 | 0.02 | < 1 | [22 empty] | 16
| 1 | 4 | 0.02 | < 1 | [22 empty] | 0.016667
| 1 | 5 | 0.01 | < 1 | [22 empty] | 10
| 1 | 6 | 0.01 | < 1
| Mean |  | 0.016667
| Median |  | 0.02
| Std Dev. |  | 0.005164
| Rel.Std.Dev. |  | 30.98%
| PDM3 |  | -2.109e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 265 (Certified Value 0.69 ppm) | [25 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.20 | 1.22 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.71 | 0.64 | 0.7 | 0.5 | 0.67 | 1.02 | 0.64 | 0.74 | 0.66892 | [15 empty] | 1
| 1 | 2 | 0.72 | 0.64 | 0.7 | 0.5 | 0.63 | 0.65 | 0.66 | 0.77 | 0.65025 | [15 empty] | 21
| 1 | 3 | 0.7 | 0.59 | 0.7 | 0.5 | 0.67 | 0.7 | 0.68 | 0.82 | 0.68332 | [15 empty] | 16
| 1 | 4 | 0.72 | 0.64 | 0.7 | 0.6 | 0.68 | 0.67 | 0.67 | 0.82 | 0.65308 | [15 empty] | 0.69122
| 1 | 5 | 0.75 | 0.69 | 0.7 | 0.5 | 0.71 | 0.71 | 0.7 | 0.77 | 0.68292 | [15 empty] | 34
| 1 | 6 | 0.74 | 0.55 | 0.7 | 0.6 | 0.65 | 0.72 | 0.69 | 0.81 | 0.68271
| Mean |  | 0.72333 | 0.625 | 0.7 | 0.53333 | 0.66833 | 0.745 | 0.67333 | 0.78833 | 0.6702
| Median |  | 0.72 | 0.64 | 0.7 | 0.5 | 0.67 | 0.705 | 0.675 | 0.79 | 0.67582
| Std Dev. |  | 0.018619 | 0.048477 | 1.2162e-16 | 0.05164 | 0.027142 | 0.13722 | 0.021602 | 0.033116 | 0.015382
| Rel.Std.Dev. |  | 2.574% | 7.756% | 1.737e-14% | 9.682% | 4.061% | 18.42% | 3.208% | 4.201% | 2.295%
| PDM3 |  | 4.646% | -9.58% | 1.27% | -22.84% | -3.311% | 7.781% | -2.588% | 14.05% | -3.041%
| Z-Score (Absolute) |  | 0.71 | 0.64 | NA | 1.9 | 0.05 | 1 | 0.02 | 1.6 | 0.02
Results from laboratory 1.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 265 (Certified Value 0.07 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.29 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.07 | 0.074 | 0.07 | 0.071748 | 0.07 | 0.067 | 0.1 | 0.077 | 0.06 | 0.073 | 0.067 | 0.07 | 0.09 | 0.09 | 0.07 | 0.1 | 0.099 | 0.074 | 0.1 | < 0.5 | 0.064 |  |  |  | 1
| 1 | 2 | 0.07 | 0.078 | 0.07 | 0.057121 | 0.07 | 0.07 | 0.07 | 0.068 | 0.07 | 0.076 | 0.066 | 0.08 | 0.09 | 0.092 | 0.07 | 0.09 | 0.106 | 0.071 | 0.09 | < 0.5 | 0.062 |  |  |  | 22
| 1 | 3 | 0.07 | 0.073 | 0.07 | 0.065621 | 0.07 | 0.069 | 0.09 | 0.069 | 0.07 | 0.071 | 0.067 | 0.07 | 0.1 | 0.108 | 0.07 | 0.07 | 0.109 | 0.074 | 0.11 | < 0.5 | 0.064 |  |  |  | 16
| 1 | 4 | 0.07 | 0.076 | 0.07 | 0.076121 | 0.07 | 0.063 | 0.08 | 0.068 | 0.07 | 0.075 | 0.066 | 0.08 | 0.08 | 0.103 | 0.07 | 0.08 | 0.1 | 0.072 | 0.09 | < 0.5 | 0.069 |  |  |  | 0.069946
| 1 | 5 | 0.07 | 0.074 | 0.07 | 0.058702 | 0.07 | 0.059 | 0.08 | 0.066 | 0.07 | 0.074 | 0.069 | 0.08 | 0.09 | 0.086 | 0.07 | 0.12 | 0.108 | 0.071 | 0.1 | < 0.5 | 0.065 |  |  |  | 35
| 1 | 6 | 0.07 | 0.071 | 0.07 | 0.082949 | 0.07 | 0.073 | 0.09 | 0.07 | 0.07 | 0.072 | 0.066 | 0.07 | 0.08 | 0.088 | 0.07 | 0.11 | 0.107 | 0.073 | 0.1 | < 0.5 | 0.066
| Mean |  | 0.07 | 0.074333 | 0.07 | 0.06871 | 0.07 | 0.066833 | 0.085 | 0.069667 | 0.068333 | 0.0735 | 0.066833 | 0.075 | 0.088333 | 0.0945 | 0.07 | 0.095 | 0.10483 | 0.0725 | 0.098333 |  | 0.065
| Median |  | 0.07 | 0.074 | 0.07 | 0.068684 | 0.07 | 0.068 | 0.085 | 0.0685 | 0.07 | 0.0735 | 0.0665 | 0.075 | 0.09 | 0.091 | 0.07 | 0.095 | 0.1065 | 0.0725 | 0.1 |  | 0.0645
| Std Dev. |  | 0 | 0.0024221 | 0 | 0.010108 | 0 | 0.0050761 | 0.010488 | 0.0038297 | 0.0040825 | 0.0018708 | 0.001169 | 0.0054772 | 0.0075277 | 0.0088938 | 0 | 0.018708 | 0.0042622 | 0.0013784 | 0.0075277 |  | 0.0023664
| Rel.Std.Dev. |  | 0% | 3.258% | 0% | 14.71% | 0% | 7.595% | 12.34% | 5.497% | 5.974% | 2.545% | 1.749% | 7.303% | 8.522% | 9.411% | 0% | 19.69% | 4.066% | 1.901% | 7.655% |  | 3.641%
| PDM3 |  | 0.07722% | 6.272% | 0.07722% | -1.766% | 0.07722% | -4.45% | 21.52% | -0.3993% | -2.306% | 5.081% | -4.45% | 7.226% | 26.29% | 35.1% | 0.07722% | 35.82% | 49.88% | 3.651% | 40.58% |  | -7.071%
| Z-Score (Absolute) |  | 0.4 | 0.3 | 0.4 | 0.61 | 0.4 | 0.92 | 2.02 | 0.46 | 0.67 | 0.16 | 0.92 | 0.4 | 2.56 | 3.56 | 0.4 | 3.64 | 5.23 | 0 | 4.18 | 28.73 | 1.21
| Analytical results for K in OREAS 265 (Certified Value 1.95 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.12 | 2.03 | 1.88 | 2.0041 | 1.9554 | 1.98 | 2 | 1.97 | 1.9003 | 1.92 | 1.96 | 1.94 | 1.75 | 1.96 | 1.93 | 2.05 | 1.9357 | 2.0866 | 1.93 | 1.951 | 1.9 | 1.8687 | 1.7384 | 2.02 | 1
| 1 | 2 | 2.04 | 2.05 | 1.87 | 1.9752 | 1.9327 | 1.93 | 1.95 | 1.98 | 1.9194 | 1.89 | 1.92 | 1.96 | 1.86 | 1.92 | 1.95 | 2.03 | 1.9306 | 2.1021 | 1.95 | 1.984 | 1.91 | 1.8555 | 1.8137 | 2.01 | #N/A
| 1 | 3 | 2 | 2.03 | 1.89 | 1.9361 | 1.9449 | 1.91 | 1.97 | 1.97 | 1.9495 | 1.91 | 1.98 | 1.96 | 1.85 | 1.98 | 1.91 | 2.02 | 1.9347 | 2.1076 | 1.95 | 1.967 | 1.87 | 1.8919 | 1.8536 | 1.99 | 16
| 1 | 4 | 2.05 | 2 | 1.87 | 1.955 | 1.9043 | 1.9 | 2 | 1.9 | 1.9012 | 1.93 | 2.02 | 1.96 | 1.72 | 1.95 | 2.02 | 2.08 | 1.9299 | 2.0518 | 1.98 | 1.934 | 1.9 | 1.892 | 1.8124 | 2.09 | 1.9491
| 1 | 5 | 1.98 | 2.06 | 1.86 | 1.9576 | 1.9205 | 1.97 | 1.91 | 1.94 | 1.9447 | 1.9 | 1.98 | 1.97 | 1.8 | 1.93 | 1.93 | 2.02 | 1.9303 | 2.0944 | 1.94 | 1.926 | 1.88 | 1.9341 | 1.8717 | 2.05 | 36
| 1 | 6 | 1.91 | 2.03 | 1.87 | 1.9945 | 1.9285 | 1.96 | 1.9 | 1.94 | 1.93 | 1.94 | 1.97 | 1.99 | 1.57 | 1.92 | 1.91 | 2.08 | 1.9248 | 2.0977 | 1.97 | 2.026 | 1.89 | 1.9199 | 1.7001 | 2.06
| Mean |  | 2.0167 | 2.0333 | 1.8733 | 1.9704 | 1.931 | 1.9417 | 1.955 | 1.95 | 1.9242 | 1.915 | 1.9717 | 1.9633 | 1.7583 | 1.9433 | 1.9417 | 2.0467 | 1.931 | 2.09 | 1.9533 | 1.9647 | 1.8917 | 1.8937 | 1.7983 | 2.0367
| Median |  | 2.02 | 2.03 | 1.87 | 1.9664 | 1.9306 | 1.945 | 1.96 | 1.955 | 1.9247 | 1.915 | 1.975 | 1.96 | 1.775 | 1.94 | 1.93 | 2.04 | 1.9304 | 2.096 | 1.95 | 1.959 | 1.895 | 1.892 | 1.8131 | 2.035
| Std Dev. |  | 0.071181 | 0.020656 | 0.010328 | 0.025755 | 0.018005 | 0.033116 | 0.043243 | 0.029665 | 0.021058 | 0.018708 | 0.032506 | 0.01633 | 0.10722 | 0.024221 | 0.041191 | 0.028048 | 0.0038886 | 0.020028 | 0.018619 | 0.036779 | 0.01472 | 0.029682 | 0.066502 | 0.036697
| Rel.Std.Dev. |  | 3.53% | 1.016% | 0.5513% | 1.307% | 0.9324% | 1.706% | 2.212% | 1.521% | 1.094% | 0.9769% | 1.649% | 0.8317% | 6.098% | 1.246% | 2.121% | 1.37% | 0.2014% | 0.9583% | 0.9532% | 1.872% | 0.7781% | 1.567% | 3.698% | 1.802%
| PDM3 |  | 3.469% | 4.324% | -3.885% | 1.096% | -0.9237% | -0.3789% | 0.3051% | 0.04861% | -1.276% | -1.747% | 1.16% | 0.7327% | -9.785% | -0.2934% | -0.3789% | 5.008% | -0.9264% | 7.233% | 0.2196% | 0.8011% | -2.944% | -2.84% | -7.734% | 4.495%
| Z-Score (Absolute) |  | 1.94 | 2.4 | 2.03 | 0.66 | 0.43 | 0.14 | 0.23 | 0.09 | 0.62 | 0.88 | 0.69 | 0.46 | 5.21 | 0.09 | 0.14 | 2.77 | 0.43 | 3.97 | 0.18 | 0.5 | 1.52 | 1.46 | 4.1 | 2.49
| Analytical results for La in OREAS 265 (Certified Value 49.8 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 54.1 | 50.8 | 52.03 | 45.98 | 52 | 52.8 | 49.8 | 51.9 | 51.5 | 51.2 | 58.4 | 47 | 47.2 | 47.3 | 50.636 | 47.9 | 50.6 | 50 | 53 | 48.8234 | 44.89 | 44.1 |  |  | 1
| 1 | 2 | 50.2 | 54.7 | 51.67 | 47.21 | 52 | 51.6 | 50.37 | 50.2 | 50.6 | 50.8 | 57.4 | 44.7 | 46.5 | 46.8 | 50.812 | 48.5 | 51.4 | 52 | 51 | 47.5214 | 45.65 | 45 |  |  | 23
| 1 | 3 | 48.2 | 52.3 | 51.02 | 46.22 | 51 | 50.8 | 49.48 | 51.2 | 52.4 | 49.7 | 56.5 | 46.4 | 46.6 | 47.6 | 50.687 | 49.2 | 51.3 | 51 | 53 | 48.7437 | 45.35 | 44.8 |  |  | 16
| 1 | 4 | 47.9 | 51.9 | 51.28 | 45.11 | 54 | 50.2 | 48.26 | 51.4 | 53.3 | 51.5 | 60.4 | 47 | 47.2 | 47.9 | 50.934 | 47.5 | 51.2 | 50 | 54 | 48.5975 | 45.254 | 46.4 |  |  | 49.801
| 1 | 5 | 47.7 | 51.8 | 51.58 | 46.41 | 48 | 50.9 | 48.77 | 52.2 | 52.2 | 53 | 58.2 | 46 | 46.1 | 47.4 | 50.208 | 48.9 | 50.1 | 51 | 49 | 49.7466 | 44.97 | 47.2 |  |  | 37
| 1 | 6 | 46.4 | 50.9 | 51.54 | 45.55 | 51 | 52.8 | 50.25 | 51.1 | 50 | 52.3 | 59.7 | 44.6 | 45.1 | 47.9 | 50.56 | 49.2 | 51.2 | 53 | 53 | 49.8926 | 45.6 | 46.5
| Mean |  | 49.0833 | 52.0667 | 51.52 | 46.08 | 51.3333 | 51.5167 | 49.4883 | 51.3333 | 51.6667 | 51.4167 | 58.4333 | 45.95 | 46.45 | 47.4833 | 50.6395 | 48.5333 | 50.9667 | 51.1667 | 52.1667 | 48.8875 | 45.2857 | 45.6667
| Median |  | 48.05 | 51.85 | 51.56 | 46.1 | 51.5 | 51.25 | 49.64 | 51.3 | 51.85 | 51.35 | 58.3 | 46.2 | 46.55 | 47.5 | 50.6615 | 48.7 | 51.2 | 51 | 53 | 48.7836 | 45.302 | 45.7
| Std Dev. |  | 2.7462 | 1.418 | 0.34473 | 0.72625 | 1.9664 | 1.0889 | 0.8342 | 0.69761 | 1.2193 | 1.1548 | 1.4376 | 1.0766 | 0.78677 | 0.41673 | 0.24951 | 0.70616 | 0.50859 | 1.169 | 1.8348 | 0.86295 | 0.31397 | 1.2028
| Rel.Std.Dev. |  | 5.595% | 2.723% | 0.6691% | 1.576% | 3.831% | 2.114% | 1.686% | 1.359% | 2.36% | 2.246% | 2.46% | 2.343% | 1.694% | 0.8776% | 0.4927% | 1.455% | 0.9979% | 2.285% | 3.517% | 1.765% | 0.6933% | 2.634%
| PDM3 |  | -1.441% | 4.549% | 3.452% | -7.472% | 3.077% | 3.445% | -0.6278% | 3.077% | 3.746% | 3.244% | 17.33% | -7.733% | -6.729% | -4.654% | 1.684% | -2.545% | 2.341% | 2.742% | 4.75% | -1.834% | -9.067% | -8.302%
| Z-Score (Absolute) |  | 0.87 | 0.64 | 0.36 | 2.38 | 0.27 | 0.36 | 0.66 | 0.27 | 0.43 | 0.31 | 3.84 | 2.44 | 2.19 | 1.67 | 0.08 | 1.14 | 0.08 | 0.18 | 0.69 | 0.96 | 2.78 | 2.59
| Analytical results for Li in OREAS 265 (Certified Value 35.6 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 50 | 37.2 | 36.3 | 34.9119 | 34.3 | 36 | 37 | 36.5 | 33.3 | 36.8 | 36.3 | 34.9 | 33.3 | 35.8 | 34 | 38.8 | 35.96 | 35 | 35 | 37 | 36.1 | 35.4779 | 39.6 |  | 1
| 1 | 2 | 50 | 38.5 | 35.9 | 34.7823 | 34 | 35 | 36 | 36.7 | 34.1 | 35.5 | 36.4 | 37.5 | 32.1 | 34.4 | 35 | 39.8 | 35.9246 | 35 | 35.8 | 40 | 36.1 | 34.6201 | 39.8 |  | #N/A
| 1 | 3 | 49 | 38 | 35.7 | 33.5151 | 35 | 34 | 36 | 36.4 | 34.2 | 36.4 | 36.7 | 36.6 | 34.4 | 36.6 | 34 | 39.4 | 35.9666 | 36 | 35.6 | 38 | 35.4 | 35.2426 | 39.3 |  | 16
| 1 | 4 | 50 | 37.1 | 35.5 | 33.6405 | 33.3 | 34 | 37 | 35 | 33.9 | 36.6 | 36.9 | 36.7 | 36.8 | 35.9 | 35 | 38.6 | 35.8806 | 34 | 36 | 38 | 37.9 | 35.873 | 41.9 |  | 35.6208
| 1 | 5 | 48 | 38 | 36.1 | 34.3454 | 33.8 | 32 | 35 | 35.4 | 34.5 | 35.9 | 37.3 | 37.5 | 33.7 | 34.5 | 34 | 39 | 35.9376 | 35 | 35.4 | 38 | 35 | 36.5722 | 42.9 |  | 38
| 1 | 6 | 45 | 37.6 | 35.9 | 34.5046 | 34 | 32 | 35 | 35.7 | 34.1 | 36.8 | 35.4 | 37.2 | 39.3 | 34.3 | 34 | 39.4 | 35.9646 | 36 | 35.7 | 39 | 34.7 | 36.1764 | 42.1
| Mean |  | 48.6667 | 37.7333 | 35.9 | 34.2833 | 34.0667 | 33.8333 | 36 | 35.95 | 34.0167 | 36.3333 | 36.5 | 36.7333 | 34.9333 | 35.25 | 34.3333 | 39.1667 | 35.939 | 35.1667 | 35.5833 | 38.3333 | 35.8667 | 35.6604 | 40.9333
| Median |  | 49.5 | 37.8 | 35.9 | 34.425 | 34 | 34 | 36 | 36.05 | 34.1 | 36.5 | 36.55 | 36.95 | 34.05 | 35.15 | 34 | 39.2 | 35.9488 | 35 | 35.65 | 38 | 35.75 | 35.6754 | 40.85
| Std Dev. |  | 1.9664 | 0.53541 | 0.28284 | 0.58314 | 0.56451 | 1.6021 | 0.89443 | 0.68337 | 0.40208 | 0.52789 | 0.64807 | 0.97707 | 2.6478 | 0.97314 | 0.5164 | 0.44572 | 0.0331 | 0.75277 | 0.34881 | 1.0328 | 1.1466 | 0.69757 | 1.5423
| Rel.Std.Dev. |  | 4.041% | 1.419% | 0.7879% | 1.701% | 1.657% | 4.735% | 2.485% | 1.901% | 1.182% | 1.453% | 1.776% | 2.66% | 7.579% | 2.761% | 1.504% | 1.138% | 0.0921% | 2.141% | 0.9803% | 2.694% | 3.197% | 1.956% | 3.768%
| PDM3 |  | 36.62% | 5.931% | 0.7838% | -3.755% | -4.363% | -5.018% | 1.065% | 0.9242% | -4.503% | 2% | 2.468% | 3.123% | -1.93% | -1.041% | -3.614% | 9.954% | 0.8933% | -1.275% | -0.1052% | 7.615% | 0.6902% | 0.1111% | 14.91%
| Z-Score (Absolute) |  | 10.33 | 1.48 | 0 | 1.31 | 1.48 | 1.67 | 0.08 | 0.04 | 1.52 | 0.35 | 0.49 | 0.67 | 0.78 | 0.53 | 1.27 | 2.64 | 0.03 | 0.59 | 0.26 | 1.97 | 0.03 | 0.19 | 4.07
| Analytical results for Lu in OREAS 265 (Certified Value 0.24 ppm) | [25 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.25 | 0.24 | 0.23 | 0.23 | 0.2 | 0.2 | 0.23 | 0.24 | 0.26 | 0.2466 | 0.24 | 0.25217 | [12 empty] | 1
| 1 | 2 | 0.24 | 0.26 | 0.25 | 0.27 | 0.2 | 0.2 | 0.23 | 0.23 | 0.26 | 0.24761 | 0.23 | 0.24355 | [12 empty] | 24
| 1 | 3 | 0.23 | 0.24 | 0.23 | 0.25 | 0.2 | 0.2 | 0.23 | 0.24 | 0.22 | 0.23545 | 0.25 | 0.25281 | [12 empty] | 16
| 1 | 4 | 0.22 | 0.23 | 0.23 | 0.24 | 0.2 | 0.2 | 0.22 | 0.24 | 0.23 | 0.24995 | 0.25 | 0.24271 | [12 empty] | 0.24036
| 1 | 5 | 0.22 | 0.28 | 0.23 | 0.23 | 0.2 | 0.2 | 0.23 | 0.24 | 0.23 | 0.24231 | 0.25 | 0.24878 | [12 empty] | 39
| 1 | 6 | 0.21 | 0.26 | 0.24 | 0.24 | 0.2 | 0.2 | 0.21 | 0.23 | 0.26 | 0.24631 | 0.25 | 0.26328
| Mean |  | 0.22833 | 0.25167 | 0.235 | 0.24333 | 0.2 | 0.2 | 0.225 | 0.23667 | 0.24333 | 0.2447 | 0.245 | 0.25055
| Median |  | 0.225 | 0.25 | 0.23 | 0.24 | 0.2 | 0.2 | 0.23 | 0.24 | 0.245 | 0.24645 | 0.25 | 0.25047
| Std Dev. |  | 0.01472 | 0.018348 | 0.0083666 | 0.015055 | 3.0405e-17 | 3.0405e-17 | 0.0083666 | 0.005164 | 0.018619 | 0.0051649 | 0.0083666 | 0.0075269
| Rel.Std.Dev. |  | 6.447% | 7.291% | 3.56% | 6.187% | 1.52e-14% | 1.52e-14% | 3.718% | 2.182% | 7.652% | 2.111% | 3.415% | 3.004%
| PDM3 |  | -5.003% | 4.705% | -2.229% | 1.238% | -16.79% | -16.79% | -6.39% | -1.536% | 1.238% | 1.808% | 1.931% | 4.24%
| Z-Score (Absolute) |  | 1.46 | 0.81 | 0.81 | 0 | NA | NA | 1.78 | 0.65 | 0 | 0.13 | 0.16 | 0.7
Results from laboratories 1.12 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 265 (Certified Value 0.841 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.89 | 0.83 | 0.831 | 0.86488 | 0.8362 | 0.831 | 0.88 | 0.87 | 0.8205 | 0.86 | 0.81 | 0.83 | 0.86 | 0.85 | 0.82 | 0.91 | 0.85452 | 0.8843 | 0.81 | 0.862 | 0.88 | 0.81638 | 0.8035 | 0.85 | 1
| 1 | 2 | 0.85 | 0.86 | 0.829 | 0.85656 | 0.832 | 0.856 | 0.81 | 0.86 | 0.8414 | 0.82 | 0.81 | 0.84 | 0.85 | 0.81 | 0.84 | 0.9 | 0.85241 | 0.8882 | 0.82 | 0.893 | 0.88 | 0.81638 | 0.821 | 0.84 | #N/A
| 1 | 3 | 0.83 | 0.86 | 0.831 | 0.85589 | 0.832 | 0.843 | 0.83 | 0.86 | 0.8407 | 0.85 | 0.81 | 0.84 | 0.85 | 0.84 | 0.82 | 0.9 | 0.85384 | 0.9012 | 0.81 | 0.862 | 0.88 | 0.83363 | 0.807 | 0.84 | 16
| 1 | 4 | 0.85 | 0.84 | 0.82 | 0.85661 | 0.8307 | 0.839 | 0.83 | 0.83 | 0.8107 | 0.85 | 0.82 | 0.84 | 0.88 | 0.85 | 0.86 | 0.93 | 0.85177 | 0.8683 | 0.82 | 0.868 | 0.89 | 0.82905 | 0.8308 | 0.88 | 0.84101
| 1 | 5 | 0.81 | 0.86 | 0.848 | 0.85659 | 0.8348 | 0.835 | 0.82 | 0.84 | 0.8425 | 0.84 | 0.82 | 0.84 | 0.86 | 0.83 | 0.81 | 0.9 | 0.85705 | 0.8968 | 0.81 | 0.868 | 0.87 | 0.85285 | 0.7923 | 0.86 | 40
| 1 | 6 | 0.8 | 0.84 | 0.844 | 0.86506 | 0.8349 | 0.831 | 0.84 | 0.85 | 0.833 | 0.86 | 0.78 | 0.83 | 0.88 | 0.81 | 0.82 | 0.93 | 0.85318 | 0.8967 | 0.82 | 0.905 | 0.85 | 0.85803 | 0.7985 | 0.87
| Mean |  | 0.83833 | 0.84833 | 0.83383 | 0.85927 | 0.83343 | 0.83917 | 0.835 | 0.85167 | 0.83147 | 0.84667 | 0.80833 | 0.83667 | 0.86333 | 0.83167 | 0.82833 | 0.91167 | 0.8538 | 0.88925 | 0.815 | 0.87633 | 0.875 | 0.83439 | 0.80885 | 0.85667
| Median |  | 0.84 | 0.85 | 0.831 | 0.8566 | 0.8334 | 0.837 | 0.83 | 0.855 | 0.83685 | 0.85 | 0.81 | 0.84 | 0.86 | 0.835 | 0.82 | 0.905 | 0.85351 | 0.89245 | 0.815 | 0.868 | 0.88 | 0.83134 | 0.80525 | 0.855
| Std Dev. |  | 0.032506 | 0.013292 | 0.010342 | 0.0044279 | 0.0021565 | 0.0094745 | 0.02429 | 0.01472 | 0.013111 | 0.015055 | 0.01472 | 0.005164 | 0.013663 | 0.018348 | 0.018348 | 0.01472 | 0.0018717 | 0.011998 | 0.0054772 | 0.018162 | 0.013784 | 0.017761 | 0.014439 | 0.01633
| Rel.Std.Dev. |  | 3.878% | 1.567% | 1.24% | 0.5153% | 0.2588% | 1.129% | 2.909% | 1.728% | 1.577% | 1.778% | 1.821% | 0.6172% | 1.583% | 2.206% | 2.215% | 1.615% | 0.2192% | 1.349% | 0.6721% | 2.073% | 1.575% | 2.129% | 1.785% | 1.906%
| PDM3 |  | -0.3188% | 0.8703% | -0.8539% | 2.17% | -0.9014% | -0.2197% | -0.7151% | 1.267% | -1.135% | 0.6721% | -3.886% | -0.517% | 2.654% | -1.111% | -1.508% | 8.401% | 1.52% | 5.735% | -3.093% | 4.2% | 4.041% | -0.7881% | -3.824% | 1.861%
| Z-Score (Absolute) |  | 0.02 | 0.55 | 0.28 | 1.19 | 0.31 | 0.02 | 0.22 | 0.75 | 0.42 | 0.46 | 1.76 | 0.12 | 1.42 | 0.41 | 0.6 | 4.21 | 0.87 | 2.92 | 1.37 | 2.18 | 2.1 | 0.25 | 1.73 | 1.04
| Analytical results for Mn in OREAS 265 (Certified Value 0.031 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.03 | 0.0317 | 0.0288 | 0.031382 | 0.0304 | 0.0315 | 0.0305 | 0.0314 | 0.0293 | 0.0309 | 0.0308 | 0.0315 | 0.0317 | 0.0303 | 0.0311 | 0.0324 | 0.029976 | 0.0329 | 0.0325 | 0.031 | 0.0307 | 0.030133 | 0.029 | 0.03 | 1
| 1 | 2 | 0.03 | 0.0322 | 0.0292 | 0.031461 | 0.0301 | 0.0308 | 0.0301 | 0.0309 | 0.0307 | 0.0302 | 0.0304 | 0.0322 | 0.0309 | 0.0289 | 0.0308 | 0.0319 | 0.029986 | 0.0331 | 0.0312 | 0.031 | 0.0314 | 0.030149 | 0.0301 | 0.0296 | 25
| 1 | 3 | 0.03 | 0.0326 | 0.0291 | 0.030899 | 0.0301 | 0.0304 | 0.0289 | 0.0305 | 0.0304 | 0.0304 | 0.0308 | 0.032 | 0.0305 | 0.0302 | 0.0308 | 0.032 | 0.02999 | 0.0334 | 0.0302 | 0.031 | 0.0306 | 0.030538 | 0.0304 | 0.0295 | 16
| 1 | 4 | 0.03 | 0.0315 | 0.0294 | 0.031095 | 0.0297 | 0.0303 | 0.0312 | 0.0296 | 0.0295 | 0.0309 | 0.0314 | 0.0324 | 0.033 | 0.03 | 0.0315 | 0.0331 | 0.030013 | 0.0325 | 0.0308 | 0.031 | 0.0309 | 0.030062 | 0.0316 | 0.0309 | 0.030778
| 1 | 5 | 0.03 | 0.0323 | 0.0287 | 0.03161 | 0.0302 | 0.031 | 0.0303 | 0.0301 | 0.0303 | 0.0306 | 0.0311 | 0.0325 | 0.0313 | 0.0309 | 0.0311 | 0.0321 | 0.030021 | 0.0332 | 0.0309 | 0.031 | 0.03 | 0.031695 | 0.03 | 0.0304 | 41
| 1 | 6 | 0.03 | 0.0316 | 0.0291 | 0.031035 | 0.0301 | 0.031 | 0.0286 | 0.0306 | 0.03 | 0.0311 | 0.0301 | 0.0315 | 0.0326 | 0.0296 | 0.0304 | 0.0333 | 0.029988 | 0.0334 | 0.0307 | 0.031 | 0.031 | 0.031492 | 0.0291 | 0.0305
| Mean |  | 0.03 | 0.031983 | 0.02905 | 0.031247 | 0.0301 | 0.030833 | 0.029933 | 0.030517 | 0.030033 | 0.030683 | 0.030767 | 0.032017 | 0.031667 | 0.029983 | 0.03095 | 0.032467 | 0.029996 | 0.033083 | 0.03105 | 0.031 | 0.030767 | 0.030678 | 0.030033 | 0.03015
| Median |  | 0.03 | 0.03195 | 0.0291 | 0.031239 | 0.0301 | 0.0309 | 0.0302 | 0.03055 | 0.03015 | 0.03075 | 0.0308 | 0.0321 | 0.0315 | 0.0301 | 0.03095 | 0.03225 | 0.029989 | 0.03315 | 0.03085 | 0.031 | 0.0308 | 0.030343 | 0.03005 | 0.0302
| Std Dev. |  | 0 | 0.0004446 | 0.00025884 | 0.00027735 | 0.00022804 | 0.00044121 | 0.00099331 | 0.00062423 | 0.00054283 | 0.00034303 | 0.00046762 | 0.00043551 | 0.00097297 | 0.00067946 | 0.00037283 | 0.00059554 | 1.7375e-05 | 0.00034303 | 0.00078166 | 0 | 0.00046762 | 0.00073134 | 0.00095219 | 0.00054681
| Rel.Std.Dev. |  | 0% | 1.39% | 0.891% | 0.8876% | 0.7576% | 1.431% | 3.318% | 2.046% | 1.807% | 1.118% | 1.52% | 1.36% | 3.073% | 2.266% | 1.205% | 1.834% | 0.05792% | 1.037% | 2.517% | 0% | 1.52% | 2.384% | 3.17% | 1.814%
| PDM3 |  | -2.528% | 3.916% | -5.615% | 1.524% | -2.203% | 0.1792% | -2.745% | -0.8496% | -2.42% | -0.3081% | -0.03736% | 4.024% | 2.887% | -2.582% | 0.5583% | 5.486% | -2.542% | 7.49% | 0.8832% | 0.7208% | -0.03736% | -0.3252% | -2.42% | -2.041%
| Z-Score (Absolute) |  | 0.74 | 1.29 | 1.72 | 0.53 | 0.64 | 0.11 | 0.81 | 0.21 | 0.71 | 0.04 | 0.04 | 1.32 | 0.96 | 0.76 | 0.23 | 1.78 | 0.75 | 2.42 | 0.33 | 0.26 | 0.04 | 0.05 | 0.71 | 0.59
| Analytical results for Mo in OREAS 265 (Certified Value 4.22 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 4.69 | 4.32 | 4.2 | 4.088 | 4.2 | 4.06 | 4.3 | 4.07 | 4.2 | 4.36 | 3.97 | 4.05 | 4.06 | 5.17 | 4.51 | 4.18 | 3.6118 | 4.3 | 4.43 | 4.29 | 4.3 | 4.0815 | 3.987 | 4.03 | 1
| 1 | 2 | 4.46 | 4.42 | 4.2 | 3.4419 | 4.2 | 4.11 | 4.2 | 3.99 | 4.2 | 4.24 | 3.94 | 4.21 | 3.99 | 5.05 | 4.48 | 4.08 | 3.6763 | 4.4 | 4.56 | 4.38 | 4.2 | 4.0475 | 4.0524 | 4.13 | 26
| 1 | 3 | 4.41 | 4.41 | 4.2 | 3.4846 | 4.3 | 4 | 4 | 3.88 | 4.2 | 4.31 | 4.01 | 4.12 | 4.14 | 4.92 | 4.66 | 4.25 | 3.618 | 4.6 | 4.6 | 4.39 | 4.2 | 4.0341 | 4.124 | 4.12 | 16
| 1 | 4 | 4.59 | 4.37 | 4.2 | 3.7736 | 4.2 | 4.23 | 4.4 | 3.86 | 4.1 | 4.33 | 4.12 | 4.18 | 3.79 | 4.63 | 4.68 | 4.29 | 3.6659 | 4.5 | 4.49 | 4.43 | 4.2 | 4.1413 | 3.9985 | 4.3 | 4.219
| 1 | 5 | 4.41 | 4.33 | 4.2 | 3.8052 | 4.4 | 4.09 | 4.1 | 3.91 | 4.4 | 4.61 | 4.03 | 4.24 | 4.23 | 4.46 | 4.57 | 4.25 | 3.6434 | 4.4 | 4.35 | 4.31 | 4.2 | 4.2438 | 4.23 | 4.31 | 42
| 1 | 6 | 4.15 | 4.31 | 4.2 | 3.9034 | 4.2 | 3.99 | 4.1 | 4 | 4.1 | 4.26 | 3.92 | 4.18 | 3.98 | 4.32 | 4.37 | 4.43 | 3.6039 | 4.6 | 4.43 | 4.47 | 4.2 | 4.233 | 4.18 | 4.31
| Mean |  | 4.4517 | 4.36 | 4.2 | 3.7495 | 4.25 | 4.08 | 4.1833 | 3.9517 | 4.2 | 4.3517 | 3.9983 | 4.1633 | 4.0317 | 4.7583 | 4.545 | 4.2467 | 3.6366 | 4.4667 | 4.4767 | 4.3783 | 4.2167 | 4.1302 | 4.0953 | 4.2
| Median |  | 4.435 | 4.35 | 4.2 | 3.7894 | 4.2 | 4.075 | 4.15 | 3.95 | 4.2 | 4.32 | 3.99 | 4.18 | 4.025 | 4.775 | 4.54 | 4.25 | 3.6307 | 4.45 | 4.46 | 4.385 | 4.2 | 4.1114 | 4.0882 | 4.215
| Std Dev. |  | 0.18465 | 0.047329 | 0 | 0.2477 | 0.083666 | 0.087636 | 0.1472 | 0.081343 | 0.10954 | 0.13408 | 0.072503 | 0.068313 | 0.15145 | 0.34008 | 0.11675 | 0.11639 | 0.030052 | 0.12111 | 0.092448 | 0.068823 | 0.040825 | 0.091691 | 0.099077 | 0.12198
| Rel.Std.Dev. |  | 4.148% | 1.086% | 0% | 6.606% | 1.969% | 2.148% | 3.519% | 2.058% | 2.608% | 3.081% | 1.813% | 1.641% | 3.756% | 7.147% | 2.569% | 2.741% | 0.8264% | 2.711% | 2.065% | 1.572% | 0.9682% | 2.22% | 2.419% | 2.904%
| PDM3 |  | 5.515% | 3.342% | -0.4504% | -11.13% | 0.7347% | -3.295% | -0.8455% | -6.337% | -0.4504% | 3.144% | -5.23% | -1.32% | -4.44% | 12.78% | 7.727% | 0.6557% | -13.81% | 5.87% | 6.107% | 3.776% | -0.05541% | -2.105% | -2.932% | -0.4504%
| Z-Score (Absolute) |  | 1.09 | 0.69 | 0 | 1.95 | 0.22 | 0.52 | 0.07 | 1.07 | 0 | 0.66 | 0.87 | 0.16 | 0.73 | 2.42 | 1.49 | 0.2 | 2.44 | 1.15 | 1.2 | 0.77 | 0.07 | 0.3 | 0.45 | 0
| Analytical results for Na in OREAS 265 (Certified Value 0.76 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.8 | 0.78 | 0.773 | 0.68371 | 0.7546 | 0.763 | 0.79 | 0.78 | 0.7521 | 0.76 | 0.77 | 0.756 | 0.758 | 0.88 | 0.9 | 0.8 | 0.73872 | 0.7853 | 0.75 | 0.779 | 0.81 | 0.71395 | 0.7015 | 0.77 | 1
| 1 | 2 | 0.75 | 0.79 | 0.748 | 0.68308 | 0.748 | 0.748 | 0.78 | 0.77 | 0.7717 | 0.73 | 0.75 | 0.784 | 0.746 | 0.84 | 0.96 | 0.79 | 0.73729 | 0.7884 | 0.75 | 0.786 | 0.81 | 0.71103 | 0.6947 | 0.76 | 44
| 1 | 3 | 0.74 | 0.79 | 0.753 | 0.64969 | 0.7371 | 0.739 | 0.78 | 0.76 | 0.7717 | 0.74 | 0.75 | 0.781 | 0.73 | 0.88 | 0.93 | 0.79 | 0.73906 | 0.7978 | 0.76 | 0.772 | 0.78 | 0.72474 | 0.6953 | 0.76 | 16
| 1 | 4 | 0.77 | 0.77 | 0.742 | 0.65969 | 0.7326 | 0.73 | 0.8 | 0.74 | 0.7423 | 0.75 | 0.78 | 0.787 | 0.815 | 0.86 | 0.95 | 0.81 | 0.73807 | 0.7688 | 0.76 | 0.772 | 0.79 | 0.72332 | 0.7094 | 0.79 | 0.76041
| 1 | 5 | 0.72 | 0.79 | 0.76 | 0.64642 | 0.7348 | 0.739 | 0.75 | 0.75 | 0.762 | 0.74 | 0.77 | 0.782 | 0.753 | 0.84 | 0.92 | 0.79 | 0.73815 | 0.7891 | 0.75 | 0.772 | 0.78 | 0.74152 | 0.7142 | 0.77 | 43
| 1 | 6 | 0.7 | 0.78 | 0.761 | 0.6613 | 0.7446 | 0.733 | 0.76 | 0.77 | 0.7617 | 0.76 | 0.77 | 0.775 | 0.804 | 0.86 | 0.9 | 0.82 | 0.73663 | 0.7896 | 0.75 | 0.809 | 0.77 | 0.73706 | 0.6985 | 0.78
| Mean |  | 0.74667 | 0.78333 | 0.75617 | 0.66398 | 0.74195 | 0.742 | 0.77667 | 0.76167 | 0.76025 | 0.74667 | 0.765 | 0.7775 | 0.76767 | 0.86 | 0.92667 | 0.8 | 0.73799 | 0.7865 | 0.75333 | 0.78167 | 0.79 | 0.72527 | 0.70227 | 0.77167
| Median |  | 0.745 | 0.785 | 0.7565 | 0.66049 | 0.74085 | 0.739 | 0.78 | 0.765 | 0.76185 | 0.745 | 0.77 | 0.7815 | 0.7555 | 0.86 | 0.925 | 0.795 | 0.73811 | 0.78875 | 0.75 | 0.7755 | 0.785 | 0.72403 | 0.7 | 0.77
| Std Dev. |  | 0.03559 | 0.008165 | 0.010944 | 0.016076 | 0.0085524 | 0.012 | 0.018619 | 0.01472 | 0.011454 | 0.012111 | 0.012247 | 0.011256 | 0.033933 | 0.017889 | 0.025033 | 0.012649 | 0.00089979 | 0.0096167 | 0.005164 | 0.014514 | 0.016733 | 0.01215 | 0.0079228 | 0.01169
| Rel.Std.Dev. |  | 4.767% | 1.042% | 1.447% | 2.421% | 1.153% | 1.617% | 2.397% | 1.933% | 1.507% | 1.622% | 1.601% | 1.448% | 4.42% | 2.08% | 2.701% | 1.581% | 0.1219% | 1.223% | 0.6855% | 1.857% | 2.118% | 1.675% | 1.128% | 1.515%
| PDM3 |  | -1.807% | 3.015% | -0.5575% | -12.68% | -2.427% | -2.421% | 2.138% | 0.1658% | -0.0205% | -1.807% | 0.6042% | 2.248% | 0.9549% | 13.1% | 21.86% | 5.207% | -2.948% | 3.432% | -0.9301% | 2.796% | 3.892% | -4.621% | -7.646% | 1.481%
| Z-Score (Absolute) |  | 0.59 | 0.71 | 0.25 | 3.52 | 0.76 | 0.76 | 0.47 | 0.06 | 0.11 | 0.59 | 0.06 | 0.5 | 0.15 | 3.42 | 5.79 | 1.3 | 0.9 | 0.82 | 0.35 | 0.64 | 0.94 | 1.35 | 2.16 | 0.3
| Analytical results for Nb in OREAS 265 (Certified Value 26.2 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 28.5 | 27.5 | 24.8 | 19.23 | 26.22 | 26.6 | 22.6 | 27.2 | 25.08 | 28.2 | 27.6 | 25.4 | 23.86 | 23.9 | 25.2 | 25.5 | 28.1072 | 25.06 | 24.2 | 28.3 | 25.7 | 26.4216 | 25.6 |  | 1
| 1 | 2 | 27.9 | 29 | 24.7 | 18.2292 | 26.97 | 26.7 | 25.5 | 26.9 | 25.36 | 26.9 | 27.6 | 26.55 | 23.2 | 23.3 | 25 | 26.5 | 28.0744 | 25.17 | 23.9 | 28.7 | 26.3 | 26.0535 | 25.9 |  | 6
| 1 | 3 | 27.6 | 28.2 | 25.3 | 18.8415 | 26.47 | 25.7 | 22.3 | 25.5 | 25.39 | 26.8 | 27.8 | 25.64 | 24.84 | 25 | 24.9 | 26.8 | 28.0557 | 26.39 | 24.6 | 27.8 | 25.7 | 26.3005 | 26 |  | 16
| 1 | 4 | 24.7 | 28.3 | 25.5 | 20.2969 | 26.72 | 27.1 | 25.1 | 26.3 | 25.07 | 27.9 | 28.3 | 26.04 | 26.45 | 24.7 | 25.3 | 26.8 | 28.1078 | 25.61 | 23.9 | 28.7 | 25.7 | 26.5105 | 26.8 |  | 26.1889
| 1 | 5 | 22.6 | 28.1 | 24.4 | 19.7814 | 28.03 | 25.7 | 25.1 | 26.5 | 25.93 | 27.8 | 27.9 | 26.03 | 23.94 | 24.4 | 25 | 26.6 | 28.0278 | 27.2 | 24.3 | 28.2 | 26 | 27.3585 | 27.9 |  | 44
| 1 | 6 | 19.3 | 27.3 | 24.7 | 20.859 | 27.53 | 26.3 | 23.8 | 27.1 | 25.85 | 27.4 | 26.8 | 26.01 | 26.81 | 23.8 | 24.2 | 27.6 | 28.9384 | 26.69 | 23.7 | 28.5 | 26.1 | 27.6769 | 27.3
| Mean |  | 25.1 | 28.0667 | 24.9 | 19.5397 | 26.99 | 26.35 | 24.0667 | 26.5833 | 25.4467 | 27.5 | 27.6667 | 25.945 | 24.85 | 24.1833 | 24.9333 | 26.6333 | 28.2185 | 26.02 | 24.1 | 28.3667 | 25.9167 | 26.7202 | 26.5833
| Median |  | 26.15 | 28.15 | 24.75 | 19.5057 | 26.845 | 26.45 | 24.45 | 26.7 | 25.375 | 27.6 | 27.7 | 26.02 | 24.39 | 24.15 | 25 | 26.7 | 28.0908 | 26 | 24.05 | 28.4 | 25.85 | 26.466 | 26.4
| Std Dev. |  | 3.6249 | 0.60882 | 0.41473 | 0.96676 | 0.67968 | 0.5648 | 1.3808 | 0.63377 | 0.36968 | 0.56569 | 0.49666 | 0.39409 | 1.4787 | 0.63061 | 0.38816 | 0.67725 | 0.35398 | 0.87068 | 0.32863 | 0.34448 | 0.25626 | 0.64443 | 0.90203
| Rel.Std.Dev. |  | 14.44% | 2.169% | 1.666% | 4.948% | 2.518% | 2.143% | 5.737% | 2.384% | 1.453% | 2.057% | 1.795% | 1.519% | 5.95% | 2.608% | 1.557% | 2.543% | 1.254% | 3.346% | 1.364% | 1.214% | 0.9888% | 2.412% | 3.393%
| PDM3 |  | -4.158% | 7.17% | -4.922% | -25.39% | 3.059% | 0.615% | -8.104% | 1.506% | -2.834% | 5.006% | 5.643% | -0.9314% | -5.113% | -7.658% | -4.794% | 1.697% | 7.75% | -0.645% | -7.976% | 8.315% | -1.04% | 2.029% | 1.506%
| Z-Score (Absolute) |  | 0.57 | 1.27 | 0.69 | 4.02 | 0.6 | 0.2 | 1.21 | 0.35 | 0.36 | 0.92 | 1.02 | 0.05 | 0.73 | 1.14 | 0.67 | 0.38 | 1.36 | 0 | 1.19 | 1.46 | 0.06 | 0.43 | 0.35
| Analytical results for Nd in OREAS 265 (Certified Value 41.1 ppm) | [25 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 1
2
| 1 | 1 | 40.75 | 42.8 | 41.5 | 44.3 | 38.9 | 39.3 | 40.7088 | 38.46 | 43.2 | 40.4517 | [14 empty] | 1
| 1 | 2 | 41.35 | 43.1 | 41.8 | 43.9 | 37 | 39.9 | 40.6617 | 39.52 | 44.3 | 39.1458 | [14 empty] | 28
| 1 | 3 | 40.99 | 41.2 | 40.7 | 44.8 | 37.9 | 39.8 | 40.5702 | 40.78 | 41.8 | 39.6363 | [14 empty] | 16
| 1 | 4 | 39.76 | 43.8 | 41.8 | 40.9 | 38.5 | 41.3 | 40.6189 | 39.57 | 43.5 | 39.7629 | [14 empty] | 41.1121
| 1 | 5 | 41.26 | 42.1 | 43.7 | 44 | 38.6 | 39 | 40.6908 | 42.04 | 43.5 | 41.1731 | [14 empty] | 45
| 1 | 6 | 41.36 | 41 | 42.6 | 40.7 | 36.9 | 40.9 | 40.6974 | 41.53 | 43.6 | 41.2582
| Mean |  | 40.9117 | 42.3333 | 42.0167 | 43.1 | 37.9667 | 40.0333 | 40.658 | 40.3167 | 43.3167 | 40.238
| Median |  | 41.125 | 42.45 | 41.8 | 43.95 | 38.2 | 39.85 | 40.6763 | 40.175 | 43.5 | 40.1073
| Std Dev. |  | 0.61206 | 1.1021 | 1.0265 | 1.81 | 0.85245 | 0.89815 | 0.053779 | 1.3635 | 0.82805 | 0.86506
| Rel.Std.Dev. |  | 1.496% | 2.603% | 2.443% | 4.199% | 2.245% | 2.243% | 0.1323% | 3.382% | 1.912% | 2.15%
| PDM3 |  | -0.4876% | 2.97% | 2.2% | 4.835% | -7.651% | -2.624% | -1.105% | -1.935% | 5.362% | -2.126%
| Z-Score (Absolute) |  | 0.09 | 1.05 | 0.84 | 1.57 | 1.92 | 0.51 | 0.09 | 0.32 | 1.72 | 0.37
| Analytical results for Ni in OREAS 265 (Certified Value 53 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 53.9 | 55.9 | 51 | 54.5428 | 54 | 56 | 54 | 54 | 54 | 57.6 | 53.4 | 52.7 | 53 | 53.5 | 51 | 52.1 | 51.7335 | 57.2 | 53.5 | 58.5 | 53.1 | 51.7658 | 52.321 | 52.2 | 1
| 1 | 2 | 52.6 | 58.9 | 50 | 53.7801 | 53 | 55 | 55 | 52.4 | 54.7 | 55.1 | 53.2 | 55.9 | 51.3 | 52.6 | 51.2 | 52.2 | 51.6361 | 59.6 | 54 | 59.9 | 52.7 | 51.2142 | 52.65 | 53.6 | 14
| 1 | 3 | 51.4 | 58.1 | 51 | 52.7378 | 53 | 53 | 51 | 52.1 | 54.6 | 55.4 | 54 | 54.1 | 51 | 53.1 | 50.7 | 53.4 | 51.8098 | 57.9 | 54.4 | 58 | 53.1 | 52.6693 | 51.98 | 52.8 | 16
| 1 | 4 | 53.3 | 57 | 50 | 53.1981 | 53 | 54 | 57 | 52.4 | 53 | 56.9 | 53.8 | 54.7 | 59.2 | 53.8 | 51.8 | 52.8 | 51.6356 | 56.3 | 53.3 | 58.6 | 52.6 | 51.7586 | 52.45 | 54.8 | 53.0533
| 1 | 5 | 51.6 | 57.7 | 51 | 53.7506 | 53 | 51 | 54 | 51.3 | 55.1 | 55.3 | 53.8 | 55.3 | 52.4 | 54.4 | 51.3 | 52.1 | 51.6667 | 57.8 | 53.1 | 59.3 | 53.2 | 53.3481 | 52.36 | 54.5 | 46
| 1 | 6 | 50.5 | 56.6 | 51 | 54.7148 | 54 | 49 | 51 | 55.3 | 54.4 | 56.2 | 52.3 | 54.9 | 58.9 | 52.1 | 50.3 | 54.8 | 51.944 | 57.9 | 54.3 | 59.7 | 53.3 | 53.8155 | 52.387 | 54.4
| Mean |  | 52.2167 | 57.3667 | 50.6667 | 53.7874 | 53.3333 | 53 | 53.6667 | 52.9167 | 54.3 | 56.0833 | 53.4167 | 54.6 | 54.3 | 53.25 | 51.05 | 52.9 | 51.7376 | 57.7833 | 53.7667 | 59 | 53 | 52.4286 | 52.358 | 53.7167
| Median |  | 52.1 | 57.35 | 51 | 53.7654 | 53 | 53.5 | 54 | 52.4 | 54.5 | 55.8 | 53.6 | 54.8 | 52.7 | 53.3 | 51.1 | 52.5 | 51.7001 | 57.85 | 53.75 | 58.95 | 53.1 | 52.2176 | 52.3735 | 54
| Std Dev. |  | 1.2766 | 1.0838 | 0.5164 | 0.7591 | 0.5164 | 2.6077 | 2.3381 | 1.4607 | 0.73212 | 1.0028 | 0.62102 | 1.1082 | 3.7513 | 0.83126 | 0.51672 | 1.0621 | 0.12128 | 1.0834 | 0.54283 | 0.74833 | 0.28284 | 1.0188 | 0.21855 | 1.04
| Rel.Std.Dev. |  | 2.445% | 1.889% | 1.019% | 1.411% | 0.9682% | 4.92% | 4.357% | 2.76% | 1.348% | 1.788% | 1.163% | 2.03% | 6.908% | 1.561% | 1.012% | 2.008% | 0.2344% | 1.875% | 1.01% | 1.268% | 0.5337% | 1.943% | 0.4174% | 1.936%
| PDM3 |  | -1.577% | 8.13% | -4.499% | 1.384% | 0.5278% | -0.1005% | 1.156% | -0.2576% | 2.35% | 5.711% | 0.6849% | 2.915% | 2.35% | 0.3707% | -3.776% | -0.289% | -2.48% | 8.916% | 1.345% | 11.21% | -0.1005% | -1.178% | -1.311% | 1.25%
| Z-Score (Absolute) |  | 0.84 | 2.91 | 1.97 | 0.3 | 0.03 | 0.27 | 0.21 | 0.33 | 0.67 | 1.97 | 0.03 | 0.89 | 0.67 | 0.09 | 1.69 | 0.35 | 1.19 | 3.21 | 0.29 | 4.1 | 0.27 | 0.69 | 0.74 | 0.25
| Analytical results for P in OREAS 265 (Certified Value 0.078 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.0759 | 0.08 | 0.079 | 0.079876 | 0.0771 | 0.08 | 0.08 | 0.079 | 0.0766 | 0.081 | 0.077 | 0.074 | 0.079 | 0.0788 | 0.0737 | 0.078 | 0.071159 | 0.0771 | 0.078 | 0.079 | 0.078 | 0.074092 | 0.0731 | 0.0804 | 1
| 1 | 2 | 0.0725 | 0.083 | 0.078 | 0.076412 | 0.0764 | 0.076 | 0.0793 | 0.078 | 0.0774 | 0.076 | 0.074 | 0.081 | 0.077 | 0.0746 | 0.0759 | 0.077 | 0.071083 | 0.0772 | 0.08 | 0.079 | 0.079 | 0.074012 | 0.0727 | 0.0782 | #N/A
| 1 | 3 | 0.0717 | 0.085 | 0.079 | 0.073104 | 0.0765 | 0.077 | 0.0785 | 0.077 | 0.0759 | 0.079 | 0.076 | 0.078 | 0.076 | 0.0785 | 0.0748 | 0.079 | 0.071237 | 0.0786 | 0.079 | 0.074 | 0.078 | 0.075905 | 0.0705 | 0.0788 | 16
| 1 | 4 | 0.0739 | 0.082 | 0.078 | 0.077121 | 0.0751 | 0.076 | 0.0797 | 0.074 | 0.0742 | 0.08 | 0.077 | 0.081 | 0.078 | 0.0777 | 0.0776 | 0.081 | 0.071389 | 0.0765 | 0.081 | 0.079 | 0.079 | 0.075804 | 0.0732 | 0.0822 | 0.077551
| 1 | 5 | 0.0705 | 0.084 | 0.078 | 0.074011 | 0.0747 | 0.079 | 0.0783 | 0.076 | 0.0743 | 0.078 | 0.079 | 0.081 | 0.078 | 0.0773 | 0.0724 | 0.079 | 0.071345 | 0.0771 | 0.079 | 0.074 | 0.078 | 0.077094 | 0.0694 | 0.0786 | 47
| 1 | 6 | 0.0669 | 0.079 | 0.079 | 0.078068 | 0.0761 | 0.078 | 0.0787 | 0.078 | 0.075 | 0.079 | 0.074 | 0.079 | 0.078 | 0.0751 | 0.0749 | 0.082 | 0.071208 | 0.0789 | 0.079 | 0.079 | 0.078 | 0.078428 | 0.0703 | 0.0802
| Mean |  | 0.0719 | 0.082167 | 0.0785 | 0.076432 | 0.075983 | 0.077667 | 0.079083 | 0.077 | 0.075567 | 0.078833 | 0.076167 | 0.079 | 0.077667 | 0.077 | 0.074883 | 0.079333 | 0.071237 | 0.077567 | 0.079333 | 0.077333 | 0.078333 | 0.075889 | 0.071533 | 0.079733
| Median |  | 0.0721 | 0.0825 | 0.0785 | 0.076767 | 0.07625 | 0.0775 | 0.079 | 0.0775 | 0.07545 | 0.079 | 0.0765 | 0.08 | 0.078 | 0.0775 | 0.07485 | 0.079 | 0.071223 | 0.07715 | 0.079 | 0.079 | 0.078 | 0.075855 | 0.0716 | 0.0795
| Std Dev. |  | 0.0030777 | 0.0023166 | 0.00054772 | 0.002528 | 0.00090866 | 0.001633 | 0.00068823 | 0.0017889 | 0.001291 | 0.0017224 | 0.0019408 | 0.0027568 | 0.0010328 | 0.0017573 | 0.0017882 | 0.0018619 | 0.00011415 | 0.00095429 | 0.0010328 | 0.002582 | 0.0005164 | 0.0017123 | 0.0016573 | 0.0015002
| Rel.Std.Dev. |  | 4.28% | 2.819% | 0.6977% | 3.308% | 1.196% | 2.103% | 0.8703% | 2.323% | 1.708% | 2.185% | 2.548% | 3.49% | 1.33% | 2.282% | 2.388% | 2.347% | 0.1602% | 1.23% | 1.302% | 3.339% | 0.6592% | 2.256% | 2.317% | 1.882%
| PDM3 |  | -7.287% | 5.951% | 1.223% | -1.444% | -2.022% | 0.1486% | 1.975% | -0.711% | -2.559% | 1.653% | -1.786% | 1.868% | 0.1486% | -0.711% | -3.44% | 2.298% | -8.142% | 0.01968% | 2.298% | -0.2812% | 1.008% | -2.143% | -7.76% | 2.814%
| Z-Score (Absolute) |  | 2.49 | 2.22 | 0.54 | 0.41 | 0.62 | 0.15 | 0.8 | 0.15 | 0.81 | 0.69 | 0.53 | 0.77 | 0.15 | 0.15 | 1.12 | 0.92 | 2.79 | 0.11 | 0.92 | 0.11 | 0.46 | 0.66 | 2.66 | 1.1
| Analytical results for Pb in OREAS 265 (Certified Value 16.1 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 16.6 | 17 | 17 | 15.102 | 17.5 | 17.8 | 16.7 | 16.2 | 16.5 | 16 | 14.8 | 17.15 | 14.43 | 13.2 | 17.1 | 15.6 | 9 | 16.6 | 15.9 | 12.3 | 16.0347 | 17.63 | 15.6 |  | 1
| 1 | 2 | 16.1 | 17.6 | 16 | 16.0512 | 17.8 | 17.9 | 17.6 | 15.6 | 15.7 | 15.4 | 14.4 | 17.11 | 14.46 | 12.4 | 16.9 | 16.2 | 9 | 16.5 | 15.9 | 13.3 | 16.0781 | 17.65 | 15.7 |  | 16
| 1 | 3 | 15.6 | 17 | 17 | 15.2658 | 16.7 | 17.5 | 16.1 | 15.6 | 15.6 | 16 | 15 | 16.94 | 14.67 | 13.3 | 16.7 | 16.1 | 9 | 15.9 | 15.5 | 12.7 | 16.3769 | 17.25 | 15.6 |  | 16
| 1 | 4 | 16 | 17 | 17 | 15.6108 | 16.5 | 17.4 | 17.8 | 15.4 | 15.6 | 16 | 15.3 | 17.33 | 13.96 | 13.6 | 16.9 | 16.2 | 9 | 15.5 | 16 | 14.3 | 16.8983 | 16.99 | 16.3 |  | 16.1234
| 1 | 5 | 15.7 | 17.5 | 16 | 15.7946 | 17.7 | 17.1 | 15.9 | 15.5 | 15.8 | 15.4 | 14.9 | 17.74 | 14.72 | 12.7 | 16.5 | 16.1 | 11 | 15.8 | 15.9 | 14 | 15.5376 | 17.414 | 16.5 |  | 48
| 1 | 6 | 15.4 | 16.6 | 17 | 16.2437 | 17.8 | 17.4 | 15.9 | 15.7 | 15.8 | 16.8 | 14.7 | 17.52 | 13.77 | 12.9 | 16.8 | 16.6 | 12 | 15.5 | 16 | 14.8 | 15.3521 | 17.6321 | 16.2
| Mean |  | 15.9 | 17.1167 | 16.6667 | 15.678 | 17.3333 | 17.5167 | 16.6667 | 15.6667 | 15.8333 | 15.9333 | 14.85 | 17.2983 | 14.335 | 13.0167 | 16.8167 | 16.1333 | 9.8333 | 15.9667 | 15.8667 | 13.5667 | 16.0463 | 17.4277 | 15.9833
| Median |  | 15.85 | 17 | 17 | 15.7027 | 17.6 | 17.45 | 16.4 | 15.6 | 15.75 | 16 | 14.85 | 17.24 | 14.445 | 13.05 | 16.85 | 16.15 | 9 | 15.85 | 15.9 | 13.65 | 16.0564 | 17.522 | 15.95
| Std Dev. |  | 0.42895 | 0.37103 | 0.5164 | 0.44251 | 0.58195 | 0.29269 | 0.85479 | 0.28048 | 0.33862 | 0.5164 | 0.30166 | 0.29363 | 0.38599 | 0.43551 | 0.20412 | 0.32042 | 1.3292 | 0.48028 | 0.18619 | 0.96678 | 0.56161 | 0.26664 | 0.39707
| Rel.Std.Dev. |  | 2.698% | 2.168% | 3.098% | 2.823% | 3.357% | 1.671% | 5.129% | 1.79% | 2.139% | 3.241% | 2.031% | 1.697% | 2.693% | 3.346% | 1.214% | 1.986% | 13.52% | 3.008% | 1.173% | 7.126% | 3.5% | 1.53% | 2.484%
| PDM3 |  | -1.386% | 6.16% | 3.369% | -2.763% | 7.504% | 8.641% | 3.369% | -2.833% | -1.799% | -1.179% | -7.898% | 7.287% | -11.09% | -19.27% | 4.3% | 0.06141% | -39.01% | -0.9723% | -1.593% | -15.86% | -0.4785% | 8.089% | -0.8689%
| Z-Score (Absolute) |  | 0.06 | 1.11 | NA | 0.28 | 1.32 | 1.49 | 0.67 | 0.29 | 0.13 | 0.03 | 1.08 | 1.28 | 1.57 | 2.84 | 0.82 | 0.16 | NA | 0 | 0.1 | 2.31 | 0.08 | 1.41 | 0.02
Results from laboratories 1.03 and 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 265 (Certified Value 11 ppm) | [25 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 1
2
| 1 | 1 | 11.23 | 11.36 | 11.3 | 10.5 | 10.63 | 11.3 | 11.1069 | 10.41 | 11.2 | 10.9311 | [14 empty] | 1
| 1 | 2 | 11.32 | 11.55 | 11.4 | 10.5 | 10.15 | 11 | 10.8212 | 10.69 | 11.4 | 10.6291 | [14 empty] | 31
| 1 | 3 | 11 | 11.33 | 11.1 | 10.5 | 10.38 | 11.2 | 10.5252 | 11.03 | 10.9 | 10.8295 | [14 empty] | 16
| 1 | 4 | 10.84 | 12.04 | 11.3 | 11.1 | 10.52 | 11.2 | 10.8132 | 10.75 | 11.2 | 10.6982 | [14 empty] | 10.9893
| 1 | 5 | 10.99 | 11.28 | 11.8 | 10.3 | 10.66 | 10.9 | 10.9373 | 11.37 | 11.2 | 11.0779 | [14 empty] | 49
| 1 | 6 | 11.16 | 11.04 | 11.6 | 11.1 | 10.14 | 11.3 | 9.0233 | 11.32 | 11.2 | 11.1836
| Mean |  | 11.09 | 11.4333 | 11.4167 | 10.6667 | 10.4133 | 11.15 | 10.5378 | 10.9283 | 11.1833 | 10.8916
| Median |  | 11.08 | 11.345 | 11.35 | 10.5 | 10.45 | 11.2 | 10.8172 | 10.89 | 11.2 | 10.8803
| Std Dev. |  | 0.17776 | 0.33939 | 0.24833 | 0.34448 | 0.22993 | 0.16432 | 0.76598 | 0.37844 | 0.16021 | 0.21526
| Rel.Std.Dev. |  | 1.603% | 2.968% | 2.175% | 3.23% | 2.208% | 1.474% | 7.269% | 3.463% | 1.433% | 1.976%
| PDM3 |  | 0.9167% | 4.041% | 3.889% | -2.936% | -5.241% | 1.463% | -4.108% | -0.5545% | 1.766% | -0.8891%
| Z-Score (Absolute) |  | 0.21 | 1.11 | 1.06 | 0.89 | 1.56 | 0.37 | 1.23 | 0.21 | 0.45 | 0.31
| Analytical results for Rb in OREAS 265 (Certified Value 100 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 103.7 | 101.5 | 100 | 100.4965 | 98.66 | 98.5 | 105 | 105 | 100.65 | 102 | 95.3 | 99.3 | 95.1 | 99.3 | 100.1 | 102 | 97.897 | 96.8 | 90.6 | 100 | 111.9 | 98.7251 | 97.2 |  | 1
| 1 | 2 | 99.3 | 104.5 | 98.6 | 99.6465 | 97 | 99.6 | 101 | 102.5 | 101.2 | 101 | 95.2 | 104.4 | 90.8 | 96.3 | 98.5 | 102 | 97.7891 | 98 | 91.3 | 103 | 112 | 97.6905 | 101.5 |  | 32
| 1 | 3 | 93.8 | 103 | 98.7 | 99.7087 | 98.91 | 97.5 | 100 | 101.5 | 99.55 | 100.5 | 95.6 | 102 | 95.5 | 97.9 | 98.9 | 103 | 97.8054 | 101.3 | 91.4 | 101 | 113.9 | 96.7474 | 99.3 |  | 16
| 1 | 4 | 96.1 | 102.5 | 98.7 | 99.9122 | 96.64 | 98.6 | 106 | 101 | 97.74 | 104 | 98.2 | 103.3 | 74.9 | 98.5 | 101.1 | 104 | 97.8306 | 100.7 | 91 | 101 | 116.2 | 98.2214 | 103.1 |  | 100.1058
| 1 | 5 | 95 | 102 | 97.4 | 100.0593 | 98.1 | 100 | 103 | 99.3 | 98.23 | 102.5 | 96.9 | 104.9 | 90.2 | 98.9 | 99.5 | 103 | 97.8532 | 105.5 | 89.4 | 102 | 109.3 | 102.0397 | 102.9 |  | 50
| 1 | 6 | 93.6 | 102 | 99.1 | 100.3508 | 99.22 | 98.4 | 100 | 102.5 | 100.91 | 102 | 98.5 | 105.5 | 72 | 98.1 | 97.5 | 107 | 98.0026 | 103.6 | 90.5 | 102 | 109.5 | 102.1083 | 102.7
| Mean |  | 96.9167 | 102.5833 | 98.75 | 100.029 | 98.0883 | 98.7667 | 102.5 | 101.9667 | 99.7133 | 102 | 96.6167 | 103.2333 | 86.4167 | 98.1667 | 99.2667 | 103.5 | 97.863 | 100.9833 | 90.7 | 101.5 | 112.1333 | 99.2554 | 101.1167
| Median |  | 95.55 | 102.25 | 98.7 | 99.9857 | 98.38 | 98.55 | 102 | 102 | 100.1 | 102 | 96.25 | 103.85 | 90.5 | 98.3 | 99.2 | 103 | 97.8419 | 101 | 90.8 | 101.5 | 111.95 | 98.4733 | 102.1
| Std Dev. |  | 3.9179 | 1.0685 | 0.84083 | 0.34221 | 1.055 | 0.90037 | 2.5884 | 1.8991 | 1.4588 | 1.2247 | 1.477 | 2.2923 | 10.3141 | 1.0482 | 1.2612 | 1.8708 | 0.078184 | 3.2811 | 0.73212 | 1.0488 | 2.6341 | 2.2795 | 2.3836
| Rel.Std.Dev. |  | 4.043% | 1.042% | 0.8515% | 0.3421% | 1.076% | 0.9116% | 2.525% | 1.862% | 1.463% | 1.201% | 1.529% | 2.221% | 11.94% | 1.068% | 1.271% | 1.808% | 0.07989% | 3.249% | 0.8072% | 1.033% | 2.349% | 2.297% | 2.357%
| PDM3 |  | -3.186% | 2.475% | -1.354% | -0.07672% | -2.015% | -1.338% | 2.392% | 1.859% | -0.392% | 1.892% | -3.485% | 3.124% | -13.67% | -1.937% | -0.8382% | 3.391% | -2.24% | 0.8766% | -9.396% | 1.393% | 12.01% | -0.8495% | 1.01%
| Z-Score (Absolute) |  | 1.02 | 1.05 | 0.35 | 0.12 | 0.59 | 0.34 | 1.02 | 0.82 | 0 | 0.83 | 1.13 | 1.28 | 4.85 | 0.56 | 0.16 | 1.38 | 0.67 | 0.46 | 3.28 | 0.65 | 4.53 | 0.17 | 0.51
| Analytical results for Re in OREAS 265 (Certified Value < 0.002 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.29 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.01 | < 0.002 | < 0.002 | < 0.002 | 0.005 | < 0.005 | 0.004 | < 0.002 | < 0.005 | < 0.01 | < 0.002 | [6 empty] | 1
| 1 | 2 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.004 | < 0.002 | < 0.002 | 0.008 | < 0.002 | < 0.002 | < 0.002 | 0.007 | < 0.005 | 0.004 | < 0.002 | < 0.005 | < 0.01 | < 0.002 | [6 empty] | 33
| 1 | 3 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.004 | < 0.002 | < 0.002 | 0.007 | < 0.002 | < 0.002 | < 0.002 | 0.006 | < 0.005 | 0.005 | < 0.002 | < 0.005 | < 0.01 | < 0.002 | [6 empty] | 16
| 1 | 4 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.005 | < 0.002 | < 0.002 | 0.005 | < 0.002 | < 0.002 | < 0.002 | 0.006 | < 0.005 | 0.005 | < 0.002 | < 0.005 | < 0.01 | < 0.002 | [6 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.005 | < 0.002 | < 0.002 | 0.002 | 0.005 | < 0.005 | 0.004 | < 0.002 | < 0.005 | < 0.01 | < 0.002 | [6 empty] | 51
| 1 | 6 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.006 | < 0.002 | < 0.002 | 0.006 | < 0.002 | < 0.002 | 0.003 | 0.007 | < 0.005 | 0.004 | < 0.002 | < 0.005 | < 0.01 | < 0.002
| Mean | [5 empty] | 0.00475 |  |  | 0.0068333 |  |  | 0.0025 | 0.006 |  | 0.0043333
| Median | [5 empty] | 0.0045 |  |  | 0.0065 |  |  | 0.0025 | 0.006 |  | 0.004
| Std Dev. | [5 empty] | 0.00095743 |  |  | 0.0019408 |  |  | 0.00070711 | 0.00089443 |  | 0.0005164
| Rel.Std.Dev. | [5 empty] | 20.16% |  |  | 28.4% |  |  | 28.28% | 14.91% |  | 11.92%
PDM3
| Z-Score (Absolute) |  | 0.67 | 15.51 | 15.51 | 0.67 | 1.01 | 0.67 | 0.67 | 3.26 | 0.67 | 0.67 | 0.34 | 2.7 | 0.34 | 1.57 | 0.67 | 0.34 | 2.02 | 0.67
| Analytical results for S in OREAS 265 (Certified Value 0.009 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.05 | 0.01 | 0.008 | 0.0079749 | 0.0075 | 0.009 | 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.04 | < 0.04 | < 0.01 | < 0.01 | 0.01 | 0.0091183 | 0.0091 | 0.01 | 0.0082 | < 0.05 | < 0.01 | < 0.1 | 0.01 | 1
| 1 | 2 | 0.04 | 0.01 | 0.008 | 0.0068215 | 0.0074 | 0.008 | 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.04 | < 0.04 | < 0.01 | < 0.01 | 0.01 | 0.0091394 | 0.0095 | 0.02 | 0.0087 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | 1
| 1 | 3 | 0.03 | 0.01 | 0.008 | 0.0062931 | 0.0073 | 0.008 | < 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.04 | < 0.04 | < 0.01 | < 0.01 | 0.01 | 0.0090416 | 0.0097 | 0.02 | 0.0075 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | 16
| 1 | 4 | 0.04 | 0.01 | 0.008 | 0.0063212 | 0.0076 | 0.008 | 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.04 | < 0.04 | < 0.01 | < 0.01 | 0.01 | 0.0090506 | 0.0095 | 0.02 | 0.007 | < 0.05 | < 0.01 | < 0.1 | 0.01 | 0.0092278
| 1 | 5 | 0.03 | 0.01 | 0.008 | 0.0076998 | 0.0078 | 0.008 | 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.04 | < 0.04 | < 0.01 | < 0.01 | 0.01 | 0.0088717 | 0.0096 | 0.01 | 0.0074 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | 52
| 1 | 6 | 0.03 | 0.01 | 0.008 | 0.0065297 | 0.0074 | 0.007 | < 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.04 | < 0.04 | < 0.01 | < 0.01 | 0.01 | 0.0090874 | 0.0095 | 0.01 | 0.0083 | < 0.05 | < 0.01 | < 0.1 | 0.01
| Mean |  | 0.036667 | 0.01 | 0.008 | 0.00694 | 0.0075 | 0.008 | 0.01 | 0.01 |  | 0.01 | 0.01 |  |  |  |  | 0.01 | 0.0090515 | 0.0094833 | 0.015 | 0.00785 |  |  |  | 0.01
| Median |  | 0.035 | 0.01 | 0.008 | 0.0066756 | 0.00745 | 0.008 | 0.01 | 0.01 |  | 0.01 | 0.01 |  |  |  |  | 0.01 | 0.009069 | 0.0095 | 0.015 | 0.00785 |  |  |  | 0.01
| Std Dev. |  | 0.008165 | 0 | 0 | 0.00072556 | 0.00017889 | 0.00063246 | 0 | 0 |  | 0 | 0 |  |  |  |  | 0 | 9.5806e-05 | 0.00020412 | 0.0054772 | 0.0006473 |  |  |  | 0
| Rel.Std.Dev. |  | 22.27% | 0% | 0% | 10.45% | 2.385% | 7.906% | 0% | 0% |  | 0% | 0% |  |  |  |  | 0% | 1.058% | 2.152% | 36.51% | 8.246% |  |  |  | 0%
| PDM3 |  | 297.3% | 8.368% | -13.31% | -24.79% | -18.72% | -13.31% | 8.368% | 8.368% |  | 8.368% | 8.368% |  |  |  |  | 8.368% | -1.91% | 2.769% | 62.55% | -14.93% |  |  |  | 8.368%
| Z-Score (Absolute) |  | 8.1 | 0.08 | 0.52 | 0.84 | 0.67 | 0.52 | 0.42 | 0.08 | 4.59 | 0.08 | 0.08 | 3.09 | 3.09 | 1.43 | 1.43 | 0.08 | 0.21 | 0.08 | 1.58 | 0.57 | 4.59 | 1.43 | 12.11 | 0.67
| Analytical results for Sb in OREAS 265 (Certified Value 2.98 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.05 | 3.19 | 3.1 | 2.8443 | 2.98 | 2.93 | 3.15 | 2.87 | 2.86 | 2.83 | 2.84 | 2.83 | 2.8 | 2.95 | 3.05 | 3.1 | 3.237 | 2.89 | 3.38 | 3.5 | 2.99 | < 5 | 2.9 |  | 1
| 1 | 2 | 3.03 | 3.34 | 3 | 2.4282 | 2.93 | 2.97 | 2.79 | 2.73 | 2.88 | 2.72 | 2.75 | 3.05 | 2.66 | 2.86 | 2.95 | 3 | 3.378 | 2.96 | 3.32 | 2.9 | 2.9206 | < 5 | 2.9 |  | 34
| 1 | 3 | 2.92 | 3.24 | 3.1 | 2.4483 | 3.05 | 2.99 | 3.28 | 2.74 | 2.94 | 2.79 | 2.89 | 2.92 | 2.8 | 3.01 | 2.91 | 3.1 | 3.3898 | 3.05 | 3.44 | 2.9 | 2.9575 | < 5 | 2.9 |  | 16
| 1 | 4 | 2.99 | 3.29 | 3 | 2.8165 | 2.92 | 3.05 | 2.95 | 2.74 | 2.99 | 2.86 | 2.86 | 2.93 | 2.84 | 2.91 | 3.05 | 3 | 3.4173 | 2.99 | 3.14 | 3.5 | 2.9248 | < 5 | 3 |  | 2.9816
| 1 | 5 | 2.62 | 3.26 | 3 | 2.4118 | 2.94 | 3 | 2.88 | 2.73 | 2.97 | 2.84 | 2.85 | 3 | 2.74 | 2.94 | 3.05 | 3 | 3.3198 | 3.12 | 3.1 | 3 | 3.0097 | < 5 | 3.1 |  | 53
| 1 | 6 | 2.38 | 3.43 | 3 | 2.6696 | 3.05 | 3 | 3.49 | 2.75 | 2.9 | 2.82 | 2.83 | 3 | 3.01 | 2.77 | 2.98 | 2.9 | 3.3961 | 3.11 | 3.17 | 2.7 | 3.0973 | < 5 | 3
| Mean |  | 2.8317 | 3.2917 | 3.0333 | 2.6031 | 2.9783 | 2.99 | 3.09 | 2.76 | 2.9233 | 2.81 | 2.8367 | 2.955 | 2.8083 | 2.9067 | 2.9983 | 3.0167 | 3.3563 | 3.02 | 3.2583 | 3.0833 | 2.9833 |  | 2.9667
| Median |  | 2.955 | 3.275 | 3 | 2.559 | 2.96 | 2.995 | 3.05 | 2.74 | 2.92 | 2.825 | 2.845 | 2.965 | 2.8 | 2.925 | 3.015 | 3 | 3.3839 | 3.02 | 3.245 | 2.95 | 2.9737 |  | 2.95
| Std Dev. |  | 0.27154 | 0.084242 | 0.05164 | 0.19964 | 0.059133 | 0.039497 | 0.26586 | 0.054406 | 0.05164 | 0.0498 | 0.047188 | 0.078166 | 0.11703 | 0.083106 | 0.0608 | 0.075277 | 0.067042 | 0.089889 | 0.14034 | 0.33714 | 0.065966 |  | 0.08165
| Rel.Std.Dev. |  | 9.59% | 2.559% | 1.702% | 7.669% | 1.985% | 1.321% | 8.604% | 1.971% | 1.766% | 1.772% | 1.663% | 2.645% | 4.167% | 2.859% | 2.028% | 2.495% | 1.997% | 2.976% | 4.307% | 10.93% | 2.211% |  | 2.752%
| PDM3 |  | -5.029% | 10.4% | 1.735% | -12.69% | -0.1095% | 0.2818% | 3.636% | -7.432% | -1.954% | -5.755% | -4.861% | -0.8921% | -5.811% | -2.513% | 0.5613% | 1.176% | 12.57% | 1.288% | 9.281% | 3.412% | 0.05681% |  | -0.5008%
| Z-Score (Absolute) |  | 0.94 | 2.01 | 0.35 | 2.41 | 0 | 0.07 | 0.72 | 1.4 | 0.35 | 1.08 | 0.91 | 0.15 | 1.09 | 0.46 | 0.13 | 0.25 | 2.43 | 0.27 | 1.8 | 0.67 | 0.03 | 3.07 | 0.07
| Analytical results for Sc in OREAS 265 (Certified Value 11.7 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 10.3 | 12.5 | 11 | 11.5525 | 11.5 | 12 | 12.3 | 13 | 11.8 | 12.2 | 11.6 | 11.5 | 11 | 11.9 | 10.9 | 11.2 | 11.0783 | 11.75 | 11.2 | 12 | 13 | 11.502 | 10.48 | 11.3 | 1
| 1 | 2 | 9.7 | 12.9 | 11 | 11.7517 | 11.5 | 12 | 12.2 | 12.7 | 12 | 11.6 | 11.4 | 11.7 | 10.9 | 11.6 | 11.3 | 11.7 | 10.9916 | 11.85 | 11.5 | 12.5 | 12.7 | 11.6728 | 10.42 | 11.4 | 35
| 1 | 3 | 9.1 | 12.7 | 11 | 11.2096 | 11.4 | 12.3 | 11 | 12.2 | 11.8 | 12.2 | 11.8 | 11.2 | 10.7 | 11.8 | 11.1 | 11.3 | 10.8941 | 11.89 | 11.4 | 12.2 | 12.7 | 11.5608 | 10.68 | 11.3 | 16
| 1 | 4 | 8.9 | 12.8 | 11 | 11.4539 | 11.3 | 12.3 | 13.2 | 12.3 | 11.5 | 11.9 | 11.8 | 11.5 | 11.6 | 11.7 | 11.6 | 11.5 | 11.1516 | 11.59 | 11.4 | 12.3 | 12.9 | 11.4517 | 10.347 | 11.8 | 11.7215
| 1 | 5 | 9.1 | 12.7 | 11 | 11.303 | 11.4 | 12.5 | 12.9 | 12.4 | 11.6 | 12 | 11.6 | 11.7 | 11 | 12 | 11 | 11.3 | 11.0866 | 11.86 | 11.2 | 12.4 | 12.2 | 11.9566 | 10.66 | 12.1 | 54
| 1 | 6 | 8.5 | 12.8 | 11 | 11.9002 | 11.4 | 12.1 | 11.2 | 12.8 | 11.9 | 12.6 | 11.4 | 11.6 | 11.5 | 11.4 | 10.9 | 11.4 | 11.0744 | 11.87 | 11.4 | 12.3 | 12.3 | 12 | 10.4524 | 11.8
| Mean |  | 9.2667 | 12.7333 | 11 | 11.5285 | 11.4167 | 12.2 | 12.1333 | 12.5667 | 11.7667 | 12.0833 | 11.6 | 11.5333 | 11.1167 | 11.7333 | 11.1333 | 11.4 | 11.0461 | 11.8017 | 11.35 | 12.2833 | 12.6333 | 11.6906 | 10.5066 | 11.6167
| Median |  | 9.1 | 12.75 | 11 | 11.5032 | 11.4 | 12.2 | 12.25 | 12.55 | 11.8 | 12.1 | 11.6 | 11.55 | 11 | 11.75 | 11.05 | 11.35 | 11.0763 | 11.855 | 11.4 | 12.3 | 12.7 | 11.6168 | 10.4662 | 11.6
| Std Dev. |  | 0.6377 | 0.13663 | 0 | 0.26331 | 0.075277 | 0.2 | 0.88468 | 0.31411 | 0.18619 | 0.33714 | 0.17889 | 0.18619 | 0.35449 | 0.21602 | 0.27325 | 0.17889 | 0.090182 | 0.11462 | 0.12247 | 0.17224 | 0.32042 | 0.23507 | 0.13433 | 0.33116
| Rel.Std.Dev. |  | 6.882% | 1.073% | 0% | 2.284% | 0.6594% | 1.639% | 7.291% | 2.5% | 1.582% | 2.79% | 1.542% | 1.614% | 3.189% | 1.841% | 2.454% | 1.569% | 0.8164% | 0.9712% | 1.079% | 1.402% | 2.536% | 2.011% | 1.278% | 2.851%
| PDM3 |  | -20.94% | 8.632% | -6.156% | -1.647% | -2.601% | 4.082% | 3.513% | 7.21% | 0.3852% | 3.087% | -1.037% | -1.605% | -5.16% | 0.1008% | -5.018% | -2.743% | -5.762% | 0.6838% | -3.17% | 4.793% | 7.779% | -0.2634% | -10.37% | -0.8945%
| Z-Score (Absolute) |  | 3.4 | 1.61 | NA | 0.13 | 0.29 | 0.84 | 0.75 | 1.37 | 0.22 | 0.67 | 0.02 | 0.12 | 0.72 | 0.17 | 0.7 | 0.31 | 0.82 | 0.27 | 0.39 | 0.96 | 1.47 | 0.11 | 1.6 | 0
Results from laboratory 1.03 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 265 (Indicative Value 0.61 ppm) | [25 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.29 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | < 2 | < 1 | 2 | < 2 | 0.5 | < 1 | 2.62 | < 1 | < 0.5 | 1 | 1 | 0.8 | 0.5 | < 1 | < 2 | 0.3 | 1.28 | 1 | < 1.5 | < 1 |  |  |  |  | 1
| 1 | 2 | < 2 | < 1 | 2 | < 2 | < 0.5 | < 1 | 2.96 | < 1 | < 0.5 | 1 | 1 | 0.3 | < 0.3 | < 1 | < 2 | < 0.3 | 0.75 | < 1 | < 1.5 | < 1 |  |  |  |  | 7
| 1 | 3 | < 2 | < 1 | 2 | < 2 | < 0.5 | < 1 | 2.26 | < 1 | < 0.5 | 1 | < 1 | < 0.3 | < 0.3 | < 1 | < 2 | < 0.3 | 0.72 | 1 | < 1.5 | < 1 |  |  |  |  | 16
| 1 | 4 | < 2 | < 1 | 2 | < 2 | < 0.5 | < 1 | 2.28 | < 1 | < 0.5 | 1 | 1 | 0.6 | 0.8 | < 1 | < 2 | 0.5 | 0.82 | 1 | < 1.5 | < 1 |  |  |  |  | 0.611
| 1 | 5 | < 2 | < 1 | 2 | < 2 | 0.5 | < 1 | 2.31 | < 1 | < 0.5 | < 1 | 1 | 0.4 | 0.5 | < 1 | < 2 | < 0.3 | 0.82 | < 1 | < 1.5 | < 1 |  |  |  |  | 9
| 1 | 6 | < 2 | < 1 | 2 | < 2 | < 0.5 | < 1 | 2.66 | < 1 | < 0.5 | < 1 | 1 | 0.4 | 1.3 | < 1 | < 2 | 0.3 | 0.83 | < 1 | < 1.5 | < 1
| Mean |  |  |  | 2 |  | 0.5 |  | 2.515 |  |  | 1 | 1 | 0.5 | 0.775 |  |  | 0.36667 | 0.87 | 1
| Median |  |  |  | 2 |  | 0.5 |  | 2.465 |  |  | 1 | 1 | 0.4 | 0.65 |  |  | 0.3 | 0.82 | 1
| Std Dev. |  |  |  | 0 |  | 0 |  | 0.28012 |  |  | 0 | 0 | 0.2 | 0.37749 |  |  | 0.11547 | 0.20572 | 0
| Rel.Std.Dev. |  |  |  | 0% |  | 0% |  | 11.14% |  |  | 0% | 0% | 40% | 48.71% |  |  | 31.49% | 23.65% | 0%
| PDM3 |  |  |  | 227.3% |  | -18.17% |  | 311.6% |  |  | 63.67% | 63.67% | -18.17% | 26.84% |  |  | -39.99% | 42.39% | 63.67%
| Z-Score (Absolute) |  | 1.08 | 0.5 | 4.22 | 1.08 | 1.02 | 0.5 | 5.85 | 0.5 | 1.29 | 0.55 | 0.81 | 0.68 | 0.29 | 0.5 | 1.08 | 1.26 | 0.67 | 0.29 | 0.29 | 0.5
| Analytical results for Sm in OREAS 265 (Certified Value 7.42 ppm) | [25 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 7.69 | 8.11 | 7.5 | 7.1 | 7.15 | 7.43 | 7.4315 | 7.11 | 7.4 | 7.4175 | [14 empty] | 1
| 1 | 2 | 7.75 | 8.69 | 7.5 | 6.9 | 6.78 | 7.13 | 7.4206 | 7.25 | 7.6 | 7.4299 | [14 empty] | 36
| 1 | 3 | 7.57 | 8.38 | 7.3 | 7.5 | 6.94 | 7.34 | 7.4068 | 7.46 | 7.3 | 7.3397 | [14 empty] | 16
| 1 | 4 | 7.66 | 8.85 | 7.6 | 7.6 | 6.96 | 7.05 | 7.3624 | 7.34 | 7.4 | 7.3751 | [14 empty] | 7.4175
| 1 | 5 | 7.67 | 7.89 | 8 | 7.1 | 7.09 | 6.84 | 7.4016 | 7.72 | 7.4 | 7.6016 | [14 empty] | 55
| 1 | 6 | 7.66 | 7.86 | 7.5 | 7.3 | 6.77 | 7.38 | 7.4688 | 7.64 | 7.5 | 7.7137
| Mean |  | 7.6667 | 8.2967 | 7.5667 | 7.25 | 6.9483 | 7.195 | 7.4153 | 7.42 | 7.4333 | 7.4796
| Median |  | 7.665 | 8.245 | 7.5 | 7.2 | 6.95 | 7.235 | 7.4137 | 7.4 | 7.4 | 7.4237
| Std Dev. |  | 0.058195 | 0.41442 | 0.23381 | 0.26646 | 0.15562 | 0.2288 | 0.035242 | 0.23298 | 0.10328 | 0.14596
| Rel.Std.Dev. |  | 0.7591% | 4.995% | 3.09% | 3.675% | 2.24% | 3.18% | 0.4753% | 3.14% | 1.389% | 1.951%
| PDM3 |  | 3.359% | 11.85% | 2.011% | -2.258% | -6.325% | -2.999% | -0.02987% | 0.03395% | 0.2137% | 0.8373%
| Z-Score (Absolute) |  | 1.02 | 3.71 | 0.6 | 0.75 | 2.04 | 0.99 | 0.05 | 0.03 | 0.03 | 0.23
| Analytical results for Sn in OREAS 265 (Certified Value 8.82 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 9.9 | 9 | 8.9 | 9.619 | 8.7 | 9.2 | 9.3 | 8.5 | 8.9 | 8.5 | 9 | 8.2 | 8.6 | 8.3 | 8.4 | 8.4 | 8.2 | 9 | 9.18 | 8.9 | 8.2953 | 7.3325 | 8.5 |  | 1
| 1 | 2 | 9.7 | 9.2 | 8.8 | 8.1005 | 8.9 | 9.2 | 9.6 | 8.2 | 9.2 | 8.4 | 8.9 | 8.3 | 8.3 | 8.1 | 8.4 | 8.8 | 8.37 | 9.1 | 9.62 | 9 | 8.1317 | 7.25 | 8.9 |  | 11
| 1 | 3 | 9.4 | 9.2 | 8.7 | 8.1353 | 9.4 | 9 | 9.9 | 8 | 9 | 8.3 | 9.1 | 8.1 | 8.6 | 8.3 | 8.3 | 9.3 | 8.67 | 9 | 9.62 | 8.9 | 8.3993 | 7.4214 | 8.6 |  | 16
| 1 | 4 | 10 | 9.2 | 8.8 | 8.7646 | 8.8 | 9.3 | 10.2 | 8.1 | 9 | 8.5 | 9.3 | 8.1 | 8.4 | 8.4 | 8.5 | 8.6 | 8.43 | 8.9 | 9.56 | 8.9 | 8.3144 | 7.1825 | 8.8 |  | 8.8218
| 1 | 5 | 9.2 | 9.1 | 8.8 | 8.7378 | 9.4 | 9.3 | 9.5 | 8.1 | 9.3 | 8.5 | 9 | 8.1 | 8.7 | 8.3 | 8.5 | 8.5 | 8.88 | 8.9 | 9.24 | 8.8 | 8.5046 | 7.3547 | 9 |  | 56
| 1 | 6 | 8.9 | 9 | 8.6 | 9.2584 | 9.2 | 9.1 | 9 | 8.3 | 9.7 | 8.5 | 8.8 | 8.5 | 8.9 | 8.2 | 8.1 | 8.9 | 8.85 | 9 | 11.6 | 9 | 8.4467 | 7.424 | 8.9
| Mean |  | 9.5167 | 9.1167 | 8.7667 | 8.7692 | 9.0667 | 9.1833 | 9.5833 | 8.2 | 9.1833 | 8.45 | 9.0167 | 8.2167 | 8.5833 | 8.2667 | 8.3667 | 8.75 | 8.5667 | 8.9833 | 9.8033 | 8.9167 | 8.3487 | 7.3275 | 8.7833
| Median |  | 9.55 | 9.15 | 8.8 | 8.7512 | 9.05 | 9.2 | 9.55 | 8.15 | 9.1 | 8.5 | 9 | 8.15 | 8.6 | 8.3 | 8.4 | 8.7 | 8.55 | 9 | 9.59 | 8.9 | 8.3569 | 7.3436 | 8.85
| Std Dev. |  | 0.42622 | 0.098319 | 0.10328 | 0.6019 | 0.30768 | 0.1169 | 0.42622 | 0.17889 | 0.29269 | 0.083666 | 0.17224 | 0.16021 | 0.2137 | 0.10328 | 0.15055 | 0.32711 | 0.27602 | 0.075277 | 0.90115 | 0.075277 | 0.13238 | 0.095839 | 0.19408
| Rel.Std.Dev. |  | 4.479% | 1.078% | 1.178% | 6.864% | 3.394% | 1.273% | 4.448% | 2.182% | 3.187% | 0.9901% | 1.91% | 1.95% | 2.49% | 1.249% | 1.799% | 3.738% | 3.222% | 0.838% | 9.192% | 0.8442% | 1.586% | 1.308% | 2.21%
| PDM3 |  | 7.877% | 3.343% | -0.6245% | -0.5953% | 2.776% | 4.099% | 8.633% | -7.048% | 4.099% | -4.214% | 2.209% | -6.859% | -2.703% | -6.292% | -5.159% | -0.8134% | -2.892% | 1.832% | 11.13% | 1.076% | -5.362% | -16.94% | -0.4355%
| Z-Score (Absolute) |  | 1.45 | 0.67 | 0 | 0 | 0.58 | 0.8 | 1.58 | 1.1 | 0.8 | 0.62 | 0.48 | 1.07 | 0.36 | 0.98 | 0.78 | 0.04 | 0.39 | 0.42 | 2.01 | 0.29 | 0.82 | 2.8 | 0.03
| Analytical results for Sr in OREAS 265 (Certified Value 261 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 240.5 | 271 | 259 | 263.0418 | 259.36 | 263 | 277 | 268 | 257.32 | 265 | 264 | 247 | 275 | 250.5 | 265.2 | 259 | 256.63 | 304 | 257 | 283 | 267 | 259.1679 | 253 | 254.4 | 1
| 1 | 2 | 229.5 | 279 | 257 | 256.7801 | 257.92 | 259 | 265 | 265 | 258.15 | 253 | 263 | 258 | 271 | 248 | 263.6 | 261 | 256.9 | 305 | 263 | 289 | 264 | 255.7011 | 236 | 248.3 | 12
| 1 | 3 | 227.5 | 275 | 259 | 246.271 | 260.99 | 256 | 266 | 263 | 258.37 | 261 | 266 | 254 | 267 | 253.3 | 264.4 | 264 | 256.52 | 308 | 262 | 284 | 262 | 259.4709 | 249 | 250.7 | 16
| 1 | 4 | 217.6 | 271 | 259 | 253.7863 | 255.83 | 254 | 272 | 255 | 256.94 | 261 | 271 | 259 | 291 | 252.9 | 266.2 | 269 | 256.45 | 298 | 263 | 284 | 268 | 259.199 | 246 | 264.4 | 260.7875
| 1 | 5 | 218 | 279 | 256 | 249.5221 | 255.8 | 261 | 259 | 256 | 260.84 | 258 | 271 | 259 | 273 | 250.5 | 265.3 | 261 | 256.83 | 306 | 261 | 283 | 259 | 265.3748 | 242 | 257.1 | 57
| 1 | 6 | 215.7 | 274 | 258 | 260.8299 | 261.1 | 260 | 257 | 267 | 258.49 | 262 | 255 | 259 | 280 | 245.2 | 264 | 273 | 256.09 | 308 | 261 | 297 | 269 | 271.9531 | 256 | 259.4
| Mean |  | 224.8 | 274.8333 | 258 | 255.0385 | 258.5 | 258.8333 | 266 | 262.3333 | 258.3517 | 260 | 265 | 256 | 276.1667 | 250.0667 | 264.7833 | 264.5 | 256.57 | 304.8333 | 261.1667 | 286.6667 | 264.8333 | 261.8111 | 247 | 255.7167
| Median |  | 222.75 | 274.5 | 258.5 | 255.2832 | 258.64 | 259.5 | 265.5 | 264 | 258.26 | 261 | 265 | 258.5 | 274 | 250.5 | 264.8 | 262.5 | 256.575 | 305.5 | 261.5 | 284 | 265.5 | 259.335 | 247.5 | 255.75
| Std Dev. |  | 9.5578 | 3.6009 | 1.2649 | 6.4729 | 2.3856 | 3.3116 | 7.5895 | 5.5737 | 1.3645 | 4.0988 | 5.9666 | 4.8166 | 8.4479 | 3.0598 | 0.96003 | 5.4314 | 0.29237 | 3.7103 | 2.2286 | 5.5377 | 3.8687 | 5.8682 | 7.3212 | 5.8786
| Rel.Std.Dev. |  | 4.252% | 1.31% | 0.4903% | 2.538% | 0.9229% | 1.279% | 2.853% | 2.125% | 0.5281% | 1.576% | 2.252% | 1.881% | 3.059% | 1.224% | 0.3626% | 2.053% | 0.114% | 1.217% | 0.8533% | 1.932% | 1.461% | 2.241% | 2.964% | 2.299%
| PDM3 |  | -13.8% | 5.386% | -1.069% | -2.204% | -0.8772% | -0.7493% | 1.999% | 0.5927% | -0.934% | -0.302% | 1.615% | -1.836% | 5.897% | -4.111% | 1.532% | 1.424% | -1.617% | 16.89% | 0.1454% | 9.923% | 1.551% | 0.3925% | -5.287% | -1.944%
| Z-Score (Absolute) |  | 5.57 | 2.22 | 0.4 | 0.86 | 0.32 | 0.27 | 0.84 | 0.27 | 0.35 | 0.09 | 0.69 | 0.71 | 2.42 | 1.64 | 0.65 | 0.61 | 0.62 | 6.89 | 0.09 | 4.06 | 0.66 | 0.19 | 2.11 | 0.76
| Analytical results for Ta in OREAS 265 (Certified Value 1.65 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.63 | 1.79 | 1.66 | 0.91896 | 1.79 | 1.63 | 1.24 | 1.7 | 1.66 | 1.71 | 1.76 | 1.7 | 1.5 | 1.49 | 1.39 | 1.69 | 1.44 | 1.82 | 1.64 | 1.6 | 1.7191 | 1.75 |  |  | 1
| 1 | 2 | 1.54 | 1.83 | 1.64 | 0.80227 | 1.82 | 1.68 | 1.35 | 1.65 | 1.64 | 1.64 | 1.74 | 1.7 | 1.5 | 1.45 | 1.48 | 1.81 | 1.53 | 1.82 | 1.7 | 1.6 | 1.7079 | 1.74 |  |  | 22
| 1 | 3 | 1.59 | 1.85 | 1.75 | 0.87103 | 1.77 | 1.66 | 1.28 | 1.62 | 1.67 | 1.65 | 1.75 | 1.6 | 1.5 | 1.47 | 1.44 | 1.79 | 1.47 | 1.8 | 1.65 | 1.6 | 1.7099 | 1.77 |  |  | 16
| 1 | 4 | 1.22 | 1.84 | 1.75 | 0.79911 | 1.79 | 1.63 | 1.66 | 1.63 | 1.69 | 1.73 | 1.78 | 1.7 | 1.6 | 1.44 | 1.46 | 1.74 | 1.42 | 1.74 | 1.67 | 1.6 | 1.7114 | 1.85 |  |  | 1.65
| 1 | 5 | 1.37 | 1.8 | 1.67 | 0.9045 | 1.96 | 1.58 | 1.66 | 1.63 | 1.79 | 1.66 | 1.75 | 1.8 | 1.5 | 1.43 | 1.28 | 1.79 | 1.5 | 1.72 | 1.64 | 1.6 | 1.7785 | 1.85 |  |  | 58
| 1 | 6 | 1.09 | 1.76 | 1.65 | 0.82213 | 1.88 | 1.62 | 2.12 | 1.66 | 1.64 | 1.68 | 1.72 | 1.7 | 1.7 | 1.38 | 1.32 | 1.86 | 1.47 | 1.73 | 1.65 | 1.6 | 1.8005 | 1.81
| Mean |  | 1.4067 | 1.8117 | 1.6867 | 0.853 | 1.835 | 1.6333 | 1.5517 | 1.6483 | 1.6817 | 1.6783 | 1.75 | 1.7 | 1.55 | 1.4433 | 1.395 | 1.78 | 1.4717 | 1.7717 | 1.6583 | 1.6 | 1.7379 | 1.795
| Median |  | 1.455 | 1.815 | 1.665 | 0.84658 | 1.805 | 1.63 | 1.505 | 1.64 | 1.665 | 1.67 | 1.75 | 1.7 | 1.5 | 1.445 | 1.415 | 1.79 | 1.47 | 1.77 | 1.65 | 1.6 | 1.7152 | 1.79
| Std Dev. |  | 0.21805 | 0.034303 | 0.050067 | 0.052456 | 0.072319 | 0.034448 | 0.33409 | 0.029269 | 0.056362 | 0.035449 | 0.02 | 0.063246 | 0.083666 | 0.037771 | 0.080436 | 0.058652 | 0.039707 | 0.046655 | 0.023166 | 2.4324e-16 | 0.040756 | 0.048888
| Rel.Std.Dev. |  | 15.5% | 1.893% | 2.968% | 6.15% | 3.941% | 2.109% | 21.53% | 1.776% | 3.352% | 2.112% | 1.143% | 3.72% | 5.398% | 2.617% | 5.766% | 3.295% | 2.698% | 2.633% | 1.397% | 1.52e-14% | 2.345% | 2.724%
| PDM3 |  | -14.75% | 9.799% | 2.224% | -48.3% | 11.21% | -1.009% | -5.958% | -0.09968% | 1.921% | 1.719% | 6.062% | 3.032% | -6.059% | -12.52% | -15.45% | 7.88% | -10.81% | 7.375% | 0.5064% | -3.029% | 5.327% | 8.789%
| Z-Score (Absolute) |  | 1.64 | 0.9 | 0.11 | 5.11 | 1.05 | 0.22 | 0.73 | 0.13 | 0.08 | 0.06 | 0.51 | 0.2 | 0.74 | 1.41 | 1.71 | 0.7 | 1.23 | 0.65 | 0.06 | 0.43 | 0.44 | 0.79
| Analytical results for Tb in OREAS 265 (Certified Value 0.76 ppm) | [25 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [11 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.81 | 0.76 | 0.76 | 0.79 | 0.7 | 0.6 | 0.76 | 0.76 | 0.79 | 0.79781 | 1.6 | 0.77 | 0.72557 | [11 empty] | 1
| 1 | 2 | 0.74 | 0.8 | 0.75 | 0.79 | 0.8 | 0.7 | 0.72 | 0.76 | 0.77 | 0.80217 | 1.5 | 0.79 | 0.72424 | [11 empty] | 23
| 1 | 3 | 0.72 | 0.79 | 0.74 | 0.74 | 0.8 | 0.6 | 0.73 | 0.74 | 0.78 | 0.79596 | 1 | 0.76 | 0.71982 | [11 empty] | 16
| 1 | 4 | 0.71 | 0.76 | 0.75 | 0.85 | 0.8 | 0.7 | 0.74 | 0.77 | 0.79 | 0.80721 | 1 | 0.81 | 0.70662 | [11 empty] | 0.76366
| 1 | 5 | 0.72 | 0.77 | 0.77 | 0.79 | 0.8 | 0.6 | 0.75 | 0.75 | 0.73 | 0.81867 | 1 | 0.77 | 0.73669 | [11 empty] | 59
| 1 | 6 | 0.7 | 0.77 | 0.75 | 0.75 | 0.8 | 0.7 | 0.74 | 0.73 | 0.82 | 0.80855 | 1.1 | 0.82 | 0.74615
| Mean |  | 0.73333 | 0.775 | 0.75333 | 0.785 | 0.78333 | 0.65 | 0.74 | 0.75167 | 0.78 | 0.80506 | 1.2 | 0.78667 | 0.72652
| Median |  | 0.72 | 0.77 | 0.75 | 0.79 | 0.8 | 0.65 | 0.74 | 0.755 | 0.785 | 0.80469 | 1.05 | 0.78 | 0.72491
| Std Dev. |  | 0.039833 | 0.016432 | 0.010328 | 0.038859 | 0.040825 | 0.054772 | 0.014142 | 0.01472 | 0.029665 | 0.0083142 | 0.27568 | 0.024221 | 0.013682
| Rel.Std.Dev. |  | 5.432% | 2.12% | 1.371% | 4.95% | 5.212% | 8.427% | 1.911% | 1.958% | 3.803% | 1.033% | 22.97% | 3.079% | 1.883%
| PDM3 |  | -3.971% | 1.485% | -1.352% | 2.795% | 2.576% | -14.88% | -3.098% | -1.57% | 2.14% | 5.422% | 57.14% | 3.013% | -4.864%
| Z-Score (Absolute) |  | 0.96 | 0.34 | 0.34 | 0.65 | NA | NA | 0.75 | 0.39 | 0.49 | 1.27 | NA | 0.7 | 1.17
Results from laboratories 1.12, 1.13 and 1.18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 265 (Certified Value 0.06 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.29 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.2 | 0.05 | 0.14 | 0.05 | < 0.2 | < 0.05 | 0.07 | 0.09 | < 0.05 | 0.07 | 0.06 | < 0.05 | 0.07 | 0.06 | < 0.1 | < 0.2 | 0.12 |  |  |  | 1
| 1 | 2 | 0.06 | < 0.05 | < 0.1 | < 0.1 | < 0.2 | 0.05 | 0.31 | 0.06 | < 0.2 | 0.05 | 0.06 | 0.06 | < 0.05 | 0.07 | 0.08 | < 0.05 | 0.05 | < 0.05 | < 0.1 | < 0.2 | 0.11 |  |  |  | 24
| 1 | 3 | 0.07 | < 0.05 | < 0.1 | < 0.1 | < 0.2 | < 0.05 | 0.25 | < 0.05 | < 0.2 | < 0.05 | 0.06 | < 0.05 | < 0.05 | 0.05 | 0.06 | < 0.05 | 0.06 | 0.05 | < 0.1 | < 0.2 | 0.13 |  |  |  | 16
| 1 | 4 | 0.07 | 0.06 | < 0.1 | < 0.1 | < 0.2 | < 0.05 | 0.3 | < 0.05 | < 0.2 | < 0.05 | 0.06 | 0.06 | 0.1 | 0.05 | 0.07 | < 0.05 | 0.05 | 0.06 | < 0.1 | < 0.2 | 0.13 |  |  |  | 0.059722
| 1 | 5 | < 0.05 | 0.05 | < 0.1 | < 0.1 | < 0.2 | < 0.05 | 0.09 | < 0.05 | < 0.2 | 0.06 | 0.05 | 0.06 | < 0.05 | 0.06 | 0.07 | 0.06 | 0.05 | 0.08 | < 0.1 | < 0.2 | 0.13 |  |  |  | 60
| 1 | 6 | 0.08 | 0.07 | < 0.1 | < 0.1 | < 0.2 | < 0.05 | 0.28 | < 0.05 | < 0.2 | 0.05 | 0.06 | 0.06 | 0.06 | 0.07 | 0.07 | < 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.2 | 0.14
| Mean |  | 0.076 | 0.0575 |  |  |  | 0.05 | 0.22833 | 0.055 |  | 0.053333 | 0.06 | 0.066 | 0.08 | 0.061667 | 0.068333 | 0.06 | 0.058333 | 0.0625 |  |  | 0.12667
| Median |  | 0.07 | 0.055 |  |  |  | 0.05 | 0.265 | 0.055 |  | 0.05 | 0.06 | 0.06 | 0.08 | 0.065 | 0.07 | 0.06 | 0.055 | 0.06 |  |  | 0.13
| Std Dev. |  | 0.015166 | 0.0095743 |  |  |  | 0 | 0.091524 | 0.0070711 |  | 0.0057735 | 0.0063246 | 0.013416 | 0.028284 | 0.0098319 | 0.0075277 |  | 0.0098319 | 0.012583 |  |  | 0.010328
| Rel.Std.Dev. |  | 19.95% | 16.65% |  |  |  | 0% | 40.08% | 12.86% |  | 10.83% | 10.54% | 20.33% | 35.36% | 15.94% | 11.02% |  | 16.85% | 20.13% |  |  | 8.154%
| PDM3 |  | 27.26% | -3.721% |  |  |  | -16.28% | 282.3% | -7.907% |  | -10.7% | 0.4651% | 10.51% | 33.95% | 3.256% | 14.42% | 0.4651% | -2.326% | 4.651% |  |  | 112.1%
| Z-Score (Absolute) |  | 0.53 | 0.67 | 0.48 | 0.48 | 2.41 | 1.44 | 9.83 | 1.35 | 2.41 | 1.11 | 0.1 | 0.05 | 0.87 | 0.19 | 0.58 | 1.59 | 0 | 0.48 | 0.48 | 2.41 | 3.95
| Analytical results for Th in OREAS 265 (Certified Value 13.6 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 13.8 | 14.8 | 13 | 14.1469 | 13.63 | 13.72 | 14.07 | 14.25 | 14.27 | 14.4 | 13.9 | 14.1 | 11.7 | 12.16 | 13.8 | 13.1 | 12.7 | 13.35 | 14.4 | 14.5 | 13.7891 | 18.35 | 12.3 |  | 1
| 1 | 2 | 12.8 | 15.15 | 13.1 | 13.8704 | 13.43 | 13.75 | 14.58 | 14 | 14.12 | 14.05 | 13.55 | 14.4 | 12 | 12.44 | 13.5 | 13.1 | 13 | 13.15 | 14.7 | 14.3 | 13.6182 | 19 | 12.8 |  | 25
| 1 | 3 | 12.1 | 14.6 | 12.9 | 13.3544 | 13.59 | 13.98 | 14.02 | 13.85 | 14.12 | 14.15 | 14.1 | 14.1 | 11.5 | 11.98 | 13.6 | 13.2 | 13.26 | 13.8 | 14.2 | 14.4 | 13.9774 | 18.553 | 12.5 |  | 16
| 1 | 4 | 12.3 | 14.05 | 12.7 | 13.5062 | 13.9 | 13.17 | 14.62 | 13.7 | 13.99 | 14.8 | 14.2 | 14.2 | 13.8 | 11.97 | 13.8 | 13.2 | 12.79 | 12.95 | 14.6 | 14.9 | 13.7703 | 18.782 | 13 |  | 13.5889
| 1 | 5 | 12.4 | 14.8 | 12.8 | 13.6759 | 13.53 | 13.35 | 13.73 | 13.8 | 13.96 | 14.15 | 14.35 | 14.9 | 12.7 | 12.1 | 13.6 | 13 | 13.57 | 12.55 | 14.6 | 14 | 14.415 | 18.02 | 13.1 |  | 61
| 1 | 6 | 11.9 | 14.4 | 12.7 | 14.1913 | 13.43 | 14.06 | 13.78 | 14.1 | 14.17 | 14.25 | 13.45 | 14.3 | 13.6 | 12.09 | 13.4 | 13.6 | 13.45 | 13.35 | 14.5 | 13.8 | 13.9145 | 18.42 | 13
| Mean |  | 12.55 | 14.6333 | 12.8667 | 13.7909 | 13.585 | 13.6717 | 14.1333 | 13.95 | 14.105 | 14.3 | 13.925 | 14.3333 | 12.55 | 12.1233 | 13.6167 | 13.2 | 13.1283 | 13.1917 | 14.5 | 14.3167 | 13.9141 | 18.5208 | 12.7833
| Median |  | 12.35 | 14.7 | 12.85 | 13.7732 | 13.56 | 13.735 | 14.045 | 13.925 | 14.12 | 14.2 | 14 | 14.25 | 12.35 | 12.095 | 13.6 | 13.15 | 13.13 | 13.25 | 14.55 | 14.35 | 13.8518 | 18.4865 | 12.9
| Std Dev. |  | 0.68337 | 0.37903 | 0.1633 | 0.34007 | 0.17456 | 0.34914 | 0.3849 | 0.20494 | 0.11502 | 0.27203 | 0.36159 | 0.30111 | 0.98133 | 0.17166 | 0.16021 | 0.20976 | 0.35504 | 0.42239 | 0.17889 | 0.38687 | 0.27522 | 0.34323 | 0.31885
| Rel.Std.Dev. |  | 5.445% | 2.59% | 1.269% | 2.466% | 1.285% | 2.554% | 2.723% | 1.469% | 0.8155% | 1.902% | 2.597% | 2.101% | 7.819% | 1.416% | 1.177% | 1.589% | 2.704% | 3.202% | 1.234% | 2.702% | 1.978% | 1.853% | 2.494%
| PDM3 |  | -7.645% | 7.686% | -5.315% | 1.486% | -0.0284% | 0.6094% | 4.007% | 2.658% | 3.798% | 5.233% | 2.474% | 5.479% | -7.645% | -10.78% | 0.2046% | -2.862% | -3.389% | -2.923% | 6.705% | 5.356% | 2.393% | 36.29% | -5.928%
| Z-Score (Absolute) |  | 1.54 | 1.05 | 1.15 | 0 | 0.26 | 0.15 | 0.43 | 0.2 | 0.39 | 0.63 | 0.17 | 0.67 | 1.54 | 2.07 | 0.22 | 0.73 | 0.82 | 0.74 | 0.88 | 0.65 | 0.15 | 5.88 | 1.25
| Analytical results for Ti in OREAS 265 (Certified Value 0.454 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.49 | 0.47 | 0.429 | 0.48991 | 0.451 | 0.486 | 0.44 | 0.465 | 0.462 | 0.461 | 0.46 | 0.47 | 0.437 | 0.44 | 0.41 | 0.45 | 0.44863 | 0.4117 | 0.423 | 0.522 | 0.433 | 0.44221 | 0.3816 | 0.46 | 1
| 1 | 2 | 0.48 | 0.481 | 0.431 | 0.45465 | 0.4608 | 0.477 | 0.44 | 0.465 | 0.4623 | 0.451 | 0.46 | 0.462 | 0.443 | 0.445 | 0.42 | 0.47 | 0.44615 | 0.4112 | 0.429 | 0.528 | 0.451 | 0.44083 | 0.3815 | 0.46 | 26
| 1 | 3 | 0.47 | 0.484 | 0.425 | 0.45688 | 0.445 | 0.466 | 0.4 | 0.452 | 0.4475 | 0.452 | 0.466 | 0.45 | 0.429 | 0.453 | 0.41 | 0.47 | 0.44776 | 0.4162 | 0.431 | 0.522 | 0.433 | 0.44973 | 0.3956 | 0.45 | 16
| 1 | 4 | 0.49 | 0.468 | 0.427 | 0.47869 | 0.4541 | 0.465 | 0.45 | 0.447 | 0.4525 | 0.458 | 0.468 | 0.453 | 0.47 | 0.456 | 0.42 | 0.48 | 0.44727 | 0.4035 | 0.428 | 0.522 | 0.43 | 0.44397 | 0.399 | 0.47 | 0.4542
| 1 | 5 | 0.45 | 0.478 | 0.431 | 0.47933 | 0.4811 | 0.44 | 0.41 | 0.451 | 0.4802 | 0.454 | 0.467 | 0.473 | 0.448 | 0.448 | 0.4 | 0.47 | 0.4468 | 0.4133 | 0.424 | 0.516 | 0.435 | 0.45757 | 0.3855 | 0.46 | 62
| 1 | 6 | 0.43 | 0.471 | 0.425 | 0.48046 | 0.473 | 0.44 | 0.42 | 0.456 | 0.4746 | 0.46 | 0.449 | 0.465 | 0.466 | 0.447 | 0.4 | 0.49 | 0.44514 | 0.4139 | 0.418 | 0.546 | 0.445 | 0.46319 | 0.3822 | 0.47
| Mean |  | 0.46833 | 0.47533 | 0.428 | 0.47332 | 0.46083 | 0.46233 | 0.42667 | 0.456 | 0.46318 | 0.456 | 0.46167 | 0.46217 | 0.44883 | 0.44817 | 0.41 | 0.47167 | 0.44696 | 0.41163 | 0.4255 | 0.526 | 0.43783 | 0.44958 | 0.38757 | 0.46167
| Median |  | 0.475 | 0.4745 | 0.428 | 0.47901 | 0.45745 | 0.4655 | 0.43 | 0.454 | 0.46215 | 0.456 | 0.463 | 0.4635 | 0.4455 | 0.4475 | 0.41 | 0.47 | 0.44703 | 0.4125 | 0.426 | 0.522 | 0.434 | 0.44685 | 0.38385 | 0.46
| Std Dev. |  | 0.024014 | 0.0065625 | 0.0027568 | 0.014212 | 0.013795 | 0.018938 | 0.019664 | 0.0075366 | 0.012508 | 0.0042426 | 0.0071181 | 0.0091524 | 0.016192 | 0.0057067 | 0.0089443 | 0.013292 | 0.0012268 | 0.0043615 | 0.0047645 | 0.010507 | 0.0082563 | 0.0090724 | 0.0077554 | 0.0075277
| Rel.Std.Dev. |  | 5.128% | 1.381% | 0.6441% | 3.003% | 2.993% | 4.096% | 4.609% | 1.653% | 2.7% | 0.9304% | 1.542% | 1.98% | 3.607% | 1.273% | 2.182% | 2.818% | 0.2745% | 1.06% | 1.12% | 1.998% | 1.886% | 2.018% | 2.001% | 1.631%
| PDM3 |  | 3.111% | 4.652% | -5.769% | 4.209% | 1.46% | 1.79% | -6.062% | 0.3958% | 1.977% | 0.3958% | 1.643% | 1.754% | -1.182% | -1.329% | -9.732% | 3.845% | -1.595% | -9.372% | -6.319% | 15.81% | -3.604% | -1.017% | -14.67% | 1.643%
| Z-Score (Absolute) |  | 0.78 | 1.22 | 1.77 | 1.09 | 0.3 | 0.4 | 1.85 | 0 | 0.45 | 0 | 0.36 | 0.39 | 0.45 | 0.49 | 2.9 | 0.99 | 0.57 | 2.8 | 1.92 | 4.39 | 1.15 | 0.4 | 4.32 | 0.36
| Analytical results for Tl in OREAS 265 (Certified Value 0.49 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.55 | 0.51 | 0.5 | 0.50862 | 0.51 | 0.5 | 0.51 | 0.47 | 0.52 | 0.44 | 0.49 | 0.5 | 0.46 | 0.38 | 0.48 | 0.47 | 0.55254 | 0.481 | 0.56 | 0.5 | 0.47218 | 0.47 |  |  | 1
| 1 | 2 | 0.53 | 0.54 | 0.48 | 0.5082 | 0.52 | 0.52 | 0.48 | 0.46 | 0.52 | 0.46 | 0.46 | 0.56 | 0.47 | 0.35 | 0.47 | 0.48 | 0.52455 | 0.484 | 0.52 | 0.5 | 0.45579 | 0.47 |  |  | 27
| 1 | 3 | 0.52 | 0.5 | 0.48 | 0.49081 | 0.53 | 0.51 | 0.51 | 0.46 | 0.5 | 0.46 | 0.48 | 0.56 | 0.43 | 0.36 | 0.48 | 0.48 | 0.52 | 0.499 | 0.55 | 0.5 | 0.46146 | 0.46 |  |  | 16
| 1 | 4 | 0.53 | 0.49 | 0.5 | 0.49449 | 0.51 | 0.51 | 0.51 | 0.47 | 0.5 | 0.46 | 0.48 | 0.55 | 0.38 | 0.37 | 0.48 | 0.48 | 0.50164 | 0.493 | 0.54 | 0.5 | 0.46376 | 0.48 |  |  | 0.49178
| 1 | 5 | 0.51 | 0.52 | 0.5 | 0.49341 | 0.5 | 0.52 | 0.48 | 0.46 | 0.48 | 0.43 | 0.48 | 0.58 | 0.45 | 0.38 | 0.45 | 0.48 | 0.50977 | 0.508 | 0.49 | 0.4 | 0.47964 | 0.49 |  |  | 63
| 1 | 6 | 0.5 | 0.51 | 0.49 | 0.51235 | 0.51 | 0.52 | 0.48 | 0.46 | 0.52 | 0.43 | 0.49 | 0.57 | 0.38 | 0.35 | 0.48 | 0.49 | 0.51812 | 0.517 | 0.49 | 0.5 | 0.471 | 0.49
| Mean |  | 0.52333 | 0.51167 | 0.49167 | 0.50131 | 0.51333 | 0.51333 | 0.495 | 0.46333 | 0.50667 | 0.44667 | 0.48 | 0.55333 | 0.42833 | 0.365 | 0.47333 | 0.48 | 0.5211 | 0.497 | 0.525 | 0.48333 | 0.4673 | 0.47667
| Median |  | 0.525 | 0.51 | 0.495 | 0.50135 | 0.51 | 0.515 | 0.495 | 0.46 | 0.51 | 0.45 | 0.48 | 0.56 | 0.44 | 0.365 | 0.48 | 0.48 | 0.51906 | 0.496 | 0.53 | 0.5 | 0.46738 | 0.475
| Std Dev. |  | 0.017512 | 0.017224 | 0.0098319 | 0.0094012 | 0.010328 | 0.008165 | 0.016432 | 0.005164 | 0.01633 | 0.015055 | 0.010954 | 0.028048 | 0.039707 | 0.013784 | 0.012111 | 0.0063246 | 0.017421 | 0.0139 | 0.030166 | 0.040825 | 0.0085885 | 0.012111
| Rel.Std.Dev. |  | 3.346% | 3.366% | 2% | 1.875% | 2.012% | 1.591% | 3.32% | 1.115% | 3.223% | 3.371% | 2.282% | 5.069% | 9.27% | 3.776% | 2.559% | 1.318% | 3.343% | 2.797% | 5.746% | 8.447% | 1.838% | 2.541%
| PDM3 |  | 6.415% | 4.043% | -0.02379% | 1.938% | 4.382% | 4.382% | 0.654% | -5.785% | 3.026% | -9.174% | -2.396% | 12.52% | -12.9% | -25.78% | -3.752% | -2.396% | 5.962% | 1.061% | 6.754% | -1.718% | -4.978% | -3.074%
| Z-Score (Absolute) |  | 1.04 | 0.61 | 0.12 | 0.23 | 0.67 | 0.67 | 0 | 1.16 | 0.43 | 1.78 | 0.55 | 2.15 | 2.45 | 4.78 | 0.8 | 0.55 | 0.96 | 0.07 | 1.1 | NA | 1.02 | 0.67
Results from laboratory 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 265 (Certified Value 0.24 ppm) | [25 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.05 | 1.07 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.22 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.26 | 0.23 | 0.3 | 0.2 | 0.25 | 0.25 | 0.24 | 0.24294 | [16 empty] | 1
| 1 | 2 | 0.25 | 0.2 | 0.2 | 0.2 | 0.25 | 0.24 | 0.24 | 0.23638 | [16 empty] | 28
| 1 | 3 | 0.26 | 0.2 | 0.2 | 0.2 | 0.27 | 0.26 | 0.24 | 0.24081 | [16 empty] | 16
| 1 | 4 | 0.29 | 0.21 | 0.2 | 0.2 | 0.24 | 0.26 | 0.24 | 0.23786 | [16 empty] | 0.24308
| 1 | 5 | 0.27 | 0.23 | 0.2 | 0.2 | 0.27 | 0.24 | 0.25 | 0.24758 | [16 empty] | 64
| 1 | 6 | 0.27 | 0.19 | 0.3 | 0.2 | 0.25 | 0.95 | 0.26 | 0.24396
| Mean |  | 0.26667 | 0.21 | 0.23333 | 0.2 | 0.255 | 0.36667 | 0.245 | 0.24159
| Median |  | 0.265 | 0.205 | 0.2 | 0.2 | 0.25 | 0.255 | 0.24 | 0.24188
| Std Dev. |  | 0.013663 | 0.016733 | 0.05164 | 3.0405e-17 | 0.012247 | 0.28591 | 0.0083666 | 0.0041236
| Rel.Std.Dev. |  | 5.123% | 7.968% | 22.13% | 1.52e-14% | 4.803% | 77.98% | 3.415% | 1.707%
| PDM3 |  | 9.701% | -13.61% | -4.011% | -17.72% | 4.902% | 50.84% | 0.7882% | -0.6154%
| Z-Score (Absolute) |  | 0.9 | 1.28 | 0.38 | 1.66 | 0.45 | 4.74 | 0.07 | 0.07
| Analytical results for U in OREAS 265 (Certified Value 2.53 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.6 | 2.6 | 2.5 | 2.7433 | 2.71 | 2.6 | 2.56 | 2.5 | 2.61 | 2.4 | 2.4 | 2.5 | 2.7 | 2.5 | 2.6 | 2.6 | 2.3885 | 2.44 | 2.6 | 2.9 | 2.3 | < 5 | < 20 | 2.4 | 1
| 1 | 2 | 2.5 | 2.7 | 2.47 | 2.6817 | 2.62 | 2.6 | 2.68 | 2.5 | 2.56 | 2.4 | 2.4 | 2.7 | 2.6 | 2.4 | 2.6 | 2.5 | 2.4097 | 2.49 | 2.6 | 2.78 | 2.3 | < 5 | < 20 | 2.4 | 29
| 1 | 3 | 2.5 | 2.9 | 2.54 | 2.6295 | 2.61 | 2.55 | 2.5 | 2.5 | 2.49 | 2.4 | 2.4 | 2.6 | 2.4 | 2.4 | 2.5 | 2.6 | 2.361 | 2.53 | 2.6 | 2.88 | 2.6 | < 5 | < 20 | 2.5 | 16
| 1 | 4 | 2.4 | 2.6 | 2.59 | 2.6253 | 2.69 | 2.54 | 2.62 | 2.4 | 2.55 | 2.5 | 2.5 | 2.6 | 2.4 | 2.5 | 2.6 | 2.5 | 2.4024 | 2.47 | 2.4 | 2.95 | 2.3 | < 5 | < 20 | 2.6 | 2.525
| 1 | 5 | 2.4 | 2.7 | 2.39 | 2.7405 | 2.72 | 2.54 | 2.57 | 2.4 | 2.61 | 2.4 | 2.5 | 2.8 | 2.4 | 2.5 | 2.5 | 2.5 | 2.3299 | 2.58 | 2.4 | 2.76 | 2.2 | < 5 | < 20 | 2.6 | 65
| 1 | 6 | 2.4 | 2.6 | 2.39 | 2.6896 | 2.64 | 2.56 | 2.48 | 2.5 | 2.6 | 2.4 | 2.4 | 2.7 | 2.4 | 2.4 | 2.5 | 2.6 | 2.3643 | 2.56 | 2.5 | 3.01 | 2.3 | < 5 | < 20 | 2.5
| Mean |  | 2.4667 | 2.6833 | 2.48 | 2.685 | 2.665 | 2.565 | 2.5683 | 2.4667 | 2.57 | 2.4167 | 2.4333 | 2.65 | 2.4833 | 2.45 | 2.55 | 2.55 | 2.376 | 2.5117 | 2.5167 | 2.88 | 2.3333 |  |  | 2.5
| Median |  | 2.45 | 2.65 | 2.485 | 2.6857 | 2.665 | 2.555 | 2.565 | 2.5 | 2.58 | 2.4 | 2.4 | 2.65 | 2.4 | 2.45 | 2.55 | 2.55 | 2.3764 | 2.51 | 2.55 | 2.89 | 2.3 |  |  | 2.5
| Std Dev. |  | 0.08165 | 0.1169 | 0.080498 | 0.051301 | 0.047645 | 0.028107 | 0.074409 | 0.05164 | 0.046904 | 0.040825 | 0.05164 | 0.10488 | 0.13292 | 0.054772 | 0.054772 | 0.054772 | 0.029917 | 0.054191 | 0.098319 | 0.09654 | 0.13663 |  |  | 0.089443
| Rel.Std.Dev. |  | 3.31% | 4.357% | 3.246% | 1.911% | 1.788% | 1.096% | 2.897% | 2.094% | 1.825% | 1.689% | 2.122% | 3.958% | 5.352% | 2.236% | 2.148% | 2.148% | 1.259% | 2.158% | 3.907% | 3.352% | 5.855% |  |  | 3.578%
| PDM3 |  | -2.312% | 6.269% | -1.784% | 6.335% | 5.543% | 1.582% | 1.714% | -2.312% | 1.78% | -4.292% | -3.632% | 4.949% | -1.652% | -2.972% | 0.9882% | 0.9882% | -5.904% | -0.5299% | -0.3319% | 14.06% | -7.593% |  |  | -0.992%
| Z-Score (Absolute) |  | 0.58 | 2.07 | 0.42 | 2.09 | 1.85 | 0.62 | 0.66 | 0.58 | 0.68 | 1.2 | 0.99 | 1.67 | 0.38 | 0.79 | 0.44 | 0.44 | 1.69 | 0.03 | 0.03 | 4.49 | 2.22 | 0.17 | 91.78 | 0.17
| Analytical results for V in OREAS 265 (Certified Value 79 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 82 | 81 | 77 | 80.4295 | 79 | 79 | 77 | 80 | 80 | 79 | 78 | 79 | 81 | 77 | 76 | 80.3 | 76.9225 | 82 | 75 | 79.8 | 77 | 77.646 | 74 | 79 | 1
| 1 | 2 | 77 | 84 | 76 | 75.1148 | 80 | 83 | 80 | 80 | 80 | 76 | 78 | 81 | 80 | 74 | 79 | 78.8 | 77.0233 | 83 | 77 | 81.9 | 77 | 77.1231 | 72 | 79 | 30
| 1 | 3 | 76 | 84 | 76 | 75.8967 | 77 | 81 | 75 | 78 | 79 | 78 | 79 | 80 | 78 | 77 | 77 | 79.2 | 77.0287 | 83 | 76 | 78.8 | 76 | 78.5387 | 76 | 78 | 16
| 1 | 4 | 80 | 81 | 76 | 78.1939 | 78 | 82 | 83 | 76 | 79 | 78 | 80 | 82 | 86 | 77 | 81 | 81.6 | 76.9126 | 81 | 77 | 78.2 | 77 | 77.2639 | 76 | 81 | 78.6616
| 1 | 5 | 76 | 82 | 77 | 77.7992 | 81 | 80 | 78 | 77 | 79 | 77 | 78 | 81 | 81 | 78 | 75 | 80.1 | 77.0712 | 83 | 76 | 78.6 | 77 | 80.0904 | 73 | 80 | 66
| 1 | 6 | 74 | 82 | 77 | 81.5795 | 78 | 80 | 75 | 78 | 80 | 80 | 75 | 80 | 86 | 74 | 78 | 82.8 | 76.9505 | 83 | 75 | 82.3 | 77 | 81.2885 | 77 | 81
| Mean |  | 77.5 | 82.3333 | 76.5 | 78.1689 | 78.8333 | 80.8333 | 78 | 78.1667 | 79.5 | 78 | 78 | 80.5 | 82 | 76.1667 | 77.6667 | 80.4667 | 76.9848 | 82.5 | 76 | 79.9333 | 76.8333 | 78.6584 | 74.6667 | 79.6667
| Median |  | 76.5 | 82 | 76.5 | 77.9965 | 78.5 | 80.5 | 77.5 | 78 | 79.5 | 78 | 78 | 80.5 | 81 | 77 | 77.5 | 80.2 | 76.9869 | 83 | 76 | 79.3 | 77 | 78.0924 | 75 | 79.5
| Std Dev. |  | 2.9496 | 1.3663 | 0.54772 | 2.5048 | 1.472 | 1.472 | 3.0984 | 1.6021 | 0.54772 | 1.4142 | 1.6733 | 1.0488 | 3.2863 | 1.7224 | 2.1602 | 1.5016 | 0.065042 | 0.83666 | 0.89443 | 1.7637 | 0.40825 | 1.6917 | 1.9664 | 1.2111
| Rel.Std.Dev. |  | 3.806% | 1.659% | 0.716% | 3.204% | 1.867% | 1.821% | 3.972% | 2.05% | 0.689% | 1.813% | 2.145% | 1.303% | 4.008% | 2.261% | 2.781% | 1.866% | 0.08449% | 1.014% | 1.177% | 2.206% | 0.5313% | 2.151% | 2.634% | 1.52%
| PDM3 |  | -1.477% | 4.668% | -2.748% | -0.6263% | 0.2183% | 2.761% | -0.8411% | -0.6292% | 1.066% | -0.8411% | -0.8411% | 2.337% | 4.244% | -3.172% | -1.265% | 2.295% | -2.132% | 4.88% | -3.384% | 1.617% | -2.324% | -0.004033% | -5.079% | 1.278%
| Z-Score (Absolute) |  | 0.32 | 1.98 | 0.79 | 0 | 0.32 | 1.27 | 0.08 | 0 | 0.63 | 0.08 | 0.08 | 1.11 | 1.82 | 0.95 | 0.24 | 1.09 | 0.56 | 2.06 | 1.03 | 0.84 | 0.64 | 0.23 | 1.67 | 0.71
| Analytical results for W in OREAS 265 (Certified Value 11.2 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 11.7 | 11.1 | 8.7 | 12.493 | 11.7 | 11.4 | 10.5 | 11 | 11.3 | 11.8 | 10.7 | 10.6 | 10.1 | 9.5 | 11.4 | 11 | 10.91 | 10 | 12.2 | 11.2 | 11.1185 | 10.5 |  |  | 1
| 1 | 2 | 11.8 | 11.8 | 9 | 4.1304 | 13.4 | 11.4 | 12.4 | 10.4 | 11.9 | 12.6 | 10.9 | 10.8 | 10.1 | 9.8 | 11.3 | 11.1 | 11.08 | 9.6 | 12.6 | 12 | 10.8671 | 11.5 |  |  | 31
| 1 | 3 | 11 | 11.6 | 9.5 | 1.8086 | 11.9 | 11.5 | 11.2 | 10.3 | 11.4 | 11.4 | 10.9 | 10.9 | 10.3 | 9.6 | 11 | 11.8 | 11.13 | 10.2 | 12.6 | 11 | 11.1284 | 10.7 |  |  | 16
| 1 | 4 | 10.9 | 11.3 | 9.6 | 8.0153 | 12.2 | 11.7 | 11 | 10.5 | 11.3 | 11.9 | 11.4 | 11.2 | 11.4 | 9.5 | 11.2 | 11.4 | 10.9 | 9.6 | 13.8 | 10.9 | 11.0173 | 11.1 |  |  | 11.2347
| 1 | 5 | 10.1 | 11 | 8.5 | 7.6669 | 12.1 | 11.4 | 11 | 11.1 | 11.7 | 11.6 | 10.9 | 11.2 | 10.2 | 9.8 | 11.1 | 11.3 | 11.46 | 9.3 | 13.6 | 11.1 | 11.5816 | 11.7 |  |  | 67
| 1 | 6 | 9 | 11.2 | 8.3 | 9.0416 | 11.7 | 11.7 | 12.3 | 10.4 | 12 | 11.4 | 11.1 | 11.1 | 11 | 9.4 | 11 | 11.7 | 11.54 | 9 | 12.8 | 11.5 | 11.1821 | 11.1
| Mean |  | 10.75 | 11.3333 | 8.9333 | 7.1926 | 12.1667 | 11.5167 | 11.4 | 10.6167 | 11.6 | 11.7833 | 10.9833 | 10.9667 | 10.5167 | 9.6 | 11.1667 | 11.3833 | 11.17 | 9.6167 | 12.9333 | 11.2833 | 11.1492 | 11.1
| Median |  | 10.95 | 11.25 | 8.85 | 7.8411 | 12 | 11.45 | 11.1 | 10.45 | 11.55 | 11.7 | 10.9 | 11 | 10.25 | 9.55 | 11.15 | 11.35 | 11.105 | 9.6 | 12.7 | 11.15 | 11.1234 | 11.1
| Std Dev. |  | 1.0559 | 0.30768 | 0.53166 | 3.7624 | 0.6377 | 0.1472 | 0.77201 | 0.34303 | 0.30984 | 0.44907 | 0.24014 | 0.24221 | 0.54924 | 0.16733 | 0.1633 | 0.31885 | 0.27247 | 0.44008 | 0.62823 | 0.40702 | 0.23935 | 0.45607
| Rel.Std.Dev. |  | 9.823% | 2.715% | 5.951% | 52.31% | 5.241% | 1.278% | 6.772% | 3.231% | 2.671% | 3.811% | 2.186% | 2.209% | 5.223% | 1.743% | 1.462% | 2.801% | 2.439% | 4.576% | 4.857% | 3.607% | 2.147% | 4.109%
| PDM3 |  | -4.314% | 0.8783% | -20.48% | -35.98% | 8.296% | 2.51% | 1.472% | -5.501% | 3.252% | 4.884% | -2.237% | -2.385% | -6.391% | -14.55% | -0.6052% | 1.323% | -0.5755% | -14.4% | 15.12% | 0.4333% | -0.761% | -1.199%
| Z-Score (Absolute) |  | 0.72 | 0.31 | 3.91 | 6.98 | 1.77 | 0.63 | 0.43 | 0.95 | 0.78 | 1.1 | 0.31 | 0.34 | 1.13 | 2.74 | 0.02 | 0.4 | 0.02 | 2.71 | 3.12 | 0.22 | 0.02 | 0.1
| Analytical results for Y in OREAS 265 (Certified Value 17.4 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 18.5 | 18.7 | 17 | 18.2402 | 17.97 | 17.3 | 17.8 | 18.9 | 17.34 | 19 | 17.2 | 17.6 | 15.7 | 17.6 | 17 | 16 | 16.9558 | 17.74 | 16.8 | 17.9 | 17.4 | 17.0655 | 15 | 16.6 | 1
| 1 | 2 | 17.5 | 19.6 | 17.2 | 18.2598 | 17.51 | 16.7 | 17.2 | 18.4 | 17.46 | 18.6 | 17 | 18.1 | 16.5 | 17.2 | 16.8 | 16.2 | 17.0463 | 17.51 | 22.4 | 18.5 | 16.5 | 16.9273 | 15 | 17 | 32
| 1 | 3 | 16.7 | 21.8 | 17.3 | 17.9788 | 17.24 | 15.8 | 16.1 | 18.4 | 17.43 | 18.9 | 17.8 | 18.6 | 16.7 | 18.1 | 16.8 | 15.8 | 16.9662 | 18.02 | 16.4 | 17.8 | 17.1 | 17.1365 | 14 | 17.1 | 16
| 1 | 4 | 16.2 | 19.3 | 17.5 | 17.9138 | 17.08 | 16.3 | 18.5 | 18.5 | 17.26 | 19.4 | 18.2 | 18.1 | 16.2 | 18 | 17.2 | 16.2 | 17.0811 | 17.32 | 16.8 | 19.6 | 17.5 | 17.147 | 14 | 17.6 | 17.4459
| 1 | 5 | 16 | 19.4 | 16.8 | 18.3006 | 21.08 | 17.5 | 17.3 | 18.2 | 17.69 | 19.2 | 17.9 | 18.2 | 16.2 | 17.5 | 16.9 | 15.4 | 16.9828 | 17.7 | 15.8 | 18.4 | 16 | 17.6946 | 14 | 18.2 | 68
| 1 | 6 | 15.8 | 18.6 | 17.1 | 18.1171 | 17.98 | 17.2 | 16.9 | 18.4 | 17.35 | 18.7 | 17.4 | 18.3 | 16.2 | 17.1 | 16.2 | 16.7 | 16.9289 | 17.39 | 16.1 | 19.1 | 17 | 17.845 | 14 | 17.7
| Mean |  | 16.7833 | 19.5667 | 17.15 | 18.1351 | 18.1433 | 16.8 | 17.3 | 18.4667 | 17.4217 | 18.9667 | 17.5833 | 18.15 | 16.25 | 17.5833 | 16.8167 | 16.05 | 16.9935 | 17.6133 | 17.3833 | 18.55 | 16.9167 | 17.3027 | 14.3333 | 17.3667
| Median |  | 16.45 | 19.35 | 17.15 | 18.1786 | 17.74 | 16.95 | 17.25 | 18.4 | 17.39 | 18.95 | 17.6 | 18.15 | 16.2 | 17.55 | 16.85 | 16.1 | 16.9745 | 17.605 | 16.6 | 18.45 | 17.05 | 17.1418 | 14 | 17.35
| Std Dev. |  | 1.0381 | 1.1639 | 0.2429 | 0.15981 | 1.4852 | 0.65727 | 0.8124 | 0.23381 | 0.14932 | 0.30111 | 0.45789 | 0.32711 | 0.33912 | 0.40702 | 0.33714 | 0.43704 | 0.058178 | 0.25905 | 2.4887 | 0.69498 | 0.57067 | 0.3733 | 0.5164 | 0.57504
| Rel.Std.Dev. |  | 6.185% | 5.948% | 1.416% | 0.8812% | 8.186% | 3.912% | 4.696% | 1.266% | 0.8571% | 1.588% | 2.604% | 1.802% | 2.087% | 2.315% | 2.005% | 2.723% | 0.3424% | 1.471% | 14.32% | 3.747% | 3.373% | 2.157% | 3.603% | 3.311%
| PDM3 |  | -3.798% | 12.16% | -1.696% | 3.95% | 3.998% | -3.702% | -0.8363% | 5.851% | -0.1389% | 8.717% | 0.7878% | 4.036% | -6.855% | 0.7878% | -3.607% | -8.001% | -2.593% | 0.9598% | -0.3586% | 6.329% | -3.034% | -0.821% | -17.84% | -0.4541%
| Z-Score (Absolute) |  | 0.7 | 2.61 | 0.27 | 0.9 | 0.91 | 0.68 | 0.09 | 1.3 | 0.06 | 1.89 | 0.25 | 0.92 | 1.34 | 0.25 | 0.66 | 1.58 | 0.45 | 0.28 | 0.01 | 1.4 | 0.55 | 0.09 | 3.62 | 0.01
| Analytical results for Yb in OREAS 265 (Certified Value 1.59 ppm) | [25 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.22 | [11 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.7 | 1.67 | 1.5 | 1.56 | 1.7 | 1.4 | 1.49 | 1.6 | 1.46 | 1.5387 | 1.52 | 1.59 | 1.5728 | [11 empty] | 1
| 1 | 2 | 1.6 | 1.69 | 1.6 | 1.5 | 1.6 | 1.4 | 1.54 | 1.5 | 1.68 | 1.5399 | 1.54 | 1.68 | 1.5622 | [11 empty] | 33
| 1 | 3 | 1.6 | 1.66 | 1.5 | 1.46 | 1.6 | 1.4 | 1.61 | 1.6 | 1.61 | 1.5331 | 1.57 | 1.61 | 1.611 | [11 empty] | 16
| 1 | 4 | 1.5 | 1.73 | 1.6 | 1.43 | 1.7 | 1.4 | 1.54 | 1.6 | 1.64 | 1.5244 | 1.52 | 1.73 | 1.5713 | [11 empty] | 1.5868
| 1 | 5 | 1.5 | 1.7 | 1.3 | 1.41 | 1.7 | 1.4 | 1.58 | 1.6 | 1.57 | 1.5021 | 1.64 | 1.67 | 1.6081 | [11 empty] | 69
| 1 | 6 | 1.6 | 1.71 | 1.5 | 1.42 | 1.7 | 1.4 | 1.62 | 1.5 | 1.63 | 1.5024 | 1.62 | 1.79 | 1.6524
| Mean |  | 1.5833 | 1.6933 | 1.5 | 1.4633 | 1.6667 | 1.4 | 1.5633 | 1.5667 | 1.5983 | 1.5234 | 1.5683 | 1.6783 | 1.5963
| Median |  | 1.6 | 1.695 | 1.5 | 1.445 | 1.7 | 1.4 | 1.56 | 1.6 | 1.62 | 1.5288 | 1.555 | 1.675 | 1.5905
| Std Dev. |  | 0.075277 | 0.02582 | 0.10954 | 0.057504 | 0.05164 | 2.4324e-16 | 0.049261 | 0.05164 | 0.07679 | 0.017295 | 0.051543 | 0.074409 | 0.034202
| Rel.Std.Dev. |  | 4.754% | 1.525% | 7.303% | 3.93% | 3.098% | 1.737e-14% | 3.151% | 3.296% | 4.804% | 1.135% | 3.286% | 4.433% | 2.143%
| PDM3 |  | -0.2175% | 6.715% | -5.469% | -7.78% | 5.034% | -11.77% | -1.478% | -1.268% | 0.7278% | -3.992% | -1.163% | 5.769% | 0.6008%
| Z-Score (Absolute) |  | 0.23 | 1.88 | 1.03 | 1.58 | 1.48 | 2.53 | 0.08 | 0.03 | 0.45 | 0.67 | 0 | 1.65 | 0.42
| Analytical results for Zn in OREAS 265 (Certified Value 61 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 57 | 64 | 61 | 61.9421 | 60 | 62 | 62 | 63 | 62 | 63 | 59 | 59.7 | 63.6 | 62 | 61 | 60 | 62.686 | 60 | 64 | 73 | 60 | 59.6293 | 60 | 61 | 1
| 1 | 2 | 56 | 66 | 60 | 61.2418 | 60 | 59 | 61 | 63 | 60 | 61 | 60 | 62.4 | 61.5 | 59 | 61 | 60 | 62.1172 | 60 | 64 | 73 | 61 | 59.3845 | 60 | 62 | 15
| 1 | 3 | 54 | 66 | 59 | 58.9897 | 60 | 59 | 59 | 62 | 60 | 62 | 60 | 61.3 | 60.8 | 62 | 62 | 60 | 62.3732 | 61 | 60 | 70 | 60 | 60.3142 | 62 | 60 | 16
| 1 | 4 | 58 | 64 | 59 | 59.3415 | 59 | 58 | 63 | 60 | 60 | 63 | 60 | 61.8 | 65.2 | 62 | 62 | 62 | 62.7548 | 59 | 61 | 73 | 62 | 60.5414 | 63 | 62 | 61.0228
| 1 | 5 | 55 | 65 | 60 | 58.9895 | 59 | 61 | 62 | 60 | 62 | 62 | 61 | 62.1 | 62.4 | 63 | 62 | 60 | 62.7244 | 61 | 60 | 72 | 60 | 62.8266 | 62 | 61 | 70
| 1 | 6 | 53 | 64 | 60 | 62.4121 | 61 | 61 | 58 | 62 | 61 | 63 | 58 | 61.5 | 65 | 59 | 60 | 63 | 62.6444 | 60 | 59 | 75 | 60 | 62.6564 | 61 | 62
| Mean |  | 55.5 | 64.8333 | 59.8333 | 60.4861 | 59.8333 | 60 | 60.8333 | 61.6667 | 60.8333 | 62.3333 | 59.6667 | 61.4667 | 63.0833 | 61.1667 | 61.3333 | 60.8333 | 62.55 | 60.1667 | 61.3333 | 72.6667 | 60.5 | 60.8921 | 61.3333 | 61.3333
| Median |  | 55.5 | 64.5 | 60 | 60.2917 | 60 | 60 | 61.5 | 62 | 60.5 | 62.5 | 60 | 61.65 | 63 | 62 | 61.5 | 60 | 62.6652 | 60 | 60.5 | 73 | 60 | 60.4278 | 61.5 | 61.5
| Std Dev. |  | 1.8708 | 0.98319 | 0.75277 | 1.5614 | 0.75277 | 1.5492 | 1.9408 | 1.3663 | 0.98319 | 0.8165 | 1.0328 | 0.95219 | 1.8225 | 1.7224 | 0.8165 | 1.3292 | 0.25233 | 0.75277 | 2.1602 | 1.633 | 0.83666 | 1.4953 | 1.2111 | 0.8165
| Rel.Std.Dev. |  | 3.371% | 1.516% | 1.258% | 2.581% | 1.258% | 2.582% | 3.19% | 2.216% | 1.616% | 1.31% | 1.731% | 1.549% | 2.889% | 2.816% | 1.331% | 2.185% | 0.4034% | 1.251% | 3.522% | 2.247% | 1.383% | 2.456% | 1.975% | 1.331%
| PDM3 |  | -9.05% | 6.244% | -1.949% | -0.8794% | -1.949% | -1.676% | -0.3104% | 1.055% | -0.3104% | 2.148% | -2.222% | 0.7274% | 3.377% | 0.2358% | 0.5089% | -0.3104% | 2.503% | -1.403% | 0.5089% | 19.08% | -0.8567% | -0.2142% | 0.5089% | 0.5089%
| Z-Score (Absolute) |  | 6.32 | 4.35 | 1.37 | 0.62 | 1.37 | 1.18 | 0.22 | 0.73 | 0.22 | 1.49 | 1.56 | 0.5 | 2.35 | 0.16 | 0.35 | 0.22 | 1.74 | 0.99 | 0.35 | 13.29 | 0.6 | 0.16 | 0.35 | 0.35
| Analytical results for Zr in OREAS 265 (Certified Value 197 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 196.3 | 204 | 184 | 189.7701 | 201.8 | 204 | 206 | 206 | 183.6 | 202 | 199 | 190.2 | 208.9 | 192.9 | 204.2 | 195 | 196.515 | 212.4 | 164.5 | 205 | 152.1 | 189.7128 | 171 | 199.3 | 1
| 1 | 2 | 189.7 | 206 | 187 | 186.3763 | 194.2 | 194 | 216 | 201 | 185.9 | 198 | 195.5 | 195.6 | 188.5 | 187.9 | 210.4 | 197 | 196.2775 | 212.2 | 168.5 | 211 | 151.7 | 187.8882 | 174 | 197.2 | 16
| 1 | 3 | 185.1 | 207 | 185 | 186.5654 | 193.8 | 188 | 200 | 196 | 182.4 | 199 | 202 | 193 | 184.8 | 195 | 206.2 | 200 | 196.2275 | 215.9 | 170 | 214 | 150.8 | 190.5526 | 168.45 | 197.7 | 16
| 1 | 4 | 189.8 | 203 | 186 | 186.5616 | 196.4 | 193 | 197 | 195 | 182.6 | 202 | 204 | 192.4 | 195.1 | 195.3 | 216.5 | 200 | 196.3525 | 206.4 | 167 | 210 | 151.3 | 191.1751 | 170.63 | 205.5 | 196.974
| 1 | 5 | 180.4 | 205 | 187 | 186.577 | 200.4 | 181 | 196 | 194 | 191.3 | 199.5 | 202 | 192.6 | 191.3 | 192.4 | 201.4 | 198 | 197.29 | 213.8 | 164 | 208 | 150.2 | 198.1512 | 166 | 202.3 | 71
| 1 | 6 | 174.3 | 202 | 186 | 189.4314 | 201.1 | 184 | 200 | 199.5 | 186.3 | 202 | 197.5 | 191.2 | 194.1 | 186.7 | 209 | 207 | 196.4025 | 213.8 | 165 | 216 | 150.3 | 198.9731 | 171.46 | 203.4
| Mean |  | 185.9333 | 204.5 | 185.8333 | 187.547 | 197.95 | 190.6667 | 202.5 | 198.5833 | 185.35 | 200.4167 | 200 | 192.5 | 193.7833 | 191.7 | 207.95 | 199.5 | 196.5108 | 212.4167 | 166.5 | 210.6667 | 151.0667 | 192.7422 | 170.2567 | 200.9
| Median |  | 187.4 | 204.5 | 186 | 186.5712 | 198.4 | 190.5 | 200 | 197.75 | 184.75 | 200.75 | 200.5 | 192.5 | 192.7 | 192.65 | 207.6 | 199 | 196.3775 | 213.1 | 166 | 210.5 | 151.05 | 190.8638 | 170.815 | 200.8
| Std Dev. |  | 7.7879 | 1.8708 | 1.169 | 1.5962 | 3.5898 | 8.2381 | 7.4766 | 4.5212 | 3.3447 | 1.8005 | 3.2094 | 1.8363 | 8.3082 | 3.6116 | 5.2959 | 4.1352 | 0.39463 | 3.2314 | 2.4083 | 3.9833 | 0.76594 | 4.6491 | 2.7399 | 3.3425
| Rel.Std.Dev. |  | 4.189% | 0.9148% | 0.6291% | 0.8511% | 1.814% | 4.321% | 3.692% | 2.277% | 1.805% | 0.8984% | 1.605% | 0.9539% | 4.287% | 1.884% | 2.547% | 2.073% | 0.2008% | 1.521% | 1.446% | 1.891% | 0.507% | 2.412% | 1.609% | 1.664%
| PDM3 |  | -5.605% | 3.821% | -5.656% | -4.786% | 0.4955% | -3.202% | 2.805% | 0.817% | -5.901% | 1.748% | 1.536% | -2.271% | -1.62% | -2.677% | 5.572% | 1.282% | -0.2351% | 7.84% | -15.47% | 6.952% | -23.31% | -2.148% | -13.56% | 1.993%
| Z-Score (Absolute) |  | 0.95 | 0.96 | 0.96 | 0.78 | 0.29 | 0.46 | 0.76 | 0.35 | 1.01 | 0.54 | 0.5 | 0.27 | 0.14 | 0.35 | 1.32 | 0.45 | 0.14 | 1.78 | 2.95 | 1.6 | 4.54 | 0.25 | 2.56 | 0.59
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 265 (Certified Value 0.629 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 0.6 | 0.6 | 0.56 | 0.60748 | 0.62 | 0.67 | 0.63 | 0.66 | 0.55 | 0.62 | 0.65 | 0.615 | 0.596 | 0.41 | 0.63 | 0.62 | 0.6 | 0.64 | 0.64 | 0.6 | 0.67104 | 0.53793 | 0.65 |  |  |  | 1
| 1 | 2 | 0.61 | 0.61 | 0.56 | 0.64061 | 0.63 | 0.67 | 0.59 | 0.63 | 0.55 | 0.61 | 0.67 | 0.617 | 0.583 | 0.39 | 0.64 | 0.65 | 0.5 | 0.63 | 0.65 | 0.6 | 0.64323 | 0.5567 | 0.65 |  |  |  | 20
| 1 | 3 | 0.42 | 0.63 | 0.61 | 0.62715 | 0.63 | 0.62 | 0.61 | 0.63 | 0.55 | 0.62 | 0.65 | 0.62 | 0.571 | 0.38 | 0.65 | 0.64 | 0.5 | 0.61 | 0.67 | 0.6 | 0.63373 | 0.57987 | 0.61 | 1 |  |  | 16
| 1 | 4 | 0.67 | 0.63 | 0.55 | 0.60114 | 0.65 | 0.63 | 0.63 | 0.62 | 0.56 | 0.62 | 0.65 | 0.632 | 0.603 | 0.4 | 0.64 | 0.63 | 0.5 | 0.63 | 0.66 | 0.6 | 0.61783 | 0.547 | 0.65 | 1 |  |  | 0.62905
| 1 | 5 | 0.47 | 0.6 | 0.57 | 0.64279 | 0.64 | 0.65 | 0.63 | 0.63 | 0.56 | 0.62 | 0.66 | 0.642 | 0.586 | 0.42 | 0.63 | 0.64 | 0.5 | 0.64 | 0.68 | 0.6 | 0.65241 | 0.55297 | 0.64 |  |  |  | 73
| 1 | 6 | 0.54 | 0.61 | 0.57 | 0.62172 | 0.66 | 0.66 | 0.6 | 0.64 | 0.55 | 0.63 | 0.65 | 0.628 | 0.597 | 0.41 | 0.64 | 0.63 | 0.5 | 0.63 | 0.68 | 0.6 | 0.6837 | 0.54326 | 0.65
| Mean |  | 0.55167 | 0.61333 | 0.57 | 0.62348 | 0.63833 | 0.65 | 0.615 | 0.635 | 0.55333 | 0.62 | 0.655 | 0.62567 | 0.58933 | 0.40167 | 0.63833 | 0.635 | 0.51667 | 0.63 | 0.66333 | 0.6 | 0.65032 | 0.55295 | 0.64167 | 1
| Median |  | 0.57 | 0.61 | 0.565 | 0.62444 | 0.635 | 0.655 | 0.62 | 0.63 | 0.55 | 0.62 | 0.65 | 0.624 | 0.591 | 0.405 | 0.64 | 0.635 | 0.5 | 0.63 | 0.665 | 0.6 | 0.64782 | 0.54998 | 0.65 | 1
| Std Dev. |  | 0.093684 | 0.013663 | 0.020976 | 0.01696 | 0.01472 | 0.020976 | 0.017607 | 0.013784 | 0.005164 | 0.0063246 | 0.0083666 | 0.010328 | 0.011639 | 0.01472 | 0.0075277 | 0.010488 | 0.040825 | 0.010954 | 0.01633 | 0 | 0.024204 | 0.014791 | 0.016021 | 0
| Rel.Std.Dev. |  | 16.98% | 2.228% | 3.68% | 2.72% | 2.306% | 3.227% | 2.863% | 2.171% | 0.9332% | 1.02% | 1.277% | 1.651% | 1.975% | 3.665% | 1.179% | 1.652% | 7.902% | 1.739% | 2.462% | 0% | 3.722% | 2.675% | 2.497% | 0%
| PDM3 |  | -12.3% | -2.498% | -9.387% | -0.8847% | 1.476% | 3.331% | -2.233% | 0.9463% | -12.04% | -1.438% | 4.126% | -0.5374% | -6.313% | -36.15% | 1.476% | 0.9463% | -17.87% | 0.1514% | 5.45% | -4.618% | 3.382% | -12.1% | 2.006% | 58.97%
| Z-Score (Absolute) |  | 3.12 | 0.52 | 2.35 | 0.09 | 0.53 | 1.03 | 0.45 | 0.39 | 3.05 | 0.24 | 1.24 | 0 | 1.53 | 9.44 | 0.53 | 0.39 | NA | 0.18 | 1.59 | NA | 1.04 | 3.06 | 0.67 | NA
Results from laboratories 1.18 and 1.21 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Al in OREAS 265 (Certified Value 1.16 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.05 | 1.17 | 1.42 | 1.2147 | 1.2921 | 1.172 | 1.22 | 1.11 | 1.1487 | 1.19 | 1.11 | 1.18 | 1.28 | 1.17 | 1.07 | 1.55 | 1.2155 | 1.16 | 0.826 | 1.2 | 1.1931 | 1.0135 | 1.13 |  |  |  | 1
| 1 | 2 | 1.04 | 1.23 | 1.46 | 1.2246 | 1.3283 | 1.165 | 1.22 | 1.1 | 1.1786 | 1.18 | 1.12 | 1.15 | 1.24 | 1.12 | 1.07 | 1.64 | 1.2288 | 1.15 | 0.82 | 1.2 | 1.2054 | 0.99177 | 1.14 |  |  |  | #N/A
| 1 | 3 | 1.03 | 1.23 | 1.47 | 1.1988 | 1.288 | 1.143 | 1.23 | 1.08 | 1.1867 | 1.19 | 1.1 | 1.15 | 1.17 | 1.14 | 1.11 | 1.59 | 1.2138 | 1.16 | 0.836 | 1.2 | 1.2542 | 0.9986 | 1.08 |  |  |  | 16
| 1 | 4 | 1.04 | 1.22 | 1.47 | 1.222 | 1.2952 | 1.116 | 1.24 | 1.07 | 1.1829 | 1.14 | 1.1 | 1.19 | 1.21 | 1.13 | 1.08 | 1.6 | 1.2201 | 1.17 | 0.852 | 1.2 | 1.2345 | 0.98776 | 1.11 |  |  |  | 1.1622
| 1 | 5 | 1.06 | 1.24 | 1.49 | 1.2115 | 1.2734 | 1.144 | 1.24 | 1.04 | 1.192 | 1.15 | 1.13 | 1.19 | 1.22 | 1.12 | 1.07 | 1.62 | 1.2182 | 1.18 | 0.841 | 1.21 | 1.2276 | 1.0044 | 1.12 |  |  |  | 74
| 1 | 6 | 1.07 | 1.23 | 1.43 | 1.1994 | 1.3123 | 1.173 | 1.22 | 1.03 | 1.2066 | 1.18 | 1.11 | 1.19 | 1.25 | 1.11 | 1.11 | 1.62 | 1.2154 | 1.14 | 0.873 | 1.21 | 1.1998 | 1.0179 | 1.11
| Mean |  | 1.0483 | 1.22 | 1.4567 | 1.2118 | 1.2982 | 1.1522 | 1.2283 | 1.0717 | 1.1826 | 1.1717 | 1.1117 | 1.175 | 1.2283 | 1.1317 | 1.085 | 1.6033 | 1.2186 | 1.16 | 0.84133 | 1.2033 | 1.2191 | 1.0023 | 1.115
| Median |  | 1.045 | 1.23 | 1.465 | 1.2131 | 1.2936 | 1.1545 | 1.225 | 1.075 | 1.1848 | 1.18 | 1.11 | 1.185 | 1.23 | 1.125 | 1.075 | 1.61 | 1.2168 | 1.16 | 0.8385 | 1.2 | 1.2165 | 1.0015 | 1.115
| Std Dev. |  | 0.01472 | 0.025298 | 0.026583 | 0.010953 | 0.01934 | 0.02214 | 0.0098319 | 0.031885 | 0.019217 | 0.02137 | 0.01169 | 0.019748 | 0.037639 | 0.02137 | 0.019748 | 0.031411 | 0.0054547 | 0.014142 | 0.019159 | 0.005164 | 0.023566 | 0.011907 | 0.020736
| Rel.Std.Dev. |  | 1.404% | 2.074% | 1.825% | 0.9039% | 1.49% | 1.922% | 0.8004% | 2.975% | 1.625% | 1.824% | 1.052% | 1.681% | 3.064% | 1.888% | 1.82% | 1.959% | 0.4476% | 1.219% | 2.277% | 0.4291% | 1.933% | 1.188% | 1.86%
| PDM3 |  | -9.801% | 4.97% | 25.33% | 4.267% | 11.7% | -0.8669% | 5.687% | -7.793% | 1.75% | 0.8109% | -4.352% | 1.098% | 5.687% | -2.631% | -6.646% | 37.95% | 4.851% | -0.1929% | -27.61% | 3.536% | 4.892% | -13.76% | -4.065%
| Z-Score (Absolute) |  | 1.6 | 0.57 | 3.56 | 0.47 | 1.56 | 0.29 | 0.67 | 1.31 | 0.1 | 0.04 | 0.8 | 0 | 0.67 | 0.55 | 1.14 | 5.42 | 0.55 | 0.19 | 4.22 | 0.36 | 0.56 | 2.18 | 0.76
| Analytical results for As in OREAS 265 (Certified Value 181 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 0
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  | 0
| 1 | 1 | 179 | 177.5 | 187 | 194.0377 | 177 | 190 | 186 | 177 | 173 | 178.5 | 199 | 179.3 | 190.4 | 184 | 172 | 178 | 189.32 | 173.9 | 179.5 | 123 | 170 | 180.7194 | 148.1 | 178.2 | 182.923 |  | 1
| 1 | 2 | 178 | 179.5 | 188 | 188.7442 | 181 | 183 | 186 | 175 | 174 | 181.5 | 189 | 182.2 | 188.6 | 180.9 | 170 | 186 | 189.6637 | 171.9 | 180 | 154 | 166.9 | 182.5815 | 143.3 | 180.5 | 181.527 |  | 10
| 1 | 3 | 177 | 179.5 | 188 | 193.3577 | 176 | 183 | 180 | 175.5 | 179 | 183 | 183 | 178.6 | 186.5 | 178.8 | 171 | 179 | 189.4325 | 175 | 177 | 139 | 165.4 | 185.0641 | 144.4 | 174.1 | 183.665 |  | 16
| 1 | 4 | 178 | 182.5 | 188 | 187.2094 | 179 | 186 | 186 | 174.5 | 176 | 176.5 | 185.5 | 183.4 | 181.7 | 180 | 170 | 188 | 188.985 | 176.2 | 179 | 146 | 169.1 | 180.6056 | 145.6 | 174.9 | 178.993 |  | 180.7574
| 1 | 5 | 181 | 185.5 | 187 | 188.4282 | 177 | 186 | 184 | 173.5 | 176 | 179 | 186 | 179.9 | 187.7 | 178.2 | 166 | 191 | 189.21 | 176 | 182.5 | 148 | 169.9 | 181.7179 | 147.9 | 177.3 | 179.234 |  | 75
| 1 | 6 | 182 | 180 | 190 | 187.1798 | 179 | 186 | 184 | 176 | 179 | 182.5 | 186.5 | 178 | 187.7 | 180.1 | 170 | 185 | 189.5163 | 176.3 | 178.5 | 159 | 169.2 | 181.6277 | 149.3 | 176.2 | 179.277
| Mean |  | 179.1667 | 180.75 | 188 | 189.8262 | 178.1667 | 185.6667 | 184.3333 | 175.25 | 176.1667 | 180.1667 | 188.1667 | 180.2333 | 187.1 | 180.3333 | 169.8333 | 184.5 | 189.3546 | 174.8833 | 179.4167 | 144.8333 | 168.4167 | 182.0527 | 146.4333 | 176.8667 | 180.9365
| Median |  | 178.5 | 179.75 | 188 | 188.5862 | 178 | 186 | 185 | 175.25 | 176 | 180.25 | 186.25 | 179.6 | 187.7 | 180.05 | 170 | 185.5 | 189.3762 | 175.5 | 179.25 | 147 | 169.15 | 181.6728 | 146.75 | 176.75 | 180.402
| Std Dev. |  | 1.9408 | 2.824 | 1.0954 | 3.072 | 1.8348 | 2.582 | 2.3381 | 1.2145 | 2.4833 | 2.5626 | 5.6451 | 2.1229 | 2.9455 | 2.0412 | 2.0412 | 5.0892 | 0.23937 | 1.7244 | 1.828 | 12.703 | 1.8541 | 1.6439 | 2.3577 | 2.3295 | 2.0576
| Rel.Std.Dev. |  | 1.083% | 1.562% | 0.5827% | 1.618% | 1.03% | 1.391% | 1.268% | 0.693% | 1.41% | 1.422% | 3% | 1.178% | 1.574% | 1.132% | 1.202% | 2.758% | 0.1264% | 0.986% | 1.019% | 8.771% | 1.101% | 0.903% | 1.61% | 1.317% | 1.137%
| PDM3 |  | -0.88% | -0.004089% | 4.007% | 5.017% | -1.433% | 2.716% | 1.978% | -3.047% | -2.54% | -0.3268% | 4.099% | -0.2899% | 3.509% | -0.2346% | -6.043% | 2.071% | 4.756% | -3.25% | -0.7417% | -19.87% | -6.827% | 0.7166% | -18.99% | -2.152% | 0.09909%
| Z-Score (Absolute) |  | 0.17 | 0.08 | 1.23 | 1.52 | 0.33 | 0.86 | 0.65 | 0.79 | 0.64 | 0.01 | 1.25 | 0 | 1.09 | 0.02 | 1.64 | 0.67 | 1.44 | 0.85 | 0.13 | 5.59 | 1.87 | 0.29 | 5.34 | 0.53 | 0.11
| Analytical results for B in OREAS 265 (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.22 | 1.23 | 1.29 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | [8 empty] | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | [8 empty] | 1
| 1 | 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | 20 | < 10 | < 10 | < 20 | < 10 | [8 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | 22 | < 10 | 20 | < 10 | < 10 | < 20 | < 10 | [8 empty] | 3
| 1 | 3 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | 20 | < 10 | < 10 | < 20 | < 10 | [8 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | 20 | < 10 | < 10 | < 20 | < 10 | [8 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | 24 | < 10 | 20 | < 10 | < 10 | < 20 | < 10 | [8 empty] | 76
| 1 | 6 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | 25 | < 10 | 30 | < 10 | < 10 | < 20 | < 10
| Mean | [11 empty] | 3 | 23.6667 |  | 21.6667
| Median | [11 empty] | 3 | 24 |  | 20
| Std Dev. | [11 empty] | 0 | 1.5275 |  | 4.0825
| Rel.Std.Dev. | [11 empty] | 0% | 6.454% |  | 18.84%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 265 (Certified Value 70 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 0
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 1
| 1 | 1 | 61 | 70 | 84 | 74.7275 | 72 | 68 | 75 | 70 | 70 | 70 | 70 | 66.3 | 78.2 | 70 | 64 | 79 | 70.5382 | 70 | 53 | 62 | 71.6618 | 73.6425 | 69 | 69.273 |  |  | 1
| 1 | 2 | 60 | 70 | 85 | 70.909 | 74 | 68 | 75 | 70 | 70 | 70 | 70 | 67 | 75.8 | 69 | 64 | 84 | 70.5119 | 70 | 56 | 63 | 71.9353 | 74.45 | 70 | 62.343 |  |  | 22
| 1 | 3 | 60 | 70 | 84 | 68.6576 | 71 | 66 | 72 | 70 | 72 | 70 | 70 | 66.6 | 72.8 | 68 | 65 | 81 | 70.6414 | 70 | 57 | 66 | 75.6973 | 73.48 | 66 | 67.483 |  |  | 16
| 1 | 4 | 61 | 70 | 83 | 73.1725 | 73 | 65 | 75 | 70 | 70 | 70 | 70 | 66.3 | 72.8 | 70 | 65 | 81 | 70.4042 | 70 | 57 | 63 | 73.8886 | 72.2925 | 68 | 61.203 |  |  | 70.0016
| 1 | 5 | 61 | 70 | 84 | 71.8895 | 72 | 67 | 74 | 60 | 71 | 70 | 70 | 69.6 | 75.1 | 69 | 64 | 82 | 70.4533 | 70 | 57 | 64 | 73.1673 | 75.84 | 68 | 63.415 |  |  | 77
| 1 | 6 | 66 | 70 | 83 | 70.3326 | 73 | 68 | 74 | 60 | 70 | 70 | 70 | 71 | 76.9 | 70 | 64 | 84 | 70.222 | 70 | 59 | 68 | 71.836 | 72.703 | 68 | 62.994
| Mean |  | 61.5 | 70 | 83.8333 | 71.6148 | 72.5 | 67 | 74.1667 | 66.6667 | 70.5 | 70 | 70 | 67.8 | 75.2667 | 69.3333 | 64.3333 | 81.8333 | 70.4618 | 70 | 56.5 | 64.3333 | 73.031 | 73.7347 | 68.1667 | 64.4518
| Median |  | 61 | 70 | 84 | 71.3993 | 72.5 | 67.5 | 74.5 | 70 | 70 | 70 | 70 | 66.8 | 75.45 | 69.5 | 64 | 81.5 | 70.4826 | 70 | 57 | 63.5 | 72.5513 | 73.5613 | 68 | 63.2045
| Std Dev. |  | 2.2583 | 0 | 0.75277 | 2.1484 | 1.0488 | 1.2649 | 1.169 | 5.164 | 0.83666 | 0 | 0 | 2.003 | 2.1796 | 0.8165 | 0.5164 | 1.9408 | 0.14248 | 0 | 1.9748 | 2.2509 | 1.5727 | 1.2776 | 1.3292 | 3.1823
| Rel.Std.Dev. |  | 3.672% | 0% | 0.8979% | 3% | 1.447% | 1.888% | 1.576% | 7.746% | 1.187% | 0% | 0% | 2.954% | 2.896% | 1.178% | 0.8027% | 2.372% | 0.2022% | 0% | 3.495% | 3.499% | 2.153% | 1.733% | 1.95% | 4.937%
| PDM3 |  | -12.14% | -0.002311% | 19.76% | 2.304% | 3.569% | -4.288% | 5.95% | -4.764% | 0.712% | -0.002311% | -0.002311% | -3.145% | 7.521% | -0.9547% | -8.097% | 16.9% | 0.6574% | -0.002311% | -19.29% | -8.097% | 4.328% | 5.333% | -2.621% | -7.928%
| Z-Score (Absolute) |  | 1.75 | NA | 2.6 | 0.22 | 0.4 | 0.67 | 0.72 | NA | 0.01 | NA | NA | 0.52 | 0.94 | 0.22 | 1.19 | 2.21 | 0 | NA | 2.72 | 1.19 | 0.5 | 0.64 | 0.45 | 1.17
Results from laboratories 1.02, 1.08, 1.10, 1.11 and 1.19 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 265 (Certified Value 1.05 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | 1.30 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.2g | [5 empty] | 3
| 1 | 1 | 1 | 0.99 | 1.09 | 1.0484 | 1.07 | 1.16 | 1.06 | 0.98 | 0.97 | 1.01 | 1.07 | 1 | 1.07 | 1.1 | 1.14 | 0.95 | 0.79 | < 1 | 1.0442 | 1.13 | 1.184 | [5 empty] | 1
| 1 | 2 | 1 | 1.01 | 1.07 | 1.1411 | 1.09 | 1.22 | 1.08 | 0.96 | 0.98 | 1.01 | 1.01 | 1.1 | 1.06 | 1.1 | 1.17 | 0.96 | 0.86 | < 1 | 1.0462 | 1.07 | 1.097 | [5 empty] | 23
| 1 | 3 | 0.9 | 1 | 1.08 | 1.0671 | 1.03 | 1.17 | 1.01 | 0.95 | 0.97 | 1.03 | 0.98 | 1.1 | 1 | 1.1 | 1.1 | 0.94 | 0.82 | < 1 | 1.0775 | 1.08 | 1.132 | [5 empty] | 16
| 1 | 4 | 0.9 | 1.01 | 1.04 | 1.1902 | 1.03 | 1.12 | 1.12 | 0.95 | 1.01 | 0.99 | 0.99 | 1.1 | 1.06 | 1.1 | 1.11 | 0.94 | 0.86 | < 1 | 1.0429 | 1.12 | 1.102 | [5 empty] | 1.0502
| 1 | 5 | 1 | 1.04 | 1.07 | 1.0152 | 1.05 | 1.17 | 1.04 | 0.94 | 1.04 | 1 | 1.01 | 1.1 | 1.04 | 1.1 | 1.17 | 0.98 | 0.88 | < 1 | 1.0379 | 1.07 | 1.132 | [5 empty] | 78
| 1 | 6 | 0.9 | 1 | 1.07 | 1.1295 | 1.08 | 1.11 | 1.12 | 0.93 | 0.99 | 1.03 | 0.99 | 1.1 | 1 | 1 | 1.16 | 0.94 | 0.91 | < 1 | 1.0233 | 1.11 | 1.147
| Mean |  | 0.95 | 1.0083 | 1.07 | 1.0986 | 1.0583 | 1.1583 | 1.0717 | 0.95167 | 0.99333 | 1.0117 | 1.0083 | 1.0833 | 1.0383 | 1.0833 | 1.1417 | 0.95167 | 0.85333 |  | 1.0453 | 1.0967 | 1.1323
| Median |  | 0.95 | 1.005 | 1.07 | 1.0983 | 1.06 | 1.165 | 1.07 | 0.95 | 0.985 | 1.01 | 1 | 1.1 | 1.05 | 1.1 | 1.15 | 0.945 | 0.86 |  | 1.0436 | 1.095 | 1.132
| Std Dev. |  | 0.054772 | 0.017224 | 0.016733 | 0.065756 | 0.025626 | 0.039707 | 0.044008 | 0.017224 | 0.027325 | 0.016021 | 0.032506 | 0.040825 | 0.031252 | 0.040825 | 0.030605 | 0.016021 | 0.04274 |  | 0.017815 | 0.026583 | 0.031791
| Rel.Std.Dev. |  | 5.766% | 1.708% | 1.564% | 5.986% | 2.421% | 3.428% | 4.106% | 1.81% | 2.751% | 1.584% | 3.224% | 3.768% | 3.01% | 3.768% | 2.681% | 1.683% | 5.009% |  | 1.704% | 2.424% | 2.808%
| PDM3 |  | -9.537% | -3.982% | 1.89% | 4.612% | 0.7789% | 10.3% | 2.049% | -9.378% | -5.411% | -3.665% | -3.982% | 3.159% | -1.126% | 3.159% | 8.714% | -9.378% | -18.74% |  | -0.4584% | 4.429% | 7.825%
| Z-Score (Absolute) |  | 1.25 | 0.49 | 0.32 | 0.7 | 0.17 | 1.48 | 0.35 | 1.23 | 0.68 | 0.44 | 0.49 | 0.5 | 0.09 | 0.5 | 1.27 | 1.23 | 2.52 | 7.16 | 0 | 0.67 | 1.14
| Analytical results for Bi in OREAS 265 (Certified Value 1.64 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.53 | 1.68 | 1.81 | 1.7349 | 1.71 | 1.66 | 1.37 | 1.33 | 1.64 | 1.57 | 1.64 | 1.74 | 1.71 | 1.75 | 1.54 | 1.85 | 1.47 | 1.83 | 1.63 | 1.7 | 1.7524 | 1.5166 | 1.77 |  |  |  | 1
| 1 | 2 | 1.63 | 1.71 | 1.83 | 1.6163 | 1.79 | 1.58 | 1.4 | 1.4 | 1.71 | 1.53 | 1.63 | 1.71 | 1.7 | 1.8 | 1.58 | 1.73 | 1.62 | 1.65 | 1.91 | 1.5 | 1.7284 | 1.5226 | 1.66 |  |  |  | 24
| 1 | 3 | 1.75 | 1.69 | 1.81 | 1.8341 | 1.7 | 1.45 | 1.41 | 1.48 | 1.71 | 1.62 | 1.56 | 1.7 | 1.73 | 1.71 | 1.59 | 2.02 | 1.71 | 1.66 | 2.18 | 1.5 | 1.6945 | 1.6279 | 1.6 |  |  |  | 16
| 1 | 4 | 1.5 | 1.68 | 1.79 | 1.6357 | 1.78 | 1.46 | 1.42 | 1.36 | 1.62 | 1.49 | 1.59 | 1.69 | 1.73 | 1.79 | 1.61 | 1.87 | 1.61 | 1.64 | 1.84 | 1.7 | 1.675 | 1.6371 | 1.57 |  |  |  | 1.6428
| 1 | 5 | 1.6 | 1.74 | 1.78 | 1.763 | 1.78 | 1.55 | 1.48 | 1.36 | 1.68 | 1.58 | 1.58 | 1.83 | 1.72 | 1.75 | 1.55 | 1.89 | 1.41 | 1.66 | 2.44 | 1.5 | 1.7349 | 1.5814 | 1.56 |  |  |  | 79
| 1 | 6 | 1.6 | 1.73 | 1.82 | 1.5961 | 1.79 | 1.57 | 1.52 | 1.38 | 1.67 | 1.47 | 1.5 | 1.75 | 1.8 | 1.79 | 1.59 | 1.91 | 1.55 | 1.62 | 1.89 | 1.6 | 1.7541 | 1.5944 | 1.62
| Mean |  | 1.6017 | 1.705 | 1.8067 | 1.6967 | 1.7583 | 1.545 | 1.4333 | 1.385 | 1.6717 | 1.5433 | 1.5833 | 1.7367 | 1.7317 | 1.765 | 1.5767 | 1.8783 | 1.5617 | 1.6767 | 1.9817 | 1.5833 | 1.7232 | 1.58 | 1.63
| Median |  | 1.6 | 1.7 | 1.81 | 1.6853 | 1.78 | 1.56 | 1.415 | 1.37 | 1.675 | 1.55 | 1.585 | 1.725 | 1.725 | 1.77 | 1.585 | 1.88 | 1.58 | 1.655 | 1.9 | 1.55 | 1.7317 | 1.5879 | 1.61
| Std Dev. |  | 0.087502 | 0.025884 | 0.018619 | 0.094919 | 0.041673 | 0.079183 | 0.055737 | 0.052058 | 0.03656 | 0.057155 | 0.050859 | 0.051251 | 0.035449 | 0.034496 | 0.026583 | 0.093897 | 0.10889 | 0.076594 | 0.2852 | 0.098319 | 0.031993 | 0.051128 | 0.07746
| Rel.Std.Dev. |  | 5.463% | 1.518% | 1.031% | 5.594% | 2.37% | 5.125% | 3.889% | 3.759% | 2.187% | 3.703% | 3.212% | 2.951% | 2.047% | 1.954% | 1.686% | 4.999% | 6.973% | 4.568% | 14.39% | 6.21% | 1.857% | 3.236% | 4.752%
| PDM3 |  | -2.506% | 3.783% | 9.972% | 3.277% | 7.03% | -5.956% | -12.75% | -15.69% | 1.754% | -6.057% | -3.622% | 5.711% | 5.407% | 7.436% | -4.028% | 14.33% | -4.941% | 2.059% | 20.62% | -3.622% | 4.893% | -3.826% | -0.7818%
| Z-Score (Absolute) |  | 0.53 | 0.25 | 1.03 | 0.19 | 0.66 | 0.97 | 1.82 | 2.19 | 0 | 0.98 | 0.67 | 0.5 | 0.46 | 0.71 | 0.73 | 1.58 | 0.84 | 0.04 | 2.37 | 0.67 | 0.39 | 0.7 | 0.32
| Analytical results for Ca in OREAS 265 (Certified Value 0.257 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  | 3
| 1 | 1 | 0.24 | 0.25 | 0.28 | 0.33986 | 0.28 | 0.268 | 0.26 | 0.25 | 0.25 | 0.26 | 0.27 | 0.25 | 0.28 | 0.26 | 0.24 | 0.31 | 0.25221 | 0.2042 | 0.26 | 0.172 | 0.26 | 0.25685 | 0.21811 | 0.26 | 0.22634 |  | 1
| 1 | 2 | 0.24 | 0.25 | 0.28 | 0.33725 | 0.28 | 0.265 | 0.26 | 0.25 | 0.26 | 0.26 | 0.25 | 0.25 | 0.27 | 0.26 | 0.24 | 0.32 | 0.25284 | 0.2017 | 0.26 | 0.172 | 0.26 | 0.25753 | 0.21252 | 0.26 | 0.21634 |  | #N/A
| 1 | 3 | 0.24 | 0.25 | 0.29 | 0.33858 | 0.28 | 0.262 | 0.26 | 0.24 | 0.26 | 0.26 | 0.24 | 0.25 | 0.26 | 0.26 | 0.25 | 0.32 | 0.25265 | 0.2044 | 0.25 | 0.179 | 0.26 | 0.26663 | 0.20141 | 0.25 | 0.22138 |  | 16
| 1 | 4 | 0.24 | 0.25 | 0.27 | 0.34309 | 0.28 | 0.257 | 0.26 | 0.24 | 0.26 | 0.25 | 0.24 | 0.25 | 0.25 | 0.26 | 0.24 | 0.32 | 0.25288 | 0.2068 | 0.26 | 0.179 | 0.25 | 0.25989 | 0.21266 | 0.25 | 0.21255 |  | 0.25726
| 1 | 5 | 0.24 | 0.26 | 0.28 | 0.33316 | 0.28 | 0.263 | 0.27 | 0.24 | 0.26 | 0.25 | 0.25 | 0.25 | 0.27 | 0.26 | 0.24 | 0.32 | 0.25102 | 0.2056 | 0.27 | 0.179 | 0.26 | 0.25713 | 0.21538 | 0.25 | 0.21626 |  | 80
| 1 | 6 | 0.25 | 0.25 | 0.28 | 0.33315 | 0.28 | 0.267 | 0.26 | 0.24 | 0.26 | 0.26 | 0.25 | 0.25 | 0.27 | 0.25 | 0.25 | 0.33 | 0.25234 | 0.2048 | 0.26 | 0.179 | 0.26 | 0.25366 | 0.21447 | 0.25 | 0.2142
| Mean |  | 0.24167 | 0.25167 | 0.28 | 0.33752 | 0.28 | 0.26367 | 0.26167 | 0.24333 | 0.25833 | 0.25667 | 0.25 | 0.25 | 0.26667 | 0.25833 | 0.24333 | 0.32 | 0.25232 | 0.20458 | 0.26 | 0.17667 | 0.25833 | 0.25861 | 0.21243 | 0.25333 | 0.21784
| Median |  | 0.24 | 0.25 | 0.28 | 0.33792 | 0.28 | 0.264 | 0.26 | 0.24 | 0.26 | 0.26 | 0.25 | 0.25 | 0.27 | 0.26 | 0.24 | 0.32 | 0.2525 | 0.2046 | 0.26 | 0.179 | 0.26 | 0.25733 | 0.21357 | 0.25 | 0.2163
| Std Dev. |  | 0.0040825 | 0.0040825 | 0.0063246 | 0.003893 | 0 | 0.0039833 | 0.0040825 | 0.005164 | 0.0040825 | 0.005164 | 0.010954 | 0 | 0.010328 | 0.0040825 | 0.005164 | 0.0063246 | 0.00069088 | 0.0017023 | 0.0063246 | 0.0036148 | 0.0040825 | 0.0044018 | 0.0057729 | 0.005164 | 0.0051159
| Rel.Std.Dev. |  | 1.689% | 1.622% | 2.259% | 1.153% | 0% | 1.511% | 1.56% | 2.122% | 1.58% | 2.012% | 4.382% | 0% | 3.873% | 1.58% | 2.122% | 1.976% | 0.2738% | 0.8321% | 2.433% | 2.046% | 1.58% | 1.702% | 2.718% | 2.038% | 2.348%
| PDM3 |  | -6.061% | -2.174% | 8.839% | 31.2% | 8.839% | 2.49% | 1.713% | -5.413% | 0.4173% | -0.2306% | -2.822% | -2.822% | 3.657% | 0.4173% | -5.413% | 24.39% | -1.919% | -20.48% | 1.065% | -31.33% | 0.4173% | 0.5262% | -17.43% | -1.526% | -15.32%
| Z-Score (Absolute) |  | 1.44 | 0.48 | 2.25 | 7.79 | 2.25 | 0.67 | 0.48 | 1.28 | 0.16 | 0 | 0.64 | 0.64 | 0.96 | 0.16 | 1.28 | 6.1 | 0.42 | 5.02 | 0.32 | 7.74 | 0.16 | 0.19 | 4.26 | 0.32 | 3.74
| Analytical results for Cd in OREAS 265 (Certified Value 0.029 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  |  | 3
| 1 | 1 | 0.02 | 0.03 | 0.01 | < 0.05 | 0.03 | 0.04 | 0.04 | 0.03 | 0.03 | 0.04 | 0.03 | 0.02 | 0.03 | 0.1 | 0.03 | 0.04 | 0.03 | 0.02 | < 0.1 | < 0.2 | < 0.1 | 0.03 | 0.675 |  |  |  | 1
| 1 | 2 | 0.03 | 0.03 | 0.02 | < 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.1 | 0.03 | 0.02 | 0.03 | 0.03 | < 0.1 | < 0.2 | < 0.1 | 0.03 | 0.71 |  |  |  | 4
| 1 | 3 | 0.03 | 0.03 | 0.01 | < 0.05 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.09 | 0.03 | 0.04 | 0.02 | 0.03 | < 0.1 | < 0.2 | < 0.1 | 0.03 | 0.67 |  |  |  | 16
| 1 | 4 | 0.02 | 0.03 | 0.01 | < 0.05 | 0.03 | 0.03 | 0.04 | 0.02 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.11 | 0.02 | 0.05 | 0.04 | 0.02 | < 0.1 | < 0.2 | < 0.1 | 0.03 | 0.634 |  |  |  | 0.029489
| 1 | 5 | 0.06 | 0.04 | 0.02 | < 0.05 | 0.04 | 0.03 | 0.04 | 0.02 | 0.02 | 0.03 | 0.02 | 0.03 | 0.03 | 0.1 | 0.03 | 0.04 | 0.03 | 0.03 | < 0.1 | < 0.2 | < 0.1 | 0.03 | 0.613 |  |  |  | 81
| 1 | 6 | 0.03 | 0.03 | 0.01 | < 0.05 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.02 | 0.01 | 0.02 | 0.04 | 0.09 | 0.03 | 0.05 | 0.02 | 0.02 | < 0.1 | < 0.2 | < 0.1 | 0.03 | 0.617
| Mean |  | 0.031667 | 0.031667 | 0.013333 |  | 0.033333 | 0.031667 | 0.04 | 0.026667 | 0.028333 | 0.03 | 0.025 | 0.025 | 0.03 | 0.098333 | 0.028333 | 0.04 | 0.028333 | 0.025 |  |  |  | 0.03 | 0.65317
| Median |  | 0.03 | 0.03 | 0.01 |  | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.025 | 0.03 | 0.1 | 0.03 | 0.04 | 0.03 | 0.025 |  |  |  | 0.03 | 0.652
| Std Dev. |  | 0.01472 | 0.0040825 | 0.005164 |  | 0.005164 | 0.0040825 | 0 | 0.005164 | 0.0040825 | 0.0063246 | 0.0083666 | 0.0054772 | 0.0063246 | 0.0075277 | 0.0040825 | 0.010954 | 0.0075277 | 0.0054772 |  |  |  | 0 | 0.038155
| Rel.Std.Dev. |  | 46.48% | 12.89% | 38.73% |  | 15.49% | 12.89% | 0% | 19.36% | 14.41% | 21.08% | 33.47% | 21.91% | 21.08% | 7.655% | 14.41% | 27.39% | 26.57% | 21.91% |  |  |  | 0% | 5.841%
| PDM3 |  | 7.385% | 7.385% | -54.79% |  | 13.04% | 7.385% | 35.64% | -9.57% | -3.919% | 1.733% | -15.22% | -15.22% | 1.733% | 233.5% | -3.919% | 35.64% | -3.919% | -15.22% |  |  |  | 1.733% | 2115%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 2.25 | 0.67 | 0.45 | 0.22 | 1.35 | 0.45 | 0.22 | 0 | 0.67 | 0.67 | 0 | 9.22 | 0.22 | 1.35 | 0.22 | 0.67 | 2.7 | 9.44 | 2.7 | 0 | 84.04
| Analytical results for Ce in OREAS 265 (Certified Value 53 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 0
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [6 empty] | 1
| 1 | 1 | 51.93 | 49.9 | 41.1 | 52.064 | 55.025 | 53.17 | 52 | 52.5 | 54.585 | 49.2 | 56.7 | 50 | 51.54 | 53.06 | 57.3 | 59.9 | 123 | 52 | 54.7047 | 52.74 | [6 empty] | 1
| 1 | 2 | 53.15 | 52.1 | 42.2 | 52.0252 | 55.285 | 50.41 | 53 | 51 | 53.999 | 48.2 | 56.8 | 49.1 | 50.38 | 52.46 | 59.6 | 57.1 | 128 | 51 | 57.2714 | 52.26 | [6 empty] | 26
| 1 | 3 | 52.02 | 52.1 | 41.4 | 50.3143 | 53.814 | 48.57 | 53 | 52 | 55.58 | 49 | 52.7 | 49.2 | 49.15 | 53.39 | 58 | 57.6 | 125 | 52 | 57.4841 | 50.52 | [6 empty] | 16
| 1 | 4 | 51.57 | 52.5 | 40.4 | 51.3658 | 55.137 | 49.13 | 53 | 51.6 | 55.185 | 47.9 | 54.4 | 50 | 49.8 | 52.76 | 59.3 | 57.4 | 130 | 52 | 56.6454 | 52.35 | [6 empty] | 53.0199
| 1 | 5 | 53.66 | 52.1 | 40.7 | 52.1312 | 54.393 | 48.38 | 54 | 53.6 | 55.523 | 49.3 | 53.5 | 51.4 | 51.37 | 51.65 | 58.6 | 57.6 | 127 | 52 | 55.9593 | 52 | [6 empty] | 82
| 1 | 6 | 54.13 | 51.5 | 41.4 | 50.5175 | 54.373 | 51.21 | 52 | 55.7 | 55.22 | 47.9 | 53.2 | 52.2 | 49.98 | 52.06 | 57.7 | 57.8 | 134 | 52 | 55.065 | 52.81
| Mean |  | 52.7433 | 51.7 | 41.2 | 51.403 | 54.6712 | 50.145 | 52.8333 | 52.7333 | 55.0153 | 48.5833 | 54.55 | 50.3167 | 50.37 | 52.5633 | 58.4167 | 57.9 | 127.8333 | 51.8333 | 56.1883 | 52.1133
| Median |  | 52.585 | 52.1 | 41.25 | 51.6955 | 54.709 | 49.77 | 53 | 52.25 | 55.2025 | 48.6 | 53.95 | 50 | 50.18 | 52.61 | 58.3 | 57.6 | 127.5 | 52 | 56.3023 | 52.305
| Std Dev. |  | 1.0478 | 0.93808 | 0.62929 | 0.81548 | 0.56921 | 1.8429 | 0.75277 | 1.699 | 0.6108 | 0.65549 | 1.7919 | 1.2368 | 0.93106 | 0.64295 | 0.91086 | 1.008 | 3.8687 | 0.40825 | 1.1469 | 0.83751
| Rel.Std.Dev. |  | 1.987% | 1.814% | 1.527% | 1.586% | 1.041% | 3.675% | 1.425% | 3.222% | 1.11% | 1.349% | 3.285% | 2.458% | 1.848% | 1.223% | 1.559% | 1.741% | 3.026% | 0.7876% | 2.041% | 1.607%
| PDM3 |  | -0.5216% | -2.489% | -22.29% | -3.05% | 3.114% | -5.422% | -0.3519% | -0.5405% | 3.764% | -8.368% | 2.886% | -5.099% | -4.998% | -0.8611% | 10.18% | 9.204% | 141.1% | -2.238% | 5.976% | -1.71%
| Z-Score (Absolute) |  | 0.03 | 0.3 | 3.59 | 0.39 | 0.63 | 0.78 | 0.06 | 0.03 | 0.74 | 1.27 | 0.6 | 0.73 | 0.71 | 0.03 | 1.81 | 1.65 | 23.57 | 0.26 | 1.11 | 0.17
| Analytical results for Co in OREAS 265 (Certified Value 8.13 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 7.6 | 8.3 | 8.7 | 9.2654 | 8.2 | 8.3 | 8.3 | 8.2 | 7.6 | 7.7 | 8.4 | 8 | 8 | 7.2 | 8.1 | 8.5 | 8.9325 | 8.2 | 5.85 | 7.8 | 8.0914 | 6.7332 | 7.8 | 7.866 |  |  | 1
| 1 | 2 | 7.8 | 8.4 | 8.6 | 8.8596 | 8.5 | 8.1 | 8.5 | 7.7 | 7.7 | 7.6 | 8.6 | 8 | 8.4 | 7 | 7.8 | 9.1 | 8.9199 | 8 | 6.11 | 7.9 | 7.9959 | 6.5565 | 7.6 | 7.849 |  |  | 14
| 1 | 3 | 7.7 | 8.4 | 8.8 | 9.4299 | 8.2 | 8.1 | 8.4 | 7.9 | 7.7 | 7.8 | 8.1 | 7.9 | 8 | 6.9 | 8 | 8.9 | 8.9802 | 7.9 | 6.1 | 7.8 | 8.2197 | 6.4031 | 7 | 7.804 |  |  | 16
| 1 | 4 | 7.7 | 8.5 | 8.5 | 8.8257 | 8.2 | 7.9 | 8.2 | 8 | 7.6 | 7.6 | 8.2 | 8.1 | 7.9 | 7.2 | 8 | 9.3 | 8.9613 | 8 | 6.06 | 7.9 | 8.0257 | 6.7053 | 7.7 | 7.529 |  |  | 8.1313
| 1 | 5 | 7.9 | 8.4 | 8.5 | 9.2879 | 8.2 | 8.2 | 8.6 | 8.4 | 7.7 | 7.9 | 8.3 | 8.3 | 8.3 | 6.9 | 7.6 | 9.1 | 8.9233 | 7.9 | 6.12 | 7.9 | 8.0841 | 6.7634 | 7.4 | 7.951 |  |  | 83
| 1 | 6 | 7.9 | 8.5 | 8.6 | 8.7206 | 8.2 | 8.2 | 8.3 | 8.4 | 7.7 | 7.7 | 8.2 | 8 | 8.2 | 6.9 | 7.8 | 9.3 | 8.9464 | 8 | 6.34 | 8 | 7.8176 | 6.4775 | 7.7 | 7.787
| Mean |  | 7.7667 | 8.4167 | 8.6167 | 9.0648 | 8.25 | 8.1333 | 8.3833 | 8.1 | 7.6667 | 7.7167 | 8.3 | 8.05 | 8.1333 | 7.0167 | 7.8833 | 9.0333 | 8.9439 | 8 | 6.0967 | 7.8833 | 8.0391 | 6.6065 | 7.5333 | 7.7977
| Median |  | 7.75 | 8.4 | 8.6 | 9.0625 | 8.2 | 8.15 | 8.35 | 8.1 | 7.7 | 7.7 | 8.25 | 8 | 8.1 | 6.95 | 7.9 | 9.1 | 8.9395 | 8 | 6.105 | 7.9 | 8.0549 | 6.6309 | 7.65 | 7.8265
| Std Dev. |  | 0.12111 | 0.075277 | 0.1169 | 0.29701 | 0.12247 | 0.13663 | 0.1472 | 0.28284 | 0.05164 | 0.1169 | 0.17889 | 0.13784 | 0.19664 | 0.1472 | 0.18348 | 0.30111 | 0.023482 | 0.10954 | 0.15629 | 0.075277 | 0.13302 | 0.14896 | 0.29439 | 0.14363
| Rel.Std.Dev. |  | 1.559% | 0.8944% | 1.357% | 3.276% | 1.485% | 1.68% | 1.756% | 3.492% | 0.6736% | 1.515% | 2.155% | 1.712% | 2.418% | 2.098% | 2.328% | 3.333% | 0.2625% | 1.369% | 2.564% | 0.9549% | 1.655% | 2.255% | 3.908% | 1.842%
| PDM3 |  | -4.485% | 3.509% | 5.969% | 11.48% | 1.459% | 0.02465% | 3.099% | -0.3853% | -5.714% | -5.1% | 2.074% | -1% | 0.02465% | -13.71% | -3.05% | 11.09% | 9.993% | -1.615% | -25.02% | -3.05% | -1.135% | -18.75% | -7.354% | -4.103%
| Z-Score (Absolute) |  | 0.62 | 0.83 | 1.27 | 2.27 | 0.46 | 0.2 | 0.75 | 0.12 | 0.84 | 0.73 | 0.57 | 0.01 | 0.2 | 2.29 | 0.36 | 2.2 | 2 | 0.1 | 4.34 | 0.36 | 0.01 | 3.2 | 1.14 | 0.55
| Analytical results for Cr in OREAS 265 (Certified Value 36 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 2
| 1 | 1 | 34 | 37 | 41 | 40.0239 | 39 | 36 | 38 | 35 | 35 | 36 | 38 | 36 | 39.6 | 33 | 35 | 40 | 39.3485 | 35 | 27 | 35 | 37.5501 | 29.67 | 35 | 32.328 |  |  | 1
| 1 | 2 | 35 | 37 | 41 | 37.7728 | 40 | 36 | 38 | 35 | 36 | 36 | 36 | 36.2 | 38.8 | 32 | 35 | 42 | 39.3525 | 35 | 29 | 35 | 38.2134 | 29.22 | 35 | 32.062 |  |  | 28
| 1 | 3 | 34 | 37 | 41 | 36.3533 | 38 | 35 | 38 | 34 | 36 | 36 | 34 | 36.4 | 37.7 | 32 | 36 | 41 | 39.5495 | 35 | 29 | 35 | 38.6854 | 28.87 | 34 | 33.652 |  |  | 16
| 1 | 4 | 34 | 38 | 40 | 38.8098 | 39 | 35 | 38 | 34 | 36 | 35 | 35 | 36.8 | 37.1 | 33 | 35 | 41 | 39.381 | 35 | 29 | 36 | 38.1828 | 29.79 | 35 | 30.84 |  |  | 35.9976
| 1 | 5 | 35 | 39 | 41 | 41.455 | 39 | 36 | 39 | 33 | 36 | 35 | 35 | 38 | 38.5 | 32 | 35 | 41 | 39.5995 | 36 | 29 | 35 | 37.8318 | 29.85 | 35 | 31.539 |  |  | 84
| 1 | 6 | 35 | 39 | 42 | 41.7696 | 39 | 36 | 38 | 33 | 37 | 36 | 35 | 37.1 | 39.1 | 32 | 36 | 43 | 39.2665 | 35 | 30 | 36 | 37.0733 | 29.51 | 35 | 31.799
| Mean |  | 34.5 | 37.8333 | 41 | 39.3641 | 39 | 35.6667 | 38.1667 | 34 | 36 | 35.6667 | 35.5 | 36.75 | 38.4667 | 32.3333 | 35.3333 | 41.3333 | 39.4162 | 35.1667 | 28.8333 | 35.3333 | 37.9228 | 29.485 | 34.8333 | 32.0367
| Median |  | 34.5 | 37.5 | 41 | 39.4168 | 39 | 36 | 38 | 34 | 36 | 36 | 35 | 36.6 | 38.65 | 32 | 35 | 41 | 39.3667 | 35 | 29 | 35 | 38.0073 | 29.59 | 35 | 31.9305
| Std Dev. |  | 0.54772 | 0.98319 | 0.63246 | 2.1212 | 0.63246 | 0.5164 | 0.40825 | 0.89443 | 0.63246 | 0.5164 | 1.3784 | 0.73144 | 0.92232 | 0.5164 | 0.5164 | 1.0328 | 0.12935 | 0.40825 | 0.98319 | 0.5164 | 0.56613 | 0.37671 | 0.40825 | 0.94138
| Rel.Std.Dev. |  | 1.588% | 2.599% | 1.543% | 5.389% | 1.622% | 1.448% | 1.07% | 2.631% | 1.757% | 1.448% | 3.883% | 1.99% | 2.398% | 1.597% | 1.462% | 2.499% | 0.3282% | 1.161% | 3.41% | 1.462% | 1.493% | 1.278% | 1.172% | 2.938%
| PDM3 |  | -4.16% | 5.1% | 13.9% | 9.352% | 8.341% | -0.9192% | 6.026% | -5.549% | 0.006778% | -0.9192% | -1.382% | 2.09% | 6.859% | -10.18% | -1.845% | 14.82% | 9.497% | -2.308% | -19.9% | -1.845% | 5.348% | -18.09% | -3.234% | -11%
| Z-Score (Absolute) |  | 0.36 | 0.66 | 1.63 | 1.13 | 1.02 | 0 | 0.76 | 0.51 | 0.1 | 0 | 0.05 | 0.33 | 0.85 | 1.02 | 0.1 | 1.73 | 1.14 | 0.15 | 2.08 | 0.1 | 0.69 | 1.88 | 0.25 | 1.11
| Analytical results for Cs in OREAS 265 (Certified Value 1.27 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.22 | 1.36 | 1.69 | 1.2366 | 1.37 | 1.33 | 1.34 | 1.07 | 1.29 | 1.12 | 1.18 | 1.26 | 1.24 | 1.36 | 1.29 | 1.5101 | 1.23 | 1 | 1.3 | 1.3262 | 1.25 | [5 empty] | 1
| 1 | 2 | 1.22 | 1.4 | 1.68 | 1.1453 | 1.39 | 1.29 | 1.43 | 1.08 | 1.29 | 1.1 | 1.18 | 1.28 | 1.19 | 1.34 | 1.43 | 1.4965 | 1.18 | 1.05 | 1.3 | 1.3254 | 1.27 | [5 empty] | 29
| 1 | 3 | 1.23 | 1.43 | 1.7 | 1.1462 | 1.37 | 1.26 | 1.35 | 1.02 | 1.33 | 1.11 | 1.09 | 1.3 | 1.14 | 1.36 | 1.33 | 1.5515 | 1.21 | 1.04 | 1.3 | 1.3819 | 1.22 | [5 empty] | 16
| 1 | 4 | 1.21 | 1.49 | 1.69 | 1.1369 | 1.37 | 1.26 | 1.44 | 1.01 | 1.3 | 1.08 | 1.14 | 1.27 | 1.21 | 1.37 | 1.34 | 1.4954 | 1.18 | 1.09 | 1.4 | 1.3162 | 1.25 | [5 empty] | 1.2666
| 1 | 5 | 1.27 | 1.46 | 1.69 | 1.2376 | 1.35 | 1.3 | 1.37 | 1.07 | 1.29 | 1.09 | 1.1 | 1.33 | 1.16 | 1.32 | 1.37 | 1.5208 | 1.22 | 1.08 | 1.3 | 1.2675 | 1.25 | [5 empty] | 85
| 1 | 6 | 1.27 | 1.39 | 1.66 | 1.0989 | 1.4 | 1.31 | 1.4 | 1.05 | 1.33 | 1.1 | 1.11 | 1.33 | 1.18 | 1.37 | 1.35 | 1.5297 | 1.22 | 1.13 | 1.3 | 1.3115 | 1.26
| Mean |  | 1.2367 | 1.4217 | 1.685 | 1.1669 | 1.375 | 1.2917 | 1.3883 | 1.05 | 1.305 | 1.1 | 1.1333 | 1.295 | 1.1867 | 1.3533 | 1.3517 | 1.5173 | 1.2067 | 1.065 | 1.3167 | 1.3215 | 1.25
| Median |  | 1.225 | 1.415 | 1.69 | 1.1457 | 1.37 | 1.295 | 1.385 | 1.06 | 1.295 | 1.1 | 1.125 | 1.29 | 1.185 | 1.36 | 1.345 | 1.5155 | 1.215 | 1.065 | 1.3 | 1.3208 | 1.25
| Std Dev. |  | 0.026583 | 0.047924 | 0.013784 | 0.057066 | 0.017607 | 0.027869 | 0.041673 | 0.028983 | 0.019748 | 0.014142 | 0.039833 | 0.030166 | 0.03559 | 0.019664 | 0.046655 | 0.021443 | 0.021602 | 0.045056 | 0.040825 | 0.036689 | 0.016733
| Rel.Std.Dev. |  | 2.15% | 3.371% | 0.818% | 4.89% | 1.28% | 2.158% | 3.002% | 2.76% | 1.513% | 1.286% | 3.515% | 2.329% | 2.999% | 1.453% | 3.452% | 1.413% | 1.79% | 4.231% | 3.101% | 2.776% | 1.339%
| PDM3 |  | -2.365% | 12.24% | 33.03% | -7.871% | 8.557% | 1.978% | 9.609% | -17.1% | 3.03% | -13.15% | -10.52% | 2.241% | -6.312% | 6.846% | 6.715% | 19.79% | -4.733% | -15.92% | 3.951% | 4.33% | -1.312%
| Z-Score (Absolute) |  | 0.45 | 0.97 | 2.98 | 0.98 | 0.61 | 0.03 | 0.71 | 1.87 | 0.08 | 1.49 | 1.23 | 0 | 0.83 | 0.45 | 0.43 | 1.7 | 0.67 | 1.76 | 0.17 | 0.2 | 0.34
| Analytical results for Cu in OREAS 265 (Certified Value 20.3 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | 1.30 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.2g | 2
| 1 | 1 | 22.6 | 19.4 | 20.2 | 20.0592 | 20.4 | 20.4 | 19.9 | 22.1 | 20.2 | 19.7 | 20.8 | 19.67 | 19.32 | 19.3 | 21.4 | 19.8 | 18.4196 | 22 | 22.1 | 14.9 | 19.3 | 20.8262 | 15.14 | 16.293 | 21.3 | 20.078 | 1
| 1 | 2 | 20.7 | 20 | 20.2 | 20.7566 | 21 | 21.1 | 20.1 | 18.6 | 20.3 | 19.6 | 21.2 | 20.2 | 19.87 | 19.2 | 21.2 | 21.4 | 18.2468 | 21 | 22.1 | 17.1 | 18.9 | 20.9215 | 14.63 | 16.446 | 20.8 | 19.258 | 16
| 1 | 3 | 20.3 | 20.5 | 19.6 | 20.2651 | 20.3 | 20.9 | 20.5 | 18.9 | 20.5 | 19.8 | 20.1 | 19.9 | 19.57 | 19.3 | 21.4 | 19.6 | 18.2644 | 21 | 21.8 | 16 | 18.7 | 20.4752 | 14.5 | 16.131 | 20 | 19.792 | 16
| 1 | 4 | 20.9 | 20.9 | 20.1 | 21.0808 | 20.7 | 20.4 | 20.4 | 19.6 | 20.4 | 19.6 | 20.2 | 20.57 | 19.37 | 18.4 | 21.4 | 21 | 18.4868 | 21 | 22 | 16.4 | 18.9 | 20.11 | 14.31 | 16.376 | 21.1 | 20.26 | 20.3415
| 1 | 5 | 20.8 | 20.8 | 19.7 | 20.0683 | 20.4 | 20.9 | 20.6 | 19.3 | 20.4 | 19.8 | 20.1 | 20.64 | 20.35 | 18.3 | 20.8 | 20.2 | 18.3828 | 21 | 23.4 | 16.7 | 19.2 | 20.5688 | 14.534 | 16.026 | 20.3 | 20.005 | 86
| 1 | 6 | 20.9 | 20.4 | 20.1 | 20.9354 | 20.6 | 21 | 20.2 | 20.6 | 20.4 | 19.6 | 20.1 | 20.12 | 19.93 | 18.9 | 21.7 | 19.9 | 18.444 | 21 | 21.7 | 17.7 | 19.2 | 20.4457 | 15.05 | 15.963 | 21.4 | 20.808
| Mean |  | 21.0333 | 20.3333 | 19.9833 | 20.5276 | 20.5667 | 20.7833 | 20.2833 | 19.85 | 20.3667 | 19.6833 | 20.4167 | 20.1833 | 19.735 | 18.9 | 21.3167 | 20.3167 | 18.3741 | 21.1667 | 22.1833 | 16.4667 | 19.0333 | 20.5579 | 14.694 | 16.2058 | 20.8167 | 20.0335
| Median |  | 20.85 | 20.45 | 20.1 | 20.5108 | 20.5 | 20.9 | 20.3 | 19.45 | 20.4 | 19.65 | 20.15 | 20.16 | 19.72 | 19.05 | 21.4 | 20.05 | 18.4012 | 21 | 22.05 | 16.55 | 19.05 | 20.522 | 14.582 | 16.212 | 20.95 | 20.0415
| Std Dev. |  | 0.79917 | 0.55737 | 0.26394 | 0.45256 | 0.2582 | 0.30605 | 0.26394 | 1.3004 | 0.10328 | 0.098319 | 0.47081 | 0.3758 | 0.3917 | 0.45166 | 0.29944 | 0.72226 | 0.097952 | 0.40825 | 0.61779 | 0.96471 | 0.23381 | 0.29131 | 0.32877 | 0.19552 | 0.56362 | 0.51258
| Rel.Std.Dev. |  | 3.8% | 2.741% | 1.321% | 2.205% | 1.255% | 1.473% | 1.301% | 6.551% | 0.5071% | 0.4995% | 2.306% | 1.862% | 1.985% | 2.39% | 1.405% | 3.555% | 0.5331% | 1.929% | 2.785% | 5.859% | 1.228% | 1.417% | 2.237% | 1.206% | 2.708% | 2.559%
| PDM3 |  | 3.401% | -0.04036% | -1.761% | 0.9145% | 1.107% | 2.172% | -0.2862% | -2.416% | 0.1235% | -3.236% | 0.3693% | -0.7778% | -2.982% | -7.087% | 4.794% | -0.1223% | -9.672% | 4.056% | 9.054% | -19.05% | -6.431% | 1.064% | -27.76% | -20.33% | 2.336% | -1.514%
| Z-Score (Absolute) |  | 0.99 | 0.04 | 0.43 | 0.31 | 0.36 | 0.65 | 0.02 | 0.61 | 0.09 | 0.83 | 0.16 | 0.16 | 0.76 | 1.89 | 1.37 | 0.02 | 2.6 | 1.17 | 2.54 | 5.17 | 1.71 | 0.35 | 7.56 | 5.52 | 0.7 | 0.36
| Analytical results for Dy in OREAS 265 (Certified Value 2.33 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.17 | 1.20 | 1.22 | [19 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | [19 empty] | 3
| 1 | 1 | 2.3 | 2.296 | 2.17 | 2.3 | 3.2006 | 2.22 | 2.2957 | [19 empty] | 1
| 1 | 2 | 2.33 | 2.303 | 2.25 | 2.6 | 3.1858 | 2.39 | 2.3434 | [19 empty] | 31
| 1 | 3 | 2.3 | 2.234 | 2.23 | 2.4 | 3.2195 | 2.35 | 2.2877 | [19 empty] | 16
| 1 | 4 | 2.24 | 2.279 | 2.17 | 2.5 | 3.2489 | 2.4 | 2.3167 | [19 empty] | 2.3257
| 1 | 5 | 2.28 | 2.317 | 2.28 | 2.4 | 3.2327 | 2.4 | 2.3622 | [19 empty] | 87
| 1 | 6 | 2.35 | 2.338 | 2.1 | 2.4 | 3.1856 | 2.48 | 2.3297
| Mean |  | 2.3 | 2.2945 | 2.2 | 2.4333 | 3.2122 | 2.3733 | 2.3225
| Median |  | 2.3 | 2.2995 | 2.2 | 2.4 | 3.21 | 2.395 | 2.3232
| Std Dev. |  | 0.038471 | 0.035703 | 0.065727 | 0.10328 | 0.025921 | 0.086178 | 0.028371
| Rel.Std.Dev. |  | 1.673% | 1.556% | 2.988% | 4.244% | 0.807% | 3.631% | 1.222%
| PDM3 |  | -1.106% | -1.343% | -5.406% | 4.627% | 38.12% | 2.047% | -0.1369%
| Z-Score (Absolute) |  | 0.3 | 0.37 | 1.63 | 1.47 | 11.81 | 0.67 | 0
| Analytical results for Er in OREAS 265 (Certified Value 0.87 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.17 | 1.20 | 1.22 | [19 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | [19 empty] | 3
| 1 | 1 | 0.93 | 0.845 | 0.83 | 0.9 | 1.2381 | 0.82 | 0.86881 | [19 empty] | 1
| 1 | 2 | 0.93 | 0.855 | 0.82 | 0.9 | 1.2442 | 0.88 | 0.84662 | [19 empty] | 32
| 1 | 3 | 0.93 | 0.833 | 0.84 | 0.9 | 1.2835 | 0.87 | 0.89511 | [19 empty] | 16
| 1 | 4 | 0.9 | 0.831 | 0.84 | 0.9 | 1.2554 | 0.89 | 0.88035 | [19 empty] | 0.87272
| 1 | 5 | 0.92 | 0.839 | 0.85 | 0.9 | 1.2803 | 0.89 | 0.88032 | [19 empty] | 88
| 1 | 6 | 0.94 | 0.831 | 0.82 | 0.9 | 1.2348 | 0.9 | 0.91039
| Mean |  | 0.925 | 0.839 | 0.83333 | 0.9 | 1.2561 | 0.875 | 0.88027
| Median |  | 0.93 | 0.836 | 0.835 | 0.9 | 1.2498 | 0.885 | 0.88034
| Std Dev. |  | 0.013784 | 0.0095499 | 0.012111 | 0 | 0.021242 | 0.02881 | 0.021867
| Rel.Std.Dev. |  | 1.49% | 1.138% | 1.453% | 0% | 1.691% | 3.293% | 2.484%
| PDM3 |  | 5.99% | -3.864% | -4.513% | 3.126% | 43.92% | 0.2612% | 0.8649%
| Z-Score (Absolute) |  | 0.77 | 0.63 | 0.72 | NA | 6.15 | 0.04 | 0.04
Results from laboratory 1.16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 265 (Certified Value 0.97 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.17 | 1.20 | 1.22 | [19 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | [19 empty] | 3
| 1 | 1 | 0.8 | 0.981 | 1.02 | 1 | 0.97378 | 0.84 | 0.9434 | [19 empty] | 1
| 1 | 2 | 0.82 | 0.984 | 1 | 1 | 0.96405 | 0.92 | 0.90967 | [19 empty] | 33
| 1 | 3 | 0.81 | 0.94 | 0.99 | 1 | 0.9833 | 0.9 | 0.94318 | [19 empty] | 16
| 1 | 4 | 0.79 | 0.979 | 0.96 | 1.1 | 0.97504 | 0.92 | 0.92517 | [19 empty] | 0.96964
| 1 | 5 | 0.82 | 0.984 | 1.01 | 1.1 | 0.96538 | 0.92 | 0.91336 | [19 empty] | 89
| 1 | 6 | 0.81 | 1.007 | 0.92 | 1 | 0.95983 | 0.96 | 0.93173
| Mean |  | 0.80833 | 0.97917 | 0.98333 | 1.0333 | 0.97023 | 0.91 | 0.92775
| Median |  | 0.81 | 0.9825 | 0.995 | 1 | 0.96958 | 0.92 | 0.92845
| Std Dev. |  | 0.01169 | 0.02172 | 0.037238 | 0.05164 | 0.0086768 | 0.039497 | 0.014419
| Rel.Std.Dev. |  | 1.446% | 2.218% | 3.787% | 4.997% | 0.8943% | 4.34% | 1.554%
| PDM3 |  | -16.64% | 0.9829% | 1.413% | 6.569% | 0.06129% | -6.15% | -4.32%
| Z-Score (Absolute) |  | 2.57 | 0.14 | 0.21 | 1 | 0 | 0.96 | 0.67
| Analytical results for Fe in OREAS 265 (Certified Value 2.53 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  | 3
| 1 | 1 | 2.43 | 2.46 | 2.79 | 2.7009 | 2.59 | 2.66 | 2.6 | 2.44 | 2.47 | 2.45 | 2.47 | 2.47 | 2.63 | 2.11 | 2.47 | 2.57 | 2.7518 | 2.0207 | 2.51 | 1.728 | 2.58 | 2.6479 | 1.8228 | 2.5 | 2.447 |  | 1
| 1 | 2 | 2.43 | 2.47 | 2.81 | 2.6433 | 2.68 | 2.67 | 2.61 | 2.4 | 2.5 | 2.46 | 2.34 | 2.45 | 2.59 | 2.08 | 2.48 | 2.71 | 2.7616 | 1.9774 | 2.5 | 1.812 | 2.56 | 2.6492 | 1.7361 | 2.55 | 2.3275 |  | 38
| 1 | 3 | 2.4 | 2.46 | 2.81 | 2.5286 | 2.6 | 2.64 | 2.65 | 2.38 | 2.53 | 2.48 | 2.28 | 2.41 | 2.53 | 2.05 | 2.56 | 2.64 | 2.7693 | 2.008 | 2.49 | 1.7 | 2.58 | 2.7197 | 1.7967 | 2.42 | 2.3111 |  | 16
| 1 | 4 | 2.43 | 2.48 | 2.89 | 2.5975 | 2.62 | 2.61 | 2.62 | 2.35 | 2.52 | 2.4 | 2.3 | 2.51 | 2.47 | 2.13 | 2.48 | 2.63 | 2.7791 | 2.0258 | 2.52 | 1.749 | 2.56 | 2.6369 | 1.8892 | 2.47 | 2.4053 |  | 2.5259
| 1 | 5 | 2.45 | 2.56 | 2.84 | 2.5779 | 2.61 | 2.63 | 2.6 | 2.3 | 2.53 | 2.4 | 2.35 | 2.53 | 2.58 | 2.1 | 2.48 | 2.66 | 2.7601 | 2.0284 | 2.58 | 1.742 | 2.58 | 2.6327 | 1.761 | 2.5 | 2.4266 |  | 90
| 1 | 6 | 2.47 | 2.48 | 2.95 | 2.4855 | 2.61 | 2.64 | 2.57 | 2.3 | 2.55 | 2.47 | 2.32 | 2.51 | 2.59 | 2.09 | 2.57 | 2.75 | 2.7559 | 2.0113 | 2.52 | 1.882 | 2.57 | 2.6035 | 1.8117 | 2.47 | 2.4047
| Mean |  | 2.435 | 2.485 | 2.8483 | 2.589 | 2.6183 | 2.6417 | 2.6083 | 2.3617 | 2.5167 | 2.4433 | 2.3433 | 2.48 | 2.565 | 2.0933 | 2.5067 | 2.66 | 2.763 | 2.0119 | 2.52 | 1.7688 | 2.5717 | 2.6483 | 1.8029 | 2.485 | 2.387
| Median |  | 2.43 | 2.475 | 2.825 | 2.5877 | 2.61 | 2.64 | 2.605 | 2.365 | 2.525 | 2.455 | 2.33 | 2.49 | 2.585 | 2.095 | 2.48 | 2.65 | 2.7608 | 2.016 | 2.515 | 1.7455 | 2.575 | 2.6424 | 1.8042 | 2.485 | 2.405
| Std Dev. |  | 0.023452 | 0.037815 | 0.0608 | 0.077503 | 0.031885 | 0.02137 | 0.026394 | 0.056006 | 0.028048 | 0.035024 | 0.067132 | 0.045166 | 0.05648 | 0.027325 | 0.045461 | 0.063246 | 0.0098555 | 0.018701 | 0.031623 | 0.066617 | 0.0098319 | 0.03866 | 0.053264 | 0.043243 | 0.054988
| Rel.Std.Dev. |  | 0.9631% | 1.522% | 2.135% | 2.994% | 1.218% | 0.8089% | 1.012% | 2.371% | 1.114% | 1.433% | 2.865% | 1.821% | 2.202% | 1.305% | 1.814% | 2.378% | 0.3567% | 0.9295% | 1.255% | 3.766% | 0.3823% | 1.46% | 2.954% | 1.74% | 2.304%
| PDM3 |  | -3.6% | -1.621% | 12.76% | 2.495% | 3.658% | 4.582% | 3.262% | -6.503% | -0.367% | -3.27% | -7.229% | -1.819% | 1.546% | -17.13% | -0.7629% | 5.307% | 9.383% | -20.35% | -0.235% | -29.97% | 1.81% | 4.845% | -28.62% | -1.621% | -5.499%
| Z-Score (Absolute) |  | 0.43 | 0.13 | 2.06 | 0.5 | 0.67 | 0.82 | 0.61 | 0.88 | 0.06 | 0.38 | 0.99 | 0.16 | 0.35 | 2.5 | 0 | 0.93 | 1.55 | 2.99 | 0.08 | 4.46 | 0.39 | 0.86 | 4.25 | 0.13 | 0.72
| Analytical results for Ga in OREAS 265 (Certified Value 4.74 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 4 | 4.67 | 4.77 | 4.4471 | 4.48 | 4.6 | 5.33 | 4.24 | 4.38 | 4.13 | 4.73 | 4.3 | 5 | 2.82 | 4.1 | 5.94 | 5.702 | 4.71 | 2.77 | 6 | 5.4291 | 4.6886 | 4.82 |  |  |  | 1
| 1 | 2 | 4.1 | 4.92 | 4.83 | 4.7357 | 4.72 | 4.6 | 5.38 | 4.15 | 4.55 | 3.96 | 4.79 | 4.2 | 5.2 | 2.82 | 4.1 | 5.83 | 5.7073 | 4.5 | 3.05 | 6 | 5.4764 | 4.7932 | 4.74 |  |  |  | 34
| 1 | 3 | 4.1 | 5.11 | 4.79 | 4.7168 | 4.52 | 4.5 | 5.09 | 4.18 | 4.49 | 4.1 | 4.43 | 4.1 | 5 | 2.77 | 4.1 | 5.53 | 5.6699 | 4.56 | 2.98 | 6 | 5.4262 | 4.9465 | 4.66 |  |  |  | 16
| 1 | 4 | 4.1 | 5.09 | 4.79 | 4.754 | 4.7 | 4.4 | 5.43 | 4.33 | 4.45 | 4.01 | 4.48 | 4.3 | 5 | 2.64 | 4.1 | 5.98 | 5.706 | 4.58 | 3 | 6 | 5.1456 | 4.5836 | 4.78 |  |  |  | 4.7417
| 1 | 5 | 4.2 | 4.93 | 4.81 | 4.4817 | 4.55 | 4.5 | 5.18 | 4.43 | 4.44 | 4.07 | 4.53 | 4.4 | 5.2 | 2.76 | 4 | 6.02 | 5.7333 | 4.58 | 3.02 | 6 | 5.5647 | 4.7296 | 4.83 |  |  |  | 91
| 1 | 6 | 4.1 | 5.04 | 4.83 | 4.5552 | 4.66 | 4.7 | 5.36 | 4.56 | 4.56 | 4.04 | 4.43 | 4.3 | 5.2 | 2.78 | 4.1 | 5.44 | 5.7564 | 4.62 | 3.17 | 6 | 5.4915 | 4.4833 | 4.79
| Mean |  | 4.1 | 4.96 | 4.8033 | 4.6151 | 4.605 | 4.55 | 5.295 | 4.315 | 4.4783 | 4.0517 | 4.565 | 4.2667 | 5.1 | 2.765 | 4.0833 | 5.79 | 5.7125 | 4.5917 | 2.9983 | 6 | 5.4223 | 4.7041 | 4.77
| Median |  | 4.1 | 4.985 | 4.8 | 4.636 | 4.605 | 4.55 | 5.345 | 4.285 | 4.47 | 4.055 | 4.505 | 4.3 | 5.1 | 2.775 | 4.1 | 5.885 | 5.7066 | 4.58 | 3.01 | 6 | 5.4528 | 4.7091 | 4.785
| Std Dev. |  | 0.063246 | 0.16273 | 0.024221 | 0.13697 | 0.10114 | 0.10488 | 0.13126 | 0.15783 | 0.069113 | 0.061779 | 0.15668 | 0.10328 | 0.10954 | 0.066257 | 0.040825 | 0.24625 | 0.029491 | 0.069976 | 0.13045 | 0 | 0.14464 | 0.16179 | 0.06261
| Rel.Std.Dev. |  | 1.543% | 3.281% | 0.5043% | 2.968% | 2.196% | 2.305% | 2.479% | 3.658% | 1.543% | 1.525% | 3.432% | 2.421% | 2.148% | 2.396% | 0.9998% | 4.253% | 0.5162% | 1.524% | 4.351% | 0% | 2.667% | 3.439% | 1.313%
| PDM3 |  | -13.53% | 4.604% | 1.3% | -2.67% | -2.883% | -4.043% | 11.67% | -8.999% | -5.555% | -14.55% | -3.727% | -10.02% | 7.556% | -41.69% | -13.88% | 22.11% | 20.47% | -3.164% | -36.77% | 26.54% | 14.35% | -0.7927% | 0.5965%
| Z-Score (Absolute) |  | 0.96 | 0.67 | 0.38 | 0.02 | 0 | 0.1 | 1.31 | 0.55 | 0.24 | 1.05 | 0.08 | 0.64 | 0.94 | 3.5 | 0.99 | 2.25 | 2.1 | 0.03 | 3.05 | 2.65 | 1.55 | 0.19 | 0.31
| Analytical results for Gd in OREAS 265 (Certified Value 3.66 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.20 | 1.22 | [20 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.5g | 0.2g | [20 empty] | 3
| 1 | 1 | 3.38 | 3.846 | 3.43 | 3.8 | 4.1 | 3.5765 | [20 empty] | 1
| 1 | 2 | 3.45 | 3.807 | 3.56 | 3.9 | 4.5 | 3.7357 | [20 empty] | 35
| 1 | 3 | 3.59 | 3.789 | 3.58 | 4 | 4.32 | 3.614 | [20 empty] | 16
| 1 | 4 | 3.27 | 3.852 | 3.58 | 3.9 | 4.51 | 3.62 | [20 empty] | 3.6635
| 1 | 5 | 3.27 | 3.84 | 3.62 | 3.9 | 4.5 | 3.6951 | [20 empty] | 92
| 1 | 6 | 3.35 | 3.755 | 3.5 | 4 | 4.61 | 3.695
| Mean |  | 3.385 | 3.8148 | 3.545 | 3.9167 | 4.4233 | 3.6561
| Median |  | 3.365 | 3.8235 | 3.57 | 3.9 | 4.5 | 3.6575
| Std Dev. |  | 0.12161 | 0.038196 | 0.068629 | 0.075277 | 0.18403 | 0.061277
| Rel.Std.Dev. |  | 3.593% | 1.001% | 1.936% | 1.922% | 4.16% | 1.676%
| PDM3 |  | -7.602% | 4.131% | -3.235% | 6.91% | 20.74% | -0.2036%
| Z-Score (Absolute) |  | 1.27 | 0.29 | 0.69 | 0.66 | 2.5 | 0.29
| Analytical results for Ge in OREAS 265 (Certified Value < 0.05 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [9 empty] | 3
|  |  | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.1 | < 0.05 | < 0.1 | 0.19 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.1 | < 0.05 | 0.06 | < 0.01 | 1.6 | < 0.2 | < 0.05 | [9 empty] | 1
| 1 | 2 | < 0.1 | < 0.05 | < 0.1 | 0.19 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.06 | < 0.1 | < 0.05 | 0.06 | < 0.01 | 1.6 | < 0.2 | < 0.05 | [9 empty] | 36
| 1 | 3 | < 0.1 | < 0.05 | < 0.1 | 0.19 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.06 | 0.1 | < 0.05 | 0.09 | < 0.01 | 1.6 | < 0.2 | < 0.05 | [9 empty] | 16
| 1 | 4 | < 0.1 | < 0.05 | < 0.1 | 0.19 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.06 | 0.1 | 0.13 | 0.07 | < 0.01 | 1.6 | < 0.2 | < 0.05 | [9 empty] | < 0.05
| 1 | 5 | < 0.1 | < 0.05 | < 0.1 | 0.18 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.06 | < 0.1 | < 0.05 | 0.07 | < 0.01 | 1.6 | < 0.2 | < 0.05 | [9 empty] | 93
| 1 | 6 | < 0.1 | < 0.05 | < 0.1 | 0.19 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.06 | 0.1 | < 0.05 | 0.08 | < 0.01 | 1.7 | < 0.2 | < 0.05
| Mean |  |  |  |  | 0.18833 |  |  | 0.05 |  |  | 0.061667 | 0.1 | 0.13 | 0.071667 |  | 1.6167
| Median |  |  |  |  | 0.19 |  |  | 0.05 |  |  | 0.06 | 0.1 | 0.13 | 0.07 |  | 1.6
| Std Dev. |  |  |  |  | 0.0040825 |  |  | 0 |  |  | 0.0040825 | 1.6997e-17 |  | 0.01169 |  | 0.040825
| Rel.Std.Dev. |  |  |  |  | 2.168% |  |  | 0% |  |  | 6.62% | 1.7e-14% |  | 16.31% |  | 2.525%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 0 | 3.73 | 0.67 | 0.67 | 0.45 | 0.67 | 0 | 0.31 | 0.67 | 0.2 | 0.58 | 1.21 | 42.26 | 1.35 | 0.67
| Analytical results for Hf in OREAS 265 (Certified Value 0.56 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.69 | 0.59 | 0.74 | 0.69566 | 0.11 | 0.69 | 0.29 | 0.38 | 0.63 | 0.41 | 0.46 | 0.21 | 0.53 | 0.78 | < 0.1 | 0.36 | 0.52 | 0.6 | 0.62173 | 0.72 | [6 empty] | 1
| 1 | 2 | 0.73 | 0.47 | 0.77 | 0.69095 | 0.13 | 0.69 | 0.32 | 0.45 | 0.64 | 0.39 | 0.45 | 0.34 | 0.55 | 0.78 | < 0.1 | 0.35 | 0.57 | 0.6 | 0.60243 | 0.73 | [6 empty] | 20
| 1 | 3 | 0.7 | 0.52 | 0.73 | 0.69902 | 0.11 | 0.64 | 0.28 | 0.4 | 0.64 | 0.4 | 0.4 | 0.34 | 0.62 | 0.78 | < 0.1 | 0.44 | 0.55 | 0.6 | 0.61223 | 0.72 | [6 empty] | 16
| 1 | 4 | 0.65 | 0.55 | 0.73 | 0.6979 | 0.11 | 0.66 | 0.32 | 0.35 | 0.67 | 0.37 | 0.42 | 0.31 | 0.57 | 0.79 | < 0.1 | 0.38 | 0.58 | 0.7 | 0.62342 | 0.63 | [6 empty] | 0.55511
| 1 | 5 | 0.7 | 0.54 | 0.74 | 0.70204 | 0.1 | 0.65 | 0.32 | 0.42 | 0.63 | 0.39 | 0.39 | 0.34 | 0.57 | 0.76 | < 0.1 | 0.39 | 0.58 | 0.7 | 0.61019 | 0.69 | [6 empty] | 94
| 1 | 6 | 0.7 | 0.56 | 0.75 | 0.68837 | 0.1 | 0.65 | 0.33 | 0.35 | 0.65 | 0.43 | 0.38 | 0.26 | 0.58 | 0.77 | < 0.1 | 0.46 | 0.62 | 0.6 | 0.62573 | 0.7
| Mean |  | 0.695 | 0.53833 | 0.74333 | 0.69566 | 0.11 | 0.66333 | 0.31 | 0.39167 | 0.64333 | 0.39833 | 0.41667 | 0.3 | 0.57 | 0.77667 |  | 0.39667 | 0.57 | 0.63333 | 0.61596 | 0.69833
| Median |  | 0.7 | 0.545 | 0.74 | 0.69678 | 0.11 | 0.655 | 0.32 | 0.39 | 0.64 | 0.395 | 0.41 | 0.325 | 0.57 | 0.78 |  | 0.385 | 0.575 | 0.6 | 0.61698 | 0.71
| Std Dev. |  | 0.025884 | 0.040702 | 0.015055 | 0.0051442 | 0.010954 | 0.021602 | 0.02 | 0.039707 | 0.015055 | 0.020412 | 0.03266 | 0.054037 | 0.030332 | 0.010328 |  | 0.044121 | 0.033466 | 0.05164 | 0.0091068 | 0.03656
| Rel.Std.Dev. |  | 3.724% | 7.561% | 2.025% | 0.7395% | 9.959% | 3.257% | 6.452% | 10.14% | 2.34% | 5.124% | 7.838% | 18.01% | 5.321% | 1.33% |  | 11.12% | 5.871% | 8.154% | 1.478% | 5.235%
| PDM3 |  | 25.2% | -3.023% | 33.91% | 25.32% | -80.18% | 19.49% | -44.16% | -29.44% | 15.89% | -28.24% | -24.94% | -45.96% | 2.682% | 39.91% |  | -28.54% | 2.682% | 14.09% | 10.96% | 25.8%
| Z-Score (Absolute) |  | 0.55 | 0.14 | 0.76 | 0.55 | 2.02 | 0.41 | 1.14 | 0.78 | 0.32 | 0.75 | 0.67 | 1.19 | 0 | 0.91 | 2.29 | 0.76 | 0 | NA | 0.2 | 0.56
Results from laboratory 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 265 (Certified Value 0.013 ppm) | [27 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | [5 empty] | 3
|  |  | 0.25g | 0.5g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.02 | 0.01 | < 0.05 | < 0.1 | 0.02 | 0.11 | 0.01 | < 0.1 | 0.01 | 0.01 | 0.009 | < 0.005 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.166 | < 1 | 0.021262 | 0.015 | [5 empty] | 1
| 1 | 2 | 0.02 | 0.01 | < 0.05 | < 0.1 | 0.04 | 0.08 | 0.01 | < 0.1 | 0.01 | 0.01 | 0.008 | < 0.005 | 0.01 | 0.01 | 0.02 | 0.02 | 0.01 | 0.123 | < 1 | 0.02 | 0.014 | [5 empty] | 4
| 1 | 3 | 0.01 | 0.01 | < 0.05 | < 0.1 | 0.02 | 0.07 | 0.01 | < 0.1 | 0.01 | 0.01 | 0.01 | < 0.005 | < 0.01 | < 0.01 | 0.03 | 0.01 | 0.01 | 0.126 | < 1 | 0.018247 | 0.011 | [5 empty] | 16
| 1 | 4 | 0.02 | 0.01 | < 0.05 | < 0.1 | 0.02 | 0.08 | 0.01 | < 0.1 | 0.01 | 0.02 | 0.01 | < 0.005 | 0.01 | 0.01 | 0.02 | 0.02 | 0.01 | 0.106 | < 1 | 0.02 | 0.014 | [5 empty] | 0.013244
| 1 | 5 | 0.02 | 0.01 | < 0.05 | < 0.1 | 0.04 | 0.1 | 0.01 | < 0.1 | 0.01 | 0.01 | 0.005 | < 0.005 | 0.01 | < 0.01 | 0.02 | 0.03 | 0.01 | 0.102 | < 1 | 0.02 | 0.014 | [5 empty] | 95
| 1 | 6 | 0.02 | 0.01 | < 0.05 | < 0.1 | 0.03 | 0.1 | 0.01 | < 0.1 | 0.01 | 0.02 | 0.006 | < 0.005 | < 0.01 | 0.01 | 0.02 | 0.01 | 0.01 | 0.103 | < 1 | 0.019299 | 0.014
| Mean |  | 0.018333 | 0.01 |  |  | 0.028333 | 0.09 | 0.01 |  | 0.01 | 0.013333 | 0.008 |  | 0.01 | 0.0125 | 0.021667 | 0.018333 | 0.01 | 0.121 |  | 0.019801 | 0.013667
| Median |  | 0.02 | 0.01 |  |  | 0.025 | 0.09 | 0.01 |  | 0.01 | 0.01 | 0.0085 |  | 0.01 | 0.01 | 0.02 | 0.02 | 0.01 | 0.1145 |  | 0.02 | 0.014
| Std Dev. |  | 0.0040825 | 0 |  |  | 0.0098319 | 0.015492 | 0 |  | 0 | 0.005164 | 0.0020976 |  | 0 | 0.005 | 0.0040825 | 0.0075277 | 0 | 0.024347 |  | 0.00099184 | 0.0013663
| Rel.Std.Dev. |  | 22.27% | 0% |  |  | 34.7% | 17.21% | 0% |  | 0% | 38.73% | 26.22% |  | 0% | 40% | 18.84% | 41.06% | 0% | 20.12% |  | 5.009% | 9.997%
| PDM3 |  | 38.43% | -24.49% |  |  | 113.9% | 579.6% | -24.49% |  | -24.49% | 0.6769% | -39.59% |  | -24.49% | -5.615% | 63.6% | 38.43% | -24.49% | 813.6% |  | 49.52% | 3.194%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.54 | 2.56 | 0.81 | 5.8 | 0.67 | 2.56 | 0.67 | 0.4 | 0.84 | 1.28 | 0.81 | 0.67 | 0.27 | 0 | 0.67 | 8.31 | 38.98 | 0.12 | 0.38
| Analytical results for Ho in OREAS 265 (Certified Value 0.36 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.20 | 1.22 | [20 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [20 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.5g | 0.2g | [20 empty] | 3
| 1 | 1 | 0.38 | 0.4 | 0.34 | 0.3 | 0.28 | 0.36013 | [20 empty] | 1
| 1 | 2 | 0.38 | 0.4 | 0.33 | 0.4 | 0.31 | 0.33034 | [20 empty] | 21
| 1 | 3 | 0.37 | 0.4 | 0.35 | 0.4 | 0.3 | 0.35472 | [20 empty] | 16
| 1 | 4 | 0.37 | 0.4 | 0.35 | 0.4 | 0.31 | 0.35721 | [20 empty] | 0.36075
| 1 | 5 | 0.37 | 0.4 | 0.35 | 0.4 | 0.31 | 0.33782 | [20 empty] | 96
| 1 | 6 | 0.39 | 0.4 | 0.36 | 0.4 | 0.32 | 0.34679
| Mean |  | 0.37667 | 0.4 | 0.34667 | 0.38333 | 0.305 | 0.34784
| Median |  | 0.375 | 0.4 | 0.35 | 0.4 | 0.31 | 0.35076
| Std Dev. |  | 0.008165 | 6.0809e-17 | 0.010328 | 0.040825 | 0.013784 | 0.011779
| Rel.Std.Dev. |  | 2.168% | 1.52e-14% | 2.979% | 10.65% | 4.519% | 3.387%
| PDM3 |  | 4.412% | 10.88% | -3.904% | 6.26% | -15.45% | -3.58%
| Z-Score (Absolute) |  | 0.53 | 1.39 | 0.57 | 0.78 | 2.11 | 0.53
| Analytical results for In in OREAS 265 (Certified Value 0.019 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.29 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [7 empty] | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.02 | 0.018 | 0.02 | < 0.05 | 0.02 | 0.018 | 0.031 | 0.017 | 0.02 | 0.017 | 0.016 | < 0.02 | 0.009 | 0.02 | < 0.02 | 0.02 | 0.01 | < 0.2 | 0.017 | [7 empty] | 1
| 1 | 2 | < 0.02 | 0.02 | 0.02 | < 0.05 | 0.02 | 0.018 | 0.024 | 0.016 | 0.02 | 0.017 | 0.02 | < 0.02 | 0.009 | 0.02 | 0.02 | 0.023 | 0.01 | < 0.2 | 0.017 | [7 empty] | 22
| 1 | 3 | < 0.02 | 0.02 | 0.02 | < 0.05 | 0.02 | 0.018 | 0.025 | 0.017 | 0.02 | 0.015 | 0.017 | < 0.02 | 0.012 | 0.02 | < 0.02 | 0.021 | 0.01 | < 0.2 | 0.017 | [7 empty] | 16
| 1 | 4 | < 0.02 | 0.019 | 0.02 | < 0.05 | 0.02 | 0.017 | 0.027 | 0.016 | 0.02 | 0.016 | 0.016 | < 0.02 | 0.012 | 0.02 | < 0.02 | 0.021 | 0.01 | < 0.2 | 0.016 | [7 empty] | 0.018897
| 1 | 5 | < 0.02 | 0.019 | 0.02 | < 0.05 | 0.02 | 0.018 | 0.027 | 0.018 | 0.02 | 0.018 | 0.018 | < 0.02 | 0.012 | < 0.02 | < 0.02 | 0.018 | 0.01 | < 0.2 | 0.016 | [7 empty] | 97
| 1 | 6 | 0.02 | 0.019 | 0.02 | < 0.05 | 0.02 | 0.02 | 0.027 | 0.018 | 0.02 | 0.018 | 0.018 | < 0.02 | 0.013 | 0.02 | < 0.02 | 0.019 | 0.01 | < 0.2 | 0.017
| Mean |  | 0.02 | 0.019167 | 0.02 |  | 0.02 | 0.018167 | 0.026833 | 0.017 | 0.02 | 0.016833 | 0.0175 |  | 0.011167 | 0.02 | 0.02 | 0.020333 | 0.01 |  | 0.016667
| Median |  | 0.02 | 0.019 | 0.02 |  | 0.02 | 0.018 | 0.027 | 0.017 | 0.02 | 0.017 | 0.0175 |  | 0.012 | 0.02 | 0.02 | 0.0205 | 0.01 |  | 0.017
| Std Dev. |  |  | 0.00075277 | 0 |  | 0 | 0.00098319 | 0.0024014 | 0.00089443 | 0 | 0.001169 | 0.0015166 |  | 0.0017224 | 0 |  | 0.0017512 | 0 |  | 0.0005164
| Rel.Std.Dev. |  |  | 3.928% | 0% |  | 0% | 5.412% | 8.949% | 5.261% | 0% | 6.945% | 8.666% |  | 15.42% | 0% |  | 8.612% | 0% |  | 3.098%
| PDM3 |  | 5.834% | 1.425% | 5.834% |  | 5.834% | -3.867% | 41.99% | -10.04% | 5.834% | -10.92% | -7.395% |  | -40.91% | 5.834% | 5.834% | 7.598% | -47.08% |  | -11.8%
| Z-Score (Absolute) |  | 2.39 | 0.37 | 0.67 | 2.51 | 0.67 | 0 | 3.19 | 0.43 | 0.67 | 0.49 | 0.25 | 3 | 2.57 | 0.06 | 2.39 | 0.8 | 3 | 30.1 | 0.55
| Analytical results for K in OREAS 265 (Certified Value 0.254 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  |  | 3
| 1 | 1 | 0.24 | 0.25 | 0.293 | 0.26638 | 0.2547 | 0.255 | 0.27 | 0.24 | 0.2549 | 0.25 | 0.24 | 0.25 | 0.27 | 0.25 | 0.24 | 0.27 | 0.26 | 0.141 | 0.27 | 0.26468 | 0.23625 | 0.25 | 0.27906 |  |  |  | 1
| 1 | 2 | 0.24 | 0.26 | 0.291 | 0.26791 | 0.2626 | 0.252 | 0.28 | 0.24 | 0.2594 | 0.26 | 0.24 | 0.25 | 0.26 | 0.24 | 0.23 | 0.28 | 0.25 | 0.149 | 0.27 | 0.26334 | 0.23769 | 0.25 | 0.21736 |  |  |  | #N/A
| 1 | 3 | 0.24 | 0.26 | 0.29 | 0.24397 | 0.2563 | 0.25 | 0.27 | 0.24 | 0.2633 | 0.26 | 0.24 | 0.24 | 0.25 | 0.24 | 0.24 | 0.27 | 0.26 | 0.158 | 0.27 | 0.27436 | 0.23932 | 0.24 | 0.25848 |  |  |  | 16
| 1 | 4 | 0.24 | 0.26 | 0.292 | 0.25279 | 0.2559 | 0.242 | 0.28 | 0.24 | 0.2612 | 0.25 | 0.24 | 0.25 | 0.24 | 0.24 | 0.24 | 0.27 | 0.26 | 0.158 | 0.27 | 0.26718 | 0.23538 | 0.25 | 0.21765 |  |  |  | 0.25383
| 1 | 5 | 0.24 | 0.26 | 0.288 | 0.25313 | 0.2519 | 0.249 | 0.27 | 0.23 | 0.2623 | 0.25 | 0.24 | 0.25 | 0.26 | 0.24 | 0.24 | 0.28 | 0.26 | 0.158 | 0.27 | 0.26631 | 0.2407 | 0.25 | 0.225 |  |  |  | 98
| 1 | 6 | 0.24 | 0.26 | 0.292 | 0.24418 | 0.2583 | 0.254 | 0.27 | 0.24 | 0.2677 | 0.26 | 0.23 | 0.25 | 0.26 | 0.24 | 0.24 | 0.28 | 0.26 | 0.158 | 0.27 | 0.26375 | 0.23767 | 0.25 | 0.21778
| Mean |  | 0.24 | 0.25833 | 0.291 | 0.25473 | 0.25662 | 0.25033 | 0.27333 | 0.23833 | 0.26147 | 0.255 | 0.23833 | 0.24833 | 0.25667 | 0.24167 | 0.23833 | 0.275 | 0.25833 | 0.15367 | 0.27 | 0.26661 | 0.23783 | 0.24833 | 0.23589
| Median |  | 0.24 | 0.26 | 0.2915 | 0.25296 | 0.2561 | 0.251 | 0.27 | 0.24 | 0.26175 | 0.255 | 0.24 | 0.25 | 0.26 | 0.24 | 0.24 | 0.275 | 0.26 | 0.158 | 0.27 | 0.2655 | 0.23768 | 0.25 | 0.22139
| Std Dev. |  | 0 | 0.0040825 | 0.0017889 | 0.010417 | 0.0036102 | 0.0046762 | 0.005164 | 0.0040825 | 0.0042514 | 0.0054772 | 0.0040825 | 0.0040825 | 0.010328 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0040825 | 0.007174 | 0 | 0.0040773 | 0.00195 | 0.0040825 | 0.026444
| Rel.Std.Dev. |  | 0% | 1.58% | 0.6147% | 4.089% | 1.407% | 1.868% | 1.889% | 1.713% | 1.626% | 2.148% | 1.713% | 1.644% | 4.024% | 1.689% | 1.713% | 1.992% | 1.58% | 4.669% | 0% | 1.529% | 0.8199% | 1.644% | 11.21%
| PDM3 |  | -5.449% | 1.774% | 14.64% | 0.3527% | 1.097% | -1.378% | 7.683% | -6.106% | 3.008% | 0.4605% | -6.106% | -2.166% | 1.117% | -4.792% | -6.106% | 8.34% | 1.774% | -39.46% | 6.37% | 5.032% | -6.302% | -2.166% | -7.069%
| Z-Score (Absolute) |  | 0.76 | 0.19 | 1.87 | 0 | 0.1 | 0.23 | 0.96 | 0.85 | 0.35 | 0.01 | 0.85 | 0.33 | 0.1 | 0.67 | 0.85 | 1.05 | 0.19 | 5.22 | 0.79 | 0.61 | 0.87 | 0.33 | 0.97
| Analytical results for La in OREAS 265 (Certified Value 27.5 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  |  | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 27.3 | 27.4 | 20.9 | 26.7187 | 28.835 | 27.5 | 28.5 | 26.8 | 27.363 | 26.4 | 29.2 | 26.2 | 26.3 | 27.5 | 27.1 | 30.4 | 30.6 | 45 | 30 | 26.4957 | 23.4625 | 27.4 |  |  |  |  | 1
| 1 | 2 | 27.5 | 28.7 | 21.2 | 26.808 | 28.959 | 27.4 | 28.8 | 26.1 | 27.165 | 25.9 | 29.1 | 26 | 25.8 | 26.7 | 27.1 | 31.4 | 29.7 | 50 | 29 | 26.6155 | 23.9319 | 27.2 |  |  |  |  | 23
| 1 | 3 | 27.1 | 28.8 | 21.2 | 25.5311 | 28.452 | 24.9 | 28.1 | 25.8 | 27.872 | 26.6 | 27.7 | 25.6 | 25.9 | 26.4 | 27.3 | 30.5 | 29.7 | 49 | 30 | 27.3815 | 23.65 | 26.3 |  |  |  |  | 16
| 1 | 4 | 27.1 | 29.2 | 20.6 | 26.4747 | 28.703 | 26.2 | 28.5 | 26.8 | 27.73 | 25.3 | 28.1 | 25.9 | 25.1 | 26.7 | 27.1 | 31.7 | 29.9 | 52 | 31 | 26.8886 | 23.7214 | 27.2 |  |  |  |  | 27.5074
| 1 | 5 | 28 | 29 | 20.5 | 27.3955 | 28.508 | 25.6 | 28.8 | 27.6 | 28.11 | 26.4 | 27.8 | 27.1 | 26.2 | 26.7 | 26.6 | 31.3 | 30.2 | 51 | 30 | 26.3043 | 24.1633 | 27.1 |  |  |  |  | 99
| 1 | 6 | 28.1 | 28.4 | 20.8 | 25.8785 | 28.879 | 26.5 | 28.1 | 27.9 | 27.72 | 26.4 | 27.6 | 27.1 | 27.5 | 26.5 | 26.7 | 31 | 30.1 | 53 | 30 | 25.9749 | 22.6814 | 27.5
| Mean |  | 27.5167 | 28.5833 | 20.8667 | 26.4678 | 28.7227 | 26.35 | 28.4667 | 26.8333 | 27.66 | 26.1667 | 28.25 | 26.3167 | 26.1333 | 26.75 | 26.9833 | 31.05 | 30.0333 | 50 | 30 | 26.6101 | 23.6017 | 27.1167
| Median |  | 27.4 | 28.75 | 20.85 | 26.5967 | 28.769 | 26.35 | 28.5 | 26.8 | 27.725 | 26.4 | 27.95 | 26.1 | 26.05 | 26.7 | 27.1 | 31.15 | 30 | 50.5 | 30 | 26.5556 | 23.6857 | 27.2
| Std Dev. |  | 0.44008 | 0.64005 | 0.29439 | 0.67291 | 0.20622 | 1.0134 | 0.31411 | 0.8165 | 0.34327 | 0.48442 | 0.71764 | 0.63692 | 0.79162 | 0.38859 | 0.27142 | 0.51672 | 0.34448 | 2.8284 | 0.63246 | 0.48634 | 0.51134 | 0.42622
| Rel.Std.Dev. |  | 1.599% | 2.239% | 1.411% | 2.542% | 0.718% | 3.846% | 1.103% | 3.043% | 1.241% | 1.851% | 2.54% | 2.42% | 3.029% | 1.453% | 1.006% | 1.664% | 1.147% | 5.657% | 2.108% | 1.828% | 2.167% | 1.572%
| PDM3 |  | 0.03356% | 3.911% | -24.14% | -3.78% | 4.418% | -4.208% | 3.487% | -2.451% | 0.5546% | -4.874% | 2.7% | -4.329% | -4.995% | -2.754% | -1.905% | 12.88% | 9.183% | 81.77% | 9.061% | -3.262% | -14.2% | -1.421%
| Z-Score (Absolute) |  | 0.35 | 1.15 | 4.63 | 0.44 | 1.25 | 0.52 | 1.06 | 0.16 | 0.46 | 0.66 | 0.9 | 0.55 | 0.69 | 0.22 | 0.05 | 3 | 2.24 | 17.19 | 2.21 | 0.33 | 2.58 | 0.05
| Analytical results for Li in OREAS 265 (Certified Value 6.19 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | 1.30 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.2g | [5 empty] | 3
| 1 | 1 | 5 | 6.1 | 7 | 6.441 | 6.7 | 6 | 6.4 | 5.6 | 6.3 | 5.9 | 4.3 | 6.2 | 3.5 | 6 | 8.1 | 6.1 | 5 | 4 | 6.2496 | 6.1 | 3.699 | [5 empty] | 1
| 1 | 2 | 5 | 6.4 | 6.9 | 5.9526 | 6.8 | 6 | 6.5 | 5.7 | 6.4 | 5.8 | 5.5 | 6.1 | 3.2 | 6 | 8.4 | 6.2 | 5 | 4 | 6.1722 | 6 | 2.235 | [5 empty] | #N/A
| 1 | 3 | 5 | 6.4 | 7 | 5.9653 | 6.6 | 6 | 6.5 | 5.6 | 6.4 | 5.9 | 5.4 | 6.2 | 3.2 | 6 | 8.2 | 6 | 5 | 5.2 | 6.2711 | 5.9 | 3.522 | [5 empty] | 16
| 1 | 4 | 5 | 6.4 | 6.8 | 6.36 | 6.6 | 6 | 6.5 | 5.6 | 6.2 | 5.7 | 5.5 | 6.4 | 3.4 | 6 | 8.3 | 6.1 | 5 | 4.5 | 6.2893 | 5.9 | 2.438 | [5 empty] | 6.191
| 1 | 5 | 5 | 6.6 | 6.8 | 6.1793 | 6.5 | 6 | 6.6 | 5.6 | 6.4 | 5.7 | 5.8 | 6.4 | 3.2 | 5 | 9.3 | 6.4 | 5 | 5.1 | 6.2495 | 5.9 | 2.287 | [5 empty] | 100
| 1 | 6 | 5 | 6.4 | 6.8 | 5.8738 | 6.7 | 6 | 6.5 | 5.5 | 6.4 | 5.9 | 5.7 | 6.3 | 3.3 | 6 | 8.2 | 6.1 | 6 | 4 | 6.0118 | 6 | 1.986
| Mean |  | 5 | 6.3833 | 6.8833 | 6.1286 | 6.65 | 6 | 6.5 | 5.6 | 6.35 | 5.8167 | 5.3667 | 6.2667 | 3.3 | 5.8333 | 8.4167 | 6.15 | 5.1667 | 4.4667 | 6.2072 | 5.9667 | 2.6945
| Median |  | 5 | 6.4 | 6.85 | 6.0723 | 6.65 | 6 | 6.5 | 5.6 | 6.4 | 5.85 | 5.5 | 6.25 | 3.25 | 6 | 8.25 | 6.1 | 5 | 4.25 | 6.2495 | 5.95 | 2.3625
| Std Dev. |  | 0 | 0.16021 | 0.098319 | 0.23507 | 0.10488 | 0 | 0.063246 | 0.063246 | 0.083666 | 0.098319 | 0.54283 | 0.12111 | 0.12649 | 0.40825 | 0.4446 | 0.13784 | 0.40825 | 0.56451 | 0.10375 | 0.08165 | 0.72646
| Rel.Std.Dev. |  | 0% | 2.51% | 1.428% | 3.836% | 1.577% | 0% | 0.973% | 1.129% | 1.318% | 1.69% | 10.11% | 1.933% | 3.833% | 6.999% | 5.282% | 2.241% | 7.902% | 12.64% | 1.671% | 1.368% | 26.96%
| PDM3 |  | -19.24% | 3.107% | 11.18% | -1.007% | 7.415% | -3.085% | 4.992% | -9.546% | 2.569% | -6.046% | -13.31% | 1.223% | -46.7% | -5.777% | 35.95% | -0.6617% | -16.55% | -27.85% | 0.2629% | -3.623% | -56.48%
| Z-Score (Absolute) |  | NA | 0.45 | 1.41 | 0.04 | 0.96 | NA | 0.67 | 1.06 | 0.39 | 0.64 | 1.51 | 0.22 | 5.49 | NA | 4.37 | 0 | NA | 3.24 | 0.11 | 0.35 | 6.66
Results from laboratories 1.01, 1.06, 1.15 and 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 265 (Certified Value 0.076 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.15 | 1.16 | 1.22 | [18 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [18 empty] | 3
|  |  | 0.25g | 01g | 25g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | [18 empty] | 3
| 1 | 1 | 0.07 | 0.08 | 0.076 | 0.08 | 0.07 | 0.08 | < 0.1 | 0.071636 | [18 empty] | 1
| 1 | 2 | 0.08 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1 | 0.076834 | [18 empty] | 24
| 1 | 3 | 0.07 | 0.08 | 0.075 | 0.08 | 0.07 | 0.08 | < 0.1 | 0.069458 | [18 empty] | 16
| 1 | 4 | 0.07 | 0.08 | 0.076 | 0.08 | 0.07 | 0.08 | < 0.1 | 0.073458 | [18 empty] | 0.076415
| 1 | 5 | 0.08 | 0.08 | 0.075 | 0.08 | 0.08 | 0.08 | < 0.1 | 0.074876 | [18 empty] | 101
| 1 | 6 | 0.07 | 0.08 | 0.077 | 0.08 | 0.07 | 0.08 | < 0.1 | 0.074179
| Mean |  | 0.073333 | 0.08 | 0.0765 | 0.08 | 0.071667 | 0.08 | 0.1 | 0.073407
| Median |  | 0.07 | 0.08 | 0.076 | 0.08 | 0.07 | 0.08 | 0.1 | 0.073818
| Std Dev. |  | 0.005164 | 0 | 0.0018708 | 0 | 0.0040825 | 0 |  | 0.0025783
| Rel.Std.Dev. |  | 7.042% | 0% | 2.446% | 0% | 5.696% | 0% |  | 3.512%
| PDM3 |  | -4.033% | 4.691% | 0.1109% | 4.691% | -6.214% | 4.691% | 30.86% | -3.937%
| Z-Score (Absolute) |  | 0.26 | 0.82 | 0.25 | 0.82 | 0.53 | 0.82 | 2.69 | 0.25
| Analytical results for Mg in OREAS 265 (Certified Value 0.416 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 0.42 | 0.4 | 0.45 | 0.43073 | 0.43 | 0.433 | 0.42 | 0.4 | 0.4 | 0.4 | 0.42 | 0.42 | 0.45 | 0.41 | 0.41 | 0.44 | 0.46332 | 0.4 | 0.356 | 0.43 | 0.4363 | 0.39446 | 0.41 | 0.41056 |  |  | 1
| 1 | 2 | 0.41 | 0.41 | 0.46 | 0.41808 | 0.44 | 0.431 | 0.42 | 0.39 | 0.41 | 0.4 | 0.38 | 0.42 | 0.43 | 0.4 | 0.41 | 0.46 | 0.46297 | 0.4 | 0.368 | 0.43 | 0.43439 | 0.39279 | 0.42 | 0.38591 |  |  | #N/A
| 1 | 3 | 0.41 | 0.4 | 0.46 | 0.40314 | 0.43 | 0.424 | 0.4 | 0.39 | 0.41 | 0.41 | 0.37 | 0.42 | 0.43 | 0.41 | 0.42 | 0.45 | 0.46195 | 0.39 | 0.368 | 0.43 | 0.45105 | 0.38554 | 0.4 | 0.39091 |  |  | 16
| 1 | 4 | 0.41 | 0.41 | 0.48 | 0.4306 | 0.43 | 0.418 | 0.44 | 0.39 | 0.41 | 0.39 | 0.37 | 0.43 | 0.41 | 0.41 | 0.41 | 0.45 | 0.46189 | 0.4 | 0.374 | 0.43 | 0.4353 | 0.39228 | 0.41 | 0.40515 |  |  | 0.4164
| 1 | 5 | 0.42 | 0.42 | 0.46 | 0.40946 | 0.43 | 0.423 | 0.44 | 0.38 | 0.42 | 0.4 | 0.38 | 0.43 | 0.44 | 0.4 | 0.41 | 0.45 | 0.4667 | 0.41 | 0.368 | 0.43 | 0.43487 | 0.39395 | 0.41 | 0.41697 |  |  | 102
| 1 | 6 | 0.42 | 0.41 | 0.45 | 0.39589 | 0.43 | 0.431 | 0.43 | 0.38 | 0.42 | 0.41 | 0.37 | 0.42 | 0.44 | 0.4 | 0.43 | 0.47 | 0.46429 | 0.4 | 0.374 | 0.43 | 0.42816 | 0.39523 | 0.41 | 0.39905
| Mean |  | 0.415 | 0.40833 | 0.46 | 0.41465 | 0.43167 | 0.42667 | 0.425 | 0.38833 | 0.41167 | 0.40167 | 0.38167 | 0.42333 | 0.43333 | 0.405 | 0.415 | 0.45333 | 0.46352 | 0.4 | 0.368 | 0.43 | 0.43668 | 0.39237 | 0.41 | 0.40142
| Median |  | 0.415 | 0.41 | 0.46 | 0.41377 | 0.43 | 0.4275 | 0.425 | 0.39 | 0.41 | 0.4 | 0.375 | 0.42 | 0.435 | 0.405 | 0.41 | 0.45 | 0.46315 | 0.4 | 0.368 | 0.43 | 0.43509 | 0.39337 | 0.41 | 0.4021
| Std Dev. |  | 0.0054772 | 0.0075277 | 0.010954 | 0.014394 | 0.0040825 | 0.0058878 | 0.015166 | 0.0075277 | 0.0075277 | 0.0075277 | 0.019408 | 0.005164 | 0.013663 | 0.0054772 | 0.0083666 | 0.010328 | 0.0017991 | 0.0063246 | 0.0065727 | 0 | 0.0076102 | 0.0035197 | 0.0063246 | 0.011798
| Rel.Std.Dev. |  | 1.32% | 1.844% | 2.381% | 3.471% | 0.9457% | 1.38% | 3.568% | 1.938% | 1.829% | 1.874% | 5.085% | 1.22% | 3.153% | 1.352% | 2.016% | 2.278% | 0.3881% | 1.581% | 1.786% | 0% | 1.743% | 0.897% | 1.543% | 2.939%
| PDM3 |  | -0.3366% | -1.938% | 10.47% | -0.4207% | 3.666% | 2.465% | 2.065% | -6.741% | -1.137% | -3.539% | -8.342% | 1.665% | 4.066% | -2.738% | -0.3366% | 8.869% | 11.32% | -3.939% | -11.62% | 3.266% | 4.87% | -5.77% | -1.537% | -3.597%
| Z-Score (Absolute) |  | 0.01 | 0.31 | 2.16 | 0.01 | 0.8 | 0.57 | 0.49 | 1.27 | 0.15 | 0.63 | 1.58 | 0.41 | 0.88 | 0.47 | 0.01 | 1.84 | 2.33 | 0.71 | 2.24 | 0.73 | 1.04 | 1.07 | 0.23 | 0.64
| Analytical results for Mn in OREAS 265 (Certified Value 0.02 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 0.0191 | 0.0198 | 0.021 | 0.020591 | 0.0199 | 0.0201 | 0.0201 | 0.019 | 0.0194 | 0.0194 | 0.0203 | 0.0195 | 0.0213 | 0.0169 | 0.0186 | 0.0214 | 0.021879 | 0.0193 | 0.015 | 0.0205 | 0.020432 | 0.015202 | 0.0191 | 0.017129 |  |  | 1
| 1 | 2 | 0.0189 | 0.0202 | 0.0209 | 0.020263 | 0.0205 | 0.02 | 0.0204 | 0.0189 | 0.0196 | 0.0194 | 0.0191 | 0.0193 | 0.0208 | 0.0165 | 0.0184 | 0.0223 | 0.021901 | 0.0191 | 0.015 | 0.0206 | 0.020811 | 0.014936 | 0.0195 | 0.016862 |  |  | 25
| 1 | 3 | 0.0187 | 0.02 | 0.0209 | 0.019137 | 0.0198 | 0.0198 | 0.0204 | 0.0185 | 0.0199 | 0.0197 | 0.0188 | 0.019 | 0.0204 | 0.0164 | 0.0188 | 0.0219 | 0.021822 | 0.0193 | 0.015 | 0.0201 | 0.021227 | 0.014223 | 0.0184 | 0.017423 |  |  | 16
| 1 | 4 | 0.0188 | 0.0202 | 0.0209 | 0.019643 | 0.0202 | 0.0195 | 0.0202 | 0.0184 | 0.0199 | 0.019 | 0.0189 | 0.0196 | 0.0197 | 0.0166 | 0.0187 | 0.0218 | 0.021868 | 0.0193 | 0.015 | 0.0207 | 0.020793 | 0.014867 | 0.0189 | 0.016709 |  |  | 0.019813
| 1 | 5 | 0.0192 | 0.0205 | 0.0206 | 0.019506 | 0.0198 | 0.0198 | 0.021 | 0.018 | 0.02 | 0.0192 | 0.0191 | 0.02 | 0.0208 | 0.0164 | 0.0182 | 0.022 | 0.021946 | 0.0195 | 0.015 | 0.0208 | 0.020466 | 0.014969 | 0.0189 | 0.017196 |  |  | 103
| 1 | 6 | 0.0192 | 0.0202 | 0.0208 | 0.020231 | 0.0201 | 0.02 | 0.02 | 0.0182 | 0.0202 | 0.0195 | 0.0186 | 0.0197 | 0.0209 | 0.0162 | 0.0184 | 0.0228 | 0.021766 | 0.0192 | 0.015 | 0.0205 | 0.020277 | 0.015128 | 0.0188 | 0.01728
| Mean |  | 0.018983 | 0.02015 | 0.02085 | 0.019895 | 0.02005 | 0.019867 | 0.02035 | 0.0185 | 0.019833 | 0.019367 | 0.019133 | 0.019517 | 0.02065 | 0.0165 | 0.018517 | 0.022033 | 0.021864 | 0.019283 | 0.015 | 0.020533 | 0.020668 | 0.014887 | 0.018933 | 0.0171
| Median |  | 0.019 | 0.0202 | 0.0209 | 0.019937 | 0.02 | 0.0199 | 0.0203 | 0.01845 | 0.0199 | 0.0194 | 0.019 | 0.01955 | 0.0208 | 0.01645 | 0.0185 | 0.02195 | 0.021874 | 0.0193 | 0.015 | 0.02055 | 0.02063 | 0.014953 | 0.0189 | 0.017162
| Std Dev. |  | 0.0002137 | 0.00023452 | 0.00013784 | 0.00055172 | 0.00027386 | 0.00021602 | 0.00035637 | 0.00038987 | 0.00028752 | 0.00024221 | 0.00060222 | 0.00034303 | 0.00054681 | 0.00023664 | 0.00022286 | 0.0004761 | 6.2518e-05 | 0.00013292 | 0 | 0.00024221 | 0.00034573 | 0.00034863 | 0.00036148 | 0.00026711
| Rel.Std.Dev. |  | 1.126% | 1.164% | 0.6611% | 2.773% | 1.366% | 1.087% | 1.751% | 2.107% | 1.45% | 1.251% | 3.147% | 1.758% | 2.648% | 1.434% | 1.204% | 2.161% | 0.2859% | 0.6893% | 0% | 1.18% | 1.673% | 2.342% | 1.909% | 1.562%
| PDM3 |  | -4.188% | 1.7% | 5.233% | 0.4135% | 1.195% | 0.2701% | 2.71% | -6.628% | 0.1019% | -2.253% | -3.431% | -1.496% | 4.224% | -16.72% | -6.544% | 11.21% | 10.35% | -2.674% | -24.29% | 3.635% | 4.313% | -24.86% | -4.441% | -13.69%
| Z-Score (Absolute) |  | 0.58 | 0.4 | 0.99 | 0.19 | 0.32 | 0.16 | 0.57 | 0.99 | 0.13 | 0.26 | 0.46 | 0.13 | 0.82 | 2.68 | 0.98 | 1.99 | 1.84 | 0.33 | 3.53 | 0.72 | 0.84 | 4.04 | 0.63 | 2.17
| Analytical results for Mo in OREAS 265 (Certified Value 3.5 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 3.55 | 3.42 | 3.6 | 3.4329 | 3.7 | 3.36 | 3.36 | 3.91 | 3.5 | 3.19 | 3.7 | 3.39 | 3.58 | 3.53 | 3.62 | 3.68 | 3.6 | 3.66 | 2.19 | 3.1 | 3.4658 | 3.2925 | 3.53 | 2.117 |  |  | 1
| 1 | 2 | 3.64 | 3.49 | 3.6 | 3.4184 | 3.8 | 3.27 | 3.26 | 3.47 | 3.6 | 3.17 | 3.68 | 3.4 | 3.67 | 3.33 | 3.56 | 3.78 | 3.4 | 3.59 | 3.01 | 3.1 | 3.5795 | 3.2251 | 3.38 | 2.269 |  |  | 26
| 1 | 3 | 3.47 | 3.55 | 3.6 | 3.4788 | 3.7 | 3.26 | 3.28 | 3.3 | 3.6 | 3.21 | 3.49 | 3.44 | 3.6 | 3.45 | 3.64 | 3.56 | 3.5 | 3.52 | 2.58 | 3 | 3.5712 | 3.3539 | 3.15 | 2.42 |  |  | 16
| 1 | 4 | 3.64 | 3.58 | 3.7 | 3.4814 | 3.7 | 3.36 | 3.28 | 3.3 | 3.6 | 3.18 | 3.57 | 3.51 | 3.73 | 3.5 | 3.6 | 3.73 | 3.6 | 3.57 | 2.83 | 3.1 | 3.3963 | 3.2126 | 3.5 | 2.225 |  |  | 3.5046
| 1 | 5 | 3.64 | 3.57 | 3.7 | 3.4548 | 3.8 | 3.22 | 3.65 | 3.4 | 3.6 | 3.26 | 3.6 | 3.46 | 3.62 | 3.76 | 3.55 | 3.65 | 3.6 | 3.56 | 2.73 | 3.1 | 3.5972 | 3.224 | 3.32 | 1.776 |  |  | 104
| 1 | 6 | 3.63 | 3.65 | 3.6 | 3.4594 | 3.8 | 3.44 | 3.25 | 3.39 | 3.6 | 3.14 | 3.55 | 3.35 | 3.75 | 3.79 | 3.52 | 3.84 | 3.6 | 3.51 | 3.11 | 3.1 | 3.5406 | 3.1093 | 3.51 | 1.962
| Mean |  | 3.595 | 3.5433 | 3.6333 | 3.4543 | 3.75 | 3.3183 | 3.3467 | 3.4617 | 3.5833 | 3.1917 | 3.5983 | 3.425 | 3.6583 | 3.56 | 3.5817 | 3.7067 | 3.55 | 3.5683 | 2.7417 | 3.0833 | 3.5251 | 3.2362 | 3.3983 | 2.1282
| Median |  | 3.635 | 3.56 | 3.6 | 3.4571 | 3.75 | 3.315 | 3.28 | 3.395 | 3.6 | 3.185 | 3.585 | 3.42 | 3.645 | 3.515 | 3.58 | 3.705 | 3.6 | 3.565 | 2.78 | 3.1 | 3.5559 | 3.2245 | 3.44 | 2.171
| Std Dev. |  | 0.07064 | 0.079415 | 0.05164 | 0.024924 | 0.054772 | 0.082077 | 0.15358 | 0.22903 | 0.040825 | 0.040702 | 0.079854 | 0.056833 | 0.070261 | 0.18022 | 0.045789 | 0.09913 | 0.083666 | 0.054191 | 0.33048 | 0.040825 | 0.078299 | 0.082412 | 0.14716 | 0.23071
| Rel.Std.Dev. |  | 1.965% | 2.241% | 1.421% | 0.7215% | 1.461% | 2.473% | 4.589% | 6.616% | 1.139% | 1.275% | 2.219% | 1.659% | 1.921% | 5.062% | 1.278% | 2.674% | 2.357% | 1.519% | 12.05% | 1.324% | 2.221% | 2.547% | 4.33% | 10.84%
| PDM3 |  | 2.58% | 1.106% | 3.674% | -1.435% | 7.003% | -5.314% | -4.506% | -1.224% | 2.247% | -8.928% | 2.675% | -2.27% | 4.388% | 1.582% | 2.2% | 5.767% | 1.296% | 1.819% | -21.77% | -12.02% | 0.5861% | -7.657% | -3.031% | -39.27%
| Z-Score (Absolute) |  | 0.39 | 0.06 | 0.64 | 0.52 | 1.4 | 1.4 | 1.21 | 0.47 | 0.32 | 2.22 | 0.42 | 0.71 | 0.8 | 0.17 | 0.31 | 1.12 | 0.1 | 0.22 | 5.13 | 2.92 | 0.06 | 1.93 | 0.88 | 9.1
| Analytical results for Na in OREAS 265 (Certified Value 0.233 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  |  | 3
| 1 | 1 | 0.22 | 0.22 | 0.23 | 0.25943 | 0.24 | 0.237 | 0.24 | 0.23 | 0.24 | 0.24 | 0.23 | 0.223 | 0.235 | 0.22 | 0.21 | 0.26 | 0.24 | 0.148 | 0.21 | 0.23768 | 0.21488 | 0.24 | 0.25871 |  |  |  | 1
| 1 | 2 | 0.22 | 0.23 | 0.23 | 0.25909 | 0.25 | 0.236 | 0.23 | 0.23 | 0.25 | 0.24 | 0.24 | 0.22 | 0.23 | 0.22 | 0.21 | 0.27 | 0.25 | 0.156 | 0.22 | 0.23728 | 0.21821 | 0.25 | 0.23902 |  |  |  | 44
| 1 | 3 | 0.22 | 0.22 | 0.23 | 0.22852 | 0.25 | 0.233 | 0.24 | 0.23 | 0.25 | 0.24 | 0.24 | 0.22 | 0.228 | 0.22 | 0.22 | 0.26 | 0.25 | 0.163 | 0.22 | 0.24433 | 0.21476 | 0.23 | 0.24298 |  |  |  | 16
| 1 | 4 | 0.22 | 0.22 | 0.22 | 0.24627 | 0.25 | 0.23 | 0.24 | 0.22 | 0.25 | 0.24 | 0.24 | 0.224 | 0.22 | 0.22 | 0.21 | 0.27 | 0.25 | 0.163 | 0.21 | 0.23665 | 0.21255 | 0.24 | 0.23995 |  |  |  | 0.23332
| 1 | 5 | 0.22 | 0.23 | 0.22 | 0.24548 | 0.24 | 0.232 | 0.24 | 0.22 | 0.25 | 0.24 | 0.24 | 0.225 | 0.229 | 0.21 | 0.21 | 0.27 | 0.25 | 0.163 | 0.22 | 0.23533 | 0.2171 | 0.24 | 0.24496 |  |  |  | 105
| 1 | 6 | 0.23 | 0.23 | 0.22 | 0.23683 | 0.24 | 0.237 | 0.23 | 0.22 | 0.25 | 0.24 | 0.24 | 0.224 | 0.231 | 0.21 | 0.22 | 0.27 | 0.25 | 0.171 | 0.22 | 0.23607 | 0.21379 | 0.24 | 0.24145
| Mean |  | 0.22167 | 0.225 | 0.225 | 0.24594 | 0.245 | 0.23417 | 0.23667 | 0.225 | 0.24833 | 0.24 | 0.23833 | 0.22267 | 0.22883 | 0.21667 | 0.21333 | 0.26667 | 0.24833 | 0.16067 | 0.21667 | 0.23789 | 0.21522 | 0.24 | 0.24451
| Median |  | 0.22 | 0.225 | 0.225 | 0.24587 | 0.245 | 0.2345 | 0.24 | 0.225 | 0.25 | 0.24 | 0.24 | 0.2235 | 0.2295 | 0.22 | 0.21 | 0.27 | 0.25 | 0.163 | 0.22 | 0.23697 | 0.21482 | 0.24 | 0.24221
| Std Dev. |  | 0.0040825 | 0.0054772 | 0.0054772 | 0.012179 | 0.0054772 | 0.0029269 | 0.005164 | 0.0054772 | 0.0040825 | 0 | 0.0040825 | 0.0021602 | 0.0049565 | 0.005164 | 0.005164 | 0.005164 | 0.0040825 | 0.0078145 | 0.005164 | 0.0032645 | 0.0020943 | 0.0063246 | 0.0072742
| Rel.Std.Dev. |  | 1.842% | 2.434% | 2.434% | 4.952% | 2.236% | 1.25% | 2.182% | 2.434% | 1.644% | 0% | 1.713% | 0.9702% | 2.166% | 2.383% | 2.421% | 1.936% | 1.644% | 4.864% | 2.383% | 1.372% | 0.9731% | 2.635% | 2.975%
| PDM3 |  | -4.995% | -3.566% | -3.566% | 5.407% | 5.006% | 0.3628% | 1.434% | -3.566% | 6.435% | 2.863% | 2.149% | -4.566% | -1.923% | -7.138% | -8.566% | 14.29% | 6.435% | -31.14% | -7.138% | 1.958% | -7.76% | 2.863% | 4.796%
| Z-Score (Absolute) |  | 0.78 | 0.57 | 0.57 | 0.73 | 0.67 | 0 | 0.16 | 0.57 | 0.88 | 0.36 | 0.26 | 0.72 | 0.33 | 1.09 | 1.3 | 2.02 | 0.88 | 4.57 | 1.09 | 0.23 | 1.18 | 0.36 | 0.64
| Analytical results for Nb in OREAS 265 (Indicative Value 0.55 ppm) | [27 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [6 empty] | 2
| 1 | 1 | 1.38 | 0.55 | 0.75 | 0.81161 | 0.52 | 0.73 | 0.53 | 0.53 | 0.44 | 0.47 | 0.59 | 0.17 | 0.96 | 1.11 | 0.35 | 0.45 | 0.05 | 0.5 | 0.63101 | 0.43 | [6 empty] | 1
| 1 | 2 | 1.35 | 0.51 | 0.74 | 0.84788 | 0.58 | 0.68 | 0.46 | 0.6 | 0.44 | 0.45 | 0.52 | 0.19 | 0.93 | 1.12 | 0.44 | 0.44 | 0.05 | 0.5 | 0.68144 | 0.44 | [6 empty] | 6
| 1 | 3 | 1.69 | 0.57 | 0.73 | 0.79241 | 0.56 | 0.75 | 0.34 | 0.51 | 0.46 | 0.47 | 0.48 | 0.2 | 0.98 | 1.11 | 0.38 | 0.48 | 0.05 | 0.4 | 0.64588 | 0.39 | [6 empty] | 16
| 1 | 4 | 1.22 | 0.62 | 0.71 | 0.79832 | 0.54 | 0.66 | 0.24 | 0.51 | 0.44 | 0.46 | 0.51 | 0.21 | 0.92 | 1.11 | 0.42 | 0.45 | 0.04 | 0.5 | 0.62986 | 0.42 | [6 empty] | 0.5541
| 1 | 5 | 1.2 | 0.58 | 0.7 | 0.83956 | 0.55 | 0.67 | 0.88 | 0.62 | 0.46 | 0.46 | 0.49 | 0.21 | 0.91 | 1.13 | 0.45 | 0.45 | 0.05 | 0.5 | 0.63333 | 0.47 | [6 empty] | 12
| 1 | 6 | 1.22 | 0.57 | 0.69 | 0.80019 | 0.6 | 0.73 | 0.2 | 0.55 | 0.45 | 0.48 | 0.46 | 0.19 | 0.91 | 1.11 | 0.55 | 0.5 | 0.07 | 0.5 | 0.66807 | 0.38
| Mean |  | 1.3433 | 0.56667 | 0.72 | 0.815 | 0.55833 | 0.70333 | 0.44167 | 0.55333 | 0.44833 | 0.465 | 0.50833 | 0.195 | 0.935 | 1.115 | 0.43167 | 0.46167 | 0.051667 | 0.48333 | 0.64827 | 0.42167
| Median |  | 1.285 | 0.57 | 0.72 | 0.8059 | 0.555 | 0.705 | 0.4 | 0.54 | 0.445 | 0.465 | 0.5 | 0.195 | 0.925 | 1.11 | 0.43 | 0.45 | 0.05 | 0.5 | 0.63961 | 0.425
| Std Dev. |  | 0.18576 | 0.036148 | 0.023664 | 0.023252 | 0.028577 | 0.037771 | 0.24887 | 0.046762 | 0.0098319 | 0.010488 | 0.04535 | 0.015166 | 0.02881 | 0.0083666 | 0.069113 | 0.023166 | 0.0098319 | 0.040825 | 0.021716 | 0.033116
| Rel.Std.Dev. |  | 13.83% | 6.379% | 3.287% | 2.853% | 5.118% | 5.37% | 56.35% | 8.451% | 2.193% | 2.256% | 8.921% | 7.777% | 3.081% | 0.7504% | 16.01% | 5.018% | 19.03% | 8.447% | 3.35% | 7.854%
| PDM3 |  | 142.4% | 2.268% | 29.94% | 47.08% | 0.7641% | 26.93% | -20.29% | -0.1383% | -19.09% | -16.08% | -8.26% | -64.81% | 68.74% | 101.2% | -22.1% | -16.68% | -90.68% | -12.77% | 16.99% | -23.9%
| Z-Score (Absolute) |  | 4.38 | 0.07 | 0.92 | 1.45 | 0.03 | 0.83 | 0.62 | 0 | 0.58 | 0.49 | 0.25 | 1.99 | 2.12 | 3.11 | 0.67 | 0.51 | 2.78 | NA | 0.53 | 0.73
Results from laboratory 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 265 (Certified Value 25.4 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.17 | 1.20 | 1.22 | [19 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 1
|  |  | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | [19 empty] | 2
| 1 | 1 | 18.6 | 25.314 | 26.4 | 25.2 | 25.8424 | 17.8 | 23.2914 | [19 empty] | 1
| 1 | 2 | 18.6 | 25.408 | 26.2 | 25.7 | 25.7848 | 19.8 | 24.8524 | [19 empty] | 28
| 1 | 3 | 18.9 | 24.941 | 25.9 | 25.5 | 25.8592 | 19.1 | 24.7199 | [19 empty] | 16
| 1 | 4 | 18.4 | 25.293 | 26.3 | 25.7 | 25.9376 | 19.7 | 23.8693 | [19 empty] | 25.3886
| 1 | 5 | 18.4 | 25.362 | 26.4 | 25.2 | 25.9184 | 19.8 | 24.3022 | [19 empty] | 106
| 1 | 6 | 18.6 | 25.005 | 26.6 | 25.3 | 25.8976 | 20.7 | 23.6604
| Mean |  | 18.5833 | 25.2205 | 26.3 | 25.4333 | 25.8733 | 19.4833 | 24.1159
| Median |  | 18.6 | 25.3035 | 26.35 | 25.4 | 25.8784 | 19.75 | 24.0858
| Std Dev. |  | 0.18348 | 0.19684 | 0.23664 | 0.23381 | 0.056116 | 0.9704 | 0.61483
| Rel.Std.Dev. |  | 0.9874% | 0.7805% | 0.8998% | 0.9193% | 0.2169% | 4.981% | 2.549%
| PDM3 |  | -26.8% | -0.6622% | 3.59% | 0.1761% | 1.909% | -23.26% | -5.013%
| Z-Score (Absolute) |  | 4.15 | 0 | 0.67 | 0.13 | 0.41 | 3.58 | 0.69
| Analytical results for Ni in OREAS 265 (Certified Value 42.4 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  |  | 2
| 1 | 1 | 40.1 | 41.4 | 45.2 | 43.4737 | 43.6 | 42 | 44.6 | 43.2 | 41.3 | 41.2 | 43.9 | 41.9 | 42.3 | 42 | 43.2 | 44.4 | 44.9 | 36.6 | 41.7 | 41.7793 | 38.24 | 42.1 | 38.962 |  |  |  | 1
| 1 | 2 | 39.9 | 42.5 | 45.9 | 42.9825 | 44.6 | 42 | 45.6 | 40.4 | 41.6 | 41.4 | 44.1 | 42.1 | 44 | 40.9 | 41.8 | 46.9 | 43.8 | 37.9 | 41.7 | 43.8126 | 39.3 | 41.8 | 37.988 |  |  |  | 14
| 1 | 3 | 38.9 | 43.7 | 45.8 | 42.0964 | 43.6 | 41 | 43.9 | 41 | 41.9 | 42.1 | 41.8 | 42.1 | 41.5 | 41.5 | 42.3 | 45.5 | 44.4 | 37.5 | 41.7 | 44.7029 | 37.37 | 38 | 38.132 |  |  |  | 16
| 1 | 4 | 39.7 | 44.1 | 46.6 | 42.4203 | 44.2 | 41 | 45.1 | 40.7 | 41.7 | 41.3 | 42.5 | 43.1 | 41.3 | 41.4 | 43.3 | 48.9 | 43.7 | 37.5 | 42.1 | 43.9788 | 38.8 | 42.1 | 38.214 |  |  |  | 42.4328
| 1 | 5 | 40.1 | 42.7 | 45 | 43.5258 | 43.7 | 41 | 44.5 | 41.1 | 41.8 | 42.1 | 42.7 | 43.3 | 43.3 | 40.8 | 40.6 | 48.1 | 44.6 | 38.5 | 42.6 | 43.8923 | 39.01 | 40.7 | 39.149 |  |  |  | 107
| 1 | 6 | 40.1 | 42.6 | 45.9 | 42.6102 | 43.9 | 42 | 43.7 | 44 | 41.9 | 41.5 | 42.6 | 41.8 | 44.8 | 40.2 | 41.8 | 47.6 | 44.1 | 39.8 | 43 | 43.4665 | 39.69 | 42.3 | 38.734
| Mean |  | 39.8 | 42.8333 | 45.7333 | 42.8515 | 43.9333 | 41.5 | 44.5667 | 41.7333 | 41.7 | 41.6 | 42.9333 | 42.3833 | 42.8667 | 41.1333 | 42.1667 | 46.9 | 44.25 | 37.9667 | 42.1333 | 43.6054 | 38.735 | 41.1667 | 38.5298
| Median |  | 40 | 42.65 | 45.85 | 42.7964 | 43.8 | 41.5 | 44.55 | 41.05 | 41.75 | 41.45 | 42.65 | 42.1 | 42.8 | 41.15 | 42.05 | 47.25 | 44.25 | 37.7 | 41.9 | 43.8524 | 38.905 | 41.95 | 38.474
| Std Dev. |  | 0.46904 | 0.95847 | 0.57155 | 0.5785 | 0.39833 | 0.54772 | 0.71461 | 1.4882 | 0.22804 | 0.4 | 0.88694 | 0.64627 | 1.4038 | 0.6314 | 1.0093 | 1.6817 | 0.46797 | 1.0912 | 0.55377 | 0.98216 | 0.82696 | 1.6537 | 0.48237
| Rel.Std.Dev. |  | 1.178% | 2.238% | 1.25% | 1.35% | 0.9067% | 1.32% | 1.603% | 3.566% | 0.5468% | 0.9615% | 2.066% | 1.525% | 3.275% | 1.535% | 2.394% | 3.586% | 1.058% | 2.874% | 1.314% | 2.252% | 2.135% | 4.017% | 1.252%
| PDM3 |  | -6.205% | 0.9439% | 7.778% | 0.9868% | 3.536% | -2.198% | 5.029% | -1.648% | -1.727% | -1.963% | 1.18% | -0.1166% | 1.023% | -3.062% | -0.6272% | 10.53% | 4.283% | -10.53% | -0.7057% | 2.763% | -8.714% | -2.984% | -9.198%
| Z-Score (Absolute) |  | 1.6 | 0.45 | 2.41 | 0.46 | 1.19 | 0.45 | 1.62 | 0.29 | 0.31 | 0.38 | 0.52 | 0.15 | 0.47 | 0.7 | 0 | 3.19 | 1.4 | 2.83 | 0.02 | 0.97 | 2.31 | 0.67 | 2.45
| Analytical results for P in OREAS 265 (Certified Value 0.047 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 0.0466 | 0.048 | 0.05 | 0.052595 | 0.0473 | 0.051 | 0.0477 | 0.047 | 0.0466 | 0.048 | 0.051 | 0.044 | 0.049 | 0.0443 | 0.047 | 0.045 | 0.045501 | 0.047 | 0.039 | 0.047 | 0.049859 | 0.045403 | 0.0466 | 0.047308 |  |  | 1
| 1 | 2 | 0.0461 | 0.049 | 0.049 | 0.05171 | 0.0466 | 0.052 | 0.0479 | 0.047 | 0.0465 | 0.048 | 0.046 | 0.044 | 0.048 | 0.0443 | 0.0483 | 0.048 | 0.04527 | 0.049 | 0.039 | 0.049 | 0.049504 | 0.043861 | 0.0477 | 0.048213 |  |  | #N/A
| 1 | 3 | 0.0458 | 0.049 | 0.05 | 0.04892 | 0.0452 | 0.051 | 0.0456 | 0.046 | 0.0481 | 0.048 | 0.046 | 0.046 | 0.048 | 0.0443 | 0.0472 | 0.047 | 0.045208 | 0.047 | 0.044 | 0.048 | 0.052181 | 0.044046 | 0.0454 | 0.048343 |  |  | 16
| 1 | 4 | 0.046 | 0.05 | 0.05 | 0.050393 | 0.0449 | 0.05 | 0.048 | 0.046 | 0.0472 | 0.046 | 0.046 | 0.046 | 0.047 | 0.0439 | 0.0465 | 0.047 | 0.04536 | 0.047 | 0.044 | 0.048 | 0.050152 | 0.045238 | 0.0461 | 0.045919 |  |  | 0.047391
| 1 | 5 | 0.0466 | 0.051 | 0.05 | 0.050496 | 0.0471 | 0.051 | 0.047 | 0.046 | 0.0478 | 0.047 | 0.046 | 0.047 | 0.048 | 0.0438 | 0.0463 | 0.048 | 0.04527 | 0.049 | 0.044 | 0.05 | 0.049508 | 0.045945 | 0.0466 | 0.04688 |  |  | 108
| 1 | 6 | 0.0478 | 0.05 | 0.05 | 0.048756 | 0.0455 | 0.051 | 0.0478 | 0.045 | 0.0485 | 0.049 | 0.046 | 0.045 | 0.048 | 0.0434 | 0.0487 | 0.05 | 0.045223 | 0.048 | 0.044 | 0.049 | 0.048589 | 0.045526 | 0.046 | 0.046823
| Mean |  | 0.046483 | 0.0495 | 0.049833 | 0.050478 | 0.0461 | 0.051 | 0.047333 | 0.046167 | 0.04745 | 0.047667 | 0.046833 | 0.045333 | 0.048 | 0.044 | 0.047333 | 0.0475 | 0.045305 | 0.047833 | 0.042333 | 0.0485 | 0.049965 | 0.045003 | 0.0464 | 0.047248
| Median |  | 0.04635 | 0.0495 | 0.05 | 0.050444 | 0.04605 | 0.051 | 0.04775 | 0.046 | 0.0475 | 0.048 | 0.046 | 0.0455 | 0.048 | 0.0441 | 0.0471 | 0.0475 | 0.04527 | 0.0475 | 0.044 | 0.0485 | 0.049684 | 0.045321 | 0.04635 | 0.047094
| Std Dev. |  | 0.00072226 | 0.0010488 | 0.00040825 | 0.0015098 | 0.0010296 | 0.00063246 | 0.00092014 | 0.00075277 | 0.0008167 | 0.0010328 | 0.0020412 | 0.0012111 | 0.00063246 | 0.00036878 | 0.00096885 | 0.0016432 | 0.00010952 | 0.00098319 | 0.002582 | 0.0010488 | 0.0012059 | 0.00084813 | 0.00077717 | 0.00091822
| Rel.Std.Dev. |  | 1.554% | 2.119% | 0.8192% | 2.991% | 2.233% | 1.24% | 1.944% | 1.631% | 1.721% | 2.167% | 4.359% | 2.671% | 1.318% | 0.8381% | 2.047% | 3.459% | 0.2417% | 2.055% | 6.099% | 2.162% | 2.414% | 1.885% | 1.675% | 1.943%
| PDM3 |  | -1.915% | 4.45% | 5.154% | 6.514% | -2.724% | 7.616% | -0.1214% | -2.583% | 0.1248% | 0.582% | -1.176% | -4.342% | 1.285% | -7.155% | -0.1214% | 0.2303% | -4.401% | 0.9336% | -10.67% | 2.34% | 5.433% | -5.038% | -2.091% | -0.302%
| Z-Score (Absolute) |  | 0.49 | 1.25 | 1.44 | 1.82 | 0.71 | 2.12 | 0 | 0.67 | 0.07 | 0.19 | 0.29 | 1.16 | 0.39 | 1.93 | 0 | 0.1 | 1.17 | 0.29 | 2.98 | 0.67 | 1.52 | 1.35 | 0.54 | 0.05
| Analytical results for Pb in OREAS 265 (Certified Value 8.44 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.2g |  | 2
| 1 | 1 | 7.8 | 9 | 9.5 | 8.467 | 9.1 | 8.7 | 8.4 | 9.6 | 8.7 | 8 | 8.4 | 8.52 | 8.2 | 7.4 | 7.4 | 10.1 | 7 | 9.7 | 5.33 | 8.6 | 9.4184 | 6.6288 | 11.648 | 8.2 | 3.688 |  | 1
| 1 | 2 | 7.7 | 9.2 | 9.7 | 8.5191 | 9.2 | 8.9 | 8.5 | 8.3 | 8.6 | 8 | 8.4 | 8.46 | 8.4 | 7 | 7.6 | 10.4 | 7 | 8.9 | 6.13 | 8.6 | 9.6892 | 6.5563 | 11.448 | 8.1 | 1.815 |  | 16
| 1 | 3 | 7.5 | 9.5 | 9.6 | 7.9387 | 9.1 | 8.6 | 8.8 | 7.9 | 8.8 | 8.1 | 8 | 8.63 | 8.25 | 6.9 | 7.3 | 10.2 | 7 | 8.8 | 5.74 | 8.5 | 9.5256 | 6.6025 | 12.048 | 7.8 | 1.262 |  | 16
| 1 | 4 | 7.5 | 9.4 | 9.7 | 8.2808 | 9.3 | 8.4 | 8.3 | 8 | 8.6 | 7.9 | 8 | 8.43 | 7.92 | 7.2 | 7.1 | 10.6 | 7 | 8.9 | 5.95 | 8.6 | 9.5746 | 6.5345 | 13.057 | 8.2 | 2.867 |  | 8.4444
| 1 | 5 | 7.9 | 9.2 | 9.5 | 8.3332 | 9.1 | 8.6 | 8.6 | 8 | 8.7 | 8 | 8.2 | 8.76 | 8.01 | 7.2 | 7.6 | 10.3 | 7 | 9 | 6 | 8.9 | 9.3146 | 6.8864 | 13.134 | 8.1 | 2.887 |  | 109
| 1 | 6 | 7.7 | 9.3 | 9.5 | 8.0203 | 9.6 | 8.8 | 8.4 | 8.2 | 8.8 | 8 | 8 | 8.84 | 8.43 | 6.8 | 7.2 | 10.4 | 7 | 8.8 | 6.43 | 8.6 | 9.3445 | 6.4214 | 13.38 | 8.4 | 2.829
| Mean |  | 7.6833 | 9.2667 | 9.5833 | 8.2598 | 9.2333 | 8.6667 | 8.5 | 8.3333 | 8.7 | 8 | 8.1667 | 8.6067 | 8.2017 | 7.0833 | 7.3667 | 10.3333 | 7 | 9.0167 | 5.93 | 8.6333 | 9.4778 | 6.605 | 12.4525 | 8.1333 | 2.558
| Median |  | 7.7 | 9.25 | 9.55 | 8.307 | 9.15 | 8.65 | 8.45 | 8.1 | 8.7 | 8 | 8.1 | 8.575 | 8.225 | 7.1 | 7.35 | 10.35 | 7 | 8.9 | 5.975 | 8.6 | 9.472 | 6.5794 | 12.5525 | 8.15 | 2.848
| Std Dev. |  | 0.16021 | 0.17512 | 0.098319 | 0.23515 | 0.19664 | 0.17512 | 0.17889 | 0.6377 | 0.089443 | 0.063246 | 0.19664 | 0.16657 | 0.20488 | 0.22286 | 0.20656 | 0.17512 | 0 | 0.34303 | 0.37186 | 0.13663 | 0.1444 | 0.15544 | 0.83785 | 0.19664 | 0.87014
| Rel.Std.Dev. |  | 2.085% | 1.89% | 1.026% | 2.847% | 2.13% | 2.021% | 2.105% | 7.652% | 1.028% | 0.7906% | 2.408% | 1.935% | 2.498% | 3.146% | 2.804% | 1.695% | 0% | 3.804% | 6.271% | 1.583% | 1.524% | 2.353% | 6.728% | 2.418% | 34.02%
| PDM3 |  | -9.012% | 9.738% | 13.49% | -2.185% | 9.343% | 2.633% | 0.659% | -1.315% | 3.027% | -5.262% | -3.288% | 1.922% | -2.874% | -16.12% | -12.76% | 22.37% | -17.1% | 6.777% | -29.78% | 2.238% | 12.24% | -21.78% | 47.47% | -3.683% | -69.71%
| Z-Score (Absolute) |  | 0.74 | 0.86 | 1.18 | 0.16 | 0.83 | 0.25 | 0.08 | 0.08 | 0.29 | 0.42 | 0.25 | 0.19 | 0.22 | 1.35 | 1.06 | 1.94 | NA | 0.61 | 2.52 | 0.22 | 1.07 | 1.83 | 4.08 | 0.29 | 5.93
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 265 (Indicative Value < 10 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.04 | 1.05 | 1.09 | 1.12 | [22 empty] | ppb
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | [22 empty] | 1
|  |  | 01g | 25g | 01g | 15g | [22 empty] | 1
| 1 | 1 | < 10 | < 10 | < 10 | < 10 | [22 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | < 10 | [22 empty] | 7
| 1 | 3 | < 10 | < 10 | < 10 | < 10 | [22 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | < 10 | [22 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 10 | < 10 | [22 empty] | 13
| 1 | 6 | < 10 | < 10 | < 10 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Pr in OREAS 265 (Certified Value 6.71 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.17 | 1.20 | 1.22 | [19 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | [19 empty] | 3
| 1 | 1 | 4.93 | 6.643 | 6.8 | 6.5 | 6.8616 | 7.25 | 6.4956 | [19 empty] | 1
| 1 | 2 | 5.01 | 6.645 | 6.87 | 6.7 | 6.8272 | 8.09 | 6.9142 | [19 empty] | 31
| 1 | 3 | 5.02 | 6.599 | 6.53 | 6.6 | 6.8574 | 7.8 | 6.7234 | [19 empty] | 16
| 1 | 4 | 4.87 | 6.713 | 6.79 | 6.9 | 6.8939 | 8.07 | 6.642 | [19 empty] | 6.7131
| 1 | 5 | 4.91 | 6.636 | 6.68 | 6.7 | 6.8861 | 8.07 | 6.691 | [19 empty] | 110
| 1 | 6 | 4.92 | 6.56 | 6.71 | 6.6 | 6.8768 | 8.44 | 6.549
| Mean |  | 4.9433 | 6.6327 | 6.73 | 6.6667 | 6.8672 | 7.9533 | 6.6692
| Median |  | 4.925 | 6.6395 | 6.75 | 6.65 | 6.8692 | 8.07 | 6.6665
| Std Dev. |  | 0.059217 | 0.051259 | 0.11916 | 0.13663 | 0.024038 | 0.40023 | 0.14751
| Rel.Std.Dev. |  | 1.198% | 0.7728% | 1.771% | 2.049% | 0.35% | 5.032% | 2.212%
| PDM3 |  | -26.36% | -1.199% | 0.2511% | -0.6923% | 2.294% | 18.47% | -0.6545%
| Z-Score (Absolute) |  | 19.14 | 0.41 | 0.67 | 0.03 | 2.2 | 14.24 | 0
| Analytical results for Pt in OREAS 265 (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.04 | 1.05 | 1.09 | 1.12 | [22 empty] | ppb
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | [22 empty] | 2
|  |  | 01g | 25g | 01g | 15g | [22 empty] | 2
| 1 | 1 | < 5 | < 5 | < 5 | < 2 | [22 empty] | 1
| 1 | 2 | < 5 | < 5 | < 5 | < 2 | [22 empty] | 6
| 1 | 3 | < 5 | < 5 | < 5 | < 2 | [22 empty] | 16
| 1 | 4 | < 5 | < 5 | < 5 | < 2 | [22 empty] | < 5
| 1 | 5 | < 5 | < 5 | < 5 | < 2 | [22 empty] | 12
| 1 | 6 | < 5 | < 5 | < 5 | < 2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Rb in OREAS 265 (Certified Value 15.1 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [6 empty] | 2
| 1 | 1 | 14.5 | 14.6 | 18.6 | 14.9467 | 14.72 | 15.66 | 16.4 | 14.2 | 15.74 | 13.9 | 14.8 | 15.1 | 14.4 | 14.8 | 14.9 | 16.5 | 10.6 | 15.2 | 15.048 | 16 | [6 empty] | 1
| 1 | 2 | 14.5 | 15.4 | 18.6 | 14.4101 | 15.45 | 15.57 | 15.8 | 13.9 | 16.19 | 13.5 | 15 | 15.4 | 14.2 | 14.4 | 15.5 | 16.1 | 11.2 | 15.2 | 16.122 | 15.9 | [6 empty] | 32
| 1 | 3 | 14.4 | 15.7 | 18.7 | 13.673 | 14.87 | 15.29 | 15.3 | 13.6 | 16.19 | 14.4 | 14.2 | 15.4 | 14 | 14.8 | 14.9 | 16.3 | 11.1 | 15.1 | 16.3404 | 14.5 | [6 empty] | 16
| 1 | 4 | 14.3 | 15.8 | 18.9 | 13.9301 | 15.09 | 14.89 | 16.2 | 13.8 | 16.24 | 13.6 | 14.2 | 15.8 | 14.5 | 14.7 | 15.5 | 16.2 | 11.3 | 15.5 | 15.6314 | 16.3 | [6 empty] | 15.096
| 1 | 5 | 14.5 | 15.6 | 18.7 | 15.1083 | 14.89 | 15.25 | 15.5 | 14.1 | 16.01 | 14.2 | 14.4 | 15.7 | 13.9 | 14.5 | 15.2 | 16.3 | 11.2 | 15.6 | 16.2172 | 15.6 | [6 empty] | 111
| 1 | 6 | 14.6 | 15.4 | 18.7 | 13.9079 | 15.3 | 16.09 | 15.9 | 14 | 16.24 | 13.5 | 14.2 | 15.6 | 14 | 14.7 | 15.5 | 16.4 | 11.6 | 15.6 | 15.7497 | 16
| Mean |  | 14.4667 | 15.4167 | 18.7 | 14.3294 | 15.0533 | 15.4583 | 15.85 | 13.9333 | 16.1017 | 13.85 | 14.4667 | 15.5 | 14.1667 | 14.65 | 15.25 | 16.3 | 11.1667 | 15.3667 | 15.8514 | 15.7167
| Median |  | 14.5 | 15.5 | 18.7 | 14.1701 | 14.99 | 15.43 | 15.85 | 13.95 | 16.19 | 13.75 | 14.3 | 15.5 | 14.1 | 14.7 | 15.35 | 16.3 | 11.2 | 15.35 | 15.9359 | 15.95
| Std Dev. |  | 0.10328 | 0.43089 | 0.10954 | 0.59368 | 0.27962 | 0.41155 | 0.41352 | 0.21602 | 0.19651 | 0.38341 | 0.35024 | 0.25298 | 0.24221 | 0.16432 | 0.29496 | 0.14142 | 0.3266 | 0.22509 | 0.47966 | 0.63692
| Rel.Std.Dev. |  | 0.7139% | 2.795% | 0.5858% | 4.143% | 1.858% | 2.662% | 2.609% | 1.55% | 1.22% | 2.768% | 2.421% | 1.632% | 1.71% | 1.122% | 1.934% | 0.8676% | 2.925% | 1.465% | 3.026% | 4.053%
| PDM3 |  | -4.169% | 2.124% | 23.87% | -5.078% | -0.2824% | 2.4% | 4.995% | -7.702% | 6.662% | -8.254% | -4.169% | 2.676% | -6.156% | -2.954% | 1.02% | 7.976% | -26.03% | 1.793% | 5.004% | 4.112%
| Z-Score (Absolute) |  | 0.78 | 0.1 | 3.15 | 0.91 | 0.24 | 0.14 | 0.5 | 1.28 | 0.74 | 1.35 | 0.78 | 0.18 | 1.06 | 0.61 | 0.05 | 0.92 | 3.85 | 0.05 | 0.5 | 0.38
| Analytical results for Re in OREAS 265 (Certified Value < 0.001 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.29 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [8 empty] | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [8 empty] | 3
| 1 | 1 | < 0.1 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.002 | < 0.002 | < 0.001 | < 0.001 | < 0.005 | < 0.01 | < 0.001 | [8 empty] | 1
| 1 | 2 | < 0.1 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.002 | < 0.002 | < 0.001 | 0.001 | < 0.005 | < 0.01 | < 0.001 | [8 empty] | 33
| 1 | 3 | < 0.1 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.002 | < 0.002 | < 0.001 | 0.001 | < 0.005 | < 0.01 | < 0.001 | [8 empty] | 16
| 1 | 4 | < 0.1 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | 0.002 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.002 | < 0.002 | < 0.001 | < 0.001 | < 0.005 | < 0.01 | < 0.001 | [8 empty] | < 0.001
| 1 | 5 | < 0.1 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.002 | < 0.002 | < 0.001 | < 0.001 | < 0.005 | < 0.01 | < 0.001 | [8 empty] | 112
| 1 | 6 | < 0.1 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.002 | < 0.002 | < 0.001 | < 0.001 | < 0.005 | < 0.01 | < 0.001
| Mean | [6 empty] | 0.002 | 0.001 | [7 empty] | 0.001
| Median | [6 empty] | 0.002 | 0.001 | [7 empty] | 0.001
| Std Dev. | [15 empty] | 0
| Rel.Std.Dev. | [15 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 800.4 | 0.67 | 0.67 | 395.82 | 0.67 | 3.37 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 7.42 | 7.42 | 0.67 | 2.02 | 31.69 | 72.15 | 0.67
| Analytical results for S in OREAS 265 (Certified Value < 0.01 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  | 3
| 1 | 1 | < 0.01 | 0.01 | 0.007 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.01 | 0.01 | < 0.02 | < 0.02 | < 0.01 | 0.01 | < 0.01 | 0.0069038 | 0.0071 | 0.01 | 0.0067 | < 0.05 | < 0.01 | 0.010902 | < 0.01 | 0.0056924 |  | 1
| 1 | 2 | < 0.01 | 0.01 | 0.007 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.02 | < 0.02 | < 0.01 | 0.01 | < 0.01 | 0.0068283 | 0.0069 | 0.01 | 0.0069 | < 0.05 | < 0.01 | 0.013831 | < 0.01 | 0.006061 |  | 1
| 1 | 3 | < 0.01 | 0.01 | 0.007 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.02 | < 0.02 | < 0.01 | 0.01 | < 0.01 | 0.0068063 | 0.007 | 0.01 | 0.0068 | < 0.05 | < 0.01 | 0.012617 | < 0.01 | 0.0055156 |  | 16
| 1 | 4 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.01 | 0.01 | < 0.02 | < 0.02 | < 0.01 | 0.01 | < 0.01 | 0.0067297 | 0.007 | 0.01 | 0.0067 | < 0.05 | < 0.01 | 0.014815 | < 0.01 | 0.0058835 |  | < 0.01
| 1 | 5 | < 0.01 | 0.01 | 0.007 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.02 | < 0.02 | < 0.01 | 0.01 | < 0.01 | 0.0067558 | 0.0071 | 0.01 | 0.0068 | < 0.05 | < 0.01 | 0.013402 | < 0.01 | 0.0058892 |  | 113
| 1 | 6 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 0.01 | 0.01 | < 0.02 | < 0.02 | < 0.01 | 0.01 | < 0.01 | 0.0068003 | 0.007 | 0.01 | 0.0071 | < 0.05 | < 0.01 | 0.011947 | < 0.01 | 0.005522
| Mean |  |  | 0.01 | 0.0066667 |  |  |  |  | 0.01 |  | 0.01 | 0.01 |  |  |  | 0.01 |  | 0.006804 | 0.0070167 | 0.01 | 0.0068333 |  |  | 0.012919 |  | 0.0057606
| Median |  |  | 0.01 | 0.007 |  |  |  |  | 0.01 |  | 0.01 | 0.01 |  |  |  | 0.01 |  | 0.0068033 | 0.007 | 0.01 | 0.0068 |  |  | 0.01301 |  | 0.005788
| Std Dev. |  |  | 0 | 0.0005164 |  |  |  |  | 0 |  | 0 | 0 |  |  |  | 0 |  | 6.0695e-05 | 7.5277e-05 | 0 | 0.00015055 |  |  | 0.0013974 |  | 0.00022068
| Rel.Std.Dev. |  |  | 0% | 7.746% |  |  |  |  | 0% |  | 0% | 0% |  |  |  | 0% |  | 0.892% | 1.073% | 0% | 2.203% |  |  | 10.82% |  | 3.831%
PDM3
| Z-Score (Absolute) |  | 0.67 | 1 | 0.12 | 0.67 | 6.01 | 0.67 | 0.67 | 0.44 | 6.01 | 1 | 1 | 1 | 1 | 0.67 | 1 | 0.67 | 0.07 | 0 | 1 | 0.06 | 6.01 | 0.67 | 1.97 | 0.67 | 0.42
| Analytical results for Sb in OREAS 265 (Certified Value 1.24 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 2
| 1 | 1 | 1.36 | 1.27 | 1.5 | 1.3716 | 1.33 | 1.28 | 1.03 | 1.31 | 1.18 | 1.12 | 1.38 | 1.06 | 1.14 | 1.43 | 1.32 | 1.19 | 1.17 | 1.24 | 0.54 | 1.09 | < 2.5 | 1.112 | 1.22 |  |  |  | 1
| 1 | 2 | 1.35 | 1.26 | 1.58 | 1.4183 | 1.39 | 1.35 | 1.79 | 1.12 | 1.17 | 1.14 | 1.42 | 1.08 | 1.15 | 1.4 | 1.28 | 1.25 | 1.25 | 1.25 | 0.95 | 1.11 | < 2.5 | 1.0826 | 1.21 | 0.616 |  |  | 34
| 1 | 3 | 1.35 | 1.29 | 1.61 | 1.4062 | 1.34 | 1.26 | 1.06 | 1.04 | 1.17 | 1.16 | 1.36 | 1.07 | 1.2 | 1.42 | 1.31 | 1.18 | 1.33 | 1.21 | 0.63 | 1.07 | < 2.5 | 1.1351 | 1.18 | 0.091 |  |  | 16
| 1 | 4 | 1.37 | 1.32 | 1.46 | 1.3471 | 1.36 | 1.3 | 1.09 | 1.02 | 1.19 | 1.13 | 1.36 | 1.11 | 1.2 | 1.34 | 1.28 | 1.28 | 1.21 | 1.19 | 0.67 | 1.11 | < 2.5 | 1.13 | 1.19 |  |  |  | 1.2433
| 1 | 5 | 1.39 | 1.29 | 1.56 | 1.3742 | 1.31 | 1.32 | 1.24 | 1.06 | 1.2 | 1.16 | 1.32 | 1.13 | 1.2 | 1.36 | 1.28 | 1.23 | 1.08 | 1.18 | 0.87 | 1.13 | < 2.5 | 1.1263 | 1.19 | 0.57 |  |  | 114
| 1 | 6 | 1.4 | 1.3 | 1.53 | 1.3323 | 1.31 | 1.35 | 0.97 | 1.03 | 1.21 | 1.12 | 1.36 | 1.08 | 1.21 | 1.36 | 1.29 | 1.28 | 1.19 | 1.23 | 0.98 | 1.09 | < 2.5 | 1.0663 | 1.23
| Mean |  | 1.37 | 1.2883 | 1.54 | 1.3749 | 1.34 | 1.31 | 1.1967 | 1.0967 | 1.1867 | 1.1383 | 1.3667 | 1.0883 | 1.1833 | 1.385 | 1.2933 | 1.235 | 1.205 | 1.2167 | 0.77333 | 1.1 |  | 1.1087 | 1.2033 | 0.42567
| Median |  | 1.365 | 1.29 | 1.545 | 1.3729 | 1.335 | 1.31 | 1.075 | 1.05 | 1.185 | 1.135 | 1.36 | 1.08 | 1.2 | 1.38 | 1.285 | 1.24 | 1.2 | 1.22 | 0.77 | 1.1 |  | 1.1192 | 1.2 | 0.57
| Std Dev. |  | 0.020976 | 0.02137 | 0.054772 | 0.033024 | 0.030984 | 0.036878 | 0.30435 | 0.11039 | 0.01633 | 0.018348 | 0.03266 | 0.026394 | 0.030111 | 0.036742 | 0.017512 | 0.043243 | 0.083367 | 0.028048 | 0.18381 | 0.020976 |  | 0.028114 | 0.019664 | 0.29074
| Rel.Std.Dev. |  | 1.531% | 1.659% | 3.557% | 2.402% | 2.312% | 2.815% | 25.43% | 10.07% | 1.376% | 1.612% | 2.39% | 2.425% | 2.545% | 2.653% | 1.354% | 3.501% | 6.918% | 2.305% | 23.77% | 1.907% |  | 2.536% | 1.634% | 68.3%
| PDM3 |  | 10.19% | 3.626% | 23.87% | 10.59% | 7.782% | 5.369% | -3.747% | -11.79% | -4.552% | -8.439% | 9.927% | -12.46% | -4.82% | 11.4% | 4.028% | -0.6639% | -3.077% | -2.138% | -37.8% | -11.52% |  | -10.82% | -3.211% | -65.76%
| Z-Score (Absolute) |  | 1.07 | 0.52 | 2.21 | 1.1 | 0.87 | 0.66 | 0.09 | 0.76 | 0.16 | 0.49 | 1.04 | 0.82 | 0.18 | 1.17 | 0.55 | 0.16 | 0.04 | 0.04 | 2.93 | 0.74 | 0.26 | 0.68 | 0.05 | 5.26
| Analytical results for Sc in OREAS 265 (Certified Value 2.82 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 2.7 | 2.6 | 2.7 | 2.5661 | 2.9 | 2.9 | 2.9 | 2.8 | 2.7 | 2.6 | 2.9 | 2.6 | 2.9 | 3 | 2.5 | 2.8 | 3.1838 | 3.1 | 2 | 2.9 | 2.8609 | 2.486 | 2.8 |  |  |  | 1
| 1 | 2 | 2.7 | 2.9 | 2.8 | 2.6668 | 2.9 | 2.9 | 3 | 2.8 | 2.7 | 2.5 | 2.9 | 2.6 | 3.2 | 3 | 2.4 | 3.1 | 3.1588 | 3 | 2.2 | 2.9 | 2.8781 | 2.42 | 2.8 |  |  |  | 35
| 1 | 3 | 2.7 | 2.9 | 2.7 | 2.8808 | 2.7 | 2.8 | 3.1 | 2.8 | 2.7 | 2.7 | 2.8 | 2.6 | 3 | 2.9 | 2.6 | 2.7 | 3.1354 | 3 | 2.1 | 3 | 2.878 | 2.387 | 2.7 |  |  |  | 16
| 1 | 4 | 2.8 | 2.9 | 2.8 | 2.5884 | 2.9 | 2.7 | 3.1 | 2.9 | 2.6 | 2.5 | 2.8 | 2.7 | 2.9 | 3 | 2.5 | 2.9 | 3.1754 | 3 | 2.1 | 3 | 2.8763 | 2.497 | 2.8 |  |  |  | 2.8183
| 1 | 5 | 2.8 | 2.8 | 2.7 | 2.8635 | 2.8 | 2.8 | 3.1 | 2.9 | 2.6 | 2.6 | 2.8 | 2.7 | 2.9 | 2.9 | 2.5 | 3.4 | 3.1094 | 3 | 2.1 | 3 | 2.8743 | 2.475 | 2.8 |  |  |  | 115
| 1 | 6 | 2.8 | 2.8 | 2.6 | 2.5534 | 2.7 | 2.9 | 2.9 | 3 | 2.6 | 2.6 | 2.8 | 2.7 | 3.2 | 2.9 | 2.5 | 3.4 | 3.1721 | 3 | 2.2 | 3 | 2.8685 | 2.464 | 2.9
| Mean |  | 2.75 | 2.8167 | 2.7167 | 2.6865 | 2.8167 | 2.8333 | 3.0167 | 2.8667 | 2.65 | 2.5833 | 2.8333 | 2.65 | 3.0167 | 2.95 | 2.5 | 3.05 | 3.1558 | 3.0167 | 2.1167 | 2.9667 | 2.8727 | 2.4548 | 2.8
| Median |  | 2.75 | 2.85 | 2.7 | 2.6276 | 2.85 | 2.85 | 3.05 | 2.85 | 2.65 | 2.6 | 2.8 | 2.65 | 2.95 | 2.95 | 2.5 | 3 | 3.1654 | 3 | 2.1 | 3 | 2.8753 | 2.4695 | 2.8
| Std Dev. |  | 0.054772 | 0.1169 | 0.075277 | 0.14921 | 0.098319 | 0.08165 | 0.098319 | 0.08165 | 0.054772 | 0.075277 | 0.05164 | 0.054772 | 0.1472 | 0.054772 | 0.063246 | 0.30166 | 0.028336 | 0.040825 | 0.075277 | 0.05164 | 0.0067755 | 0.042555 | 0.063246
| Rel.Std.Dev. |  | 1.992% | 4.15% | 2.771% | 5.554% | 3.491% | 2.882% | 3.259% | 2.848% | 2.067% | 2.914% | 1.823% | 2.067% | 4.879% | 1.857% | 2.53% | 9.891% | 0.8979% | 1.353% | 3.556% | 1.741% | 0.2359% | 1.734% | 2.259%
| PDM3 |  | -2.424% | -0.05889% | -3.607% | -4.677% | -0.05889% | 0.5325% | 7.038% | 1.715% | -5.973% | -8.338% | 0.5325% | -5.973% | 7.038% | 4.672% | -11.29% | 8.22% | 11.97% | 7.038% | -24.9% | 5.263% | 1.929% | -12.9% | -0.6503%
| Z-Score (Absolute) |  | 0.3 | 0 | 0.45 | 0.59 | 0 | 0.07 | 0.9 | 0.22 | 0.75 | 1.05 | 0.07 | 0.75 | 0.9 | 0.6 | 1.42 | 1.05 | 1.52 | 0.9 | 3.15 | 0.67 | 0.25 | 1.63 | 0.07
| Analytical results for Se in OREAS 265 (Indicative Value 0.31 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | < 1 | 0.2 | 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | < 0.2 | 0.2 | 0.2 | 0.2 | 0.5 | < 1 | 1.3 | 0.52 | 0.2 | < 0.1 | 1.5 | < 1.5 | 1.2462 | < 0.2 |  |  |  | 1
| 1 | 2 | < 1 | 0.2 | 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.6 | 0.3 | 0.4 | 0.5 | < 1 | 1 | 0.55 | 0.2 | < 0.1 | 1.7 | < 1.5 | 1.3791 | 0.2 |  |  |  | 7
| 1 | 3 | < 1 | 0.3 | 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.4 | 0.3 | < 0.1 | 0.6 | < 1 | 0.3 | 0.5 | < 0.2 | 0.3 | 1.7 | < 1.5 | 1.2978 | 0.2 |  |  |  | 16
| 1 | 4 | < 1 | 0.2 | 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.5 | 0.2 | 0.3 | 0.5 | < 1 | 0.2 | 0.57 | < 0.2 | 0.5 | 1.7 | < 1.5 | 1.2933 | < 0.2 |  |  |  | 0.31083
| 1 | 5 | < 1 | 0.2 | 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.3 | 0.3 | 0.3 | 0.5 | < 1 | 1 | 0.48 | 0.2 | 0.2 | 1.7 | < 1.5 | 1.4171 | 0.2 |  |  |  | 13
| 1 | 6 | < 1 | 0.2 | 1 | < 1 | < 1 | < 1 | < 1 | 0.2 | < 1 | < 0.2 | 0.3 | 0.2 | 0.4 | 0.5 | < 1 | 0.4 | 0.54 | 0.2 | < 0.1 | 1.8 | < 1.5 | 1.33 | 0.2
| Mean |  |  | 0.21667 | 1 |  |  |  |  | 0.28333 |  | 0.2 | 0.38333 | 0.25 | 0.32 | 0.51667 |  | 0.7 | 0.52667 | 0.2 | 0.33333 | 1.6833 |  | 1.3272 | 0.2
| Median |  |  | 0.2 | 1 |  |  |  |  | 0.3 |  | 0.2 | 0.35 | 0.25 | 0.3 | 0.5 |  | 0.7 | 0.53 | 0.2 | 0.3 | 1.7 |  | 1.3139 | 0.2
| Std Dev. |  |  | 0.040825 | 0 |  |  |  |  | 0.040825 |  | 0 | 0.1472 | 0.054772 | 0.083666 | 0.040825 |  | 0.45607 | 0.033267 | 0 | 0.15275 | 0.098319 |  | 0.062207 | 0
| Rel.Std.Dev. |  |  | 18.84% | 0% |  |  |  |  | 14.41% |  | 0% | 38.4% | 21.91% | 26.15% | 7.902% |  | 65.15% | 6.316% | 0% | 45.83% | 5.841% |  | 4.687% | 0%
| PDM3 |  |  | -30.29% | 221.7% |  |  |  |  | -8.847% |  | -35.66% | 23.32% | -19.57% | 2.949% | 66.22% |  | 125.2% | 69.44% | -35.66% | 7.239% | 441.6% |  | 327% | -35.66%
| Z-Score (Absolute) |  | 0 | 0.88 | 1.56 | 0 | 0 | 0 | 0 | 0.67 | 0 | 1.04 | 0.36 | 0.78 | 0.7 | 0.05 | 0 | 0.62 | 0.08 | 1.04 | 0.96 | 3.68 | 0.78 | 2.57 | 1.04
| Analytical results for Sm in OREAS 265 (Certified Value 4.42 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.17 | 1.20 | 1.22 | [19 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | [19 empty] | 3
| 1 | 1 | 3.87 | 4.672 | 4.3 | 4.9 | 4.3077 | 5.48 | 4.1943 | [19 empty] | 1
| 1 | 2 | 3.87 | 4.836 | 4.19 | 5.2 | 4.2967 | 6.19 | 4.2275 | [19 empty] | 36
| 1 | 3 | 3.99 | 4.78 | 4.67 | 4.8 | 4.3129 | 5.96 | 4.2793 | [19 empty] | 16
| 1 | 4 | 3.76 | 4.745 | 4.49 | 5 | 4.2881 | 6.14 | 4.1815 | [19 empty] | 4.4207
| 1 | 5 | 3.86 | 4.758 | 4.7 | 4.8 | 4.316 | 6.14 | 4.1656 | [19 empty] | 116
| 1 | 6 | 3.92 | 4.725 | 4.48 | 4.9 | 4.2636 | 6.38 | 4.0945
| Mean |  | 3.8783 | 4.7527 | 4.4717 | 4.9333 | 4.2975 | 6.0483 | 4.1905
| Median |  | 3.87 | 4.7515 | 4.485 | 4.9 | 4.3022 | 6.14 | 4.1879
| Std Dev. |  | 0.075741 | 0.054866 | 0.20034 | 0.15055 | 0.019611 | 0.30909 | 0.061935
| Rel.Std.Dev. |  | 1.953% | 1.154% | 4.48% | 3.052% | 0.4563% | 5.11% | 1.478%
| PDM3 |  | -12.27% | 7.51% | 1.154% | 11.6% | -2.786% | 36.82% | -5.207%
| Z-Score (Absolute) |  | 1.42 | 0.67 | 0 | 1.11 | 0.42 | 3.78 | 0.67
| Analytical results for Sn in OREAS 265 (Certified Value 1.4 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | 1.30 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.2g | [5 empty] | 3
| 1 | 1 | 1.4 | 1.6 | 1.71 | 1.4353 | 1.44 | 1.5 | 1.6 | 1.3 | 1.42 | 1.3 | 1.4 | 1.4 | 1.4 | 1.3 | 1.63 | 1.6 | 0.85 | 1.4 | 1.5034 | 1.4 | 1.808 | [5 empty] | 1
| 1 | 2 | 1.4 | 1.7 | 1.71 | 1.3578 | 1.51 | 1.4 | 1.6 | 1.3 | 1.4 | 1.2 | 1.4 | 1.4 | 1.4 | 1.3 | 1.82 | 1.5 | 0.97 | 1.4 | 1.469 | 1.4 | [6 empty] | 11
| 1 | 3 | 1.4 | 1.7 | 1.68 | 1.4095 | 1.51 | 1.4 | 1.3 | 1.3 | 1.44 | 1.3 | 1.3 | 1.4 | 1.4 | 1.4 | 1.64 | 1.5 | 0.94 | 1.4 | 1.5716 | 1.2 | [6 empty] | 16
| 1 | 4 | 1.4 | 1.7 | 1.73 | 1.3767 | 1.47 | 1.4 | 1.5 | 1.3 | 1.46 | 1.3 | 1.4 | 1.4 | 1.4 | 1.4 | 1.7 | 1.5 | 1.33 | 1.4 | 1.47 | 1.4 | 1.297 | [5 empty] | 1.4004
| 1 | 5 | 1.4 | 1.7 | 1.7 | 1.2892 | 1.46 | 1.4 | 1.4 | 1.4 | 1.43 | 1.3 | 1.4 | 1.4 | 1.3 | 1.3 | 1.57 | 1.6 | 0.96 | 1.4 | 1.5116 | 1.3 | 1.239 | [5 empty] | 117
| 1 | 6 | 1.4 | 1.7 | 1.68 | 1.2985 | 1.43 | 1.5 | 1.5 | 1.4 | 1.46 | 1.3 | 1.3 | 1.4 | 1.4 | 1.3 | 1.68 | 1.5 | 1.09 | 1.4 | 1.456 | 1.4 | 1.495
| Mean |  | 1.4 | 1.6833 | 1.7017 | 1.3612 | 1.47 | 1.4333 | 1.4833 | 1.3333 | 1.435 | 1.2833 | 1.3667 | 1.4 | 1.3833 | 1.3333 | 1.6733 | 1.5333 | 1.0233 | 1.4 | 1.4969 | 1.35 | 1.4598
| Median |  | 1.4 | 1.7 | 1.705 | 1.3673 | 1.465 | 1.4 | 1.5 | 1.3 | 1.435 | 1.3 | 1.4 | 1.4 | 1.4 | 1.3 | 1.66 | 1.5 | 0.965 | 1.4 | 1.4867 | 1.4 | 1.396
| Std Dev. |  | 2.4324e-16 | 0.040825 | 0.019408 | 0.058648 | 0.034059 | 0.05164 | 0.1169 | 0.05164 | 0.023452 | 0.040825 | 0.05164 | 2.4324e-16 | 0.040825 | 0.05164 | 0.084774 | 0.05164 | 0.16872 | 2.4324e-16 | 0.042439 | 0.083666 | 0.25674
| Rel.Std.Dev. |  | 1.737e-14% | 2.425% | 1.141% | 4.309% | 2.317% | 3.603% | 7.881% | 3.873% | 1.634% | 3.181% | 3.779% | 1.737e-14% | 2.951% | 3.873% | 5.066% | 3.368% | 16.49% | 1.737e-14% | 2.835% | 6.197% | 17.59%
| PDM3 |  | -0.02852% | 20.2% | 21.51% | -2.8% | 4.97% | 2.352% | 5.922% | -4.789% | 2.471% | -8.359% | -2.409% | -0.02852% | -1.219% | -4.789% | 19.49% | 9.493% | -26.93% | -0.02852% | 6.893% | -3.599% | 4.238%
| Z-Score (Absolute) |  | 0 | 2.87 | 3.05 | 0.39 | 0.71 | 0.34 | 0.84 | 0.67 | 0.35 | 1.18 | 0.34 | 0 | 0.17 | 0.67 | 2.76 | 1.35 | 3.81 | 0 | 0.98 | 0.51 | 0.6
| Analytical results for Sr in OREAS 265 (Certified Value 39.8 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 2
| 1 | 1 | 36.4 | 38.8 | 39 | 40.5352 | 40.32 | 40.3 | 39.5 | 40.8 | 42.41 | 40.3 | 44.1 | 37.6 | 44 | 40.9 | 37.3 | 49 | 40.9308 | 44.5 | 32 | 40 | 39.4692 | 35.78 | 38 | 36.9298 |  |  | 1
| 1 | 2 | 37.1 | 40.6 | 39.6 | 41.1675 | 41.47 | 40 | 40.5 | 40 | 43.13 | 38.8 | 44.5 | 37.6 | 42.8 | 40.7 | 36.6 | 53.2 | 41.036 | 43.8 | 34 | 40 | 40.5874 | 35.6377 | 37.8 | 34.2636 |  |  | 12
| 1 | 3 | 36.1 | 40.8 | 38.9 | 38.7385 | 40.3 | 39.5 | 40.9 | 40 | 43.33 | 40.6 | 42 | 37.5 | 41.4 | 40.5 | 37.4 | 50.6 | 40.906 | 43.4 | 34 | 40 | 41.8703 | 36.3913 | 33.9 | 35.1776 |  |  | 16
| 1 | 4 | 37.1 | 41.9 | 37.3 | 39.9357 | 41.03 | 38.7 | 40.9 | 40.7 | 43.03 | 39.3 | 43.2 | 38.1 | 40 | 40.7 | 37.2 | 53.8 | 40.961 | 43.9 | 35 | 40 | 40.4987 | 36.973 | 38 | 33.6797 |  |  | 39.8064
| 1 | 5 | 37.2 | 40.5 | 39.1 | 41.5342 | 39.98 | 39.8 | 41.7 | 41.7 | 43.29 | 39.5 | 43 | 38.6 | 42.3 | 40.3 | 35.9 | 52 | 40.9905 | 43.8 | 34 | 41 | 40.0607 | 35.6549 | 36.6 | 34.414 |  |  | 118
| 1 | 6 | 37.5 | 40.6 | 39.5 | 38.8972 | 40.85 | 40.2 | 40.2 | 42.4 | 43.13 | 39.7 | 42.8 | 38.5 | 43.2 | 40.4 | 37.1 | 51.5 | 41.0517 | 43.9 | 35 | 40 | 40.3528 | 35.619 | 37.9 | 34.2399
| Mean |  | 36.9 | 40.5333 | 38.9 | 40.1347 | 40.6583 | 39.75 | 40.6167 | 40.9333 | 43.0533 | 39.7 | 43.2667 | 37.9833 | 42.2833 | 40.5833 | 36.9167 | 51.6833 | 40.9793 | 43.8833 | 34 | 40.1667 | 40.4732 | 36.0093 | 37.0333 | 34.7841
| Median |  | 37.1 | 40.6 | 39.05 | 40.2355 | 40.585 | 39.9 | 40.7 | 40.75 | 43.13 | 39.6 | 43.1 | 37.85 | 42.55 | 40.6 | 37.15 | 51.75 | 40.9757 | 43.85 | 34 | 40 | 40.4257 | 35.7174 | 37.85 | 34.3388
| Std Dev. |  | 0.53292 | 0.99532 | 0.83187 | 1.1581 | 0.55438 | 0.58907 | 0.74409 | 0.95429 | 0.33429 | 0.6603 | 0.90701 | 0.48751 | 1.4176 | 0.22286 | 0.57067 | 1.7486 | 0.057692 | 0.35449 | 1.0954 | 0.40825 | 0.79477 | 0.55557 | 1.6256 | 1.1561
| Rel.Std.Dev. |  | 1.444% | 2.456% | 2.138% | 2.886% | 1.364% | 1.482% | 1.832% | 2.331% | 0.7764% | 1.663% | 2.096% | 1.283% | 3.353% | 0.5491% | 1.546% | 3.383% | 0.1408% | 0.8078% | 3.222% | 1.016% | 1.964% | 1.543% | 4.39% | 3.324%
| PDM3 |  | -7.301% | 1.826% | -2.277% | 0.8249% | 2.14% | -0.1416% | 2.036% | 2.831% | 8.157% | -0.2672% | 8.693% | -4.58% | 6.223% | 1.952% | -7.259% | 29.84% | 2.947% | 10.24% | -14.59% | 0.9052% | 1.675% | -9.539% | -6.966% | -12.62%
| Z-Score (Absolute) |  | 1.36 | 0.09 | 0.57 | 0.07 | 0.13 | 0.23 | 0.12 | 0.24 | 1.09 | 0.25 | 1.17 | 0.93 | 0.78 | 0.1 | 1.36 | 4.53 | 0.26 | 1.42 | 2.52 | 0.06 | 0.06 | 1.72 | 1.31 | 2.21
| Analytical results for Ta in OREAS 265 (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [6 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | [6 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | [6 empty] | 22
| 1 | 3 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | [6 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | [6 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | [6 empty] | 119
| 1 | 6 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.1 | < 0.01 | < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 265 (Certified Value 0.45 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [16 empty] | 2
|  |  | 0.25g | 01g | 25g | 0.25g | 0.5g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | [16 empty] | 3
| 1 | 1 | 0.42 | 0.42 | 0.468 | 0.48 | 0.43 | 0.46 | 0.5 | 0.62213 | 0.42 | 0.46389 | [16 empty] | 1
| 1 | 2 | 0.43 | 0.44 | 0.485 | 0.47 | 0.42 | 0.47 | 0.5 | 0.6139 | 0.45 | 0.42907 | [16 empty] | 23
| 1 | 3 | 0.41 | 0.43 | 0.469 | 0.47 | 0.44 | 0.47 | 0.5 | 0.6198 | 0.44 | 0.46946 | [16 empty] | 16
| 1 | 4 | 0.42 | 0.42 | 0.465 | 0.46 | 0.44 | 0.46 | 0.5 | 0.62145 | 0.46 | 0.45666 | [16 empty] | 0.44915
| 1 | 5 | 0.43 | 0.44 | 0.463 | 0.47 | 0.45 | 0.45 | 0.5 | 0.61098 | 0.45 | 0.43752 | [16 empty] | 120
| 1 | 6 | 0.43 | 0.45 | 0.477 | 0.48 | 0.43 | 0.47 | 0.5 | 0.6188 | 0.47 | 0.42556
| Mean |  | 0.42333 | 0.43333 | 0.47117 | 0.47167 | 0.435 | 0.46333 | 0.5 | 0.61784 | 0.44833 | 0.44703
| Median |  | 0.425 | 0.435 | 0.4685 | 0.47 | 0.435 | 0.465 | 0.5 | 0.6193 | 0.45 | 0.44709
| Std Dev. |  | 0.008165 | 0.012111 | 0.0083046 | 0.0075277 | 0.010488 | 0.008165 | 0 | 0.0044454 | 0.017224 | 0.01873
| Rel.Std.Dev. |  | 1.929% | 2.795% | 1.763% | 1.596% | 2.411% | 1.762% | 0% | 0.7195% | 3.842% | 4.19%
| PDM3 |  | -5.748% | -3.521% | 4.902% | 5.013% | -3.15% | 3.158% | 11.32% | 37.56% | -0.1816% | -0.4726%
| Z-Score (Absolute) |  | 1.12 | 0.67 | 1.03 | 1.05 | 0.6 | 0.67 | NA | 7.62 | 0 | 0.06
Results from laboratory 1.16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 265 (Certified Value 0.034 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | < 0.05 | 0.04 | 0.02 | < 0.05 | < 0.1 | < 0.05 | 0.05 | 0.03 | < 0.1 | 0.04 | 0.04 | 0.03 | < 0.02 | 0.05 | < 0.05 | 0.04 | 0.05 | 0.04 | < 0.01 | < 0.1 | < 0.2 | < 0.2 | 0.03 | 9.1027 |  |  | 1
| 1 | 2 | < 0.05 | 0.03 | 0.02 | < 0.05 | < 0.1 | < 0.05 | 0.04 | 0.03 | < 0.1 | 0.04 | 0.04 | 0.03 | < 0.02 | 0.05 | < 0.05 | 0.04 | 0.04 | 0.04 | < 0.01 | < 0.1 | < 0.2 | < 0.2 | 0.03 | 3.3576 |  |  | 24
| 1 | 3 | < 0.05 | 0.04 | 0.02 | < 0.05 | < 0.1 | < 0.05 | < 0.01 | 0.02 | < 0.1 | 0.04 | 0.04 | 0.03 | 0.03 | 0.04 | < 0.05 | < 0.01 | 0.04 | 0.04 | < 0.01 | < 0.1 | < 0.2 | < 0.2 | 0.03 | 4.6776 |  |  | 16
| 1 | 4 | < 0.05 | 0.04 | 0.03 | < 0.05 | < 0.1 | < 0.05 | 0.08 | 0.03 | < 0.1 | 0.04 | 0.04 | 0.04 | 0.03 | 0.05 | < 0.05 | < 0.01 | 0.04 | 0.03 | < 0.01 | < 0.1 | < 0.2 | < 0.2 | 0.04 | 6.5392 |  |  | 0.034015
| 1 | 5 | < 0.05 | 0.03 | 0.03 | < 0.05 | < 0.1 | < 0.05 | 0.08 | 0.03 | < 0.1 | 0.02 | 0.03 | 0.04 | < 0.02 | 0.05 | < 0.05 | < 0.01 | < 0.03 | 0.04 | < 0.01 | < 0.1 | < 0.2 | < 0.2 | 0.03 | 7.7064 |  |  | 121
| 1 | 6 | < 0.05 | 0.03 | 0.02 | < 0.05 | < 0.1 | < 0.05 | 0.05 | 0.03 | < 0.1 | 0.03 | 0.02 | 0.03 | < 0.02 | 0.06 | < 0.05 | < 0.01 | < 0.03 | 0.04 | < 0.01 | < 0.1 | < 0.2 | < 0.2 | 0.04 | 8.2101
| Mean |  |  | 0.035 | 0.023333 |  |  |  | 0.06 | 0.028333 |  | 0.035 | 0.035 | 0.033333 | 0.03 | 0.05 |  | 0.04 | 0.0425 | 0.038333 |  |  |  |  | 0.033333 | 6.5989
| Median |  |  | 0.035 | 0.02 |  |  |  | 0.05 | 0.03 |  | 0.04 | 0.04 | 0.03 | 0.03 | 0.05 |  | 0.04 | 0.04 | 0.04 |  |  |  |  | 0.03 | 7.1228
| Std Dev. |  |  | 0.0054772 | 0.005164 |  |  |  | 0.018708 | 0.0040825 |  | 0.0083666 | 0.0083666 | 0.005164 | 0 | 0.0063246 |  | 0 | 0.005 | 0.0040825 |  |  |  |  | 0.005164 | 2.2042
| Rel.Std.Dev. |  |  | 15.65% | 22.13% |  |  |  | 31.18% | 14.41% |  | 23.9% | 23.9% | 15.49% | 0% | 12.65% |  | 0% | 11.76% | 10.65% |  |  |  |  | 15.49% | 33.4%
| PDM3 |  |  | 2.895% | -31.4% |  |  |  | 76.39% | -16.7% |  | 2.895% | 2.895% | -2.004% | -11.8% | 46.99% |  | 17.59% | 24.94% | 12.69% |  |  |  |  | -2.004% | 19300%
| Z-Score (Absolute) |  | 0.62 | 0.06 | 0.73 | 0.62 | 1.07 | 0.62 | 1.12 | 0.39 | 1.07 | 0.06 | 0.06 | 0.06 | 1.18 | 1.07 | 0.62 | 1.18 | 0.06 | 0.28 | 1.97 | 1.07 | 4.44 | 4.44 | 0.06 | 442.67
| Analytical results for Th in OREAS 265 (Certified Value 8.76 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 8.5 | 8.2 | 8.26 | 9.0014 | 8.88 | 8.99 | 8.12 | 8.5 | 9.14 | 8.3 | 9.2 | 8.5 | 8.1 | 8.6 | 7.6 | 9.3 | 10.3618 | 10 | 8.65 | 7.6 | 8.9478 | 7.6177 | 9.6 | 10.9032 |  |  | 1
| 1 | 2 | 8.4 | 8.4 | 8.4 | 8.6523 | 8.91 | 8.89 | 8.07 | 8.6 | 9.1 | 8.2 | 9.3 | 8.5 | 8.2 | 9 | 7.2 | 9.4 | 10.4009 | 9.6 | 9.7 | 7.6 | 9.1729 | 7.7422 | 9.3 | 9.7132 |  |  | 25
| 1 | 3 | 8.3 | 8.3 | 8.36 | 9.14 | 8.82 | 8.73 | 8.33 | 8.5 | 9.17 | 8.2 | 9 | 8.5 | 8.3 | 8.7 | 7.5 | 9.1 | 10.3844 | 9.4 | 9.27 | 7.6 | 9.0462 | 7.8418 | 8.5 | 9.7903 |  |  | 16
| 1 | 4 | 8.2 | 8.4 | 8.01 | 8.8265 | 8.94 | 8.47 | 7.99 | 8.3 | 9.07 | 8.2 | 9 | 8.4 | 8.5 | 8.6 | 7.2 | 9.5 | 10.4048 | 9.5 | 9.5 | 7.8 | 9.1552 | 8.0437 | 9.4 | 10.0904 |  |  | 8.7598
| 1 | 5 | 8.4 | 8.2 | 8.19 | 9.0972 | 8.84 | 8.65 | 8.47 | 8.8 | 9.15 | 8.4 | 9.1 | 8.8 | 8.1 | 8.7 | 7.4 | 9.4 | 10.364 | 9.5 | 9.77 | 7.9 | 9.2333 | 7.6388 | 9.1 | 10.2068 |  |  | 122
| 1 | 6 | 8.4 | 8.3 | 8.22 | 8.6897 | 8.89 | 8.78 | 8.08 | 8.9 | 9.19 | 8.1 | 8.8 | 8.8 | 8.1 | 8.6 | 7.2 | 9.4 | 10.3719 | 9.4 | 10.2 | 7.8 | 8.8449 | 7.8614 | 9.4 | 10.5843
| Mean |  | 8.3667 | 8.3 | 8.24 | 8.9012 | 8.88 | 8.7517 | 8.1767 | 8.6 | 9.1367 | 8.2333 | 9.0667 | 8.5833 | 8.2167 | 8.7 | 7.35 | 9.35 | 10.3813 | 9.5667 | 9.515 | 7.7167 | 9.0667 | 7.7909 | 9.2167 | 10.2147
| Median |  | 8.4 | 8.3 | 8.24 | 8.9139 | 8.885 | 8.755 | 8.1 | 8.55 | 9.145 | 8.2 | 9.05 | 8.5 | 8.15 | 8.65 | 7.3 | 9.4 | 10.3782 | 9.5 | 9.6 | 7.7 | 9.1007 | 7.792 | 9.35 | 10.1486
| Std Dev. |  | 0.10328 | 0.089443 | 0.13871 | 0.20869 | 0.044272 | 0.18269 | 0.18348 | 0.21909 | 0.044572 | 0.10328 | 0.17512 | 0.17224 | 0.16021 | 0.15492 | 0.17607 | 0.13784 | 0.018509 | 0.22509 | 0.52478 | 0.13292 | 0.14874 | 0.15937 | 0.38687 | 0.45994
| Rel.Std.Dev. |  | 1.234% | 1.078% | 1.683% | 2.345% | 0.4986% | 2.088% | 2.244% | 2.548% | 0.4878% | 1.254% | 1.931% | 2.007% | 1.95% | 1.781% | 2.395% | 1.474% | 0.1783% | 2.353% | 5.515% | 1.722% | 1.641% | 2.046% | 4.197% | 4.503%
| PDM3 |  | -4.488% | -5.249% | -5.934% | 1.614% | 1.372% | -0.09268% | -6.657% | -1.824% | 4.302% | -6.01% | 3.503% | -2.014% | -6.2% | -0.6825% | -16.09% | 6.738% | 18.51% | 9.211% | 8.621% | -11.91% | 3.504% | -11.06% | 5.216% | 16.61%
| Z-Score (Absolute) |  | 0.5 | 0.59 | 0.67 | 0.24 | 0.21 | 0.04 | 0.76 | 0.17 | 0.57 | 0.68 | 0.47 | 0.2 | 0.7 | 0.04 | 1.9 | 0.86 | 2.29 | 1.16 | 1.09 | 1.39 | 0.47 | 1.29 | 0.68 | 2.06
| Analytical results for Ti in OREAS 265 (Certified Value 0.072 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 3
| 1 | 1 | 0.07 | 0.077 | 0.0875 | 0.087145 | 0.0696 | 0.076 | 0.07 | 0.069 | 0.07 | 0.075 | 0.074 | 0.062 | 0.075 | 0.053 | 0.07 | 0.076 | 0.07238 | 0.073 | 0.018 | 0.072 | 0.073558 | 0.034559 | 0.068 | 0.060737 |  |  | 1
| 1 | 2 | 0.07 | 0.078 | 0.0861 | 0.088659 | 0.073 | 0.076 | 0.07 | 0.069 | 0.0712 | 0.073 | 0.071 | 0.06 | 0.073 | 0.051 | 0.07 | 0.082 | 0.07227 | 0.072 | 0.018 | 0.072 | 0.074268 | 0.033831 | 0.069 | 0.052915 |  |  | 26
| 1 | 3 | 0.07 | 0.078 | 0.0888 | 0.073718 | 0.0699 | 0.075 | 0.07 | 0.066 | 0.073 | 0.074 | 0.068 | 0.059 | 0.07 | 0.051 | 0.07 | 0.079 | 0.072282 | 0.074 | 0.018 | 0.072 | 0.077651 | 0.033 | 0.063 | 0.059249 |  |  | 16
| 1 | 4 | 0.07 | 0.079 | 0.0895 | 0.081667 | 0.0702 | 0.072 | 0.07 | 0.066 | 0.0721 | 0.071 | 0.069 | 0.064 | 0.07 | 0.052 | 0.07 | 0.084 | 0.072205 | 0.074 | 0.018 | 0.072 | 0.076766 | 0.03361 | 0.067 | 0.048783 |  |  | 0.072088
| 1 | 5 | 0.07 | 0.079 | 0.0865 | 0.078642 | 0.0692 | 0.073 | 0.07 | 0.065 | 0.0716 | 0.071 | 0.071 | 0.066 | 0.072 | 0.052 | 0.07 | 0.089 | 0.072156 | 0.074 | 0.018 | 0.072 | 0.07574 | 0.032577 | 0.067 | 0.053796 |  |  | 123
| 1 | 6 | 0.07 | 0.078 | 0.0864 | 0.07789 | 0.0706 | 0.075 | 0.07 | 0.065 | 0.072 | 0.073 | 0.067 | 0.063 | 0.074 | 0.053 | 0.07 | 0.086 | 0.072278 | 0.073 | 0.024 | 0.073 | 0.073818 | 0.034043 | 0.065 | 0.053992
| Mean |  | 0.07 | 0.078167 | 0.087467 | 0.081287 | 0.070417 | 0.0745 | 0.07 | 0.066667 | 0.07165 | 0.072833 | 0.07 | 0.062333 | 0.072333 | 0.052 | 0.07 | 0.082667 | 0.072262 | 0.073333 | 0.019 | 0.072167 | 0.0753 | 0.033603 | 0.0665 | 0.054912
| Median |  | 0.07 | 0.078 | 0.087 | 0.080155 | 0.07005 | 0.075 | 0.07 | 0.066 | 0.0718 | 0.073 | 0.07 | 0.0625 | 0.0725 | 0.052 | 0.07 | 0.083 | 0.072274 | 0.0735 | 0.018 | 0.072 | 0.075004 | 0.033721 | 0.067 | 0.053894
| Std Dev. |  | 0 | 0.00075277 | 0.0014038 | 0.0057382 | 0.0013541 | 0.0016432 | 0 | 0.0018619 | 0.0010075 | 0.0016021 | 0.0025298 | 0.002582 | 0.0020656 | 0.00089443 | 0 | 0.0047188 | 7.6134e-05 | 0.0008165 | 0.0024495 | 0.00040825 | 0.001683 | 0.00071775 | 0.0021679 | 0.0043905
| Rel.Std.Dev. |  | 0% | 0.963% | 1.605% | 7.059% | 1.923% | 2.206% | 0% | 2.793% | 1.406% | 2.2% | 3.614% | 4.142% | 2.856% | 1.72% | 0% | 5.708% | 0.1054% | 1.113% | 12.89% | 0.5657% | 2.235% | 2.136% | 3.26% | 7.996%
| PDM3 |  | -2.897% | 8.432% | 21.33% | 12.76% | -2.319% | 3.346% | -2.897% | -7.521% | -0.6078% | 1.034% | -2.897% | -13.53% | 0.3401% | -27.87% | -2.897% | 14.67% | 0.2411% | 1.727% | -73.64% | 0.1089% | 4.456% | -53.39% | -7.752% | -23.83%
| Z-Score (Absolute) |  | 0.18 | 1.24 | 2.87 | 1.79 | 0.11 | 0.6 | 0.18 | 0.76 | 0.11 | 0.31 | 0.18 | 1.52 | 0.23 | 3.32 | 0.18 | 2.03 | 0.21 | 0.4 | 9.08 | 0.2 | 0.74 | 6.53 | 0.79 | 2.81
| Analytical results for Tl in OREAS 265 (Certified Value 0.08 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.29 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.08 | 0.08 | 0.1 | 0.090412 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.05 | 0.09 | 0.1 | 0.1 | 0.09 | 0.08 | < 0.2 | 0.081745 | 0.08 |  |  |  |  | 1
| 1 | 2 | 0.08 | 0.08 | 0.1 | 0.084861 | 0.09 | 0.08 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | 0.08 | 0.06 | 0.05 | 0.09 | 0.1 | 0.09 | 0.08 | 0.08 | < 0.2 | 0.080778 | 0.08 |  |  |  |  | 27
| 1 | 3 | 0.07 | 0.09 | 0.1 | 0.086595 | 0.08 | 0.07 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.04 | 0.08 | 0.09 | 0.09 | 0.08 | 0.08 | < 0.2 | 0.087327 | 0.07 |  |  |  |  | 16
| 1 | 4 | 0.08 | 0.09 | 0.1 | 0.081717 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.06 | 0.04 | 0.08 | 0.09 | 0.09 | 0.08 | 0.09 | < 0.2 | 0.08862 | 0.07 |  |  |  |  | 0.080391
| 1 | 5 | 0.08 | 0.08 | 0.1 | 0.09064 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.04 | 0.08 | 0.1 | < 0.08 | 0.08 | 0.08 | < 0.2 | 0.078575 | 0.07 |  |  |  |  | 124
| 1 | 6 | 0.07 | 0.08 | 0.09 | 0.082721 | 0.09 | 0.08 | 0.08 | 0.07 | 0.09 | 0.07 | 0.07 | 0.09 | 0.08 | 0.04 | 0.08 | 0.1 | 0.09 | 0.07 | 0.08 | < 0.2 | 0.078547 | 0.08
| Mean |  | 0.076667 | 0.083333 | 0.098333 | 0.086158 | 0.085 | 0.078333 | 0.08 | 0.073333 | 0.081667 | 0.07 | 0.071667 | 0.081667 | 0.068333 | 0.043333 | 0.083333 | 0.096667 | 0.092 | 0.08 | 0.081667 |  | 0.082599 | 0.075
| Median |  | 0.08 | 0.08 | 0.1 | 0.085728 | 0.085 | 0.08 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.04 | 0.08 | 0.1 | 0.09 | 0.08 | 0.08 |  | 0.081262 | 0.075
| Std Dev. |  | 0.005164 | 0.005164 | 0.0040825 | 0.0037842 | 0.0054772 | 0.0040825 | 0 | 0.005164 | 0.0040825 | 0 | 0.0040825 | 0.0040825 | 0.0075277 | 0.005164 | 0.005164 | 0.005164 | 0.0044721 | 0.0063246 | 0.0040825 |  | 0.0043651 | 0.0054772
| Rel.Std.Dev. |  | 6.736% | 6.197% | 4.152% | 4.392% | 6.444% | 5.212% | 0% | 7.042% | 4.999% | 0% | 5.696% | 4.999% | 11.02% | 11.92% | 6.197% | 5.342% | 4.861% | 7.906% | 4.999% |  | 5.285% | 7.303%
| PDM3 |  | -4.632% | 3.66% | 22.32% | 7.174% | 5.734% | -2.559% | -0.486% | -8.779% | 1.587% | -12.93% | -10.85% | 1.587% | -15% | -46.1% | 3.66% | 20.25% | 14.44% | -0.486% | 1.587% |  | 2.747% | -6.706%
| Z-Score (Absolute) |  | 0.86 | 0.29 | 2.87 | 0.77 | 0.57 | 0.57 | 0.29 | 1.44 | 0 | 2.01 | 1.72 | 0 | 2.3 | 6.61 | 0.29 | 2.59 | 0.29 | 0.29 | 0 | 3.16 | 0.16 | 1.15
| Analytical results for Tm in OREAS 265 (Certified Value 0.1 ppm) | [27 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.03 | 1.05 | 1.07 | 1.16 | 1.22 | [21 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 2
|  |  | 01g | 25g | 0.5g | 0.5g | 0.2g | [21 empty] | 3
| 1 | 1 | 0.11 | 0.1 | 0.1 | < 0.1 | 0.10794 | [21 empty] | 1
| 1 | 2 | 0.11 | 0.097 | 0.09 | 0.1 | 0.10682 | [21 empty] | 28
| 1 | 3 | 0.11 | 0.098 | 0.1 | 0.1 | 0.10485 | [21 empty] | 16
| 1 | 4 | 0.11 | 0.099 | 0.09 | 0.1 | 0.11405 | [21 empty] | 0.10151
| 1 | 5 | 0.11 | 0.095 | 0.1 | 0.1 | 0.10341 | [21 empty] | 125
| 1 | 6 | 0.1 | 0.098 | 0.09 | 0.1 | 0.10111
| Mean |  | 0.10833 | 0.097833 | 0.095 | 0.1 | 0.10636
| Median |  | 0.11 | 0.098 | 0.095 | 0.1 | 0.10583
| Std Dev. |  | 0.0040825 | 0.0017224 | 0.0054772 | 0 | 0.0044805
| Rel.Std.Dev. |  | 3.768% | 1.761% | 5.766% | 0% | 4.212%
| PDM3 |  | 6.726% | -3.618% | -6.409% | -1.483% | 4.785%
| Z-Score (Absolute) |  | 1.15 | 0 | 0.31 | 0.67 | 0.93
| Analytical results for U in OREAS 265 (Certified Value 1.23 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.14 | 1.18 | 1.25 | 1.331 | 1.26 | 1.32 | 1.22 | 1.13 | 1.34 | 1.14 | 1.33 | 1.2 | 0.96 | 1.24 | 1.27 | 1.23 | 1.4558 | 1.47 | 1.3 | 0.93 | 1.1 | 1.1965 | 1.1982 | 1.31 |  |  | 1
| 1 | 2 | 1.18 | 1.28 | 1.25 | 1.26 | 1.23 | 1.33 | 1.26 | 1.16 | 1.33 | 1.13 | 1.33 | 1.2 | 0.9 | 1.25 | 1.27 | 1.25 | 1.4524 | 1.57 | 1.3 | 1.04 | 1.1 | 1.1753 | 1.2074 | 1.28 |  |  | 29
| 1 | 3 | 1.12 | 1.26 | 1.26 | 1.2737 | 1.21 | 1.25 | 1.21 | 1.12 | 1.35 | 1.14 | 1.28 | 1.2 | 0.98 | 1.23 | 1.28 | 1.2 | 1.4772 | 1.57 | 1.25 | 0.99 | 1.1 | 1.2075 | 1.2312 | 1.25 |  |  | 16
| 1 | 4 | 1.15 | 1.26 | 1.22 | 1.2574 | 1.25 | 1.27 | 1.24 | 1.14 | 1.35 | 1.11 | 1.27 | 1.2 | 0.95 | 1.28 | 1.26 | 1.28 | 1.4662 | 1.53 | 1.26 | 1.21 | 1.1 | 1.262 | 1.2359 | 1.3 |  |  | 1.2328
| 1 | 5 | 1.16 | 1.21 | 1.23 | 1.3342 | 1.24 | 1.28 | 1.25 | 1.17 | 1.36 | 1.13 | 1.26 | 1.2 | 0.97 | 1.24 | 1.27 | 1.26 | 1.4905 | 1.36 | 1.26 | 1.03 | 1.1 | 1.1589 | 1.2133 | 1.28 |  |  | 126
| 1 | 6 | 1.17 | 1.25 | 1.26 | 1.2626 | 1.25 | 1.31 | 1.27 | 1.19 | 1.34 | 1.14 | 1.26 | 1.2 | 0.94 | 1.27 | 1.26 | 1.25 | 1.4143 | 1.53 | 1.24 | 1.11 | 1.1 | 1.2147 | 1.2192 | 1.3
| Mean |  | 1.1533 | 1.24 | 1.245 | 1.2865 | 1.24 | 1.2933 | 1.2417 | 1.1517 | 1.345 | 1.1317 | 1.2883 | 1.2 | 0.95 | 1.2517 | 1.2683 | 1.245 | 1.4594 | 1.505 | 1.2683 | 1.0517 | 1.1 | 1.2025 | 1.2175 | 1.2867
| Median |  | 1.155 | 1.255 | 1.25 | 1.2681 | 1.245 | 1.295 | 1.245 | 1.15 | 1.345 | 1.135 | 1.275 | 1.2 | 0.955 | 1.245 | 1.27 | 1.25 | 1.461 | 1.53 | 1.26 | 1.035 | 1.1 | 1.202 | 1.2162 | 1.29
| Std Dev. |  | 0.021602 | 0.037417 | 0.016432 | 0.03617 | 0.017889 | 0.031411 | 0.023166 | 0.026394 | 0.010488 | 0.01169 | 0.033116 | 0 | 0.028284 | 0.019408 | 0.0075277 | 0.027386 | 0.026163 | 0.079937 | 0.025626 | 0.097656 | 2.4324e-16 | 0.035744 | 0.014275 | 0.021602
| Rel.Std.Dev. |  | 1.873% | 3.017% | 1.32% | 2.812% | 1.443% | 2.429% | 1.866% | 2.292% | 0.7798% | 1.033% | 2.57% | 0% | 2.977% | 1.551% | 0.5935% | 2.2% | 1.793% | 5.311% | 2.02% | 9.286% | 2.211e-14% | 2.973% | 1.172% | 1.679%
| PDM3 |  | -6.448% | 0.582% | 0.9876% | 4.351% | 0.582% | 4.908% | 0.7172% | -6.583% | 9.099% | -8.205% | 4.503% | -2.663% | -22.94% | 1.528% | 2.88% | 0.9876% | 18.38% | 22.08% | 2.88% | -14.69% | -10.77% | -2.461% | -1.24% | 4.367%
| Z-Score (Absolute) |  | 1.4 | 0.05 | 0.03 | 0.67 | 0.05 | 0.78 | 0.03 | 1.43 | 1.58 | 1.74 | 0.7 | 0.67 | 4.56 | 0.13 | 0.39 | 0.03 | 3.36 | 4.07 | 0.39 | 2.98 | 2.23 | 0.64 | 0.4 | 0.67
| Analytical results for V in OREAS 265 (Certified Value 27.3 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  | 2
| 1 | 1 | 27 | 27 | 30 | 30.1413 | 28 | 28 | 28 | 27 | 27 | 27 | 28 | 27 | 29 | 26 | 28 | 29 | 24.0205 | 27 | 18.6 | 25 | 27.8317 | 22.83 | 27 | 24.2288 |  |  | 1
| 1 | 2 | 27 | 28 | 29 | 27.7162 | 29 | 28 | 28 | 27 | 27 | 28 | 27 | 26 | 28 | 25 | 28 | 30 | 24.2615 | 27 | 21.5 | 25 | 28.5334 | 22.72 | 28 | 23.9925 |  |  | 30
| 1 | 3 | 27 | 28 | 30 | 27.8335 | 28 | 28 | 28 | 26 | 27 | 28 | 26 | 27 | 28 | 25 | 29 | 30 | 24.0897 | 26 | 20.4 | 25 | 28.5911 | 22.62 | 26 | 26.3059 |  |  | 16
| 1 | 4 | 27 | 28 | 30 | 28.9232 | 29 | 27 | 28 | 26 | 27 | 27 | 26 | 27 | 27 | 25 | 28 | 30 | 24.0732 | 27 | 21 | 25 | 28.2497 | 22.66 | 27 | 23.2383 |  |  | 27.2565
| 1 | 5 | 27 | 29 | 30 | 29.9897 | 28 | 28 | 29 | 26 | 28 | 27 | 27 | 27 | 28 | 25 | 28 | 30 | 24.0431 | 27 | 21.2 | 26 | 28.2967 | 23.02 | 27 | 24.2972 |  |  | 127
| 1 | 6 | 28 | 28 | 30 | 27.4254 | 28 | 28 | 28 | 26 | 28 | 28 | 26 | 27 | 29 | 25 | 28 | 31 | 23.9287 | 27 | 22.6 | 26 | 27.8664 | 23.18 | 27 | 24.2762
| Mean |  | 27.1667 | 28 | 29.8333 | 28.6715 | 28.3333 | 27.8333 | 28.1667 | 26.3333 | 27.3333 | 27.5 | 26.6667 | 26.8333 | 28.1667 | 25.1667 | 28.1667 | 30 | 24.0694 | 26.8333 | 20.8833 | 25.3333 | 28.2282 | 22.8383 | 27 | 24.3898
| Median |  | 27 | 28 | 30 | 28.3783 | 28 | 28 | 28 | 26 | 27 | 27.5 | 26.5 | 27 | 28 | 25 | 28 | 30 | 24.0581 | 27 | 21.1 | 25 | 28.2732 | 22.775 | 27 | 24.2525
| Std Dev. |  | 0.40825 | 0.63246 | 0.40825 | 1.1941 | 0.5164 | 0.40825 | 0.40825 | 0.5164 | 0.5164 | 0.54772 | 0.8165 | 0.40825 | 0.75277 | 0.40825 | 0.40825 | 0.63246 | 0.10972 | 0.40825 | 1.3333 | 0.5164 | 0.32188 | 0.22058 | 0.63246 | 1.0201
| Rel.Std.Dev. |  | 1.503% | 2.259% | 1.368% | 4.165% | 1.823% | 1.467% | 1.449% | 1.961% | 1.889% | 1.992% | 3.062% | 1.521% | 2.673% | 1.622% | 1.449% | 2.108% | 0.4558% | 1.521% | 6.384% | 2.038% | 1.14% | 0.9658% | 2.342% | 4.182%
| PDM3 |  | -0.3295% | 2.728% | 9.454% | 5.192% | 3.951% | 2.116% | 3.339% | -3.387% | 0.282% | 0.8935% | -2.164% | -1.552% | 3.339% | -7.667% | 3.339% | 10.07% | -11.69% | -1.552% | -23.38% | -7.056% | 3.565% | -16.21% | -0.941% | -10.52%
| Z-Score (Absolute) |  | 0.06 | 0.55 | 1.9 | 1.05 | 0.8 | 0.43 | 0.67 | 0.67 | 0.06 | 0.18 | 0.43 | 0.31 | 0.67 | 1.53 | 0.67 | 2.02 | 2.34 | 0.31 | 4.68 | 1.41 | 0.72 | 3.25 | 0.18 | 2.1
| Analytical results for W in OREAS 265 (Certified Value 2.03 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.29 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.2g |  |  |  | 2
| 1 | 1 | 2.8 | 2.39 | 2.14 | 2.5628 | 1.83 | 1.92 | 1.98 | 2.09 | 2.48 | 1.82 | 2.36 | 1.5 | 1.61 | 1.32 | 2.6 | 1.37 | 2.22 | 0.4 | 1.8 | 2.057 | 2.04 | 1.85 | 4.4721 |  |  |  | 1
| 1 | 2 | 2.5 | 2.22 | 2.17 | 2.3088 | 1.97 | 1.95 | 2.1 | 2.16 | 2.47 | 1.82 | 2.38 | 1.6 | 1.6 | 1.33 | 2.7 | 1.6 | 2.16 | 0.4 | 1.8 | 2.1779 | 1.9691 | 1.78 | 4.5777 |  |  |  | 31
| 1 | 3 | 2.5 | 2.25 | 2.1 | 2.3485 | 1.91 | 1.97 | 2.1 | 1.98 | 2.58 | 1.89 | 2.16 | 1.6 | 1.73 | 1.31 | 2.7 | 1.3 | 2.09 | 0.4 | 1.9 | 2.1165 | 2.0506 | 1.8 | 5.3777 |  |  |  | 16
| 1 | 4 | 2.4 | 2.25 | 2.19 | 2.3764 | 1.85 | 1.92 | 2.02 | 1.96 | 2.55 | 1.86 | 2.22 | 1.6 | 1.65 | 1.29 | 2.8 | 1.5 | 2.09 | 0.7 | 1.9 | 2.0677 | 1.9588 | 1.84 | 3.403 |  |  |  | 2.0302
| 1 | 5 | 2.5 | 2.32 | 2.02 | 2.6198 | 1.98 | 1.95 | 2.2 | 2.05 | 2.46 | 1.98 | 2.12 | 1.8 | 1.78 | 1.29 | 2.8 | 1.68 | 2.13 | 0.4 | 1.7 | 2.2222 | 1.9726 | 1.86 | 4.8716 |  |  |  | 128
| 1 | 6 | 2.8 | 2.31 | 1.98 | 2.2711 | 1.9 | 2.06 | 2.1 | 1.94 | 2.49 | 1.84 | 2.17 | 1.7 | 1.64 | 1.29 | 2.6 | 1.24 | 2.42 | 0.4 | 1.9 | 2.1333 | 1.9878 | 1.7 | 4.1662
| Mean |  | 2.5833 | 2.29 | 2.1 | 2.4146 | 1.9067 | 1.9617 | 2.0833 | 2.03 | 2.505 | 1.8683 | 2.235 | 1.6333 | 1.6683 | 1.305 | 2.7 | 1.4483 | 2.185 | 0.45 | 1.8333 | 2.1291 | 1.9965 | 1.805 | 4.4781
| Median |  | 2.5 | 2.28 | 2.12 | 2.3625 | 1.905 | 1.95 | 2.1 | 2.015 | 2.485 | 1.85 | 2.195 | 1.6 | 1.645 | 1.3 | 2.7 | 1.435 | 2.145 | 0.4 | 1.85 | 2.1249 | 1.9802 | 1.82 | 4.5249
| Std Dev. |  | 0.17224 | 0.06229 | 0.084143 | 0.1426 | 0.060882 | 0.051929 | 0.076333 | 0.085323 | 0.048477 | 0.0608 | 0.1095 | 0.10328 | 0.071391 | 0.017607 | 0.089443 | 0.17371 | 0.12502 | 0.12247 | 0.08165 | 0.063545 | 0.039061 | 0.059917 | 0.66742
| Rel.Std.Dev. |  | 6.667% | 2.72% | 4.007% | 5.906% | 3.193% | 2.647% | 3.664% | 4.203% | 1.935% | 3.254% | 4.899% | 6.323% | 4.279% | 1.349% | 3.313% | 11.99% | 5.722% | 27.22% | 4.454% | 2.985% | 1.956% | 3.319% | 14.9%
| PDM3 |  | 27.24% | 12.8% | 3.437% | 18.93% | -6.086% | -3.377% | 2.616% | -0.01128% | 23.39% | -7.974% | 10.09% | -19.55% | -17.83% | -35.72% | 32.99% | -28.66% | 7.623% | -77.84% | -9.698% | 4.87% | -1.663% | -11.09% | 120.6%
| Z-Score (Absolute) |  | 1.66 | 0.78 | 0.21 | 1.15 | 0.37 | 0.2 | 0.16 | 0 | 1.42 | 0.48 | 0.61 | 1.19 | 1.08 | 2.17 | 2.01 | 1.74 | 0.46 | 4.74 | 0.59 | 0.3 | 0.1 | 0.67 | 7.34
| Analytical results for Y in OREAS 265 (Certified Value 8.7 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.2g |  |  |  |  | 3
| 1 | 1 | 8.25 | 8.42 | 9.06 | 8.8342 | 8.72 | 9.22 | 8.8 | 8.67 | 8.42 | 8.14 | 9.33 | 8.57 | 8.91 | 8.64 | 8.59 | 8.4153 | 9.79 | 7.77 | 8.6 | 8.5864 | 9.03 | 9.3899 |  |  |  |  | 1
| 1 | 2 | 8.37 | 8.79 | 9.03 | 8.6053 | 9.02 | 9.29 | 8.9 | 8.38 | 8.42 | 7.99 | 9.27 | 8.35 | 8.9 | 8.47 | 9.01 | 8.3857 | 9.6 | 8.38 | 8.5 | 8.9665 | 8.9 | 9.3386 |  |  |  |  | 32
| 1 | 3 | 8.21 | 8.78 | 8.5 | 8.2022 | 8.66 | 8.95 | 8.8 | 8.38 | 8.56 | 8.51 | 8.8 | 8.47 | 8.71 | 8.72 | 8.79 | 8.4123 | 9.57 | 8.14 | 8.5 | 8.9289 | 8.83 | 9.2505 |  |  |  |  | 16
| 1 | 4 | 8.31 | 8.86 | 8.66 | 8.5901 | 8.82 | 9.01 | 8.8 | 8.3 | 8.5 | 8.19 | 8.99 | 8.7 | 9.02 | 8.61 | 9.06 | 8.362 | 9.51 | 8.28 | 8.6 | 8.7315 | 9.01 | 8.885 |  |  |  |  | 8.7041
| 1 | 5 | 8.44 | 8.87 | 8.88 | 8.8896 | 8.75 | 8.8 | 9.1 | 8.86 | 8.57 | 8.4 | 9.11 | 8.73 | 8.78 | 8.61 | 8.87 | 8.3797 | 9.55 | 8.41 | 8.8 | 8.77 | 8.95 | 9.1647 |  |  |  |  | 129
| 1 | 6 | 8.47 | 8.76 | 9.13 | 8.1257 | 8.9 | 9.3 | 8.7 | 8.79 | 8.53 | 8.29 | 9.01 | 8.66 | 8.91 | 8.46 | 8.9 | 8.3505 | 9.57 | 8.51 | 8.7 | 8.5886 | 9.01 | 8.8878
| Mean |  | 8.3417 | 8.7467 | 8.8767 | 8.5412 | 8.8117 | 9.095 | 8.85 | 8.5633 | 8.5 | 8.2533 | 9.085 | 8.58 | 8.8717 | 8.585 | 8.87 | 8.3843 | 9.5983 | 8.2483 | 8.6167 | 8.762 | 8.955 | 9.1528
| Median |  | 8.34 | 8.785 | 8.955 | 8.5977 | 8.785 | 9.115 | 8.8 | 8.525 | 8.515 | 8.24 | 9.06 | 8.615 | 8.905 | 8.61 | 8.885 | 8.3827 | 9.57 | 8.33 | 8.6 | 8.7507 | 8.98 | 9.2076
| Std Dev. |  | 0.10362 | 0.16609 | 0.24905 | 0.31663 | 0.13152 | 0.2054 | 0.13784 | 0.23972 | 0.066633 | 0.18705 | 0.19532 | 0.14724 | 0.1098 | 0.10134 | 0.16817 | 0.026083 | 0.098472 | 0.26589 | 0.1169 | 0.16222 | 0.07791 | 0.22016
| Rel.Std.Dev. |  | 1.242% | 1.899% | 2.806% | 3.707% | 1.493% | 2.258% | 1.558% | 2.799% | 0.7839% | 2.266% | 2.15% | 1.716% | 1.238% | 1.18% | 1.896% | 0.3111% | 1.026% | 3.224% | 1.357% | 1.851% | 0.87% | 2.405%
| PDM3 |  | -4.164% | 0.4892% | 1.983% | -1.872% | 1.236% | 4.491% | 1.676% | -1.617% | -2.345% | -5.179% | 4.376% | -1.426% | 1.925% | -1.368% | 1.906% | -3.675% | 10.27% | -5.236% | -1.004% | 0.665% | 2.883% | 5.155%
| Z-Score (Absolute) |  | 1.42 | 0.03 | 0.42 | 0.73 | 0.2 | 1.17 | 0.33 | 0.66 | 0.88 | 1.73 | 1.14 | 0.6 | 0.4 | 0.58 | 0.4 | 1.27 | 2.91 | 1.74 | 0.47 | 0.03 | 0.69 | 1.37
| Analytical results for Yb in OREAS 265 (Certified Value 0.58 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.15 | 1.16 | 1.20 | 1.22 | [17 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [17 empty] | 2
|  |  | 0.25g | 01g | 25g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.2g | [17 empty] | 3
| 1 | 1 | 0.5 | 0.63 | 0.577 | 0.6 | 0.59 | 0.6 | 0.6 | 0.52 | 0.53823 | [17 empty] | 1
| 1 | 2 | 0.5 | 0.62 | 0.577 | 0.6 | 0.67 | 0.6 | 0.6 | 0.54 | 0.55858 | [17 empty] | 33
| 1 | 3 | 0.5 | 0.63 | 0.572 | 0.6 | 0.61 | 0.6 | 0.6 | 0.53 | 0.55763 | [17 empty] | 16
| 1 | 4 | 0.5 | 0.65 | 0.584 | 0.6 | 0.62 | 0.6 | 0.6 | 0.54 | 0.56922 | [17 empty] | 0.58151
| 1 | 5 | 0.5 | 0.61 | 0.582 | 0.6 | 0.62 | 0.6 | 0.6 | 0.54 | 0.57691 | [17 empty] | 130
| 1 | 6 | 0.6 | 0.62 | 0.565 | 0.6 | 0.58 | 0.6 | 0.6 | 0.56 | 0.5641
| Mean |  | 0.51667 | 0.62667 | 0.57617 | 0.6 | 0.615 | 0.6 | 0.6 | 0.53833 | 0.56078
| Median |  | 0.5 | 0.625 | 0.577 | 0.6 | 0.615 | 0.6 | 0.6 | 0.54 | 0.56134
| Std Dev. |  | 0.040825 | 0.013663 | 0.0069113 | 0 | 0.031464 | 0 | 0 | 0.013292 | 0.013157
| Rel.Std.Dev. |  | 7.902% | 2.18% | 1.2% | 0% | 5.116% | 0% | 0% | 2.469% | 2.346%
| PDM3 |  | -11.15% | 7.765% | -0.9193% | 3.179% | 5.759% | 3.179% | 3.179% | -7.425% | -3.565%
| Z-Score (Absolute) |  | 2.36 | 0.75 | 0.67 | 0 | 0.42 | 0 | 0 | 1.74 | 1.11
| Analytical results for Zn in OREAS 265 (Certified Value 36.3 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.2g |  | 2
| 1 | 1 | 34 | 37 | 39 | 35.9348 | 36 | 38 | 38 | 38 | 35 | 36 | 37 | 36.1 | 36.4 | 31 | 36 | 38 | 31 | 36 | 23.9 | 35 | 36.7659 | 25.09 | 51.343 | 35 | 33.2252 |  | 1
| 1 | 2 | 33 | 37 | 39 | 37.9631 | 38 | 38 | 39 | 35 | 36 | 36 | 35 | 35.4 | 37.7 | 30 | 36 | 40 | 30 | 36 | 26.6 | 36 | 37.4441 | 24.62 | 52.446 | 35 | 31.5566 |  | 15
| 1 | 3 | 33 | 37 | 40 | 35.7832 | 36 | 37 | 39 | 35 | 36 | 36 | 34 | 35.4 | 37.1 | 30 | 37 | 39 | 30 | 35 | 25.4 | 36 | 38.3269 | 25.041 | 52.084 | 34 | 30.3478 |  | 16
| 1 | 4 | 33 | 37 | 40 | 39.1427 | 37 | 36 | 39 | 35 | 36 | 35 | 34 | 36.7 | 36.7 | 30 | 36 | 39 | 31 | 36 | 26.2 | 36 | 37.7918 | 24.39 | 52.18 | 34 | 31.9493 |  | 36.3123
| 1 | 5 | 34 | 38 | 40 | 34.1812 | 37 | 37 | 39 | 35 | 36 | 35 | 35 | 37.3 | 37.5 | 30 | 35 | 39 | 31 | 37 | 25.8 | 37 | 37.4782 | 25.03 | 51.804 | 34 | 33.4554 |  | 131
| 1 | 6 | 34 | 37 | 39 | 37.22 | 38 | 37 | 39 | 35 | 36 | 36 | 34 | 36.2 | 37.5 | 30 | 36 | 40 | 31 | 36 | 27.7 | 37 | 36.991 | 24.65 | 51.924 | 35 | 31.8596
| Mean |  | 33.5 | 37.1667 | 39.5 | 36.7042 | 37 | 37.1667 | 38.8333 | 35.5 | 35.8333 | 35.6667 | 34.8333 | 36.1833 | 37.15 | 30.1667 | 36 | 39.1667 | 30.6667 | 36 | 25.9333 | 36.1667 | 37.4663 | 24.8035 | 51.9635 | 34.5 | 32.0656
| Median |  | 33.5 | 37 | 39.5 | 36.5774 | 37 | 37 | 39 | 35 | 36 | 36 | 34.5 | 36.15 | 37.3 | 30 | 36 | 39 | 31 | 36 | 26 | 36 | 37.4612 | 24.84 | 52.004 | 34.5 | 31.9044
| Std Dev. |  | 0.54772 | 0.40825 | 0.54772 | 1.7665 | 0.89443 | 0.75277 | 0.40825 | 1.2247 | 0.40825 | 0.5164 | 1.169 | 0.7414 | 0.51284 | 0.40825 | 0.63246 | 0.75277 | 0.5164 | 0.63246 | 1.2707 | 0.75277 | 0.55926 | 0.28914 | 0.37582 | 0.54772 | 1.1439
| Rel.Std.Dev. |  | 1.635% | 1.098% | 1.387% | 4.813% | 2.417% | 2.025% | 1.051% | 3.45% | 1.139% | 1.448% | 3.356% | 2.049% | 1.38% | 1.353% | 1.757% | 1.922% | 1.684% | 1.757% | 4.9% | 2.081% | 1.493% | 1.166% | 0.7232% | 1.588% | 3.567%
| PDM3 |  | -7.745% | 2.353% | 8.779% | 1.079% | 1.894% | 2.353% | 6.943% | -2.237% | -1.319% | -1.778% | -4.073% | -0.3552% | 2.307% | -16.92% | -0.8601% | 7.861% | -15.55% | -0.8601% | -28.58% | -0.4011% | 3.178% | -31.69% | 43.1% | -4.991% | -11.69%
| Z-Score (Absolute) |  | 1.44 | 0.67 | 2.02 | 0.41 | 0.58 | 0.67 | 1.64 | 0.29 | 0.1 | 0.19 | 0.67 | 0.11 | 0.66 | 3.37 | 0 | 1.83 | 3.08 | 0 | 5.82 | 0.1 | 0.85 | 6.47 | 9.23 | 0.87 | 2.27
| Analytical results for Zr in OREAS 265 (Certified Value 31.1 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.29 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 0.25g | 0.5g | 01g | 01g | 25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.2g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | [5 empty] | 1
| 1 | 1 | 34.6 | 29.4 | 35.5 | 37.2835 | 9.6 | 34.7 | 29 | 24.2 | 31.9 | 23 | 27.9 | 13.8 | 32.1 | 34.8 | 3.6 | 38.63 | 22.2 | 30 | 30.8 | 30.8885 | 33.4 | [5 empty] | 1
| 1 | 2 | 34.3 | 25.3 | 35.4 | 34.0788 | 11.4 | 34.7 | 30 | 25.9 | 33.3 | 21.7 | 26.8 | 19.8 | 32.3 | 34.3 | 3.9 | 38.39 | 22.1 | 32.6 | 33 | 30.3514 | 33.2 | [5 empty] | 16
| 1 | 3 | 34.3 | 27.5 | 33.7 | 35.1719 | 9.4 | 35 | 28 | 23.5 | 33.2 | 22.3 | 24.4 | 21.4 | 33.5 | 34.9 | 3.6 | 38.64 | 25.9 | 31.7 | 33.5 | 31.5673 | 28.9 | [5 empty] | 16
| 1 | 4 | 33.5 | 29.1 | 33.6 | 34.6504 | 9 | 33.6 | 30 | 22.2 | 34.1 | 20.7 | 26 | 19.7 | 32 | 34.7 | 2.3 | 38.61 | 24 | 32.4 | 34.6 | 31.4003 | 32.6 | [5 empty] | 31.0582
| 1 | 5 | 35 | 28 | 34.3 | 37.1037 | 9.2 | 33.5 | 30 | 24.8 | 32.7 | 22.4 | 24.3 | 20.6 | 33 | 34.5 | 3.3 | 38.29 | 23.3 | 32.5 | 34 | 30.95 | 30.8 | [5 empty] | 132
| 1 | 6 | 34.6 | 28.9 | 35 | 33.315 | 9 | 34.1 | 31 | 22 | 33.3 | 23.8 | 23.9 | 16.2 | 32.2 | 34.9 | 1.5 | 38.7 | 27.7 | 35.4 | 34.2 | 30.3611 | 30
| Mean |  | 34.3833 | 28.0333 | 34.5833 | 35.2672 | 9.6 | 34.2667 | 29.6667 | 23.7667 | 33.0833 | 22.3167 | 25.55 | 18.5833 | 32.5167 | 34.6833 | 3.0333 | 38.5433 | 24.2 | 32.4333 | 33.35 | 30.9198 | 31.4833
| Median |  | 34.45 | 28.45 | 34.65 | 34.9111 | 9.3 | 34.4 | 30 | 23.85 | 33.25 | 22.35 | 25.2 | 19.75 | 32.25 | 34.75 | 3.45 | 38.62 | 23.65 | 32.45 | 33.75 | 30.9192 | 31.7
| Std Dev. |  | 0.50365 | 1.5175 | 0.83766 | 1.6156 | 0.91214 | 0.62823 | 1.0328 | 1.5135 | 0.73326 | 1.0647 | 1.6059 | 2.9424 | 0.59805 | 0.24014 | 0.9331 | 0.16342 | 2.2091 | 1.7489 | 1.3678 | 0.5074 | 1.8552
| Rel.Std.Dev. |  | 1.465% | 5.413% | 2.422% | 4.581% | 9.501% | 1.833% | 3.481% | 6.368% | 2.216% | 4.771% | 6.285% | 15.83% | 1.839% | 0.6924% | 30.76% | 0.424% | 9.128% | 5.392% | 4.101% | 1.641% | 5.893%
| PDM3 |  | 10.71% | -9.739% | 11.35% | 13.55% | -69.09% | 10.33% | -4.48% | -23.48% | 6.521% | -28.15% | -17.73% | -40.17% | 4.696% | 11.67% | -90.23% | 24.1% | -22.08% | 4.428% | 7.379% | -0.4456% | 1.369%
| Z-Score (Absolute) |  | 0.61 | 0.73 | 0.65 | 0.8 | 4.61 | 0.59 | 0.38 | 1.63 | 0.34 | 1.93 | 1.25 | 2.72 | 0.22 | 0.67 | 6 | 1.49 | 1.53 | 0.2 | 0.39 | 0.12 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C870: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C906: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C942: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C961: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1015: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1033: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1051: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1069: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1087: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1105: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1123: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1159: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 265 (Indicative Value 12.86 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.85 | 1
| 1 | 2 | 12.87 | 9
| Mean |  | 12.86 | 16
| Median |  | 12.86 | 12.86
| Std Dev. |  | 0.014142 | 15
| Rel.Std.Dev. |  | 0.11%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 265 (Indicative Value 200 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 200 | 1
| 1 | 2 | 200 | 10
| Mean |  | 200 | 16
| Median |  | 200 | 200
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 265 (Indicative Value 610 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 590 | 1
| 1 | 2 | 630 | 11
| Mean |  | 610 | 16
| Median |  | 610 | 610
| Std Dev. |  | 28.2843 | 17
| Rel.Std.Dev. |  | 4.637%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 265 (Indicative Value 1.21 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.21 | 12
| Mean |  | 1.205 | 16
| Median |  | 1.205 | 1.205
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.5868%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 265 (Indicative Value 175 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 190 | 1
| 1 | 2 | 160 | 13
| Mean |  | 175 | 16
| Median |  | 175 | 175
| Std Dev. |  | 21.2132 | 19
| Rel.Std.Dev. |  | 12.12%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 265 (Indicative Value 20 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 14
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 265 (Indicative Value 165 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 160 | 15
| Mean |  | 165 | 16
| Median |  | 165 | 165
| Std Dev. |  | 7.0711 | 21
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 265 (Indicative Value 30 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 16
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 265 (Indicative Value 5 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.99 | 1
| 1 | 2 | 5.01 | 9
| Mean |  | 5 | 16
| Median |  | 5 | 5
| Std Dev. |  | 0.014142 | 15
| Rel.Std.Dev. |  | 0.2828%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 265 (Indicative Value 2.36 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.35 | 1
| 1 | 2 | 2.36 | 10
| Mean |  | 2.355 | 16
| Median |  | 2.355 | 2.355
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.3003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 265 (Indicative Value 1.47 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.47 | 1
| 1 | 2 | 1.47 | 11
| Mean |  | 1.47 | 16
| Median |  | 1.47 | 1.47
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 265 (Indicative Value 0.04 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.039 | 1
| 1 | 2 | 0.04 | 12
| Mean |  | 0.0395 | 16
| Median |  | 0.0395 | 0.0395
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 265 (Indicative Value 1.03 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 1.03 | 13
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 265 (Indicative Value 60 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 60 | 14
| Mean |  | 60 | 16
| Median |  | 60 | 60
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 265 (Indicative Value 0.175 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.175 | 1
| 1 | 2 | 0.175 | 15
| Mean |  | 0.175 | 16
| Median |  | 0.175 | 0.175
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 265 (Indicative Value 45 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 50 | 16
| Mean |  | 45 | 16
| Median |  | 45 | 45
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 265 (Indicative Value 0.008 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0084 | 1
| 1 | 2 | 0.0084 | 1
| Mean |  | 0.0084 | 16
| Median |  | 0.0084 | 0.0084095
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.1124%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 265 (Indicative Value 71.93 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 71.95 | 1
| 1 | 2 | 71.9 | 10
| Mean |  | 71.925 | 16
| Median |  | 71.925 | 71.925
| Std Dev. |  | 0.035355 | 16
| Rel.Std.Dev. |  | 0.04916%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 265 (Indicative Value 20 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 11
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 265 (Indicative Value 211 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 250 | 12
| Mean |  | 210 | 16
| Median |  | 210 | 211.3986
| Std Dev. |  | 56.5685 | 18
| Rel.Std.Dev. |  | 26.94%
| PDM3 |  | -0.6616%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 265 (Indicative Value 0.85 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.847 | 1
| 1 | 2 | 0.852 | 13
| Mean |  | 0.8495 | 16
| Median |  | 0.8495 | 0.8495
| Std Dev. |  | 0.0035355 | 19
| Rel.Std.Dev. |  | 0.4162%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 265 (Indicative Value 155 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 160 | 1
| 1 | 2 | 150 | 14
| Mean |  | 155 | 16
| Median |  | 155 | 155
| Std Dev. |  | 7.0711 | 20
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 265 (Indicative Value 65 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 70 | 15
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 7.0711 | 21
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 265 (Indicative Value 267 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 259 | 1
| 1 | 2 | 274 | 16
| Mean |  | 266.5 | 16
| Median |  | 266.5 | 266.5159
| Std Dev. |  | 10.6066 | 22
| Rel.Std.Dev. |  | 3.98%
| PDM3 |  | -0.005973%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 265 (Indicative Value 2.8 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.82 | 1
| 1 | 2 | 2.77 | 18
| Mean |  | 2.795 | 16
| Median |  | 2.795 | 2.795
| Std Dev. |  | 0.035355 | 24
| Rel.Std.Dev. |  | 1.265%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 265 (Indicative Value 0.65 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.7 | 20
| Mean |  | 0.65 | 16
| Median |  | 0.65 | 0.65
| Std Dev. |  | 0.070711 | 26
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 265 (Indicative Value 199 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 197 | 1
| 1 | 2 | 201 | 10
| Mean |  | 199 | 16
| Median |  | 199 | 199
| Std Dev. |  | 2.8284 | 27
| Rel.Std.Dev. |  | 1.421%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 265 (Indicative Value 524 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 525 | 1
| 1 | 2 | 522 | 22
| Mean |  | 523.5 | 16
| Median |  | 523.5 | 523.5
| Std Dev. |  | 2.1213 | 28
| Rel.Std.Dev. |  | 0.4052%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 265 (Indicative Value 2.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.4 | 23
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0.28284 | 29
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 265 (Indicative Value 2.16 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2.22 | 24
| Mean |  | 2.16 | 16
| Median |  | 2.16 | 2.16
| Std Dev. |  | 0.084853 | 30
| Rel.Std.Dev. |  | 3.928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 265 (Indicative Value 0.075 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | < 0.1 | 4
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.075
| Std Dev. |  |  | 31
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 265 (Indicative Value 93 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 92.1 | 1
| 1 | 2 | 92.9 | 26
| Mean |  | 92.5 | 16
| Median |  | 92.5 | 92.5
| Std Dev. |  | 0.56569 | 32
| Rel.Std.Dev. |  | 0.6116%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 265 (Indicative Value 11.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.8 | 1
| 1 | 2 | 11.4 | 14
| Mean |  | 11.1 | 16
| Median |  | 11.1 | 11.1
| Std Dev. |  | 0.42426 | 33
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 265 (Indicative Value 113 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 113 | 1
| 1 | 2 | 113 | 28
| Mean |  | 113 | 16
| Median |  | 113 | 113
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 265 (Indicative Value 6.08 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.98 | 1
| 1 | 2 | 6.18 | 29
| Mean |  | 6.08 | 16
| Median |  | 6.08 | 6.08
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 2.326%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 265 (Indicative Value 24 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 24 | 1
| 1 | 2 | 24 | 16
| Mean |  | 24 | 16
| Median |  | 24 | 24
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 265 (Indicative Value 5.16 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.11 | 1
| 1 | 2 | 5.21 | 31
| Mean |  | 5.16 | 16
| Median |  | 5.16 | 5.16
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 1.37%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 265 (Indicative Value 2.73 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.77 | 1
| 1 | 2 | 2.68 | 32
| Mean |  | 2.725 | 16
| Median |  | 2.725 | 2.725
| Std Dev. |  | 0.06364 | 38
| Rel.Std.Dev. |  | 2.335%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 265 (Indicative Value 1.59 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.56 | 1
| 1 | 2 | 1.62 | 33
| Mean |  | 1.59 | 16
| Median |  | 1.59 | 1.59
| Std Dev. |  | 0.042426 | 39
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 265 (Indicative Value 19.1 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.7 | 1
| 1 | 2 | 19.4 | 34
| Mean |  | 19.05 | 16
| Median |  | 19.05 | 19.05
| Std Dev. |  | 0.49497 | 40
| Rel.Std.Dev. |  | 2.598%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 265 (Indicative Value 5.97 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.94 | 1
| 1 | 2 | 5.99 | 35
| Mean |  | 5.965 | 16
| Median |  | 5.965 | 5.965
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 0.5927%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 265 (Indicative Value 1.4 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.4 | 36
| Mean |  | 1.4 | 16
| Median |  | 1.4 | 1.4
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 265 (Indicative Value 7.26 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.24 | 1
| 1 | 2 | 7.28 | 20
| Mean |  | 7.26 | 16
| Median |  | 7.26 | 7.26
| Std Dev. |  | 0.028284 | 26
| Rel.Std.Dev. |  | 0.3896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 265 (Indicative Value 1 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 0.99 | 21
| Mean |  | 0.995 | 16
| Median |  | 0.995 | 0.995
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 0.7107%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 265 (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 22
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 265 (Indicative Value 50 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 49.6 | 1
| 1 | 2 | 50.6 | 23
| Mean |  | 50.1 | 16
| Median |  | 50.1 | 50.1
| Std Dev. |  | 0.70711 | 29
| Rel.Std.Dev. |  | 1.411%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 265 (Indicative Value 0.35 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.33 | 1
| 1 | 2 | 0.36 | 24
| Mean |  | 0.345 | 16
| Median |  | 0.345 | 0.345
| Std Dev. |  | 0.021213 | 30
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 265 (Indicative Value 0.032 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0313 | 1
| 1 | 2 | 0.0318 | 25
| Mean |  | 0.03155 | 16
| Median |  | 0.03155 | 0.03155
| Std Dev. |  | 0.00035355 | 31
| Rel.Std.Dev. |  | 1.121%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 265 (Indicative Value 4.1 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4 | 26
| Mean |  | 4.1 | 16
| Median |  | 4.1 | 4.1
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 265 (Indicative Value 30.3 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30.5 | 6
| Mean |  | 30.25 | 16
| Median |  | 30.25 | 30.25
| Std Dev. |  | 0.35355 | 33
| Rel.Std.Dev. |  | 1.169%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 265 (Indicative Value 41.5 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.8 | 1
| 1 | 2 | 42.1 | 28
| Mean |  | 41.45 | 16
| Median |  | 41.45 | 41.45
| Std Dev. |  | 0.91924 | 34
| Rel.Std.Dev. |  | 2.218%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 265 (Indicative Value 55 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 56 | 1
| 1 | 2 | 54 | 14
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 1.4142 | 35
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 265 (Indicative Value 15.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15 | 1
| 1 | 2 | 16 | 16
| Mean |  | 15.5 | 16
| Median |  | 15.5 | 15.5
| Std Dev. |  | 0.70711 | 36
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 265 (Indicative Value 11.5 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.4 | 1
| 1 | 2 | 11.6 | 31
| Mean |  | 11.5 | 16
| Median |  | 11.5 | 11.5
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 1.23%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 265 (Indicative Value 97 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 97.4 | 1
| 1 | 2 | 97.2 | 32
| Mean |  | 97.3 | 16
| Median |  | 97.3 | 97.3
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 0.1453%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 265 (Indicative Value 0.013 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | < 0.01 | 33
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.0125
| Std Dev. |  |  | 39
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 265 (Indicative Value 2.95 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 2.9 | 34
| Mean |  | 2.95 | 16
| Median |  | 2.95 | 2.95
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 265 (Indicative Value 11 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.2 | 1
| 1 | 2 | 10.7 | 35
| Mean |  | 10.95 | 16
| Median |  | 10.95 | 10.95
| Std Dev. |  | 0.35355 | 41
| Rel.Std.Dev. |  | 3.229%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 265 (Indicative Value 7.76 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.8 | 1
| 1 | 2 | 7.71 | 36
| Mean |  | 7.755 | 16
| Median |  | 7.755 | 7.755
| Std Dev. |  | 0.06364 | 42
| Rel.Std.Dev. |  | 0.8206%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 265 (Indicative Value 11 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.2 | 1
| 1 | 2 | 10.8 | 11
| Mean |  | 11 | 16
| Median |  | 11 | 11
| Std Dev. |  | 0.28284 | 26
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 265 (Indicative Value 254 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 253 | 1
| 1 | 2 | 255 | 12
| Mean |  | 254 | 16
| Median |  | 254 | 254
| Std Dev. |  | 1.4142 | 27
| Rel.Std.Dev. |  | 0.5568%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 265 (Indicative Value 2.09 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.09 | 1
| 1 | 2 | 2.08 | 22
| Mean |  | 2.085 | 16
| Median |  | 2.085 | 2.085
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.3391%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 265 (Indicative Value 0.93 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.92 | 1
| 1 | 2 | 0.94 | 23
| Mean |  | 0.93 | 16
| Median |  | 0.93 | 0.93
| Std Dev. |  | 0.014142 | 29
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 265 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 24
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 265 (Indicative Value 13.7 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.5 | 1
| 1 | 2 | 13.9 | 25
| Mean |  | 13.7 | 16
| Median |  | 13.7 | 13.7
| Std Dev. |  | 0.28284 | 31
| Rel.Std.Dev. |  | 2.065%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 265 (Indicative Value 0.526 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.523 | 1
| 1 | 2 | 0.528 | 26
| Mean |  | 0.5255 | 16
| Median |  | 0.5255 | 0.5255
| Std Dev. |  | 0.0035355 | 32
| Rel.Std.Dev. |  | 0.6728%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 265 (Indicative Value 0.5 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.4 | 27
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 265 (Indicative Value 0.38 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.39 | 28
| Mean |  | 0.38 | 16
| Median |  | 0.38 | 0.38
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 3.722%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 265 (Indicative Value 2.82 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.82 | 1
| 1 | 2 | 2.81 | 29
| Mean |  | 2.815 | 16
| Median |  | 2.815 | 2.815
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.2512%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 265 (Indicative Value 84 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 82.8 | 1
| 1 | 2 | 85.1 | 30
| Mean |  | 83.95 | 16
| Median |  | 83.95 | 83.95
| Std Dev. |  | 1.6263 | 36
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 265 (Indicative Value 12.3 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.5 | 1
| 1 | 2 | 12 | 31
| Mean |  | 12.25 | 16
| Median |  | 12.25 | 12.25
| Std Dev. |  | 0.35355 | 37
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 265 (Indicative Value 25.7 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25.5 | 1
| 1 | 2 | 25.8 | 32
| Mean |  | 25.65 | 16
| Median |  | 25.65 | 25.65
| Std Dev. |  | 0.21213 | 38
| Rel.Std.Dev. |  | 0.827%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 265 (Indicative Value 2.49 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.47 | 1
| 1 | 2 | 2.5 | 33
| Mean |  | 2.485 | 16
| Median |  | 2.485 | 2.485
| Std Dev. |  | 0.021213 | 39
| Rel.Std.Dev. |  | 0.8537%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 265 (Indicative Value 55 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 55 | 1
| 1 | 2 | 55 | 15
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 265 (Indicative Value 279 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 278 | 1
| 1 | 2 | 280 | 16
| Mean |  | 279 | 16
| Median |  | 279 | 279
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 0.5069%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## INAA
| Analytical results for Ag in OREAS 265 (Indicative Value < 5 ppm) |  |  | Indicative
Ag | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 5 | 1
| Mean |  |  | 20
| Median |  |  | 16
| Std Dev. |  |  | < 5
| Rel.Std.Dev. |  |  | 44
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 265 (Indicative Value 229 ppm) |  |  | Indicative
As | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | 229 | 1
| Mean |  | 229 | 10
| Median |  | 229 | 16
| Std Dev. |  |  | 229
| Rel.Std.Dev. |  |  | 45
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Au in OREAS 265 (Indicative Value 1.19 ppm) |  |  | Indicative
Au | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | 1.19 | 1
| Mean |  | 1.19 | 40
| Median |  | 1.19 | 16
| Std Dev. |  |  | 1.19
| Rel.Std.Dev. |  |  | 46
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 265 (Indicative Value 600 ppm) |  |  | Indicative
Ba | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | 600 | 1
| Mean |  | 600 | 22
| Median |  | 600 | 16
| Std Dev. |  |  | 600
| Rel.Std.Dev. |  |  | 47
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Br in OREAS 265 (Indicative Value < 1 ppm) |  |  | Indicative
Br | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 1 | 1
| Mean |  |  | 42
| Median |  |  | 16
| Std Dev. |  |  | < 1
| Rel.Std.Dev. |  |  | 48
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 265 (Indicative Value < 10 ppm) |  |  | Indicative
Cd | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | < 10 | 1
| Mean |  |  | 4
| Median |  |  | 16
| Std Dev. |  |  | < 10
| Rel.Std.Dev. |  |  | 49
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 265 (Indicative Value < 10 ppm) |  |  | Indicative
Co | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | < 10 | 1
| Mean |  |  | 14
| Median |  |  | 16
| Std Dev. |  |  | < 10
| Rel.Std.Dev. |  |  | 50
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 265 (Indicative Value 120 ppm) |  |  | Indicative
Cr | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | 120 | 1
| Mean |  | 120 | 28
| Median |  | 120 | 16
| Std Dev. |  |  | 120
| Rel.Std.Dev. |  |  | 51
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 265 (Indicative Value < 10000 ppm) |  |  | Indicative
Cu | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | < 10000 | 1
| Mean |  |  | 16
| Median |  |  | 16
| Std Dev. |  |  | < 10000
| Rel.Std.Dev. |  |  | 52
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 265 (Indicative Value 3.6 wt.%) |  |  | Indicative
Fe | Round | Replicate | ALS | 1
| No | No | Malaga | wt.%
|  |  | INAA | 2
2
| 1 | 1 | 3.6 | 1
| Mean |  | 3.6 | 38
| Median |  | 3.6 | 16
| Std Dev. |  |  | 3.6
| Rel.Std.Dev. |  |  | 44
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hg in OREAS 265 (Indicative Value < 10 ppm) |  |  | Indicative
Hg | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | < 10 | 1
| Mean |  |  | 4
| Median |  |  | 16
| Std Dev. |  |  | < 10
| Rel.Std.Dev. |  |  | 45
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ir in OREAS 265 (Indicative Value < 0.1 ppm) |  |  | Indicative
Ir | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | < 0.1 | 1
| Mean |  |  | 40
| Median |  |  | 16
| Std Dev. |  |  | < 0.1
| Rel.Std.Dev. |  |  | 46
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 265 (Indicative Value 55 ppm) |  |  | Indicative
La | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | 55 | 1
| Mean |  | 55 | 23
| Median |  | 55 | 16
| Std Dev. |  |  | 55
| Rel.Std.Dev. |  |  | 47
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 265 (Indicative Value 0.031 wt.%) |  |  | Indicative
Mn | Round | Replicate | ALS | 1
| No | No | Malaga | wt.%
|  |  | INAA | 3
3
| 1 | 1 | 0.0312 | 1
| Mean |  | 0.0312 | 25
| Median |  | 0.0312 | 16
| Std Dev. |  |  | 0.0312
| Rel.Std.Dev. |  |  | 48
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 265 (Indicative Value < 10 ppm) |  |  | Indicative
Mo | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | < 10 | 1
| Mean |  |  | 26
| Median |  |  | 16
| Std Dev. |  |  | < 10
| Rel.Std.Dev. |  |  | 49
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Na in OREAS 265 (Indicative Value 0.75 wt.%) |  |  | Indicative
Na | Round | Replicate | ALS | 1
| No | No | Malaga | wt.%
|  |  | INAA | 3
3
| 1 | 1 | 0.75 | 1
| Mean |  | 0.75 | 44
| Median |  | 0.75 | 16
| Std Dev. |  |  | 0.75
| Rel.Std.Dev. |  |  | 50
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 265 (Indicative Value < 100 ppm) |  |  | Indicative
Ni | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | < 100 | 1
| Mean |  |  | 14
| Median |  |  | 16
| Std Dev. |  |  | < 100
| Rel.Std.Dev. |  |  | 51
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 265 (Indicative Value 100 ppm) |  |  | Indicative
Rb | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | 100 | 1
| Mean |  | 100 | 32
| Median |  | 100 | 16
| Std Dev. |  |  | 100
| Rel.Std.Dev. |  |  | 52
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 265 (Indicative Value 3.1 ppm) |  |  | Indicative
Sb | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | 3.1 | 1
| Mean |  | 3.1 | 34
| Median |  | 3.1 | 16
| Std Dev. |  |  | 3.1
| Rel.Std.Dev. |  |  | 44
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 265 (Indicative Value 12.7 ppm) |  |  | Indicative
Sc | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | 12.7 | 1
| Mean |  | 12.7 | 35
| Median |  | 12.7 | 16
| Std Dev. |  |  | 12.7
| Rel.Std.Dev. |  |  | 45
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 265 (Indicative Value < 10 ppm) |  |  | Indicative
Se | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | < 10 | 1
| Mean |  |  | 7
| Median |  |  | 16
| Std Dev. |  |  | < 10
| Rel.Std.Dev. |  |  | 46
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 265 (Indicative Value 7.9 ppm) |  |  | Indicative
Sm | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | 7.9 | 1
| Mean |  | 7.9 | 36
| Median |  | 7.9 | 16
| Std Dev. |  |  | 7.9
| Rel.Std.Dev. |  |  | 47
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 265 (Indicative Value 2 ppm) |  |  | Indicative
Ta | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | 2 | 1
| Mean |  | 2 | 22
| Median |  | 2 | 16
| Std Dev. |  |  | 2
| Rel.Std.Dev. |  |  | 48
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 265 (Indicative Value 13.4 ppm) |  |  | Indicative
Th | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | 13.4 | 1
| Mean |  | 13.4 | 25
| Median |  | 13.4 | 16
| Std Dev. |  |  | 13.4
| Rel.Std.Dev. |  |  | 49
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 265 (Indicative Value 3 ppm) |  |  | Indicative
U | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 2
2
| 1 | 1 | 3 | 1
| Mean |  | 3 | 29
| Median |  | 3 | 16
| Std Dev. |  |  | 3
| Rel.Std.Dev. |  |  | 50
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 265 (Indicative Value 16 ppm) |  |  | Indicative
W | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 1
1
| 1 | 1 | 16 | 1
| Mean |  | 16 | 31
| Median |  | 16 | 16
| Std Dev. |  |  | 16
| Rel.Std.Dev. |  |  | 51
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 265 (Indicative Value < 200 ppm) |  |  | Indicative
Zn | Round | Replicate | ALS | 1
| No | No | Malaga | ppm
|  |  | INAA | 0
0
| 1 | 1 | < 200 | 1
| Mean |  |  | 15
| Median |  |  | 16
| Std Dev. |  |  | < 200
| Rel.Std.Dev. |  |  | 52
PDM3
| Z-Score (Absolute) |  | NA
Note on C12: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C25: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C38: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C51: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C64: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C77: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C116: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C142: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C155: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C194: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C207: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C246: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C259: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C285: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C298: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C337: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C350: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.