## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 702
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 9.6988 | 9.5828 | 9.8148 | 9.6288 | 9.7688
As, Arsenic (ppm) | < 100 | IND | IND | IND | IND
BaO, Barium oxide (ppm) | 738.1906 | 668.3531 | 808.028 | 673.0162 | 803.3649
Bi, Bismuth (ppm) | 286.6877 | 237.7718 | 335.6036 | 270.3597 | 303.0157
CaO, Calcium oxide (wt.%) | 22.5383 | 22.3484 | 22.7282 | 22.419 | 22.6575
Co, Cobalt (ppm) | < 100 | IND | IND | IND | IND
Cr2O3, Chromium(III) oxide (ppm) | < 150 | IND | IND | IND | IND
Cu, Copper (ppm) | 352.1244 | 328.9627 | 375.2862 | 334.122 | 370.1268
Fe2O3, Iron(III) oxide (wt.%) | 3.3575 | 3.3075 | 3.4076 | 3.3254 | 3.3897
HfO2, Hafnium dioxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 2.2682 | 2.2365 | 2.3 | 2.2487 | 2.2878
MgO, Magnesium oxide (wt.%) | 1.0285 | 1.0013 | 1.0556 | 1.0114 | 1.0456
MnO, Manganese oxide (wt.%) | 0.038364 | 0.034737 | 0.041991 | 0.036147 | 0.040581
Mo, Molybdenum (ppm) | 625.2316 | 584.852 | 665.6112 | 605.8345 | 644.6287
Na2O, Sodium oxide (wt.%) | 1.7681 | 1.7156 | 1.8205 | 1.7515 | 1.7846
Ni, Nickel (ppm) | < 50 | IND | IND | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.1289 | 0.12221 | 0.1356 | 0.12601 | 0.13179
Pb, Lead (ppm) | < 50 | IND | IND | IND | IND
SiO2, Silicon dioxide (wt.%) | 41.4064 | 41.0759 | 41.7369 | 41.1932 | 41.6196
SO3, Sulphur trioxide (wt.%) | 0.90646 | 0.87915 | 0.93377 | 0.88759 | 0.92533
Sr, Strontium (ppm) | 195.9718 | 186.3573 | 205.5864 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.38611 | 0.37378 | 0.39843 | 0.37653 | 0.39568
WO3, Tungsten trioxide (wt.%) | 0.27664 | 0.27064 | 0.28265 | 0.27225 | 0.28104
Zn, Zinc (ppm) | 60.015 | 46.5595 | 73.4705 | IND | IND
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 16.0603 | 15.9714 | 16.1491 | 16.013 | 16.1075
Ion Selective Electrode
F, Fluorine (wt.%) | 0.29949 | 0.27098 | 0.32799 | 0.28489 | 0.31408
Infrared Combustion
S, Sulphur (wt.%) | 0.34852 | 0.33423 | 0.36281 | 0.33802 | 0.35902
Borate / Peroxide Fusion ICP
Al2O3, Aluminium(III) oxide (wt.%) | 9.5629 | 9.3198 | 9.806 | 9.445 | 9.6808
Ba, Barium (ppm) | 649.9487 | 631.1389 | 668.7586 | 636.6922 | 663.2053
Be, Beryllium (ppm) | 5.6298 | 4.7307 | 6.529 | 5.2475 | 6.0122
Bi, Bismuth (ppm) | 275.1158 | 263.1453 | 287.0864 | 269.274 | 280.9577
CaO, Calcium oxide (wt.%) | 22.2132 | 21.6786 | 22.7479 | 21.8963 | 22.5301
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (ppm) | 55.7091 | 52.6884 | 58.7299 | 54.4599 | 56.9584
Co, Cobalt (ppm) | 6.7415 | 5.8994 | 7.5835 | 6.3212 | 7.1617
Cr, Chromium (ppm) | 29.4662 | 24.5053 | 34.4271 | 27.3401 | 31.5923
Cs, Caesium (ppm) | 8.5304 | 8.0536 | 9.0072 | 8.3459 | 8.7149
Cu, Copper (ppm) | 362.1484 | 341.0609 | 383.2359 | 353.391 | 370.9058
Dy, Dysprosium (ppm) | 4.4564 | 3.995 | 4.9179 | 4.2639 | 4.6489
Er, Erbium (ppm) | 2.2622 | 2.0462 | 2.4782 | 2.1082 | 2.4162
Eu, Europium (ppm) | 1.0105 | 0.89051 | 1.1306 | 0.9416 | 1.0795
Fe2O3, Iron(III) oxide (wt.%) | 3.3839 | 3.2889 | 3.4789 | 3.3172 | 3.4507
Ga, Gallium (ppm) | 14.2344 | 13.4788 | 14.9901 | 13.9285 | 14.5404
Gd, Gadolinium (ppm) | 5.0726 | 4.7352 | 5.41 | 4.8871 | 5.2581
Hf, Hafnium (ppm) | 4.5206 | 4.0751 | 4.9661 | 4.3387 | 4.7024
Ho, Holmium (ppm) | 0.84143 | 0.76087 | 0.92198 | 0.80231 | 0.88054
K2O, Potassium oxide (wt.%) | 2.3156 | 2.2436 | 2.3876 | 2.2564 | 2.3748
La, Lanthanum (ppm) | 27.8141 | 26.2917 | 29.3364 | 26.985 | 28.6432
Li, Lithium (ppm) | 37.9487 | 34.6881 | 41.2092 | 36.8535 | 39.0439
Lu, Lutetium (ppm) | 0.28664 | 0.25173 | 0.32156 | 0.25743 | 0.31586
MgO, Magnesium oxide (wt.%) | 1.0051 | 0.96653 | 1.0437 | 0.98541 | 1.0248
MnO, Manganese oxide (wt.%) | 0.040068 | 0.039399 | 0.040738 | 0.039051 | 0.041086
Mo, Molybdenum (ppm) | 587.5244 | 558.3044 | 616.7443 | 577.2809 | 597.7678
Na2O, Sodium oxide (wt.%) | 1.7683 | 1.6989 | 1.8376 | 1.713 | 1.8235
Nb, Niobium (ppm) | 8.7778 | 8.023 | 9.5327 | 8.3565 | 9.1992
Nd, Neodymium (ppm) | 25.6724 | 23.8395 | 27.5054 | 24.8976 | 26.4473
P2O5, Phosphorus(V) oxide (wt.%) | 0.13294 | 0.12508 | 0.1408 | 0.12797 | 0.13791
Pb, Lead (ppm) | 17.5514 | 15.0259 | 20.0769 | 16.4111 | 18.6917
Pr, Praseodymium (ppm) | 6.714 | 6.2353 | 7.1927 | 6.4973 | 6.9307
Rb, Rubidium (ppm) | 119.5442 | 115.2245 | 123.8639 | 117.6532 | 121.4352
S, Sulphur (wt.%) | 0.34622 | 0.31516 | 0.37727 | 0.32967 | 0.36277
Sc, Scandium (ppm) | 5.8088 | 5.3015 | 6.316 | 5.4229 | 6.1947
SiO2, Silicon dioxide (wt.%) | 42.0143 | 40.874 | 43.1547 | 41.2193 | 42.8094
Sm, Samarium (ppm) | 5.5985 | 5.1234 | 6.0735 | 5.1004 | 6.0965
Sn, Tin (ppm) | 4.3564 | 3.8402 | 4.8726 | 3.9893 | 4.7235
Sr, Strontium (ppm) | 194.7788 | 187.9938 | 201.5638 | 191.3336 | 198.224
Ta, Tantalum (ppm) | 0.78373 | 0.65192 | 0.91553 | IND | IND
Tb, Terbium (ppm) | 0.78335 | 0.71662 | 0.85008 | 0.72185 | 0.84485
Th, Thorium (ppm) | 9.8764 | 9.1421 | 10.6107 | 9.3066 | 10.4463
Ti, Titanium (wt.%) | 0.23741 | 0.22993 | 0.24489 | 0.23264 | 0.24218
Tl, Thallium (ppm) | 0.70067 | 0.59726 | 0.80407 | 0.64852 | 0.75281
Tm, Thulium (ppm) | 0.31169 | 0.28634 | 0.33704 | 0.27 | 0.35338
U, Uranium (ppm) | 4.1971 | 3.6634 | 4.7308 | 3.8886 | 4.5055
V, Vanadium (ppm) | 41.6191 | 39.2578 | 43.9805 | 40.275 | 42.9633
WO3, Tungsten trioxide (wt.%) | 0.28601 | 0.2754 | 0.29663 | 0.27995 | 0.29208
Y, Yttrium (ppm) | 22.8092 | 21.3117 | 24.3068 | 22.0067 | 23.6118
Yb, Ytterbium (ppm) | 1.9907 | 1.7357 | 2.2456 | 1.8835 | 2.0978
Zn, Zinc (ppm) | 59.2041 | 53.4145 | 64.9938 | 56.312 | 62.0962
Zr, Zirconium (ppm) | 167.3616 | 157.7708 | 176.9524 | 159.0352 | 175.688
4-Acid Digestion
Ag, Silver (ppm) | 0.17836 | 0.15571 | 0.201 | IND | IND
Al, Aluminium (wt.%) | 5.1021 | 4.9883 | 5.2159 | 5.0154 | 5.1887
As, Arsenic (ppm) | 3.5448 | 2.7028 | 4.3868 | 3.1148 | 3.9748
Ba, Barium (ppm) | 656.8759 | 634.4214 | 679.3304 | 644.379 | 669.3728
Be, Beryllium (ppm) | 5.2428 | 5.0036 | 5.4821 | 5.1073 | 5.3784
Bi, Bismuth (ppm) | 277.4113 | 263.6944 | 291.1282 | 271.5983 | 283.2244
Ca, Calcium (wt.%) | 15.6337 | 15.0965 | 16.171 | 15.4162 | 15.8513
Ce, Cerium (ppm) | 59.2081 | 55.7553 | 62.6609 | 57.4756 | 60.9406
Co, Cobalt (ppm) | 6.4747 | 6.0242 | 6.9253 | 6.2845 | 6.6649
Cr, Chromium (ppm) | 21.3867 | 19.1922 | 23.5813 | 20.1024 | 22.6711
Cs, Caesium (ppm) | 8.7725 | 8.2686 | 9.2764 | 8.4992 | 9.0458
Cu, Copper (ppm) | 357.6129 | 345.5317 | 369.6942 | 351.3174 | 363.9085
Dy, Dysprosium (ppm) | 3.1816 | 2.8909 | 3.4724 | 3.057 | 3.3062
Er, Erbium (ppm) | 1.2626 | 1.1668 | 1.3585 | 1.2066 | 1.3187
Eu, Europium (ppm) | 1.055 | 0.9392 | 1.1709 | 0.98569 | 1.1244
Fe, Iron (wt.%) | 2.3355 | 2.2728 | 2.3982 | 2.2878 | 2.3832
Ga, Gallium (ppm) | 13.9892 | 13.181 | 14.7974 | 13.6171 | 14.3613
Gd, Gadolinium (ppm) | 4.5163 | 4.0401 | 4.9926 | 4.3709 | 4.6617
Hf, Hafnium (ppm) | 1.2612 | 1.0939 | 1.4286 | 1.1466 | 1.3759
Ho, Holmium (ppm) | 0.50436 | 0.44196 | 0.56676 | 0.475 | 0.53372
In, Indium (ppm) | 0.097278 | 0.086781 | 0.10778 | 0.092826 | 0.10173
K, Potassium (wt.%) | 1.9104 | 1.8592 | 1.9616 | 1.8713 | 1.9495
La, Lanthanum (ppm) | 28.7942 | 27.0123 | 30.5762 | 28.055 | 29.5335
Li, Lithium (ppm) | 36.9187 | 35.2073 | 38.6301 | 36.1695 | 37.6679
Lu, Lutetium (ppm) | 0.14038 | 0.11834 | 0.16241 | IND | IND
Mg, Magnesium (wt.%) | 0.58653 | 0.56522 | 0.60783 | 0.57445 | 0.5986
Mn, Manganese (wt.%) | 0.030685 | 0.029652 | 0.031718 | 0.030143 | 0.031227
Mo, Molybdenum (ppm) | 576.9731 | 556.6676 | 597.2786 | 568.2836 | 585.6626
Na, Sodium (wt.%) | 1.3211 | 1.2818 | 1.3604 | 1.296 | 1.3462
Nb, Niobium (ppm) | 8.6839 | 8.2281 | 9.1398 | 8.3277 | 9.0402
Nd, Neodymium (ppm) | 24.8154 | 21.804 | 27.8267 | 24.0415 | 25.5892
Ni, Nickel (ppm) | 11.2795 | 10.5695 | 11.9895 | 10.8499 | 11.7092
P, Phosphorus (wt.%) | 0.05756 | 0.055252 | 0.059868 | 0.056396 | 0.058724
Pb, Lead (ppm) | 17.2722 | 16.0503 | 18.4941 | 16.1201 | 18.4243
Pr, Praseodymium (ppm) | 6.892 | 6.3355 | 7.4485 | 6.6637 | 7.1204
Rb, Rubidium (ppm) | 123.4713 | 117.5031 | 129.4395 | 121.079 | 125.8635
Re, Rhenium (ppm) | 0.035125 | 0.028009 | 0.042241 | 0.032006 | 0.038244
S, Sulphur (wt.%) | 0.34935 | 0.33382 | 0.36488 | 0.33975 | 0.35895
Sb, Antimony (ppm) | 0.36513 | 0.30862 | 0.42163 | 0.32797 | 0.40228
Sc, Scandium (ppm) | 5.6188 | 5.2793 | 5.9582 | 5.4687 | 5.7689
Sm, Samarium (ppm) | 5.1734 | 4.4731 | 5.8737 | 4.9641 | 5.3827
Sn, Tin (ppm) | 4.1031 | 3.9136 | 4.2926 | 3.9436 | 4.2626
Sr, Strontium (ppm) | 197.5478 | 191.4672 | 203.6284 | 193.2952 | 201.8004
Ta, Tantalum (ppm) | 0.72409 | 0.65515 | 0.79302 | 0.67004 | 0.77813
Tb, Terbium (ppm) | 0.65995 | 0.58104 | 0.73885 | 0.62002 | 0.69988
Te, Tellurium (ppm) | 1.1282 | 0.96344 | 1.2929 | 1.0106 | 1.2457
Th, Thorium (ppm) | 10.0105 | 9.3379 | 10.683 | 9.5344 | 10.4866
Ti, Titanium (wt.%) | 0.23015 | 0.2221 | 0.23821 | 0.22652 | 0.23379
Tl, Thallium (ppm) | 0.70667 | 0.65878 | 0.75455 | 0.68385 | 0.72949
Tm, Thulium (ppm) | 0.16657 | 0.13758 | 0.19557 | IND | IND
U, Uranium (ppm) | 3.931 | 3.4183 | 4.4437 | 3.6209 | 4.2411
V, Vanadium (ppm) | 38.4563 | 36.8886 | 40.024 | 37.4392 | 39.4734
W, Tungsten (wt.%) | 0.21492 | 0.20583 | 0.22402 | 0.20836 | 0.22148
Y, Yttrium (ppm) | 13.0604 | 12.1717 | 13.9491 | 12.7415 | 13.3793
Yb, Ytterbium (ppm) | 1.028 | 0.92036 | 1.1356 | 0.95761 | 1.0983
Zn, Zinc (ppm) | 57.156 | 53.9327 | 60.3794 | 55.5664 | 58.7456
Zr, Zirconium (ppm) | 40.0151 | 36.5057 | 43.5245 | 38.3395 | 41.6907
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 702
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Cd | ppm | 65.657 | Nb | ppm | < 50 | Th | ppm | 87.8812
Ce | ppm | 65.5776 | Rb | ppm | 116.5833 | V2O5 | ppm | < 180
Cl | ppm | 194.2708 | Sb | ppm | 137.5 | Y2O3 | ppm | 52.9167
F | wt.% | 0.23601 | Sc | ppm | 13.4 | Zr | ppm | 155.1747
La2O3 | ppm | < 100 | SnO2 | ppm | < 63
Thermogravimetry
H2O- | wt.% | 0.31167
Infrared Combustion
C | wt.% | 4.3889
Borate / Peroxide Fusion ICP
Ag | ppm | < 1 | In | ppm | 0.13921 | Se | ppm | < 10
As | ppm | 5.8487 | Ni | ppm | 18.7573 | Te | ppm | 1.1096
B | ppm | 29.4675 | Re | ppm | 0.036556
Ge | ppm | 0.96871 | Sb | ppm | 0.40867
4-Acid Digestion
B | ppm | 22.0533 | Ge | ppm | 0.098024 | Se | ppm | 1.0333
Cd | ppm | 0.20553 | Hg | ppm | 1.015
Aqua Regia Digestion
Ag | ppm | 0.11667 | Cu | ppm | 327.95 | Sb | ppm | 0.2
As | ppm | 3.1 | Hg | ppm | < 0.01 | Se | ppm | < 0.5
Au | ppm | 0.056583 | Mo | ppm | 595.4833 | Tl | ppm | 0.5
Bi | ppm | 267.7667 | Ni | ppm | 10.75 | Zn | ppm | 51.5
Cd | ppm | < 0.1 | Pb | ppm | 4.4333
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 702
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 9.6988 | 0.1245 | 9.4498 | 9.9478 | 9.3253 | 10.0723 | 1.284% | 2.567% | 3.851% | 9.2139 | 10.1837
As, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
BaO, ppm | 738.1906 | 49.2197 | 639.7512 | 836.6299 | 590.5315 | 885.8496 | 6.668% | 13.34% | 20% | 701.281 | 775.1001
Bi, ppm | 286.6877 | 53.2247 | 180.2383 | 393.1371 | 127.0136 | 446.3618 | 18.57% | 37.13% | 55.7% | 272.3533 | 301.0221
CaO, wt.% | 22.5383 | 0.19466 | 22.149 | 22.9276 | 21.9543 | 23.1222 | 0.8637% | 1.727% | 2.591% | 21.4114 | 23.6652
Co, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cr2O3, ppm | < 150 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cu, ppm | 352.1244 | 17.7782 | 316.568 | 387.6808 | 298.7898 | 405.459 | 5.049% | 10.1% | 15.15% | 334.5182 | 369.7306
Fe2O3, wt.% | 3.3575 | 0.081366 | 3.1948 | 3.5203 | 3.1134 | 3.6016 | 2.423% | 4.847% | 7.27% | 3.1897 | 3.5254
HfO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 2.2682 | 0.032385 | 2.2035 | 2.333 | 2.1711 | 2.3654 | 1.428% | 2.855% | 4.283% | 2.1548 | 2.3817
MgO, wt.% | 1.0285 | 0.028871 | 0.97072 | 1.0862 | 0.94185 | 1.1151 | 2.807% | 5.614% | 8.422% | 0.97704 | 1.0799
MnO, wt.% | 0.038364 | 0.0057288 | 0.026906 | 0.049822 | 0.021178 | 0.055551 | 14.93% | 29.87% | 44.8% | 0.036446 | 0.040282
Mo, ppm | 625.2316 | 35.1065 | 555.0185 | 695.4446 | 519.912 | 730.5512 | 5.615% | 11.23% | 16.84% | 593.97 | 656.4932
Na2O, wt.% | 1.7681 | 0.072651 | 1.6228 | 1.9134 | 1.5501 | 1.986 | 4.109% | 8.218% | 12.33% | 1.6797 | 1.8565
Ni, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
P2O5, wt.% | 0.1289 | 0.0057613 | 0.11738 | 0.14043 | 0.11162 | 0.14619 | 4.469% | 8.939% | 13.41% | 0.12246 | 0.13535
Pb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SiO2, wt.% | 41.4064 | 0.34455 | 40.7173 | 42.0955 | 40.3727 | 42.44 | 0.8321% | 1.664% | 2.496% | 39.3361 | 43.4767
SO3, wt.% | 0.90646 | 0.019232 | 0.868 | 0.94492 | 0.84876 | 0.96416 | 2.122% | 4.243% | 6.365% | 0.86114 | 0.95178
Sr, ppm | 195.9718 | 7.4431 | 181.0856 | 210.8581 | 173.6425 | 218.3012 | 3.798% | 7.596% | 11.39% | 186.1732 | 205.7704
TiO2, wt.% | 0.38611 | 0.013429 | 0.35925 | 0.41296 | 0.34582 | 0.42639 | 3.478% | 6.956% | 10.43% | 0.3668 | 0.40541
WO3, wt.% | 0.27664 | 0.0071897 | 0.26226 | 0.29102 | 0.25507 | 0.29821 | 2.599% | 5.198% | 7.797% | 0.26281 | 0.29048
Zn, ppm | 60.015 | 6.444 | 47.1269 | 72.9031 | 40.6829 | 79.3471 | 10.74% | 21.47% | 32.21% | 57.0142 | 63.0157
Thermogravimetry
LOI1000, wt.% | 16.0603 | 0.073432 | 15.9134 | 16.2071 | 15.84 | 16.2805 | 0.4572% | 0.9145% | 1.372% | 15.2572 | 16.8633
Ion Selective Electrode
F, wt.% | 0.29949 | 0.038344 | 0.2228 | 0.37617 | 0.18445 | 0.41452 | 12.8% | 25.61% | 38.41% | 0.28451 | 0.31446
Infrared Combustion
S, wt.% | 0.34852 | 0.01374 | 0.32104 | 0.376 | 0.3073 | 0.38974 | 3.942% | 7.885% | 11.83% | 0.33109 | 0.36594
Borate / Peroxide Fusion ICP
Al2O3, wt.% | 9.5629 | 0.1523 | 9.2583 | 9.8675 | 9.106 | 10.0198 | 1.593% | 3.185% | 4.778% | 9.0848 | 10.041
Ba, ppm | 649.9487 | 17.8809 | 614.187 | 685.7105 | 596.3061 | 703.5914 | 2.751% | 5.502% | 8.253% | 617.4513 | 682.4462
Be, ppm | 5.6298 | 0.779 | 4.0719 | 7.1878 | 3.2929 | 7.9668 | 13.84% | 27.67% | 41.51% | 5.3484 | 5.9113
Bi, ppm | 275.1158 | 11.6369 | 251.8419 | 298.3897 | 240.205 | 310.0266 | 4.23% | 8.46% | 12.69% | 261.36 | 288.8716
CaO, wt.% | 22.2132 | 0.69464 | 20.824 | 23.6025 | 20.1293 | 24.2972 | 3.127% | 6.254% | 9.381% | 21.1026 | 23.3239
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 55.7091 | 2.5992 | 50.5107 | 60.9075 | 47.9115 | 63.5067 | 4.666% | 9.331% | 14% | 52.9237 | 58.4946
Co, ppm | 6.7415 | 0.71422 | 5.313 | 8.1699 | 4.5988 | 8.8841 | 10.59% | 21.19% | 31.78% | 6.4044 | 7.0785
Cr, ppm | 29.4662 | 3.9197 | 21.6269 | 37.3055 | 17.7072 | 41.2252 | 13.3% | 26.6% | 39.91% | 27.9929 | 30.9395
Cs, ppm | 8.5304 | 0.51357 | 7.5033 | 9.5575 | 6.9897 | 10.0711 | 6.02% | 12.04% | 18.06% | 8.1039 | 8.9569
Cu, ppm | 362.1484 | 22.7185 | 316.7113 | 407.5855 | 293.9928 | 430.304 | 6.273% | 12.55% | 18.82% | 344.041 | 380.2558
Dy, ppm | 4.4564 | 0.29424 | 3.868 | 5.0449 | 3.5737 | 5.3391 | 6.603% | 13.21% | 19.81% | 4.2336 | 4.6792
Er, ppm | 2.2622 | 0.16134 | 1.9395 | 2.5849 | 1.7782 | 2.7462 | 7.132% | 14.26% | 21.4% | 2.1491 | 2.3753
Eu, ppm | 1.0105 | 0.073087 | 0.86437 | 1.1567 | 0.79128 | 1.2298 | 7.232% | 14.46% | 21.7% | 0.96002 | 1.0611
Fe2O3, wt.% | 3.3839 | 0.087935 | 3.208 | 3.5598 | 3.1201 | 3.6477 | 2.599% | 5.197% | 7.796% | 3.2147 | 3.5531
Ga, ppm | 14.2344 | 0.59577 | 13.0429 | 15.426 | 12.4471 | 16.0217 | 4.185% | 8.371% | 12.56% | 13.5227 | 14.9462
Gd, ppm | 5.0726 | 0.30911 | 4.4544 | 5.6908 | 4.1453 | 6 | 6.094% | 12.19% | 18.28% | 4.819 | 5.3263
Hf, ppm | 4.5206 | 0.30784 | 3.9049 | 5.1362 | 3.597 | 5.4441 | 6.81% | 13.62% | 20.43% | 4.2945 | 4.7466
Ho, ppm | 0.84143 | 0.073687 | 0.69405 | 0.9888 | 0.62037 | 1.0625 | 8.757% | 17.51% | 26.27% | 0.79935 | 0.8835
K2O, wt.% | 2.3156 | 0.079574 | 2.1565 | 2.4748 | 2.0769 | 2.5543 | 3.436% | 6.873% | 10.31% | 2.1998 | 2.4314
La, ppm | 27.8141 | 1.1314 | 25.5513 | 30.0768 | 24.4199 | 31.2082 | 4.068% | 8.135% | 12.2% | 26.4234 | 29.2048
Li, ppm | 37.9487 | 3.3834 | 31.1818 | 44.7155 | 27.7984 | 48.099 | 8.916% | 17.83% | 26.75% | 36.0513 | 39.8461
Lu, ppm | 0.28664 | 0.022552 | 0.24154 | 0.33175 | 0.21899 | 0.3543 | 7.867% | 15.73% | 23.6% | 0.27231 | 0.30097
MgO, wt.% | 1.0051 | 0.044977 | 0.91517 | 1.0951 | 0.87019 | 1.1401 | 4.475% | 8.949% | 13.42% | 0.95487 | 1.0554
MnO, wt.% | 0.040068 | 0.00083407 | 0.0384 | 0.041737 | 0.037566 | 0.042571 | 2.082% | 4.163% | 6.245% | 0.038065 | 0.042072
Mo, ppm | 587.5244 | 40.701 | 506.1223 | 668.9264 | 465.4213 | 709.6274 | 6.928% | 13.86% | 20.78% | 558.1481 | 616.9006
Na2O, wt.% | 1.7683 | 0.042799 | 1.6827 | 1.8539 | 1.6399 | 1.8967 | 2.42% | 4.841% | 7.261% | 1.6798 | 1.8567
Nb, ppm | 8.7778 | 0.46153 | 7.8548 | 9.7009 | 7.3933 | 10.1624 | 5.258% | 10.52% | 15.77% | 8.339 | 9.2167
Nd, ppm | 25.6724 | 1.3456 | 22.9812 | 28.3636 | 21.6357 | 29.7092 | 5.241% | 10.48% | 15.72% | 24.3888 | 26.956
P2O5, wt.% | 0.13294 | 0.0053847 | 0.12217 | 0.14371 | 0.11679 | 0.14909 | 4.05% | 8.101% | 12.15% | 0.12629 | 0.13959
Pb, ppm | 17.5514 | 2.9052 | 11.741 | 23.3618 | 8.8357 | 26.267 | 16.55% | 33.11% | 49.66% | 16.6738 | 18.429
Pr, ppm | 6.714 | 0.36276 | 5.9885 | 7.4395 | 5.6257 | 7.8023 | 5.403% | 10.81% | 16.21% | 6.3783 | 7.0497
Rb, ppm | 119.5442 | 4.1687 | 111.2067 | 127.8816 | 107.038 | 132.0503 | 3.487% | 6.974% | 10.46% | 113.567 | 125.5214
S, wt.% | 0.34622 | 0.028538 | 0.28914 | 0.40329 | 0.2606 | 0.43183 | 8.243% | 16.49% | 24.73% | 0.32891 | 0.36353
Sc, ppm | 5.8088 | 0.43939 | 4.93 | 6.6876 | 4.4906 | 7.127 | 7.564% | 15.13% | 22.69% | 5.5184 | 6.0992
SiO2, wt.% | 42.0143 | 0.8759 | 40.2625 | 43.7662 | 39.3866 | 44.6421 | 2.085% | 4.17% | 6.254% | 39.9136 | 44.1151
Sm, ppm | 5.5985 | 0.37991 | 4.8387 | 6.3583 | 4.4587 | 6.7382 | 6.786% | 13.57% | 20.36% | 5.3185 | 5.8784
Sn, ppm | 4.3564 | 0.29847 | 3.7595 | 4.9534 | 3.461 | 5.2518 | 6.851% | 13.7% | 20.55% | 4.1386 | 4.5742
Sr, ppm | 194.7788 | 9.6268 | 175.5252 | 214.0324 | 165.8984 | 223.6592 | 4.942% | 9.885% | 14.83% | 185.0399 | 204.5177
Ta, ppm | 0.78373 | 0.16039 | 0.46296 | 1.1045 | 0.30257 | 1.2649 | 20.46% | 40.93% | 61.39% | 0.74454 | 0.82291
Tb, ppm | 0.78335 | 0.050137 | 0.68307 | 0.88362 | 0.63294 | 0.93376 | 6.4% | 12.8% | 19.2% | 0.74418 | 0.82252
Th, ppm | 9.8764 | 0.50197 | 8.8725 | 10.8804 | 8.3705 | 11.3823 | 5.083% | 10.17% | 15.25% | 9.3826 | 10.3702
Ti, wt.% | 0.23741 | 0.0089014 | 0.21961 | 0.25521 | 0.21071 | 0.26411 | 3.749% | 7.499% | 11.25% | 0.22554 | 0.24928
Tl, ppm | 0.70067 | 0.070648 | 0.55937 | 0.84196 | 0.48872 | 0.91261 | 10.08% | 20.17% | 30.25% | 0.66563 | 0.7357
Tm, ppm | 0.31169 | 0.01911 | 0.27347 | 0.34991 | 0.25436 | 0.36902 | 6.131% | 12.26% | 18.39% | 0.29611 | 0.32727
U, ppm | 4.1971 | 0.43606 | 3.3249 | 5.0692 | 2.8889 | 5.5052 | 10.39% | 20.78% | 31.17% | 3.9872 | 4.4069
V, ppm | 41.6191 | 2.8338 | 35.9515 | 47.2867 | 33.1177 | 50.1205 | 6.809% | 13.62% | 20.43% | 39.5382 | 43.7001
WO3, wt.% | 0.28601 | 0.011044 | 0.26393 | 0.3081 | 0.25288 | 0.31914 | 3.861% | 7.722% | 11.58% | 0.27171 | 0.30031
Y, ppm | 22.8092 | 1.6731 | 19.4631 | 26.1554 | 17.7901 | 27.8284 | 7.335% | 14.67% | 22.01% | 21.6688 | 23.9497
Yb, ppm | 1.9907 | 0.16086 | 1.6689 | 2.3124 | 1.5081 | 2.4732 | 8.081% | 16.16% | 24.24% | 1.8911 | 2.0902
Zn, ppm | 59.2041 | 6.4002 | 46.4038 | 72.0045 | 40.0036 | 78.4047 | 10.81% | 21.62% | 32.43% | 56.2439 | 62.1643
Zr, ppm | 167.3616 | 7.6516 | 152.0584 | 182.6647 | 144.4068 | 190.3163 | 4.572% | 9.144% | 13.72% | 158.9935 | 175.7296
4-Acid Digestion
Ag, ppm | 0.17836 | 0.016396 | 0.14556 | 0.21115 | 0.12917 | 0.22755 | 9.193% | 18.39% | 27.58% | 0.16944 | 0.18728
Al, wt.% | 5.1021 | 0.11058 | 4.8809 | 5.3232 | 4.7703 | 5.4338 | 2.167% | 4.335% | 6.502% | 4.847 | 5.3572
As, ppm | 3.5448 | 0.74665 | 2.0515 | 5.0381 | 1.3048 | 5.7848 | 21.06% | 42.13% | 63.19% | 3.3676 | 3.722
Ba, ppm | 656.8759 | 14.7541 | 627.3677 | 686.384 | 612.6137 | 701.1381 | 2.246% | 4.492% | 6.738% | 624.0321 | 689.7197
Be, ppm | 5.2428 | 0.37464 | 4.4935 | 5.9921 | 4.1189 | 6.3667 | 7.146% | 14.29% | 21.44% | 4.9807 | 5.505
Bi, ppm | 277.4113 | 20.7588 | 235.8938 | 318.9289 | 215.1351 | 339.6876 | 7.483% | 14.97% | 22.45% | 263.5408 | 291.2819
Ca, wt.% | 15.6337 | 0.60802 | 14.4177 | 16.8498 | 13.8097 | 17.4578 | 3.889% | 7.778% | 11.67% | 14.8521 | 16.4154
Ce, ppm | 59.2081 | 3.1089 | 52.9903 | 65.4259 | 49.8815 | 68.5347 | 5.251% | 10.5% | 15.75% | 56.2477 | 62.1685
Co, ppm | 6.4747 | 0.45019 | 5.5744 | 7.3751 | 5.1242 | 7.8253 | 6.953% | 13.91% | 20.86% | 6.151 | 6.7985
Cr, ppm | 21.3867 | 3.0163 | 15.354 | 27.4194 | 12.3377 | 30.4357 | 14.1% | 28.21% | 42.31% | 20.3174 | 22.4561
Cs, ppm | 8.7725 | 0.51301 | 7.7465 | 9.7985 | 7.2335 | 10.3115 | 5.848% | 11.7% | 17.54% | 8.3339 | 9.2111
Cu, ppm | 357.6129 | 16.0456 | 325.5218 | 389.7041 | 309.4763 | 405.7496 | 4.487% | 8.974% | 13.46% | 339.7323 | 375.4936
Dy, ppm | 3.1816 | 0.20058 | 2.7804 | 3.5828 | 2.5799 | 3.7833 | 6.304% | 12.61% | 18.91% | 3.0225 | 3.3407
Er, ppm | 1.2626 | 0.053899 | 1.1548 | 1.3704 | 1.1009 | 1.4243 | 4.269% | 8.538% | 12.81% | 1.1995 | 1.3257
Eu, ppm | 1.055 | 0.067774 | 0.91949 | 1.1906 | 0.85171 | 1.2584 | 6.424% | 12.85% | 19.27% | 1.0023 | 1.1078
Fe, wt.% | 2.3355 | 0.061259 | 2.213 | 2.458 | 2.1517 | 2.5193 | 2.623% | 5.246% | 7.869% | 2.2187 | 2.4523
Ga, ppm | 13.9892 | 0.77641 | 12.4364 | 15.542 | 11.66 | 16.3184 | 5.55% | 11.1% | 16.65% | 13.2897 | 14.6887
Gd, ppm | 4.5163 | 0.37413 | 3.7681 | 5.2646 | 3.3939 | 5.6387 | 8.284% | 16.57% | 24.85% | 4.2905 | 4.7421
Hf, ppm | 1.2612 | 0.15869 | 0.94387 | 1.5786 | 0.78519 | 1.7373 | 12.58% | 25.16% | 37.75% | 1.1982 | 1.3243
Ho, ppm | 0.50436 | 0.054613 | 0.39514 | 0.61359 | 0.34052 | 0.6682 | 10.83% | 21.66% | 32.48% | 0.47914 | 0.52958
In, ppm | 0.097278 | 0.0079694 | 0.081339 | 0.11322 | 0.073369 | 0.12119 | 8.192% | 16.38% | 24.58% | 0.092414 | 0.10214
K, wt.% | 1.9104 | 0.049425 | 1.8115 | 2.0092 | 1.7621 | 2.0587 | 2.587% | 5.174% | 7.761% | 1.8149 | 2.0059
La, ppm | 28.7942 | 2.0544 | 24.6855 | 32.903 | 22.6312 | 34.9573 | 7.135% | 14.27% | 21.4% | 27.3545 | 30.234
Li, ppm | 36.9187 | 1.636 | 33.6467 | 40.1907 | 32.0107 | 41.8267 | 4.431% | 8.863% | 13.29% | 35.0728 | 38.7646
Lu, ppm | 0.14038 | 0.012747 | 0.11488 | 0.16587 | 0.10214 | 0.17862 | 9.08% | 18.16% | 27.24% | 0.13336 | 0.14739
Mg, wt.% | 0.58653 | 0.02505 | 0.53643 | 0.63663 | 0.51138 | 0.66168 | 4.271% | 8.542% | 12.81% | 0.5572 | 0.61585
Mn, wt.% | 0.030685 | 0.00087821 | 0.028929 | 0.032442 | 0.028051 | 0.03332 | 2.862% | 5.724% | 8.586% | 0.029151 | 0.03222
Mo, ppm | 576.9731 | 32.1532 | 512.6666 | 641.2796 | 480.5133 | 673.4328 | 5.573% | 11.15% | 16.72% | 548.1244 | 605.8217
Na, wt.% | 1.3211 | 0.023294 | 1.2745 | 1.3677 | 1.2512 | 1.391 | 1.763% | 3.526% | 5.29% | 1.255 | 1.3871
Nb, ppm | 8.6839 | 0.38302 | 7.9179 | 9.45 | 7.5349 | 9.833 | 4.411% | 8.821% | 13.23% | 8.2497 | 9.1181
Nd, ppm | 24.8154 | 3.235 | 18.3454 | 31.2853 | 15.1104 | 34.5203 | 13.04% | 26.07% | 39.11% | 23.5746 | 26.0561
Ni, ppm | 11.2795 | 0.7553 | 9.7689 | 12.7901 | 9.0136 | 13.5455 | 6.696% | 13.39% | 20.09% | 10.7156 | 11.8435
P, wt.% | 0.05756 | 0.0018182 | 0.053924 | 0.061197 | 0.052106 | 0.063015 | 3.159% | 6.318% | 9.477% | 0.054682 | 0.060438
Pb, ppm | 17.2722 | 1.488 | 14.2962 | 20.2482 | 12.8082 | 21.7362 | 8.615% | 17.23% | 25.85% | 16.4086 | 18.1358
Pr, ppm | 6.892 | 0.3548 | 6.1824 | 7.6016 | 5.8276 | 7.9564 | 5.148% | 10.3% | 15.44% | 6.5474 | 7.2366
Rb, ppm | 123.4713 | 5.4907 | 112.4898 | 134.4527 | 106.9991 | 139.9434 | 4.447% | 8.894% | 13.34% | 117.2977 | 129.6448
Re, ppm | 0.035125 | 0.0075276 | 0.02007 | 0.05018 | 0.012542 | 0.057708 | 21.43% | 42.86% | 64.29% | 0.033369 | 0.036881
S, wt.% | 0.34935 | 0.020804 | 0.30774 | 0.39096 | 0.28694 | 0.41176 | 5.955% | 11.91% | 17.86% | 0.33188 | 0.36682
Sb, ppm | 0.36513 | 0.049203 | 0.26672 | 0.46353 | 0.21752 | 0.51273 | 13.48% | 26.95% | 40.43% | 0.34687 | 0.38338
Sc, ppm | 5.6188 | 0.32554 | 4.9677 | 6.2698 | 4.6422 | 6.5954 | 5.794% | 11.59% | 17.38% | 5.3378 | 5.8997
Sm, ppm | 5.1734 | 0.66127 | 3.8509 | 6.4959 | 3.1896 | 7.1572 | 12.78% | 25.56% | 38.35% | 4.9147 | 5.4321
Sn, ppm | 4.1031 | 0.10927 | 3.8845 | 4.3216 | 3.7753 | 4.4309 | 2.663% | 5.326% | 7.989% | 3.8979 | 4.3082
Sr, ppm | 197.5478 | 8.2893 | 180.9692 | 214.1264 | 172.6799 | 222.4157 | 4.196% | 8.392% | 12.59% | 187.6704 | 207.4252
Ta, ppm | 0.72409 | 0.070129 | 0.58383 | 0.86435 | 0.5137 | 0.93447 | 9.685% | 19.37% | 29.06% | 0.68788 | 0.76029
Tb, ppm | 0.65995 | 0.05046 | 0.55903 | 0.76087 | 0.50857 | 0.81133 | 7.646% | 15.29% | 22.94% | 0.62695 | 0.69295
Te, ppm | 1.1282 | 0.12854 | 0.87108 | 1.3853 | 0.74253 | 1.5138 | 11.39% | 22.79% | 34.18% | 1.0718 | 1.1846
Th, ppm | 10.0105 | 0.49502 | 9.0205 | 11.0005 | 8.5254 | 11.4955 | 4.945% | 9.89% | 14.83% | 9.51 | 10.511
Ti, wt.% | 0.23015 | 0.0086262 | 0.2129 | 0.24741 | 0.20427 | 0.25603 | 3.748% | 7.496% | 11.24% | 0.21865 | 0.24166
Tl, ppm | 0.70667 | 0.037734 | 0.6312 | 0.78213 | 0.59346 | 0.81987 | 5.34% | 10.68% | 16.02% | 0.67133 | 0.742
Tm, ppm | 0.16657 | 0.013343 | 0.13989 | 0.19326 | 0.12655 | 0.2066 | 8.01% | 16.02% | 24.03% | 0.15825 | 0.1749
U, ppm | 3.931 | 0.37203 | 3.1869 | 4.6751 | 2.8149 | 5.0471 | 9.464% | 18.93% | 28.39% | 3.7345 | 4.1276
V, ppm | 38.4563 | 1.5535 | 35.3493 | 41.5633 | 33.7957 | 43.1168 | 4.04% | 8.079% | 12.12% | 36.5335 | 40.3791
W, wt.% | 0.21492 | 0.013559 | 0.1878 | 0.24204 | 0.17424 | 0.2556 | 6.309% | 12.62% | 18.93% | 0.20418 | 0.22567
Y, ppm | 13.0604 | 1.3175 | 10.4253 | 15.6955 | 9.1078 | 17.013 | 10.09% | 20.18% | 30.26% | 12.4074 | 13.7134
Yb, ppm | 1.028 | 0.050922 | 0.92612 | 1.1298 | 0.8752 | 1.1807 | 4.954% | 9.907% | 14.86% | 0.97657 | 1.0794
Zn, ppm | 57.156 | 2.7963 | 51.5635 | 62.7485 | 48.7673 | 65.5448 | 4.892% | 9.785% | 14.68% | 54.2982 | 60.0138
Zr, ppm | 40.0151 | 3.9231 | 32.1689 | 47.8614 | 28.2458 | 51.7845 | 9.804% | 19.61% | 29.41% | 38.0144 | 42.0159
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 702
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
ISE | ion specific electrode
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PPP*XRF | pressed powder pellet with X-ray fluorescence spectroscopy
Text Values:
NA | Not Applicable (Lab 12)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 702
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
ESAN Istanbul, Istanbul, Turkey
Intertek, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
Nagrom, Perth, WA, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan
UIS Analytical Services, Centurion, South Africa

## Fusion XRF
| Analytical results for Al2O3 in OREAS 702 (Certified Value 9.7 wt.%) | [20 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 | 24 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 9.64 | 9.62 | 9.71 | 9.76 | 9.75 | 9.91 | 9.66 | 9.26 | 9.75 | 9.0546 | 9.72 | 9.693 | 9.7 | 8.18 | 9.5449 | 9.5848 | 9.7712 | 9.9 | 9.636 | 1
| 1 | 2 | 9.64 | 9.57 | 9.68 | 9.82 | 9.88 | 9.4 | 9.61 | 9.23 | 9.73 | 9.07 | 9.66 | 9.75 | 9.55 | 8.15 | 9.5953 | 9.5559 | 9.7402 | 9.72 | 9.618 | #N/A
| 1 | 3 | 9.66 | 9.59 | 9.65 | 9.96 | 9.82 | 9.44 | 9.66 | 9.25 | 9.77 | 9.1995 | 9.7 | 9.712 | 9.5 | 8.21 | 9.4252 | 9.5542 | 9.7639 | 9.9 | 9.636 | 16
| 1 | 4 | 9.69 | 9.61 | 9.78 | 9.87 | 9.83 | 9.94 | 9.68 | 9.21 | 9.74 | 9.0454 | 9.72 | 9.769 | 9.58 | 8.23 | 9.4733 | 9.6016 | 9.7317 | 9.83 | 9.599 | 9.6988
| 1 | 5 | 9.61 | 9.57 | 9.74 | 9.83 | 9.96 | 10.06 | 9.59 | 9.23 | 9.73 | 9.0064 | 9.72 | 9.788 | 9.72 | 2.26 | 9.562 | 9.5952 | 9.74 | 9.76 | 9.618 | 7
| 1 | 6 | 9.67 | 9.63 | 9.67 | 9.71 | 9.84 | 9.94 | 9.74 | 9.19 | 9.74 | 9.153 | 9.81 | 9.731 | 9.69 | 8.19 | 9.4742 | 9.604 | 9.73 | 9.86 | 9.599
| Mean |  | 9.6517 | 9.5983 | 9.705 | 9.825 | 9.8467 | 9.7817 | 9.6567 | 9.2283 | 9.7433 | 9.0882 | 9.7217 | 9.7405 | 9.6233 | 7.2033 | 9.5125 | 9.5826 | 9.7462 | 9.8283 | 9.6177
| Median |  | 9.65 | 9.6 | 9.695 | 9.825 | 9.835 | 9.925 | 9.66 | 9.23 | 9.74 | 9.0623 | 9.72 | 9.7405 | 9.635 | 8.185 | 9.5095 | 9.59 | 9.7401 | 9.845 | 9.618
| Std Dev. |  | 0.027869 | 0.025626 | 0.048477 | 0.086891 | 0.069761 | 0.28513 | 0.053166 | 0.025626 | 0.015055 | 0.07289 | 0.04916 | 0.035546 | 0.091797 | 2.4219 | 0.064794 | 0.022361 | 0.017242 | 0.074409 | 0.016549
| Rel.Std.Dev. |  | 0.2887% | 0.267% | 0.4995% | 0.8844% | 0.7085% | 2.915% | 0.5506% | 0.2777% | 0.1545% | 0.802% | 0.5057% | 0.3649% | 0.9539% | 33.62% | 0.6812% | 0.2333% | 0.1769% | 0.7571% | 0.1721%
| PDM3 |  | -0.486% | -1.036% | 0.06389% | 1.301% | 1.525% | 0.8544% | -0.4345% | -4.851% | 0.4591% | -6.296% | 0.2357% | 0.4299% | -0.7781% | -25.73% | -1.921% | -1.198% | 0.4883% | 1.336% | -0.8366%
| Z-Score (Absolute) |  | 0.04 | 0.45 | 0.38 | 1.31 | 1.48 | 0.97 | 0 | 3.33 | 0.67 | 4.42 | 0.51 | 0.65 | 0.26 | 19.09 | 1.12 | 0.58 | 0.7 | 1.34 | 0.3
| Analytical results for As in OREAS 702 (Certified Value < 100 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 12 | 14 | 20 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | 100 | < 100 | < 100 | < 50 | < 100 | < 10 | 416.572 | [11 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 50 | < 100 | < 10 | 408.998 | [11 empty] | 12
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 50 | < 100 | < 10 | 424.146 | [11 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 50 | < 100 | < 10 | 439.294 | [11 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 50 | < 100 | < 10 | 424.146 | [11 empty] | 8
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 50 | < 100 | < 10 | 416.572
| Mean |  |  | 100 | [5 empty] | 421.6213
| Median |  |  | 100 | [5 empty] | 420.359
| Std Dev. | [8 empty] | 10.3481
| Rel.Std.Dev. | [8 empty] | 2.454%
PDM3
| Z-Score (Absolute) |  | 0 | 1.35 | 0 | 0 | 4.05 | 0 | 7.28 | 60.14
| Analytical results for BaO in OREAS 702 (Certified Value 738 ppm) | [20 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 09 | 12 | 13 | 14 | 16 | 19 | 20 | 22 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 700 | 700 | 700 | 800 | 700 | 869.75 | 670 | < 11 | 810 | 743.1333 | 720 | 830 | [7 empty] | 1
| 1 | 2 | 700 | 700 | 700 | 800 | 700 | 907.16 | 670 | < 11 | 820 | 757.0333 | 710 | 740 | [7 empty] | #N/A
| 1 | 3 | 700 | 800 | 700 | 800 | 700 | 967.45 | 560 | < 11 | 790 | 719.8667 | 766.15 | 790 | [7 empty] | 16
| 1 | 4 | 700 | 700 | 800 | 800 | 700 | 955.17 | 670 | < 11 | 800 | 736.5667 | 740 | 810 | [7 empty] | 738.1906
| 1 | 5 | 800 | 700 | 700 | 800 | 700 | 839.05 | 670 | < 11 | 770 | 719.6333 | 713.55 | 810 | [7 empty] | 9
| 1 | 6 | 700 | 700 | 700 | 800 | 800 | 978.05 | 670 | < 11 | 770 | 706.1 | 710 | 790
| Mean |  | 716.6667 | 716.6667 | 716.6667 | 800 | 716.6667 | 919.4383 | 651.6667 |  | 793.3333 | 730.3889 | 726.6167 | 795
| Median |  | 700 | 700 | 700 | 800 | 700 | 931.165 | 670 |  | 795 | 728.2167 | 716.775 | 800
| Std Dev. |  | 40.8248 | 40.8248 | 40.8248 | 0 | 40.8248 | 56.7436 | 44.9073 |  | 20.6559 | 18.5842 | 22.3997 | 30.8221
| Rel.Std.Dev. |  | 5.696% | 5.696% | 5.696% | 0% | 5.696% | 6.172% | 6.891% |  | 2.604% | 2.544% | 3.083% | 3.877%
| PDM3 |  | -2.916% | -2.916% | -2.916% | 8.373% | -2.916% | 24.55% | -11.72% |  | 7.47% | -1.057% | -1.568% | 7.696%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 0.08 | 1.34 | 0.08 | 3.37 | 1.2 | 12.21 | 1.22 | 0.15 | 0.08 | 1.25
| Analytical results for Bi in OREAS 702 (Certified Value 287 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 09 | 11 | 14 | 16 | 20 | 24 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 428 | 641 | 390 | 282.241 | 300 | [7 empty] | 1
| 1 | 2 | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 423 | 590 | 300 | 322.92 | 300 | [7 empty] | 22
| 1 | 3 | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 432 | 606 | 400 | 348.081 | 300 | [7 empty] | 16
| 1 | 4 | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 429 | 630 | 400 | 322.92 | 300 | [7 empty] | 286.6877
| 1 | 5 | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 429 | 606 | 400 | 315.52 | 300 | [7 empty] | 10
| 1 | 6 | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 430 | 609 | 400 | 304.98 | 300
| Mean |  | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 428.5 | 613.6667 | 381.6667 | 316.1103 | 300
| Median |  | 300 | 200 | 200 | 300 | 300 | 300 | 270 | 429 | 607.5 | 400 | 319.22 | 300
| Std Dev. |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.0166 | 18.5113 | 40.2078 | 21.846 | 0
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0.704% | 3.017% | 10.53% | 6.911% | 0%
| PDM3 |  | 4.643% | -30.24% | -30.24% | 4.643% | 4.643% | 4.643% | -5.821% | 49.47% | 114.1% | 33.13% | 10.26% | 4.643%
| Z-Score (Absolute) |  | 0 | 2.87 | 2.87 | 0 | 0 | 0 | 0.89 | 3.69 | 9 | 2.34 | 0.46 | 0
| Analytical results for CaO in OREAS 702 (Certified Value 22.54 wt.%) | [20 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 | 24 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 22.5 | 22.2 | 22.6 | 22.4 | 22.3 | 22.83 | 22.77 | 22.28 | 20.79 | 18.1134 | 22.53 | 23.21 | 22.6 | 21.4 | 22.8445 | 22.4351 | 22.4216 | 22.74 | 22.303 | 1
| 1 | 2 | 22.4 | 22.3 | 22.6 | 22.3 | 22.5 | 22.63 | 22.81 | 22.53 | 20.75 | 18.1554 | 22.5 | 23.15 | 22.3 | 21.42 | 22.8144 | 22.4076 | 22.5797 | 22.76 | 22.331 | #N/A
| 1 | 3 | 22.5 | 22.3 | 22.6 | 22.3 | 22.5 | 22.64 | 22.82 | 22.48 | 20.76 | 18.2759 | 22.47 | 23.17 | 22.1 | 21.01 | 22.6415 | 22.4181 | 22.6025 | 22.81 | 22.275 | 16
| 1 | 4 | 22.6 | 22.1 | 22.8 | 22.4 | 22.5 | 22.79 | 22.87 | 22.36 | 20.78 | 18.1821 | 22.55 | 23.2 | 22.5 | 21.46 | 22.8519 | 22.5295 | 22.6457 | 22.71 | 22.289 | 22.5383
| 1 | 5 | 22.5 | 22.3 | 22.8 | 22.5 | 22.6 | 22.83 | 22.74 | 22.4 | 20.78 | 18.2262 | 22.47 | 23.26 | 22.8 | 21.16 | 22.8597 | 22.5618 | 22.7103 | 22.72 | 22.359 | 11
| 1 | 6 | 22.5 | 22.2 | 22.6 | 22.3 | 22.4 | 22.77 | 22.75 | 22.4 | 20.82 | 18.2431 | 22.71 | 23.17 | 22.6 | 21.2 | 22.7472 | 22.4783 | 22.5621 | 22.53 | 22.275
| Mean |  | 22.5 | 22.2333 | 22.6667 | 22.3667 | 22.4667 | 22.7483 | 22.7933 | 22.4083 | 20.78 | 18.1993 | 22.5383 | 23.1933 | 22.4833 | 21.275 | 22.7932 | 22.4717 | 22.587 | 22.7117 | 22.3053
| Median |  | 22.5 | 22.25 | 22.6 | 22.35 | 22.5 | 22.78 | 22.79 | 22.4 | 20.78 | 18.2042 | 22.515 | 23.185 | 22.55 | 21.3 | 22.8294 | 22.4567 | 22.5911 | 22.73 | 22.296
| Std Dev. |  | 0.063246 | 0.08165 | 0.10328 | 0.08165 | 0.10328 | 0.090866 | 0.049261 | 0.088185 | 0.024495 | 0.060181 | 0.089981 | 0.039328 | 0.24833 | 0.17886 | 0.08494 | 0.062967 | 0.096884 | 0.095795 | 0.033619
| Rel.Std.Dev. |  | 0.2811% | 0.3672% | 0.4556% | 0.3651% | 0.4597% | 0.3994% | 0.2161% | 0.3935% | 0.1179% | 0.3307% | 0.3992% | 0.1696% | 1.104% | 0.8407% | 0.3727% | 0.2802% | 0.4289% | 0.4218% | 0.1507%
| PDM3 |  | -0.1698% | -1.353% | 0.5697% | -0.7614% | -0.3177% | 0.932% | 1.132% | -0.5765% | -7.801% | -19.25% | 0.0002566% | 2.906% | -0.2438% | -5.605% | 1.131% | -0.2953% | 0.2161% | 0.7693% | -1.034%
| Z-Score (Absolute) |  | 0.06 | 0.92 | 0.67 | 0.43 | 0.06 | 0.97 | 1.14 | 0.28 | 6.26 | 15.76 | 0.2 | 2.61 | 0 | 4.44 | 1.14 | 0.04 | 0.38 | 0.84 | 0.65
| Analytical results for Cd in OREAS 702 (Indicative Value 66 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 14 | 20 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | [17 empty] | 0
0
| 1 | 1 | < 10 | 78.788 | [17 empty] | 1
| 1 | 2 | < 10 | 61.28 | [17 empty] | 1
| 1 | 3 | < 10 | 61.28 | [17 empty] | 16
| 1 | 4 | < 10 | 61.28 | [17 empty] | 65.657
| 1 | 5 | < 10 | 70.034 | [17 empty] | 7
| 1 | 6 | < 10 | 61.28
| Mean |  |  | 65.657
| Median |  |  | 61.28
| Std Dev. |  |  | 7.3241
| Rel.Std.Dev. |  |  | 11.16%
| PDM3 |  |  | -1.187e-05%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ce in OREAS 702 (Indicative Value 66 ppm) | [20 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [15 empty] | 0
0
| 1 | 1 | 80 | < 80 | < 80 | < 80 | [15 empty] | 1
| 1 | 2 | 80 | < 80 | < 80 | 80 | [15 empty] | 2
| 1 | 3 | 80 | < 80 | < 80 | 80 | [15 empty] | 16
| 1 | 4 | 80 | 80 | < 80 | 80 | [15 empty] | 65.5776
| 1 | 5 | < 80 | < 80 | < 80 | 80 | [15 empty] | 8
| 1 | 6 | 80 | < 80 | < 80 | 80
| Mean |  | 80 | 80 |  | 80
| Median |  | 80 | 80 |  | 80
| Std Dev. |  | 0 |  |  | 0
| Rel.Std.Dev. |  | 0% |  |  | 0%
| PDM3 |  | 21.99% | 21.99% |  | 21.99%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 1.01 | 0.67
| Analytical results for Cl in OREAS 702 (Indicative Value 194 ppm) | [20 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 12 | 20 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | [17 empty] | 0
0
| 1 | 1 | 232.45 | 131.95 | [17 empty] | 1
| 1 | 2 | 248 | 140 | [17 empty] | 3
| 1 | 3 | 212 | 179.1 | [17 empty] | 16
| 1 | 4 | 238.5 | 160 | [17 empty] | 194.2708
| 1 | 5 | 256.5 | 140.7 | [17 empty] | 9
| 1 | 6 | 242.05 | 150
| Mean |  | 238.25 | 150.2917
| Median |  | 240.275 | 145.35
| Std Dev. |  | 15.2665 | 17.0798
| Rel.Std.Dev. |  | 6.408% | 11.36%
| PDM3 |  | 22.64% | -22.64%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 702 (Certified Value < 100 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 12 | 14 | 20 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 37 | < 100 | < 50 | 7.9741 | [11 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 37 | < 100 | < 50 | [12 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 37 | < 100 | < 50 | [12 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 37 | < 100 | < 50 | [12 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 37 | < 100 | < 50 | 15.791 | [11 empty] | 12
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 37 | < 100 | < 50
| Mean | [8 empty] | 11.8825
| Median | [8 empty] | 11.8825
| Std Dev. | [8 empty] | 5.5274
| Rel.Std.Dev. | [8 empty] | 46.52%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cr2O3 in OREAS 702 (Certified Value < 150 ppm) | [20 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 12 | 13 | 14 | 19 | 22 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 150 | < 150 | < 150 | < 150 | < 100 | < 50 | < 150 | < 100 | < 100 | 830 | < 100 | [8 empty] | 1
| 1 | 2 | < 150 | < 150 | < 150 | < 150 | 100 | < 50 | < 150 | < 100 | < 100 | 576.6667 | < 100 | [8 empty] | #N/A
| 1 | 3 | < 150 | < 150 | < 150 | < 150 | 200 | < 50 | < 150 | < 100 | < 100 | 583.3333 | < 100 | [8 empty] | 16
| 1 | 4 | < 150 | < 150 | < 150 | < 150 | 100 | 60 | < 150 | < 100 | < 100 | 600 | < 100 | [8 empty] | < 150
| 1 | 5 | < 150 | < 150 | < 150 | < 150 | 100 | 50 | < 150 | < 100 | < 100 | 566.6667 | < 100 | [8 empty] | 13
| 1 | 6 | < 150 | < 150 | < 150 | < 150 | < 100 | < 50 | < 150 | < 100 | < 100 | 480 | < 100
| Mean | [5 empty] | 125 | 55 |  |  |  | 606.1111
| Median | [5 empty] | 100 | 55 |  |  |  | 580
| Std Dev. | [5 empty] | 50 | 7.0711 |  |  |  | 117.4813
| Rel.Std.Dev. | [5 empty] | 40% | 12.86% |  |  |  | 19.38%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 0 | 0 | 0.79 | 1.11 | 0 | 0.67 | 0.67 | 15.58 | 0.67
| Analytical results for Cu in OREAS 702 (Certified Value 352 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 12 | 14 | 16 | 19 | 20 | 22 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 360 | 390 | 450 | 360 | 359 | 360 | 338 | 700 | 421 | < 160 | 447.3558 | 328 | [7 empty] | 1
| 1 | 2 | 350 | 350 | 390 | 370 | 375 | 360 | 333 | 800 | 450 | < 160 | 439.3673 | 328 | [7 empty] | 19
| 1 | 3 | 360 | 360 | 330 | 380 | 359 | 350 | 312.5 | 700 | 400 | < 160 | 408.9531 | 336 | [7 empty] | 16
| 1 | 4 | 360 | 350 | 340 | 360 | 359 | 360 | 329 | 700 | 420 | < 160 | 419.3961 | 328 | [7 empty] | 352.1244
| 1 | 5 | 360 | 360 | 350 | 360 | 375 | 360 | 326 | 700 | 430 | < 160 | 455.3443 | 344 | [7 empty] | 14
| 1 | 6 | 370 | 350 | 330 | 460 | 359 | 360 | 323.5 | 800 | 390 | < 160 | 444.5598 | 336
| Mean |  | 360 | 360 | 365 | 381.6667 | 364.3333 | 358.3333 | 327 | 733.3333 | 418.5 |  | 435.8294 | 333.3333
| Median |  | 360 | 355 | 345 | 365 | 359 | 360 | 327.5 | 700 | 420.5 |  | 441.9636 | 332
| Std Dev. |  | 6.3246 | 15.4919 | 47.2229 | 39.2003 | 8.2624 | 4.0825 | 8.7693 | 51.6398 | 21.3892 |  | 17.8595 | 6.532
| Rel.Std.Dev. |  | 1.757% | 4.303% | 12.94% | 10.27% | 2.268% | 1.139% | 2.682% | 7.042% | 5.111% |  | 4.098% | 1.96%
| PDM3 |  | 2.237% | 2.237% | 3.657% | 8.39% | 3.467% | 1.763% | -7.135% | 108.3% | 18.85% |  | 23.77% | -5.336%
| Z-Score (Absolute) |  | 0.07 | 0.07 | 0.07 | 0.53 | 0.07 | 0.11 | 0.98 | 10.25 | 1.55 | 7.81 | 2.03 | 0.82
| Analytical results for F in OREAS 702 (Indicative Value 0.236 wt.%) | [20 empty] | Indicative
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 09 | 11 | 16 | 17 | [11 empty] | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | OxBF*XRF | [11 empty] | 3
3
| 1 | 1 | 0.3 | 0.3 | 0.1 | 0.3 | 0.13 | 0.2448 | 0.22 | < 1 | [11 empty] | 1
| 1 | 2 | 0.3 | 0.4 | 0.1 | 0.3 | 0.11 | 0.2388 | 0.29 | < 1 | [11 empty] | 4
| 1 | 3 | 0.3 | 0.3 | 0.1 | 0.3 | 0.13 | 0.2466 | 0.24 | < 1 | [11 empty] | 16
| 1 | 4 | 0.3 | 0.4 | 0.1 | 0.3 | 0.14 | 0.2448 | 0.26 | < 1 | [11 empty] | 0.23601
| 1 | 5 | 0.3 | 0.3 | 0.1 | 0.3 | 0.09 | 0.2502 | 0.24 | < 1 | [11 empty] | 10
| 1 | 6 | 0.3 | 0.4 | 0.1 | 0.3 | 0.07 | 0.2472 | 0.22 | < 1
| Mean |  | 0.3 | 0.35 | 0.1 | 0.3 | 0.11167 | 0.2454 | 0.245
| Median |  | 0.3 | 0.35 | 0.1 | 0.3 | 0.12 | 0.2457 | 0.24
| Std Dev. |  | 0 | 0.054772 | 1.5202e-17 | 0 | 0.027142 | 0.0037947 | 0.026646
| Rel.Std.Dev. |  | 0% | 15.65% | 1.52e-14% | 0% | 24.31% | 1.546% | 10.88%
| PDM3 |  | 27.11% | 48.3% | -57.63% | 27.11% | -52.69% | 3.979% | 3.809%
| Z-Score (Absolute) |  | 0.35 | 0.99 | 2.22 | 0.35 | 2.07 | 0.35 | 0.36 | 2.92
| Analytical results for Fe2O3 in OREAS 702 (Certified Value 3.36 wt.%) | [20 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 | 24 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 3.388 | 3.303 | 3.503 | 3.39 | 3.345 | 3.45 | 3.37 | 3.44 | 3.43 | 3.2451 | 3.38 | 3.203 | 3.43 | 3.2 | 3.2352 | 3.3927 | 3.302 | 3.39 | 3.317 | 1
| 1 | 2 | 3.36 | 3.274 | 3.46 | 3.36 | 3.374 | 3.43 | 3.41 | 3.43 | 3.46 | 3.2544 | 3.37 | 3.203 | 3.39 | 3.5 | 3.2265 | 3.4244 | 3.3464 | 3.39 | 3.317 | #N/A
| 1 | 3 | 3.374 | 3.303 | 3.474 | 3.35 | 3.431 | 3.43 | 3.39 | 3.48 | 3.43 | 3.3105 | 3.35 | 3.188 | 3.36 | 3.2 | 3.2083 | 3.3724 | 3.3296 | 3.38 | 3.317 | 16
| 1 | 4 | 3.36 | 3.26 | 3.474 | 3.34 | 3.374 | 3.44 | 3.39 | 3.44 | 3.44 | 3.2577 | 3.37 | 3.203 | 3.44 | 3.23 | 3.2423 | 3.3866 | 3.3516 | 3.38 | 3.317 | 3.3575
| 1 | 5 | 3.36 | 3.288 | 3.474 | 3.36 | 3.431 | 3.44 | 3.36 | 3.44 | 3.45 | 3.2574 | 3.35 | 3.203 | 3.44 | 3.23 | 3.2768 | 3.4135 | 3.3399 | 3.37 | 3.317 | 15
| 1 | 6 | 3.388 | 3.288 | 3.46 | 3.38 | 3.488 | 3.44 | 3.38 | 3.45 | 3.44 | 3.3455 | 3.4 | 3.217 | 3.41 | 3.23 | 3.2058 | 3.39 | 3.318 | 3.36 | 3.303
| Mean |  | 3.3717 | 3.286 | 3.4742 | 3.3633 | 3.4072 | 3.4383 | 3.3833 | 3.4467 | 3.4417 | 3.2784 | 3.37 | 3.2028 | 3.4117 | 3.265 | 3.2325 | 3.3966 | 3.3312 | 3.3783 | 3.3147
| Median |  | 3.367 | 3.288 | 3.474 | 3.36 | 3.4025 | 3.44 | 3.385 | 3.44 | 3.44 | 3.2576 | 3.37 | 3.203 | 3.42 | 3.23 | 3.2309 | 3.3913 | 3.3347 | 3.38 | 3.317
| Std Dev. |  | 0.013765 | 0.016769 | 0.015702 | 0.018619 | 0.052412 | 0.0075277 | 0.017512 | 0.017512 | 0.01169 | 0.040205 | 0.018974 | 0.0091742 | 0.031885 | 0.11606 | 0.026071 | 0.018974 | 0.018709 | 0.01169 | 0.0057155
| Rel.Std.Dev. |  | 0.4082% | 0.5103% | 0.452% | 0.5536% | 1.538% | 0.2189% | 0.5176% | 0.5081% | 0.3397% | 1.226% | 0.563% | 0.2864% | 0.9346% | 3.555% | 0.8065% | 0.5586% | 0.5616% | 0.346% | 0.1724%
| PDM3 |  | 0.4205% | -2.131% | 3.473% | 0.1723% | 1.478% | 2.406% | 0.768% | 2.654% | 2.505% | -2.356% | 0.3709% | -4.608% | 1.612% | -2.756% | -3.724% | 1.163% | -0.7836% | 0.6191% | -1.277%
| Z-Score (Absolute) |  | 0 | 1.01 | 1.21 | 0.1 | 0.42 | 0.79 | 0.14 | 0.88 | 0.82 | 1.1 | 0.02 | 1.99 | 0.47 | 1.26 | 1.64 | 0.29 | 0.48 | 0.08 | 0.67
| Analytical results for HfO2 in OREAS 702 (Certified Value < 100 ppm) | [20 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 14 | 19 | 22 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 200 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 200 | < 100 | [12 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 200 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 200 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 200 | < 100 | [12 empty] | 16
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 200 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for K2O in OREAS 702 (Certified Value 2.27 wt.%) | [20 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 | 24 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.26 | 2.29 | 2.21 | 2.25 | 2.25 | 2.19 | 2.25 | 2.266 | 2.42 | 2.3018 | 2.27 | 2.22 | 2.3 | 2.14 | 2.2669 | 2.2655 | 2.2536 | 2.31 | 2.302 | 1
| 1 | 2 | 2.27 | 2.3 | 2.21 | 2.25 | 2.25 | 2.17 | 2.26 | 2.281 | 2.44 | 2.3235 | 2.25 | 2.21 | 2.28 | 2.13 | 2.2349 | 2.2558 | 2.2588 | 2.32 | 2.2887 | #N/A
| 1 | 3 | 2.27 | 2.29 | 2.2 | 2.26 | 2.25 | 2.19 | 2.24 | 2.279 | 2.42 | 2.3268 | 2.25 | 2.22 | 2.25 | 2.11 | 2.2231 | 2.2602 | 2.2613 | 2.34 | 2.2839 | 16
| 1 | 4 | 2.27 | 2.28 | 2.23 | 2.26 | 2.26 | 2.17 | 2.26 | 2.256 | 2.43 | 2.3135 | 2.25 | 2.22 | 2.29 | 2.12 | 2.2418 | 2.2695 | 2.2491 | 2.33 | 2.2911 | 2.2682
| 1 | 5 | 2.26 | 2.29 | 2.22 | 2.27 | 2.28 | 2.19 | 2.24 | 2.263 | 2.42 | 2.3454 | 2.24 | 2.23 | 2.3 | 2.13 | 2.2458 | 2.2772 | 2.2711 | 2.33 | 2.2875 | 17
| 1 | 6 | 2.27 | 2.3 | 2.21 | 2.31 | 2.29 | 2.19 | 2.25 | 2.27 | 2.41 | 2.3527 | 2.29 | 2.21 | 2.29 | 2.11 | 2.2505 | 2.2679 | 2.2855 | 2.29 | 2.2803
| Mean |  | 2.2667 | 2.2917 | 2.2133 | 2.2667 | 2.2633 | 2.1833 | 2.25 | 2.2692 | 2.4233 | 2.3273 | 2.2583 | 2.2183 | 2.285 | 2.1233 | 2.2438 | 2.266 | 2.2632 | 2.32 | 2.2889
| Median |  | 2.27 | 2.29 | 2.21 | 2.26 | 2.255 | 2.19 | 2.25 | 2.268 | 2.42 | 2.3251 | 2.25 | 2.22 | 2.29 | 2.125 | 2.2438 | 2.2667 | 2.26 | 2.325 | 2.2881
| Std Dev. |  | 0.005164 | 0.0075277 | 0.010328 | 0.022509 | 0.017512 | 0.010328 | 0.0089443 | 0.0095795 | 0.010328 | 0.01914 | 0.018348 | 0.0075277 | 0.018708 | 0.012111 | 0.014788 | 0.0074804 | 0.01323 | 0.017889 | 0.0074486
| Rel.Std.Dev. |  | 0.2278% | 0.3285% | 0.4666% | 0.9931% | 0.7737% | 0.473% | 0.3975% | 0.4222% | 0.4262% | 0.8224% | 0.8125% | 0.3393% | 0.8187% | 0.5704% | 0.6591% | 0.3301% | 0.5845% | 0.7711% | 0.3254%
| PDM3 |  | -0.06929% | 1.033% | -2.421% | -0.06929% | -0.2162% | -3.743% | -0.8041% | 0.04093% | 6.838% | 2.603% | -0.4367% | -2.2% | 0.739% | -6.388% | -1.076% | -0.09782% | -0.2204% | 2.282% | 0.9117%
| Z-Score (Absolute) |  | 0.02 | 0.78 | 1.6 | 0.02 | 0.08 | 2.51 | 0.49 | 0.1 | 4.78 | 1.86 | 0.23 | 1.45 | 0.58 | 4.34 | 0.67 | 0 | 0.08 | 1.64 | 0.7
| Analytical results for La2O3 in OREAS 702 (Indicative Value < 100 ppm) | [20 empty] | Indicative
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 14 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | 100 | < 100 | [14 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | 100 | < 100 | [14 empty] | 5
| 1 | 3 | 100 | < 100 | < 100 | 100 | < 100 | [14 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | 100 | < 100 | [14 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | 100 | < 100 | [14 empty] | 11
| 1 | 6 | 100 | < 100 | < 100 | 100 | < 100
| Mean |  | 100 |  |  | 100
| Median |  | 100 |  |  | 100
| Std Dev. |  | 0 |  |  | 0
| Rel.Std.Dev. |  | 0% |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for MgO in OREAS 702 (Certified Value 1.03 wt.%) | [20 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 | 24 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.08 | 0.99 | 1.04 | 1.08 | 1.02 | 1.05 | 1 | 1 | 1.05 | 1.0083 | 0.99 | 0.929 | 1.03 | 1.11 | 1.0299 | 1.0205 | 1.013 | 0.94 | 0.995 | 1
| 1 | 2 | 1.04 | 0.99 | 1.03 | 1.02 | 1.04 | 1.02 | 1.02 | 1.01 | 1.06 | 0.99245 | 0.98 | 0.912 | 1.01 | 1.1 | 1.0307 | 1.0146 | 1.0281 | 0.94 | 0.995 | #N/A
| 1 | 3 | 1.06 | 1 | 1.02 | 1.06 | 1.05 | 1.01 | 1.01 | 1.01 | 1.04 | 1.025 | 0.99 | 0.912 | 1 | 1.08 | 1.0254 | 1.0149 | 1.015 | 0.93 | 1.012 | 16
| 1 | 4 | 1.06 | 0.99 | 1.06 | 1.04 | 1.03 | 1.04 | 1 | 1 | 1.05 | 1.0148 | 1 | 0.929 | 1.03 | 1.1 | 1.051 | 1.0216 | 1.0235 | 0.94 | 0.995 | 1.0285
| 1 | 5 | 1.06 | 1 | 1.04 | 1.02 | 1.08 | 1.05 | 1.01 | 1.01 | 1.06 | 1.0351 | 0.99 | 0.945 | 1.03 | 1.09 | 1.0361 | 1.0268 | 1.0182 | 0.93 | 0.995 | 18
| 1 | 6 | 1.06 | 1 | 1.02 | 1.06 | 1.04 | 1.07 | 1 | 1 | 1.05 | 1.0301 | 1.01 | 0.929 | 1.03 | 1.1 | 1.0364 | 1.0247 | 1.0109 | 0.92 | 0.995
| Mean |  | 1.06 | 0.995 | 1.035 | 1.0467 | 1.0433 | 1.04 | 1.0067 | 1.005 | 1.0517 | 1.0176 | 0.99333 | 0.926 | 1.0217 | 1.0967 | 1.0349 | 1.0205 | 1.0181 | 0.93333 | 0.99783
| Median |  | 1.06 | 0.995 | 1.035 | 1.05 | 1.04 | 1.045 | 1.005 | 1.005 | 1.05 | 1.0199 | 0.99 | 0.929 | 1.03 | 1.1 | 1.0334 | 1.021 | 1.0166 | 0.935 | 0.995
| Std Dev. |  | 0.012649 | 0.0054772 | 0.015166 | 0.024221 | 0.020656 | 0.021909 | 0.008165 | 0.0054772 | 0.0075277 | 0.015768 | 0.010328 | 0.01249 | 0.013292 | 0.010328 | 0.0089052 | 0.0050128 | 0.0065933 | 0.008165 | 0.0069402
| Rel.Std.Dev. |  | 1.193% | 0.5505% | 1.465% | 2.314% | 1.98% | 2.107% | 0.8111% | 0.545% | 0.7158% | 1.549% | 1.04% | 1.349% | 1.301% | 0.9418% | 0.8605% | 0.4912% | 0.6476% | 0.8748% | 0.6955%
| PDM3 |  | 3.066% | -3.254% | 0.6357% | 1.77% | 1.446% | 1.122% | -2.119% | -2.281% | 2.256% | -1.053% | -3.416% | -9.963% | -0.6608% | 6.632% | 0.6282% | -0.7735% | -1.006% | -9.25% | -2.978%
| Z-Score (Absolute) |  | 1.17 | 0.75 | 0.43 | 0.77 | 0.67 | 0.58 | 0.41 | 0.46 | 0.92 | 0.08 | 0.8 | 2.8 | 0.03 | 2.25 | 0.43 | 0 | 0.07 | 2.58 | 0.67
| Analytical results for MnO in OREAS 702 (Certified Value 0.038 wt.%) | [20 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 |  |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.039 | 0.039 | 0.039 | 0.039 | 0.047 | 0.03 | 0.021 | 0.06 | 0.0467 | 0.04 | 0.0387 | 0.03 | 0.023 | 0.035594 | 0.04405 | 0.038263 | 0.042 |  |  | 1
| 1 | 2 | 0.039 | 0.039 | 0.039 | 0.039 | 0.05 | 0.04 | 0.021 | 0.06 | 0.04465 | 0.04 | 0.0387 | 0.03 | 0.023 | 0.037742 | 0.044057 | 0.039493 | 0.048 |  |  | #N/A
| 1 | 3 | 0.039 | 0.039 | 0.039 | 0.039 | 0.046 | 0.04 | 0.026 | 0.06 | 0.04515 | 0.04 | 0.0387 | 0.03 | 0.025 | 0.036106 | 0.043373 | 0.039295 | 0.044 |  |  | 16
| 1 | 4 | 0.039 | 0.039 | 0.039 | 0.039 | 0.044 | 0.04 | 0.027 | 0.06 | 0.0439 | 0.04 | 0.0387 | 0.03 | 0.027 | 0.037482 | 0.04371 | 0.039166 | 0.045 |  |  | 0.038364
| 1 | 5 | 0.039 | 0.039 | 0.039 | 0.039 | 0.045 | 0.03 | 0.022 | 0.07 | 0.0432 | 0.04 | 0.0387 | 0.03 | 0.028 | 0.036254 | 0.0427 | 0.037886 | 0.044 |  |  | 19
| 1 | 6 | 0.039 | 0.039 | 0.026 | 0.039 | 0.046 | 0.03 | 0.021 | 0.06 | 0.0452 | 0.04 | 0.0387 | 0.03 | 0.025 | 0.03605 | 0.04406 | 0.037698 | 0.044
| Mean |  | 0.039 | 0.039 | 0.036833 | 0.039 | 0.046333 | 0.035 | 0.023 | 0.061667 | 0.0448 | 0.04 | 0.0387 | 0.03 | 0.025167 | 0.036538 | 0.043658 | 0.038633 | 0.0445
| Median |  | 0.039 | 0.039 | 0.039 | 0.039 | 0.046 | 0.035 | 0.0215 | 0.06 | 0.0449 | 0.04 | 0.0387 | 0.03 | 0.025 | 0.03618 | 0.04388 | 0.038715 | 0.044
| Std Dev. |  | 0 | 0 | 0.0053072 | 0 | 0.0020656 | 0.0054772 | 0.0027568 | 0.0040825 | 0.0012071 | 0 | 7.6012e-18 | 0 | 0.0020412 | 0.00086456 | 0.00054327 | 0.0007787 | 0.0019748
| Rel.Std.Dev. |  | 0% | 0% | 14.41% | 0% | 4.458% | 15.65% | 11.99% | 6.62% | 2.694% | 0% | 1.964e-14% | 0% | 8.111% | 2.366% | 1.244% | 2.016% | 4.438%
| PDM3 |  | 1.657% | 1.657% | -3.99% | 1.657% | 20.77% | -8.769% | -40.05% | 60.74% | 16.78% | 4.264% | 0.8755% | -21.8% | -34.4% | -4.76% | 13.8% | 0.7021% | 15.99%
| Z-Score (Absolute) |  | 0 | 0 | 0.39 | 0 | 1.37 | 0.67 | 2.84 | 4.14 | 1.09 | 0.23 | 0 | 1.58 | 2.45 | 0.4 | 0.89 | 0.02 | 1.04
| Analytical results for Mo in OREAS 702 (Certified Value 625 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 09 | 11 | 14 | 16 | 20 | 24 | [7 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 640 | 670 | 610 | 610 | 630 | 600 | 670 | 626 | 558 | 550 | 669.766 | 600 | [7 empty] | 1
| 1 | 2 | 620 | 640 | 590 | 610 | 620 | 600 | 670 | 656 | 569 | 600 | 679.864 | 600 | [7 empty] | 21
| 1 | 3 | 630 | 640 | 600 | 610 | 610 | 530 | 670 | 623 | 560 | 600 | 676.465 | 600 | [7 empty] | 16
| 1 | 4 | 620 | 640 | 590 | 590 | 620 | 530 | 670 | 660 | 560 | 630 | 666.533 | 600 | [7 empty] | 625.2316
| 1 | 5 | 600 | 620 | 620 | 580 | 650 | 600 | 670 | 663 | 565 | 620 | 666.533 | 600 | [7 empty] | 20
| 1 | 6 | 620 | 640 | 610 | 630 | 660 | 530 | 670 | 664 | 565 | 600 | 666.533 | 600
| Mean |  | 621.6667 | 641.6667 | 603.3333 | 605 | 631.6667 | 565 | 670 | 648.6667 | 562.8333 | 600 | 670.949 | 600
| Median |  | 620 | 640 | 605 | 610 | 625 | 565 | 670 | 658 | 562.5 | 600 | 668.1495 | 600
| Std Dev. |  | 13.2916 | 16.0208 | 12.1106 | 17.6068 | 19.4079 | 38.3406 | 0 | 18.9491 | 4.1673 | 27.5681 | 5.8276 | 0
| Rel.Std.Dev. |  | 2.138% | 2.497% | 2.007% | 2.91% | 3.072% | 6.786% | 0% | 2.921% | 0.7404% | 4.595% | 0.8686% | 0%
| PDM3 |  | -0.5702% | 2.629% | -3.502% | -3.236% | 1.029% | -9.633% | 7.16% | 3.748% | -9.98% | -4.036% | 7.312% | -4.036%
| Z-Score (Absolute) |  | 0.15 | 0.45 | 0.69 | 0.64 | 0.15 | NA | 1.19 | 0.65 | 1.9 | 0.79 | 1.32 | NA
Results from laboratories 7 and 24 were removed due to their 100 ppm reading resolution.
| Analytical results for Na2O in OREAS 702 (Certified Value 1.77 wt.%) | [20 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 24 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 2
3
| 1 | 1 | 1.77 | 1.84 | 1.78 | 2.06 | 1.8614 | 1.84 | 1.69 | 1.78 | 1.6304 | 1.7792 | 1.837 | 1.68 | 1.752 | [6 empty] | 1
| 1 | 2 | 1.75 | 1.86 | 1.78 | 2.07 | 1.8443 | 1.84 | 1.66 | 1.76 | 1.6235 | 1.7809 | 1.852 | 1.68 | 1.739 | [6 empty] | #N/A
| 1 | 3 | 1.75 | 1.87 | 1.78 | 2.08 | 1.8132 | 1.85 | 1.68 | 1.72 | 1.6325 | 1.7722 | 1.8427 | 1.68 | 1.752 | [6 empty] | 16
| 1 | 4 | 1.75 | 1.86 | 1.78 | 2.06 | 1.8235 | 1.86 | 1.67 | 1.78 | 1.6176 | 1.7856 | 1.841 | 1.68 | 1.752 | [6 empty] | 1.7681
| 1 | 5 | 1.76 | 1.84 | 1.79 | 2.07 | 1.8106 | 1.85 | 1.69 | 1.78 | 1.6293 | 1.7852 | 1.8492 | 1.68 | 1.752 | [6 empty] | 21
| 1 | 6 | 1.77 | 1.87 | 1.79 | 2.05 | 1.8509 | 1.88 | 1.67 | 1.79 | 1.6455 | 1.7696 | 1.838 | 1.69 | 1.766
| Mean |  | 1.7583 | 1.8567 | 1.7833 | 2.065 | 1.834 | 1.8533 | 1.6767 | 1.7683 | 1.6298 | 1.7788 | 1.8433 | 1.6817 | 1.7522
| Median |  | 1.755 | 1.86 | 1.78 | 2.065 | 1.8339 | 1.85 | 1.675 | 1.78 | 1.6299 | 1.7801 | 1.8419 | 1.68 | 1.752
| Std Dev. |  | 0.0098319 | 0.013663 | 0.005164 | 0.010488 | 0.021137 | 0.015055 | 0.012111 | 0.025626 | 0.009402 | 0.0066486 | 0.0060643 | 0.0040825 | 0.0085421
| Rel.Std.Dev. |  | 0.5592% | 0.7359% | 0.2896% | 0.5079% | 1.153% | 0.8123% | 0.7223% | 1.449% | 0.5769% | 0.3738% | 0.329% | 0.2428% | 0.4875%
| PDM3 |  | -0.5497% | 5.012% | 0.8643% | 16.8% | 3.729% | 4.823% | -5.169% | 0.01587% | -7.819% | 0.6067% | 4.257% | -4.886% | -0.8985%
| Z-Score (Absolute) |  | 0.21 | 0.81 | 0.05 | 2.99 | 0.58 | 0.78 | 1.07 | 0.11 | 1.56 | 0 | 0.67 | 1.01 | 0.28
| Analytical results for Nb in OREAS 702 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [15 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | < 50 | [15 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 50 | [15 empty] | 1
| 1 | 3 | < 50 | < 50 | < 50 | < 50 | [15 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 50 | [15 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | < 50 | [15 empty] | 7
| 1 | 6 | < 50 | < 50 | < 50 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Ni in OREAS 702 (Certified Value < 50 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 12 | 14 | 19 | 22 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | < 50 | < 24 | < 100 | < 10 | 484.6103 | < 47 | [10 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 50 | < 24 | < 100 | < 10 | 235.7564 | < 47 | [10 empty] | 12
| 1 | 3 | < 50 | < 50 | < 50 | < 50 | 24 | < 100 | < 10 | 311.7223 | < 47 | [10 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 50 | < 24 | < 100 | < 10 | 259.332 | < 47 | [10 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | < 50 | 24 | < 100 | < 10 | 188.6051 | < 47 | [10 empty] | 22
| 1 | 6 | < 50 | < 50 | < 50 | < 50 | 39 | < 100 | < 10 | 366.7322 | < 47
| Mean | [5 empty] | 29 |  |  | 307.7931
| Median | [5 empty] | 24 |  |  | 285.5272
| Std Dev. | [5 empty] | 8.6603 |  |  | 106.2988
| Rel.Std.Dev. | [5 empty] | 29.86% |  |  | 34.54%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 0 | 0 | 2.22 | 11.84 | 9.47 | 133.88 | 0.67
| Analytical results for P2O5 in OREAS 702 (Certified Value 0.129 wt.%) | [20 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 24 |  |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.126 | 0.12 | 0.15 | 0.13142 | 0.14 | 0.12 | 0.15 | 0.12106 | 0.13241 | 0.156 | 0.13 | 0.1306 |  |  | 1
| 1 | 2 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.124 | 0.122 | 0.15 | 0.132 | 0.13 | 0.12 | 0.14 | 0.12126 | 0.13241 | 0.1596 | 0.13 | 0.1306 |  |  | #N/A
| 1 | 3 | 0.13 | 0.13 | 0.13 | 0.12 | 0.13 | 0.127 | 0.122 | 0.15 | 0.13326 | 0.13 | 0.12 | 0.15 | 0.12184 | 0.13094 | 0.16298 | 0.14 | 0.1306 |  |  | 16
| 1 | 4 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.128 | 0.119 | 0.15 | 0.12902 | 0.13 | 0.12 | 0.15 | 0.12398 | 0.13204 | 0.16065 | 0.14 | 0.1306 |  |  | 0.1289
| 1 | 5 | 0.13 | 0.13 | 0.14 | 0.13 | 0.13 | 0.128 | 0.118 | 0.16 | 0.1273 | 0.14 | 0.12 | 0.15 | 0.1252 | 0.13203 | 0.155 | 0.14 | 0.1306 |  |  | 23
| 1 | 6 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.127 | 0.114 | 0.15 | 0.13131 | 0.14 | 0.12 | 0.15 | 0.12079 | 0.13315 | 0.1533 | 0.14 | 0.1283
| Mean |  | 0.13 | 0.13167 | 0.13167 | 0.12833 | 0.13 | 0.12667 | 0.11917 | 0.15167 | 0.13072 | 0.135 | 0.12 | 0.14833 | 0.12235 | 0.13217 | 0.15792 | 0.13667 | 0.13022
| Median |  | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.127 | 0.1195 | 0.15 | 0.13137 | 0.135 | 0.12 | 0.15 | 0.12155 | 0.13223 | 0.1578 | 0.14 | 0.1306
| Std Dev. |  | 0 | 0.0040825 | 0.0040825 | 0.0040825 | 0 | 0.0015055 | 0.0029944 | 0.0040825 | 0.0021687 | 0.0054772 | 0 | 0.0040825 | 0.0018069 | 0.00072432 | 0.0037268 | 0.005164 | 0.00093897
| Rel.Std.Dev. |  | 0% | 3.101% | 3.101% | 3.181% | 0% | 1.189% | 2.513% | 2.692% | 1.659% | 4.057% | 0% | 2.752% | 1.477% | 0.548% | 2.36% | 3.779% | 0.7211%
| PDM3 |  | 0.851% | 2.144% | 2.144% | -0.442% | 0.851% | -1.735% | -7.553% | 17.66% | 1.407% | 4.73% | -6.907% | 15.07% | -5.08% | 2.531% | 22.51% | 6.023% | 1.019%
| Z-Score (Absolute) |  | 0.12 | 0.16 | 0.16 | 0.4 | 0.12 | 0.67 | 1.92 | 3.49 | 0 | 0.71 | 1.78 | 2.93 | 1.39 | 0.24 | 4.53 | 0.99 | 0.08
| Analytical results for Pb in OREAS 702 (Certified Value < 50 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 16 | 19 | 22 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | 60 | < 100 | < 10 | < 100 | < 190 | < 74 | [10 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | 50 | < 100 | < 10 | < 100 | < 190 | < 74 | [10 empty] | 23
| 1 | 3 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | < 100 | < 190 | < 74 | [10 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | < 100 | < 190 | < 74 | [10 empty] | < 50
| 1 | 5 | < 50 | < 50 | 50 | 70 | < 100 | < 10 | < 100 | < 190 | < 74 | [10 empty] | 24
| 1 | 6 | < 50 | < 50 | < 50 | 90 | < 100 | < 10 | < 100 | < 190 | < 74
| Mean |  |  |  | 50 | 67.5
| Median |  |  |  | 50 | 65
| Std Dev. |  |  |  |  | 17.0783
| Rel.Std.Dev. |  |  |  |  | 25.3%
PDM3
| Z-Score (Absolute) |  | 0.64 | 0.64 | 0.42 | 0.85 | 0.67 | 1.68 | 0.67 | 2.92 | 0
| Analytical results for Rb in OREAS 702 (Indicative Value 117 ppm) | [20 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [15 empty] | 0
0
| 1 | 1 | 120 | 100 | 120 | 130 | [15 empty] | 1
| 1 | 2 | 110 | 120 | 120 | 110 | [15 empty] | 2
| 1 | 3 | 120 | 100 | 130 | 120 | [15 empty] | 16
| 1 | 4 | 110 | 100 | 130 | 120 | [15 empty] | 116.5833
| 1 | 5 | 110 | 80 | 120 | 110 | [15 empty] | 8
| 1 | 6 | 120 | 120 | 130 | 120
| Mean |  | 115 | 103.3333 | 125 | 118.3333
| Median |  | 115 | 100 | 125 | 120
| Std Dev. |  | 5.4772 | 15.0555 | 5.4772 | 7.5277
| Rel.Std.Dev. |  | 4.763% | 14.57% | 4.382% | 6.361%
| PDM3 |  | -1.358% | -11.37% | 7.219% | 1.501%
| Z-Score (Absolute) |  | 0.22 | 1.8 | 1.12 | 0.22
| Analytical results for Sb in OREAS 702 (Indicative Value 138 ppm) | [20 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 14 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 50 | 100 | 250 | 60 | < 10 | [14 empty] | 1
| 1 | 2 | < 50 | 100 | 200 | < 50 | < 10 | [14 empty] | 3
| 1 | 3 | < 50 | 120 | 250 | 70 | < 10 | [14 empty] | 16
| 1 | 4 | < 50 | 120 | 230 | 50 | < 10 | [14 empty] | 137.5
| 1 | 5 | < 50 | 130 | 270 | 50 | < 10 | [14 empty] | 9
| 1 | 6 | < 50 | 140 | 240 | 70 | < 10
| Mean |  |  | 118.3333 | 240 | 60
| Median |  |  | 120 | 245 | 60
| Std Dev. |  |  | 16.0208 | 23.6643 | 10
| Rel.Std.Dev. |  |  | 13.54% | 9.86% | 16.67%
| PDM3 |  |  | -13.94% | 74.55% | -56.36%
| Z-Score (Absolute) |  | 0.4 | 0.88 | 2.55 | 0 | 0.67
| Analytical results for Sc in OREAS 702 (Indicative Value 13.4 ppm) | [20 empty] | Indicative
Sc | Round | Replicate | Lab | [18 empty] | 1
| No | No | 14 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 1
1
| 1 | 1 | 13 | [18 empty] | 1
| 1 | 2 | 13 | [18 empty] | 4
| 1 | 3 | < 10 | [18 empty] | 16
| 1 | 4 | 12 | [18 empty] | 13.4
| 1 | 5 | 17 | [18 empty] | 10
| 1 | 6 | 12
| Mean |  | 13.4
| Median |  | 13
| Std Dev. |  | 2.0736
| Rel.Std.Dev. |  | 15.47%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 702 (Certified Value 41.41 wt.%) | [20 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 22 | 24 | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 41.3 | 41.1 | 41.6 | 41.5 | 41.5 | 41.84 | 41.63 | 39.07 | 42.56 | 40.641 | 41.62 | 41.53 | 42.4 | 39.69 | 41.0372 | 41.1077 | 41.2546 | 41.44 | 41.394 | 1
| 1 | 2 | 41.2 | 41.2 | 41.9 | 41.2 | 41.7 | 41.26 | 41.73 | 38.98 | 42.55 | 40.7428 | 41.48 | 41.28 | 41.7 | 39.68 | 41.1096 | 41.0477 | 41.2574 | 41.53 | 41.287 | #N/A
| 1 | 3 | 41.4 | 41.1 | 41.6 | 41.2 | 41.8 | 41.2 | 41.75 | 39.16 | 42.59 | 41.3214 | 41.32 | 41.34 | 41.4 | 38.94 | 40.858 | 41.0548 | 41.312 | 41.33 | 41.351 | 16
| 1 | 4 | 41.5 | 41 | 42.2 | 41.2 | 41.6 | 41.47 | 41.9 | 38.83 | 42.56 | 40.6525 | 41.42 | 41.47 | 42.1 | 39.58 | 40.9394 | 41.2394 | 41.3795 | 41.46 | 41.415 | 41.4064
| 1 | 5 | 41.3 | 41.1 | 42.1 | 41 | 42 | 41.56 | 41.73 | 38.84 | 42.53 | 40.4604 | 41.42 | 41.52 | 42.3 | 38.36 | 41.1497 | 41.3495 | 41.2093 | 41.52 | 41.458 | 25
| 1 | 6 | 41.5 | 41.2 | 41.6 | 41.4 | 41.8 | 41.52 | 41.76 | 38.91 | 42.55 | 41.2432 | 41.74 | 41.34 | 42 | 39.29 | 41.048 | 41.1643 | 41.2461 | 41.21 | 41.265
| Mean |  | 41.3667 | 41.1167 | 41.8333 | 41.25 | 41.7333 | 41.475 | 41.75 | 38.965 | 42.5567 | 40.8435 | 41.5 | 41.4133 | 41.9833 | 39.2567 | 41.0236 | 41.1606 | 41.2765 | 41.415 | 41.3617
| Median |  | 41.35 | 41.1 | 41.75 | 41.2 | 41.75 | 41.495 | 41.74 | 38.945 | 42.555 | 40.6977 | 41.45 | 41.405 | 42.05 | 39.435 | 41.0426 | 41.136 | 41.256 | 41.45 | 41.3725
| Std Dev. |  | 0.12111 | 0.075277 | 0.27325 | 0.17607 | 0.17512 | 0.22993 | 0.086948 | 0.13126 | 0.019664 | 0.35283 | 0.15336 | 0.10652 | 0.37639 | 0.52477 | 0.10825 | 0.11718 | 0.060229 | 0.12341 | 0.075104
| Rel.Std.Dev. |  | 0.2928% | 0.1831% | 0.6532% | 0.4268% | 0.4196% | 0.5544% | 0.2083% | 0.3369% | 0.04621% | 0.8638% | 0.3695% | 0.2572% | 0.8965% | 1.337% | 0.2639% | 0.2847% | 0.1459% | 0.298% | 0.1816%
| PDM3 |  | -0.09595% | -0.6997% | 1.031% | -0.3777% | 0.7896% | 0.1657% | 0.8298% | -5.896% | 2.778% | -1.359% | 0.2261% | 0.01675% | 1.393% | -5.192% | -0.9244% | -0.5937% | -0.3137% | 0.02077% | -0.108%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.26 | 0.31 | 0.99 | 0.29 | 1.03 | 6.48 | 3.21 | 1.41 | 0.36 | 0.13 | 1.66 | 5.69 | 0.93 | 0.56 | 0.24 | 0.13 | 0.01
| Analytical results for SnO2 in OREAS 702 (Indicative Value < 63 ppm) | [20 empty] | Indicative
SnO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 12 | 14 | 19 | 20 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 63 | < 63 | < 63 | < 63 | < 100 | < 130 | < 100 | < 200 | [11 empty] | 1
| 1 | 2 | < 63 | < 63 | < 63 | 63 | < 100 | < 130 | < 100 | < 200 | [11 empty] | 5
| 1 | 3 | < 63 | < 63 | < 63 | < 63 | < 100 | < 130 | < 100 | < 200 | 49.75 | [10 empty] | 16
| 1 | 4 | < 63 | < 63 | < 63 | < 63 | < 100 | < 130 | < 100 | < 200 | [11 empty] | < 63
| 1 | 5 | < 63 | < 63 | < 63 | < 63 | < 100 | < 130 | < 100 | < 200 | 10.05 | [10 empty] | 11
| 1 | 6 | < 63 | < 63 | < 63 | 63 | < 100 | < 130 | < 100 | < 200
| Mean |  |  |  |  | 63 |  |  |  |  | 29.9
| Median |  |  |  |  | 63 |  |  |  |  | 29.9
| Std Dev. |  |  |  |  | 0 |  |  |  |  | 28.0721
| Rel.Std.Dev. |  |  |  |  | 0% |  |  |  |  | 93.89%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 0 | 0.49 | 1.35 | 0.49 | 3.68 | 0.79
| Analytical results for SO3 in OREAS 702 (Certified Value 0.906 wt.%) | [20 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 11 | 12 | 13 | 14 | 20 | 22 | 24 | [6 empty] | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.899 | 0.924 | 0.924 | 0.874 | 0.932 | 0.8 | 0.75 | 0.79145 | 0.804 | 0.924 | 0.90102 | 0.89 | 0.9165 | [6 empty] | 1
| 1 | 2 | 0.899 | 0.949 | 0.899 | 0.899 | 0.921 | 0.79 | 0.76 | 0.79662 | 0.758 | 0.924 | 0.9163 | 0.88 | 0.914 | [6 empty] | #N/A
| 1 | 3 | 0.899 | 0.949 | 0.874 | 0.874 | 0.918 | 0.78 | 0.74 | 0.801 | 0.743 | 0.874 | 0.91938 | 0.88 | 0.9115 | [6 empty] | 16
| 1 | 4 | 0.899 | 0.924 | 0.899 | 0.899 | 0.934 | 0.78 | 0.75 | 0.79201 | 0.742 | 0.899 | 0.911 | 0.9 | 0.919 | [6 empty] | 0.90646
| 1 | 5 | 0.899 | 0.924 | 0.899 | 0.874 | 0.905 | 0.79 | 0.74 | 0.80347 | 0.739 | 0.874 | 0.90651 | 0.9 | 0.919 | [6 empty] | 26
| 1 | 6 | 0.924 | 0.924 | 0.899 | 0.899 | 0.924 | 0.78 | 0.75 | 0.80206 | 0.804 | 0.899 | 0.9273 | 0.87 | 0.914
| Mean |  | 0.90317 | 0.93233 | 0.899 | 0.8865 | 0.92233 | 0.78667 | 0.74833 | 0.79777 | 0.765 | 0.899 | 0.91358 | 0.88667 | 0.91567
| Median |  | 0.899 | 0.924 | 0.899 | 0.8865 | 0.9225 | 0.785 | 0.75 | 0.79881 | 0.7505 | 0.899 | 0.91365 | 0.885 | 0.91525
| Std Dev. |  | 0.010206 | 0.01291 | 0.015811 | 0.013693 | 0.01052 | 0.008165 | 0.0075277 | 0.0052122 | 0.030919 | 0.022361 | 0.0094211 | 0.012111 | 0.0030277
| Rel.Std.Dev. |  | 1.13% | 1.385% | 1.759% | 1.545% | 1.141% | 1.038% | 1.006% | 0.6533% | 4.042% | 2.487% | 1.031% | 1.366% | 0.3306%
| PDM3 |  | -0.3633% | 2.854% | -0.8229% | -2.202% | 1.751% | -13.22% | -17.44% | -11.99% | -15.61% | -0.8229% | 0.786% | -2.184% | 1.016%
| Z-Score (Absolute) |  | 0.17 | 1.35 | 0 | 0.51 | 0.95 | 4.55 | 6.1 | 4.1 | 5.43 | 0 | 0.59 | 0.5 | 0.67
| Analytical results for Sr in OREAS 702 (Certified Value 196 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 09 | 11 | 12 | 13 | 14 | 19 | 22 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 200 | 200 | 200 | 200 | 178 | 250 | < 100 | 190 | 200 | 149.3884 | < 80 | [8 empty] | 1
| 1 | 2 | 200 | 200 | 200 | 200 | 178 | 250 | < 100 | 190 | 200 | 155.0257 | < 80 | [8 empty] | #N/A
| 1 | 3 | 200 | 200 | 200 | 200 | 186 | 250 | < 100 | 180 | 200 | 160.6629 | < 80 | [8 empty] | 16
| 1 | 4 | 200 | 100 | 200 | 200 | 178 | 250 | < 100 | 200 | 200 | 138.1138 | < 80 | [8 empty] | 195.9718
| 1 | 5 | 200 | 200 | 200 | 200 | 186 | 250 | < 100 | 190 | 200 | 155.0257 | < 80 | [8 empty] | 27
| 1 | 6 | 200 | 200 | 200 | 200 | 186 | 250 | < 100 | 230 | 200 | 157.8443 | < 80
| Mean |  | 200 | 183.3333 | 200 | 200 | 182 | 250 |  | 196.6667 | 200 | 152.6768
| Median |  | 200 | 200 | 200 | 200 | 182 | 250 |  | 190 | 200 | 155.0257
| Std Dev. |  | 0 | 40.8248 | 0 | 0 | 4.3818 | 0 |  | 17.5119 | 0 | 8.055
| Rel.Std.Dev. |  | 0% | 22.27% | 0% | 0% | 2.408% | 0% |  | 8.904% | 0% | 5.276%
| PDM3 |  | 2.055% | -6.449% | 2.055% | 2.055% | -7.13% | 27.57% |  | 0.3546% | 2.055% | -22.09%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 0.17 | 0.17 | 0.75 | 2.88 | 7.42 | 0 | 0.17 | 2.22 | 7.81
| Analytical results for Th in OREAS 702 (Indicative Value 88 ppm) | [20 empty] | Indicative
Th | Round | Replicate | Lab | [18 empty] | 1
| No | No | 14 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | 90 | [18 empty] | 1
| 1 | 2 | 90 | [18 empty] | 1
| 1 | 3 | 90 | [18 empty] | 16
| 1 | 4 | 90 | [18 empty] | 87.8812
| 1 | 5 | 90 | [18 empty] | 7
| 1 | 6 | 90
| Mean |  | 90
| Median |  | 90
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.411%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 702 (Certified Value 0.386 wt.%) | [20 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 24 |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.42 | 0.37 | 0.4 | 0.39 | 0.39 | 0.39 | 0.4 | 0.39 | 0.38 | 0.373 | 0.4 | 0.367 | 0.37 | 0.38932 | 0.39717 | 0.366 | 0.38 | 0.3853 |  | 1
| 1 | 2 | 0.42 | 0.38 | 0.39 | 0.4 | 0.4 | 0.39 | 0.4 | 0.39 | 0.39 | 0.38451 | 0.4 | 0.367 | 0.36 | 0.38487 | 0.40772 | 0.3465 | 0.39 | 0.3887 |  | #N/A
| 1 | 3 | 0.42 | 0.37 | 0.4 | 0.39 | 0.4 | 0.38 | 0.4 | 0.39 | 0.38 | 0.39144 | 0.4 | 0.367 | 0.36 | 0.39626 | 0.40644 | 0.35352 | 0.38 | 0.387 |  | 16
| 1 | 4 | 0.42 | 0.38 | 0.39 | 0.4 | 0.41 | 0.39 | 0.4 | 0.39 | 0.37 | 0.38029 | 0.4 | 0.367 | 0.38 | 0.38971 | 0.40637 | 0.3707 | 0.38 | 0.3853 |  | 0.38611
| 1 | 5 | 0.42 | 0.37 | 0.39 | 0.4 | 0.41 | 0.39 | 0.4 | 0.39 | 0.39 | 0.38087 | 0.4 | 0.367 | 0.38 | 0.38649 | 0.40895 | 0.39687 | 0.38 | 0.3887 |  | 28
| 1 | 6 | 0.42 | 0.37 | 0.39 | 0.4 | 0.4 | 0.39 | 0.4 | 0.39 | 0.38 | 0.3944 | 0.41 | 0.367 | 0.37 | 0.39096 | 0.39239 | 0.363 | 0.38 | 0.3887
| Mean |  | 0.42 | 0.37333 | 0.39333 | 0.39667 | 0.40167 | 0.38833 | 0.4 | 0.39 | 0.38167 | 0.38409 | 0.40167 | 0.367 | 0.37 | 0.3896 | 0.40317 | 0.3661 | 0.38167 | 0.38728
| Median |  | 0.42 | 0.37 | 0.39 | 0.4 | 0.4 | 0.39 | 0.4 | 0.39 | 0.38 | 0.38269 | 0.4 | 0.367 | 0.37 | 0.38952 | 0.40641 | 0.3645 | 0.38 | 0.38785
| Std Dev. |  | 0 | 0.005164 | 0.005164 | 0.005164 | 0.0075277 | 0.0040825 | 6.0809e-17 | 6.0809e-17 | 0.0075277 | 0.0078535 | 0.0040825 | 0 | 0.0089443 | 0.0039583 | 0.0067415 | 0.017437 | 0.0040825 | 0.0016714
| Rel.Std.Dev. |  | 0% | 1.383% | 1.313% | 1.302% | 1.874% | 1.051% | 1.52e-14% | 1.559e-14% | 1.972% | 2.045% | 1.016% | 0% | 2.417% | 1.016% | 1.672% | 4.763% | 1.07% | 0.4316%
| PDM3 |  | 8.778% | -3.308% | 1.872% | 2.735% | 4.03% | 0.5766% | 3.598% | 1.008% | -1.15% | -0.5236% | 4.03% | -4.949% | -4.172% | 0.9051% | 4.42% | -5.182% | -1.15% | 0.3047%
| Z-Score (Absolute) |  | 2.56 | 1.21 | 0.4 | 0.67 | 1.08 | 0 | NA | 0.13 | 0.54 | 0.34 | 1.08 | 1.73 | 1.48 | 0.1 | 1.2 | 1.8 | 0.54 | 0.09
Results from laboratory 7 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for V2O5 in OREAS 702 (Indicative Value < 180 ppm) | [20 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 12 | 14 | 19 | 22 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | 180 | < 180 | < 180 | < 180 | 60 | 80 | < 180 | < 62.5 | < 200 | < 100 | [9 empty] | 1
| 1 | 2 | < 180 | < 180 | < 180 | 180 | 60 | 80 | < 180 | < 62.5 | < 200 | < 100 | [9 empty] | 2
| 1 | 3 | < 180 | < 180 | < 180 | < 180 | 70 | 80 | < 180 | < 62.5 | < 200 | < 100 | [9 empty] | 16
| 1 | 4 | < 180 | < 180 | < 180 | < 180 | 70 | < 50 | < 180 | < 62.5 | < 200 | < 100 | [9 empty] | < 180
| 1 | 5 | < 180 | 180 | < 180 | < 180 | 60 | 80 | < 180 | < 62.5 | < 200 | < 100 | [9 empty] | 8
| 1 | 6 | < 180 | < 180 | < 180 | < 180 | 80 | 80 | < 180 | < 62.5 | < 200 | < 100
| Mean |  | 180 | 180 |  | 180 | 66.6667 | 80
| Median |  | 180 | 180 |  | 180 | 65 | 80
| Std Dev. | [5 empty] | 8.165 | 0
| Rel.Std.Dev. | [5 empty] | 12.25% | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0 | 0.67 | 1.02 | 0.84 | 0 | 2.63 | 0.49 | 1.78
| Analytical results for WO3 in OREAS 702 (Certified Value 0.277 wt.%) | [20 empty] | Certified
WO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 20 | 24 |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.2787 | 0.2862 | 0.2736 | 0.2787 | 0.2724 | 0.2774 | 0.27 | 0.28 | 0.328 | 0.29795 | NA | 0.26 | 0.29457 | 0.298 | 0.28 | 0.27503 | 0.271 | 0.2762 |  | 1
| 1 | 2 | 0.2787 | 0.2812 | 0.2686 | 0.2787 | 0.2724 | 0.2774 | 0.27 | 0.28 | 0.315 | 0.30127 | 0.301 | 0.26 | 0.28864 | 0.298 | 0.25 | 0.27993 | 0.27405 | 0.2762 |  | #N/A
| 1 | 3 | 0.2799 | 0.285 | 0.2711 | 0.2799 | 0.2774 | 0.2724 | 0.27 | 0.28 | 0.315 | 0.29669 | 0.304 | 0.26 | 0.28978 | 0.297 | 0.27 | 0.27988 | 0.26955 | 0.2736 |  | 16
| 1 | 4 | 0.2787 | 0.2825 | 0.2774 | 0.2787 | 0.2774 | 0.2787 | 0.27 | 0.27 | 0.328 | 0.29787 | 0.305 | 0.27 | 0.29041 | 0.298 | 0.29 | 0.27877 | 0.271 | 0.2774 |  | 0.27664
| 1 | 5 | 0.2774 | 0.2837 | 0.2736 | 0.2762 | 0.2799 | 0.2774 | 0.27 | 0.27 | 0.315 | 0.29862 | 0.297 | 0.26 | 0.29116 | 0.301 | 0.28 | 0.27877 | 0.27331 | 0.2774 |  | 29
| 1 | 6 | 0.2787 | 0.285 | 0.2736 | 0.2812 | 0.2799 | 0.2787 | 0.27 | 0.28 | 0.328 | 0.30062 | 0.296 | 0.26 | 0.29167 | 0.297 | 0.25 | 0.27767 | 0.26985 | 0.2736
| Mean |  | 0.27868 | 0.28393 | 0.27298 | 0.2789 | 0.27657 | 0.277 | 0.27 | 0.27667 | 0.3215 | 0.29884 | 0.3006 | 0.26167 | 0.29104 | 0.29817 | 0.27 | 0.27834 | 0.27146 | 0.27573
| Median |  | 0.2787 | 0.28435 | 0.2736 | 0.2787 | 0.2774 | 0.2774 | 0.27 | 0.28 | 0.3215 | 0.29828 | 0.301 | 0.26 | 0.29078 | 0.298 | 0.275 | 0.27877 | 0.271 | 0.2762
| Std Dev. |  | 0.00079099 | 0.0018435 | 0.0029465 | 0.0016553 | 0.0034157 | 0.0023418 | 0 | 0.005164 | 0.0071204 | 0.001757 | 0.0040373 | 0.0040825 | 0.0020306 | 0.001472 | 0.016733 | 0.0018262 | 0.0018336 | 0.0017374
| Rel.Std.Dev. |  | 0.2838% | 0.6493% | 1.079% | 0.5935% | 1.235% | 0.8454% | 0% | 1.866% | 2.215% | 0.588% | 1.343% | 1.56% | 0.6977% | 0.4937% | 6.197% | 0.6561% | 0.6755% | 0.6301%
| PDM3 |  | 0.7375% | 2.635% | -1.323% | 0.8158% | -0.02766% | 0.129% | -2.401% | 0.00849% | 16.21% | 8.022% | 8.66% | -5.414% | 5.204% | 7.78% | -2.401% | 0.614% | -1.874% | -0.3289%
| Z-Score (Absolute) |  | 0.2 | 0.84 | 0.49 | 0.23 | 0.05 | 0 | 0.85 | 0.04 | NA | 2.66 | 2.87 | 1.87 | 1.71 | 2.58 | 0.85 | 0.16 | 0.67 | 0.15
| Analytical results for Y2O3 in OREAS 702 (Indicative Value 53 ppm) | [20 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 14 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 50 | < 50 | 50 | 60 | < 100 | [14 empty] | 1
| 1 | 2 | < 50 | < 50 | 50 | < 50 | < 100 | [14 empty] | 3
| 1 | 3 | < 50 | < 50 | 60 | 70 | < 100 | [14 empty] | 16
| 1 | 4 | < 50 | < 50 | 50 | 50 | < 100 | [14 empty] | 52.9167
| 1 | 5 | < 50 | < 50 | 60 | 60 | < 100 | [14 empty] | 9
| 1 | 6 | < 50 | < 50 | 50 | 50 | < 100
| Mean |  |  |  | 53.3333 | 58
| Median |  |  |  | 50 | 60
| Std Dev. |  |  |  | 5.164 | 8.3666
| Rel.Std.Dev. |  |  |  | 9.682% | 14.43%
| PDM3 |  |  |  | 0.7874% | 9.606%
| Z-Score (Absolute) |  | 5.06 | 5.06 | 0.67 | 0.51 | 0
| Analytical results for Zn in OREAS 702 (Certified Value 60 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 12 | 14 | 16 | 19 | 20 | 22 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 70 | < 50 | 50 | 60 | 80 | < 100 | < 50 | < 100 | < 160 | 65.231 | 48 | [8 empty] | 1
| 1 | 2 | 60 | < 50 | 60 | 70 | 80 | < 100 | < 50 | < 100 | < 160 | 64.267 | 56 | [8 empty] | 22
| 1 | 3 | 70 | < 50 | 60 | 60 | 90 | < 100 | < 50 | < 100 | < 160 | 71.939 | 56 | [8 empty] | 16
| 1 | 4 | 60 | < 50 | 60 | 70 | 80 | < 100 | < 50 | < 100 | < 160 | 64.267 | 48 | [8 empty] | 60.015
| 1 | 5 | 60 | < 50 | 60 | 60 | 80 | < 100 | < 50 | < 100 | < 160 | 64.589 | 48 | [8 empty] | 30
| 1 | 6 | 60 | < 50 | 50 | 60 | 80 | < 100 | < 50 | < 100 | < 160 | 64.267 | 56
| Mean |  | 63.3333 |  | 56.6667 | 63.3333 | 81.6667 |  |  |  |  | 65.76 | 52
| Median |  | 60 |  | 60 | 60 | 80 |  |  |  |  | 64.428 | 52
| Std Dev. |  | 5.164 |  | 5.164 | 5.164 | 4.0825 |  |  |  |  | 3.0502 | 4.3818
| Rel.Std.Dev. |  | 8.154% |  | 9.113% | 8.154% | 4.999% |  |  |  |  | 4.638% | 8.427%
| PDM3 |  | 5.529% |  | -5.579% | 5.529% | 36.08% |  |  |  |  | 9.573% | -13.35%
| Z-Score (Absolute) |  | 0.67 | 3.2 | 0 | 0.67 | 2.53 | 0.67 | 3.2 | 0.67 | 2.39 | 0.92 | 0.45
| Analytical results for Zr in OREAS 702 (Indicative Value 155 ppm) | [20 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 11 | 12 | 14 | 19 | 22 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 200 | 100 | 200 | 100 | 150 | < 100 | 150 | 170.2741 | 200 | [10 empty] | 1
| 1 | 2 | 200 | 100 | 200 | 100 | 150 | < 100 | 150 | 165.3386 | 192 | [10 empty] | 4
| 1 | 3 | 200 | 100 | 100 | 100 | 150 | < 100 | 150 | 170.2741 | 192 | [10 empty] | 16
| 1 | 4 | 200 | 100 | 200 | 100 | 150 | < 100 | 150 | 167.8063 | 192 | [10 empty] | 155.1747
| 1 | 5 | 200 | 100 | 200 | 100 | 150 | < 100 | 150 | 162.8708 | 207 | [10 empty] | 10
| 1 | 6 | 200 | 100 | 200 | 100 | 150 | < 100 | 150 | 165.3386 | 185
| Mean |  | 200 | 100 | 183.3333 | 100 | 150 |  | 150 | 166.9837 | 194.6667
| Median |  | 200 | 100 | 200 | 100 | 150 |  | 150 | 166.5725 | 192
| Std Dev. |  | 0 | 0 | 40.8248 | 0 | 0 |  | 0 | 2.9886 | 7.6855
| Rel.Std.Dev. |  | 0% | 0% | 22.27% | 0% | 0% |  | 0% | 1.79% | 3.948%
| PDM3 |  | 28.89% | -35.56% | 18.15% | -35.56% | -3.335% |  | -3.335% | 7.61% | 25.45%
| Z-Score (Absolute) |  | 0.75 | 0.69 | 0.51 | 0.69 | 0 | 1.41 | 0 | 0.27 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C558: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C576: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C649: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C667: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 702 (Indicative Value 0.312 wt.%) | [22 empty] | Indicative
H2O- | Round | Replicate | Lab | [20 empty] | 1
| No | No | 07 | [20 empty] | wt.%
|  |  | LOI*TGA | [20 empty] | 3
3
| 1 | 1 | 0.31 | [20 empty] | 1
| 1 | 2 | 0.3 | [20 empty] | 7
| 1 | 3 | 0.32 | [20 empty] | 16
| 1 | 4 | 0.3 | [20 empty] | 0.31167
| 1 | 5 | 0.32 | [20 empty] | 13
| 1 | 6 | 0.32
| Mean |  | 0.31167
| Median |  | 0.315
| Std Dev. |  | 0.0098319
| Rel.Std.Dev. |  | 3.155%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 702 (Certified Value 16.06 wt.%) | [22 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | BF*OES/MS | BF*XRF | LOI*Furnace | LOI*Furnace | OxBF*XRF | PF*OES/MS | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | 2
3
| 1 | 1 | 15.88 | 16.07 | 16.1 | 16.01 | 16.02 | 16.28 | 16 | 16.03 | 16.14 | 15.98 | 16.1 | 15.91 | 16.45 | 16.11 | 16.0888 | 16.63 | 16.0196 | 16.0478 | 16.11 | 16.163 | 16.04 | 1
| 1 | 2 | 15.76 | 16.08 | 16.08 | 16 | 16.05 | 16.25 | 16 | 16.16 | 16.13 | 16.02 | 16 | 15.93 | 16.48 | 16.08 | 16.154 | 16.72 | 16.0179 | 16.0525 | 16.11 | 16.179 | 16.06 | #N/A
| 1 | 3 | 15.74 | 16.02 | 16.13 | 16.01 | 16.03 | 16.32 | 16 | 16.02 | 16.12 | 15.97 | 16.1 | 15.95 | 16.4 | 16.08 | 16.2146 | 16.75 | 16.0374 | 16.1159 | 16.12 | 16.16 | 16.05 | 16
| 1 | 4 | 15.54 | 16.03 | 16.11 | 16.04 | 16.05 | 16.46 | 16 | 15.97 | 16.15 | 16.01 | 16.1 | 15.9 | 16.59 | 16.13 | 16.205 | 16.73 | 16.0247 | 16.0976 | 16.13 | 16.169 | 16 | 16.0603
| 1 | 5 | 15.53 | 16.01 | 16.15 | 16 | 16.04 | 16.41 | 16.1 | 15.95 | 16.13 | 15.925 | 16.1 | 15.86 | 16.45 | 16.11 | 16.1383 | 16.58 | 16.033 | 16.0998 | 16.13 | 16.133 | 15.99 | 32
| 1 | 6 | 15.53 | 15.95 | 15.99 | 16 | 16.06 | 16.26 | 16 | 16.08 | 16.14 | 16.07 | 16.1 | 15.82 | 16.43 | 16.07 | 16.0243 | 16.57 | 16.0375 | 16.2052 | 16.12 | 16.233 | 15.97
| Mean |  | 15.6633 | 16.0267 | 16.0933 | 16.01 | 16.0417 | 16.33 | 16.0167 | 16.035 | 16.135 | 15.9958 | 16.0833 | 15.895 | 16.4667 | 16.0967 | 16.1375 | 16.6633 | 16.0284 | 16.1031 | 16.12 | 16.1728 | 16.0183
| Median |  | 15.64 | 16.025 | 16.105 | 16.005 | 16.045 | 16.3 | 16 | 16.025 | 16.135 | 15.995 | 16.1 | 15.905 | 16.45 | 16.095 | 16.1461 | 16.675 | 16.0288 | 16.0987 | 16.12 | 16.166 | 16.02
| Std Dev. |  | 0.15029 | 0.046762 | 0.056095 | 0.015492 | 0.01472 | 0.086255 | 0.040825 | 0.076616 | 0.010488 | 0.049439 | 0.040825 | 0.047645 | 0.065929 | 0.023381 | 0.072055 | 0.079917 | 0.0088056 | 0.056997 | 0.0089443 | 0.033229 | 0.03656
| Rel.Std.Dev. |  | 0.9595% | 0.2918% | 0.3486% | 0.09676% | 0.09176% | 0.5282% | 0.2549% | 0.4778% | 0.065% | 0.3091% | 0.2538% | 0.2997% | 0.4004% | 0.1453% | 0.4465% | 0.4796% | 0.05494% | 0.3539% | 0.05549% | 0.2055% | 0.2282%
| PDM3 |  | -2.471% | -0.2091% | 0.206% | -0.3129% | -0.1157% | 1.68% | -0.2714% | -0.1572% | 0.4654% | -0.4011% | 0.1437% | -1.029% | 2.531% | 0.2267% | 0.481% | 3.755% | -0.1986% | 0.267% | 0.372% | 0.701% | -0.261%
| Z-Score (Absolute) |  | 5 | 0.67 | 0.12 | 0.87 | 0.5 | 2.94 | 0.79 | 0.58 | 0.61 | 1.04 | 0 | 2.24 | 4.56 | 0.16 | 0.64 | 6.9 | 0.65 | 0.24 | 0.44 | 1.07 | 0.77
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ISE
| Analytical results for F in OREAS 702 (Certified Value 0.299 wt.%) | [18 empty] | Certified
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 21 | 22 | 23 | 24 | wt.%
|  |  | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | 3
3
| 1 | 1 | 0.341 | 0.312 | 0.369 | 0.26 | 0.3189 | 0.31 | 0.3233 | 0.2856 | 0.334 | 0.2268 | 0.2849 | 0.33794 | 0.1959 | 0.31 | 0.25957 | 0.33 | 0.28 | 1
| 1 | 2 | 0.343 | 0.312 | 0.356 | 0.26 | 0.3375 | 0.22 | 0.3019 | 0.2786 | 0.3256 | 0.2275 | 0.3079 | 0.31838 | 0.2009 | 0.3 | 0.2534 | 0.31 | 0.32 | 4
| 1 | 3 | 0.349 | 0.31 | 0.344 | 0.26 | 0.3272 | 0.22 | 0.3102 | 0.2877 | 0.3282 | 0.2286 | 0.26986 | 0.30266 | 0.1989 | 0.31 | 0.26083 | 0.33 | 0.26 | 16
| 1 | 4 | 0.348 | 0.295 | 0.336 | 0.26 | 0.3279 | 0.29 | 0.3152 | 0.2856 | 0.3244 | 0.2304 | 0.31654 | 0.35342 | 0.2191 | 0.3 | 0.25216 | 0.32 | 0.3 | 0.29949
| 1 | 5 | 0.347 | 0.307 | 0.351 | 0.26 | 0.3202 | 0.21 | 0.3095 | 0.2919 | 0.3005 | 0.2259 | 0.31978 | 0.35142 | 0.2491 | 0.3 | 0.25957 | 0.32 | 0.34 | 34
| 1 | 6 | 0.348 | 0.316 | 0.333 | 0.26 | 0.3321 | 0.22 | 0.3258 | 0.2884 | 0.3173 | 0.2299 | 0.27661 | 0.33858 | 0.1975 | 0.29 | 0.23345 | 0.32 | 0.3
| Mean |  | 0.346 | 0.30867 | 0.34817 | 0.26 | 0.3273 | 0.245 | 0.31432 | 0.2863 | 0.32167 | 0.22818 | 0.29593 | 0.33373 | 0.21023 | 0.30167 | 0.25316 | 0.32167 | 0.3
| Median |  | 0.3475 | 0.311 | 0.3475 | 0.26 | 0.32755 | 0.22 | 0.3127 | 0.28665 | 0.325 | 0.22805 | 0.2964 | 0.33826 | 0.1999 | 0.3 | 0.25648 | 0.32 | 0.3
| Std Dev. |  | 0.0032249 | 0.0073121 | 0.013408 | 0 | 0.0070487 | 0.043243 | 0.0090296 | 0.0044272 | 0.011703 | 0.001768 | 0.021501 | 0.019712 | 0.020844 | 0.0075277 | 0.010302 | 0.0075277 | 0.028284
| Rel.Std.Dev. |  | 0.9321% | 2.369% | 3.851% | 0% | 2.154% | 17.65% | 2.873% | 1.546% | 3.638% | 0.7748% | 7.265% | 5.906% | 9.915% | 2.495% | 4.069% | 2.34% | 9.428%
| PDM3 |  | 15.53% | 3.066% | 16.26% | -13.18% | 9.288% | -18.19% | 4.952% | -4.403% | 7.407% | -23.81% | -1.186% | 11.44% | -29.8% | 0.7285% | -15.47% | 7.407% | 0.1719%
| Z-Score (Absolute) |  | 1.17 | 0.18 | 1.22 | 1.1 | 0.67 | 1.49 | 0.33 | 0.4 | 0.53 | 1.93 | 0.15 | 0.84 | 2.41 | 0 | 1.28 | 0.53 | 0.04
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 702 (Indicative Value 4.39 wt.%) | [25 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 09 | 13 | 16 | [21 empty] | wt.%
|  |  | IRC | IRC | IRC | [21 empty] | 2
2
| 1 | 1 | 4.35 | 4.65 | 4.24 | [21 empty] | 1
| 1 | 2 | 4.23 | 4.68 | 4.21 | [21 empty] | 9
| 1 | 3 | 4.23 | 4.66 | 4.25 | [21 empty] | 16
| 1 | 4 | 4.28 | 4.54 | 4.3 | [21 empty] | 4.3889
| 1 | 5 | 4.3 | 4.6 | 4.22 | [21 empty] | 15
| 1 | 6 | 4.3 | 4.73 | 4.23
| Mean |  | 4.2817 | 4.6433 | 4.2417
| Median |  | 4.29 | 4.655 | 4.235
| Std Dev. |  | 0.046224 | 0.065929 | 0.031885
| Rel.Std.Dev. |  | 1.08% | 1.42% | 0.7517%
| PDM3 |  | -2.443% | 5.797% | -3.354%
| Z-Score (Absolute) |  | 0 | 6.1 | 0.67
| Analytical results for S in OREAS 702 (Certified Value 0.349 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.35 | 0.38 | 0.34 | 0.34 | 0.35 | 0.35 | 0.35 | 0.32 | 0.23 | 0.35413 | 0.397 | 0.35 | 0.37 | 0.327 | 0.35 | 0.295 | 0.36 | 0.35084 | 0.39 | 0.346 | 0.36 | 0.3 | 0.3 | 0.3636 | 1
| 1 | 2 | 0.34 | 0.4 | 0.34 | 0.35 | 0.34 | 0.36 | 0.35 | 0.33 | 0.23 | 0.36076 | 0.392 | 0.37 | 0.36 | 0.326 | 0.38 | 0.29 | 0.362 | 0.35732 | 0.4 | 0.357 | 0.35 | 0.4 | 0.33 | 0.3606 | #N/A
| 1 | 3 | 0.35 | 0.39 | 0.34 | 0.35 | 0.35 | 0.35 | 0.35 | 0.32 | 0.24 | 0.35254 | 0.387 | 0.35 | 0.36 | 0.308 | 0.35 | 0.324 | 0.361 | 0.3496 | 0.4 | 0.334 | 0.35 | 0.3 | 0.34 | 0.3519 | 16
| 1 | 4 | 0.34 | 0.39 | 0.34 | 0.34 | 0.34 | 0.37 | 0.35 | 0.32 | 0.24 | 0.35844 | 0.397 | 0.34 | 0.36 | 0.337 | 0.37 | 0.295 | 0.36 | 0.35155 | 0.4 | 0.347 | 0.35 | 0.4 | 0.33 | 0.3627 | 0.34852
| 1 | 5 | 0.35 | 0.38 | 0.35 | 0.35 | 0.35 | 0.36 | 0.36 | 0.31 | 0.24 | 0.35836 | 0.4002 | 0.36 | 0.37 | 0.327 | 0.38 | 0.276 | 0.356 | 0.34281 | 0.41 | 0.345 | 0.35 | 0.3 | 0.34 | 0.354 | 36
| 1 | 6 | 0.35 | 0.38 | 0.33 | 0.33 | 0.34 | 0.35 | 0.36 | 0.33 | 0.23 | 0.353 | 0.403 | 0.35 | 0.37 | 0.316 | 0.36 | 0.29 | 0.358 | 0.32636 | 0.4 | 0.341 | 0.35 | 0.3 | 0.35 | 0.3614
| Mean |  | 0.34667 | 0.38667 | 0.34 | 0.34333 | 0.345 | 0.35667 | 0.35333 | 0.32167 | 0.235 | 0.3562 | 0.39603 | 0.35333 | 0.365 | 0.3235 | 0.365 | 0.295 | 0.3595 | 0.34641 | 0.4 | 0.345 | 0.35167 | 0.33333 | 0.33167 | 0.35903
| Median |  | 0.35 | 0.385 | 0.34 | 0.345 | 0.345 | 0.355 | 0.35 | 0.32 | 0.235 | 0.35624 | 0.397 | 0.35 | 0.365 | 0.3265 | 0.365 | 0.2925 | 0.36 | 0.35022 | 0.4 | 0.3455 | 0.35 | 0.3 | 0.335 | 0.361
| Std Dev. |  | 0.005164 | 0.008165 | 0.0063246 | 0.008165 | 0.0054772 | 0.008165 | 0.005164 | 0.0075277 | 0.0054772 | 0.0034172 | 0.0057556 | 0.010328 | 0.0054772 | 0.010095 | 0.013784 | 0.015824 | 0.0021679 | 0.010865 | 0.0063246 | 0.0075631 | 0.0040825 | 0.05164 | 0.017224 | 0.0048698
| Rel.Std.Dev. |  | 1.49% | 2.112% | 1.86% | 2.378% | 1.588% | 2.289% | 1.462% | 2.34% | 2.331% | 0.9593% | 1.453% | 2.923% | 1.501% | 3.12% | 3.776% | 5.364% | 0.603% | 3.136% | 1.581% | 2.192% | 1.161% | 15.49% | 5.193% | 1.356%
| PDM3 |  | -0.5313% | 10.95% | -2.444% | -1.488% | -1.009% | 2.338% | 1.382% | -7.705% | -32.57% | 2.206% | 13.63% | 1.382% | 4.729% | -7.178% | 4.729% | -15.36% | 3.151% | -0.6041% | 14.77% | -1.009% | 0.9034% | -4.357% | -4.835% | 3.017%
| Z-Score (Absolute) |  | 0.4 | 2.83 | 0.94 | 0.67 | 0.54 | 0.4 | 0.13 | 2.43 | 9.44 | 0.37 | 3.59 | 0.13 | 1.08 | 2.28 | 1.08 | 4.59 | 0.63 | 0.43 | 3.91 | 0.54 | 0 | NA | 1.62 | 0.6
Results from laboratory 24 were removed due to their 0.1 wt.% reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 702 (Indicative Value < 1 ppm) | [22 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 06 | 17 | 18 | 25 | [16 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | [16 empty] | 2
2
| 1 | 1 | < 5 | < 0.5 | < 1 | 0.15 | < 10 | [16 empty] | 1
| 1 | 2 | < 5 | < 0.5 | < 1 | 0.11 | < 10 | [16 empty] | 11
| 1 | 3 | < 5 | < 0.5 | < 1 | 0.14 | < 10 | [16 empty] | 16
| 1 | 4 | < 5 | < 0.5 | < 1 | 0.11 | < 10 | [16 empty] | < 1
| 1 | 5 | < 5 | < 0.5 | < 1 | 0.13 | < 10 | [16 empty] | 17
| 1 | 6 | < 5 | < 0.5 | < 1 | 0.13 | < 10
| Mean |  |  |  |  | 0.12833
| Median |  |  |  |  | 0.13
| Std Dev. |  |  |  |  | 0.016021
| Rel.Std.Dev. |  |  |  |  | 12.48%
PDM3
| Z-Score (Absolute) |  | 3.63 | 0.45 | 0 | 0.67 | 8.16
| Analytical results for Al2O3 in OREAS 702 (Certified Value 9.56 wt.%) | [22 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 25 | 26 | [5 empty] | wt.%
|  |  | BF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | [5 empty] | 2
3
| 1 | 1 | 9.5 | 9.51 | 10.07 | 9.3122 | 9.788 | 9.5201 | 9.82 | 9.523 | 9.7196 | 9.9635 | 9.693 | 9.5282 | 9.33 | 9.372 | 9.429 | 9.4409 | [5 empty] | 1
| 1 | 2 | 9.36 | 9.45 | 10.05 | 9.302 | 9.769 | 9.562 | 9.67 | 9.542 | 9.7045 | 9.87 | 9.75 | 9.5317 | 9.32 | 9.882 | 9.863 | 9.5676 | [5 empty] | #N/A
| 1 | 3 | 9.44 | 9.56 | 9.71 | 9.559 | 9.75 | 9.512 | 9.73 | 9.901 | 9.5008 | 10.0517 | 9.693 | 9.5751 | 9.6 | 10.128 | 9.24 | 9.5278 | [5 empty] | 16
| 1 | 4 | 9.59 | 9.6 | 10.13 | 9.7634 | 9.693 | 9.5357 | 9.66 | 9.372 | 9.7016 | 9.8616 | 9.655 | 9.5177 | 9.16 | 10.166 | 9.636 | 9.5223 | [5 empty] | 9.5629
| 1 | 5 | 9.43 | 9.62 | 9.9 | 9.5112 | 9.674 | 9.5573 | 9.66 | 9.58 | 9.7032 | 9.9561 | 9.674 | 9.5396 | 8.99 | 10.128 | 9.391 | 9.4078 | [5 empty] | 38
| 1 | 6 | 9.28 | 9.66 | 10.01 | 9.5518 | 9.693 | 9.5241 | 9.71 | 9.523 | 9.7904 | 10.2911 | 9.712 | 9.4718 | 9.34 | 10.033 | 9.391 | 9.6601
| Mean |  | 9.4333 | 9.5667 | 9.9783 | 9.4999 | 9.7278 | 9.5352 | 9.7083 | 9.5735 | 9.6867 | 9.999 | 9.6962 | 9.5274 | 9.29 | 9.9515 | 9.4917 | 9.5211
| Median |  | 9.435 | 9.58 | 10.03 | 9.5315 | 9.7215 | 9.5299 | 9.69 | 9.5325 | 9.7038 | 9.9598 | 9.693 | 9.5299 | 9.325 | 10.0805 | 9.41 | 9.5251
| Std Dev. |  | 0.10764 | 0.076855 | 0.15211 | 0.17329 | 0.047182 | 0.020466 | 0.061779 | 0.1755 | 0.09718 | 0.15918 | 0.032726 | 0.033578 | 0.20396 | 0.30186 | 0.22192 | 0.090341
| Rel.Std.Dev. |  | 1.141% | 0.8034% | 1.524% | 1.824% | 0.485% | 0.2146% | 0.6364% | 1.833% | 1.003% | 1.592% | 0.3375% | 0.3524% | 2.195% | 3.033% | 2.338% | 0.9489%
| PDM3 |  | -1.355% | 0.03939% | 4.344% | -0.6584% | 1.725% | -0.2897% | 1.521% | 0.1108% | 1.294% | 4.56% | 1.394% | -0.3717% | -2.854% | 4.064% | -0.7449% | -0.4372%
| Z-Score (Absolute) |  | 0.76 | 0.02 | 2.27 | 0.39 | 0.88 | 0.19 | 0.77 | 0.02 | 0.65 | 2.38 | 0.7 | 0.24 | 1.56 | 2.12 | 0.44 | 0.27
| Analytical results for As in OREAS 702 (Indicative Value 5.85 ppm) | [22 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 06 | 07 | 10 | 14 | 15 | 17 | 18 | 22 | 23 | 24 | 25 | 26 | [8 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [8 empty] | 2
2
| 1 | 1 | 5 | < 5 | 3.9 | < 20 | 11.75 | < 30 | < 10 | 2.7 | 23.85 | < 30 | < 100 | 16 | < 100 | [8 empty] | 1
| 1 | 2 | 6 | < 5 | 3.9 | < 20 | 11.02 | < 30 | < 10 | 3.1 | 24.51 | < 30 | < 100 | 11 | < 100 | [8 empty] | 12
| 1 | 3 | 5 | < 5 | 2.9 | < 20 | 11.05 | < 30 | < 10 | 3.3 | 22.8 | < 30 | < 100 | < 10 | < 100 | [8 empty] | 16
| 1 | 4 | 5 | < 5 | 4 | < 20 | 10.68 | < 30 | < 10 | 2.7 | 21.84 | < 30 | < 100 | < 10 | < 100 | [8 empty] | 5.8487
| 1 | 5 | 6 | < 5 | 3.2 | < 20 | 11.13 | < 30 | < 10 | 2.8 | 23.08 | < 30 | < 100 | < 10 | < 100 | [8 empty] | 18
| 1 | 6 | 6 | < 5 | 3.1 | < 20 | 11.03 | < 30 | < 10 | 3 | 22.27 | < 30 | < 100 | < 10 | < 100
| Mean |  | 5.5 |  | 3.5 |  | 11.11 |  |  | 2.9333 | 23.0583 |  |  | 13.5
| Median |  | 5.5 |  | 3.55 |  | 11.04 |  |  | 2.9 | 22.94 |  |  | 13.5
| Std Dev. |  | 0.54772 |  | 0.4858 |  | 0.35014 |  |  | 0.24221 | 0.99066 |  |  | 3.5355
| Rel.Std.Dev. |  | 9.959% |  | 13.88% |  | 3.152% |  |  | 8.257% | 4.296% |  |  | 26.19%
| PDM3 |  | -5.961% |  | -40.16% |  | 89.96% |  |  | -49.85% | 294.2% |  |  | 130.8%
| Z-Score (Absolute) |  | 0.61 | 1.01 | 0.88 | 0 | 0.15 | 0.67 | 0.67 | 0.95 | 1.76 | 0.67 | 5.39 | 0.29 | 5.39
| Analytical results for B in OREAS 702 (Indicative Value 29.5 ppm) | [22 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 09 | 10 | 14 | 22 | 23 | 24 | [15 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | [15 empty] | 1
1
| 1 | 1 | < 30 | < 50 | 56 | 7.16 | < 20 | < 50 | [15 empty] | 1
| 1 | 2 | < 30 | < 50 | 57 | 5.78 | < 20 | < 50 | [15 empty] | 13
| 1 | 3 | < 30 | < 50 | 46 | 5.68 | < 20 | < 50 | [15 empty] | 16
| 1 | 4 | < 30 | < 50 | 55 | 5.47 | < 20 | < 50 | [15 empty] | 29.4675
| 1 | 5 | < 30 | < 50 | 48 | 5.49 | < 20 | < 50 | [15 empty] | 19
| 1 | 6 | < 30 | < 50 | 56 | 6.03 | < 20 | < 50
| Mean |  |  |  | 53 | 5.935
| Median |  |  |  | 55.5 | 5.73
| Std Dev. |  |  |  | 4.7329 | 0.6345
| Rel.Std.Dev. |  |  |  | 8.93% | 10.69%
| PDM3 |  |  |  | 79.86% | -79.86%
| Z-Score (Absolute) |  | 0.45 | 0.45 | 2.97 | 1.26 | 0.9 | 0.45
| Analytical results for Ba in OREAS 702 (Certified Value 650 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  | 0
0
| 1 | 1 | 648 | 638 | 658 | 680 | 630 | 663 | 655 | 641 | 688 | 648 | 646 | 626 | 655.2467 | 635 | 660.5221 | 735.71 | 691 | 650 | 668 | 646.58 |  | 1
| 1 | 2 | 645 | 623 | 665 | 671 | 630 | 656 | 628.2 | 648 | 681 | 645 | 667 | 626 | 647.0515 | 628 | 647.4554 | 726.75 | 706 | 650 | 666 | 641.44 |  | #N/A
| 1 | 3 | 657 | 613 | 669 | 678 | 630 | 659 | 646.8 | 652 | 645 | 649 | 674 | 621 | 647.2288 | 638 | 651.2463 | 733.38 | 652 | 650 | 652 | 641.46 |  | 16
| 1 | 4 | 644 | 641 | 661 | 652 | 630 | 664 | 641.3 | 676 | 675 | 659 | 641 | 619 | 637.6471 | 629 | 644.1095 | 750.23 | 652 | 650 | 660 | 635.67 |  | 649.9487
| 1 | 5 | 669 | 615 | 654 | 689 | 630 | 666 | 651.7 | 647 | 674 | 649 | 642 | 617 | 662.0174 | 636 | 647.1076 | 724.78 | 691 | 650 | 724 | 647.1 |  | 39
| 1 | 6 | 643 | 607 | 650 | 676 | 630 | 659 | 641.1 | 647 | 673 | 651 | 647 | 607 | 653.7868 | 635 | 653.4424 | 733.3 | 682 | 650 | 660 | 645.39
| Mean |  | 651 | 622.8333 | 659.5 | 674.3333 | 630 | 661.1667 | 644.0167 | 651.8333 | 672.6667 | 650.1667 | 652.8333 | 619.3333 | 650.4964 | 633.5 | 650.6472 | 734.025 | 679 | 650 | 671.6667 | 642.94
| Median |  | 646.5 | 619 | 659.5 | 677 | 630 | 661 | 644.05 | 647.5 | 674.5 | 649 | 646.5 | 620 | 650.5078 | 635 | 649.3508 | 733.34 | 686.5 | 650 | 663 | 643.425
| Std Dev. |  | 10.1784 | 13.92 | 7.0071 | 12.4365 | 0 | 3.7639 | 9.5233 | 12.3518 | 14.6788 | 4.7504 | 14.0487 | 7.0616 | 8.4116 | 4.0373 | 5.8505 | 8.9975 | 22.289 | 0 | 26.2425 | 4.3325
| Rel.Std.Dev. |  | 1.564% | 2.235% | 1.062% | 1.844% | 0% | 0.5693% | 1.479% | 1.895% | 2.182% | 0.7306% | 2.152% | 1.14% | 1.293% | 0.6373% | 0.8992% | 1.226% | 3.283% | 0% | 3.907% | 0.6739%
| PDM3 |  | 0.1617% | -4.172% | 1.47% | 3.752% | -3.069% | 1.726% | -0.9127% | 0.29% | 3.495% | 0.03353% | 0.4438% | -4.71% | 0.08426% | -2.531% | 0.1075% | 12.94% | 4.47% | 0.007886% | 3.341% | -1.078%
| Z-Score (Absolute) |  | 0.01 | 1.98 | 0.62 | 1.67 | 1.69 | 0.73 | 0.48 | 0.07 | 1.55 | 0.05 | 0.14 | 2.23 | 0.02 | 1.23 | 0.01 | 5.9 | 2 | 0.06 | 1.48 | 0.56
| Analytical results for Be in OREAS 702 (Certified Value 5.63 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [5 empty] | 2
2
| 1 | 1 | 4.1 | 6 | 3.8 | 3.6 | 6 | 5.2267 | 6 | 6.4 | 6 | 5.4449 | 5.66 | 6.11 | 6 | 6 | 7 | < 10 | [5 empty] | 1
| 1 | 2 | 4.2 | 6 | 3.6 | 4.3 | 5 | 5.2417 | 6 | 6.5 | 5 | 5.2152 | 5.69 | 6.14 | 6 | 6 | 7 | < 10 | [5 empty] | #N/A
| 1 | 3 | 4.5 | 6 | 6.2 | 4.4 | 6 | 5.2634 | 6 | 6.7 | 6 | 5.035 | 5.75 | 6.33 | 6 | 7 | 7 | < 10 | [5 empty] | 16
| 1 | 4 | 4.5 | 6 | 6.2 | 5 | 5 | 5.2586 | 6 | 6.2 | 5 | 5.1251 | 5.77 | 6.13 | 6 | 6 | 7 | < 10 | [5 empty] | 5.6298
| 1 | 5 | 5.1 | 6 | 8.8 | 5 | 6 | 5.2393 | 6 | 6 | 5 | 5.1315 | 5.64 | 5.93 | 6 | 6 | 7 | < 10 | [5 empty] | 40
| 1 | 6 | 4.4 | 5 | 4.7 | 4.8 | 6 | 5.2884 | 6 | 6.3 | 7 | 5.1061 | 5.62 | 5.94 | 6 | 6 | 5 | < 10
| Mean |  | 4.4667 | 5.8333 | 5.55 | 4.5167 | 5.6667 | 5.253 | 6 | 6.35 | 5.6667 | 5.1763 | 5.6883 | 6.0967 | 6 | 6.1667 | 6.6667
| Median |  | 4.45 | 6 | 5.45 | 4.6 | 6 | 5.2501 | 6 | 6.35 | 5.5 | 5.1283 | 5.675 | 6.12 | 6 | 6 | 7
| Std Dev. |  | 0.35024 | 0.40825 | 1.9491 | 0.53821 | 0.5164 | 0.021933 | 0 | 0.2429 | 0.8165 | 0.14367 | 0.06047 | 0.14828 | 0 | 0.40825 | 0.8165
| Rel.Std.Dev. |  | 7.841% | 6.999% | 35.12% | 11.92% | 9.113% | 0.4175% | 0% | 3.825% | 14.41% | 2.776% | 1.063% | 2.432% | 0% | 6.62% | 12.25%
| PDM3 |  | -20.66% | 3.614% | -1.418% | -19.77% | 0.6541% | -6.694% | 6.575% | 12.79% | 0.6541% | -8.056% | 1.039% | 8.292% | 6.575% | 9.535% | 18.42%
| Z-Score (Absolute) |  | 1.94 | 0.25 | 0.2 | 1.86 | 0.02 | 0.68 | 0.52 | 1.08 | 0.02 | 0.8 | 0.02 | 0.67 | 0.52 | 0.78 | 1.58 | 1.08
| Analytical results for Bi in OREAS 702 (Certified Value 275 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 06 | 07 | 09 | 10 | 11 | 14 | 15 | 17 | 18 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  |  |  | 0
0
| 1 | 1 | 227 | 275 | 264 | 281 | 277.61 | 273 | 280 | 313.5644 | 261 | 268.5 | 273.4452 | 291 | 317.142 | 253 | 278 | 296.2 | 267.0273 |  |  |  |  | 1
| 1 | 2 | 227 | 267 | 254 | 279 | 272.87 | 255 | 282 | 312.8894 | 277 | 269.7 | 272.5893 | 291.62 | 316.1466 | 257 | 269.5 | 280.3 | 262.9211 |  |  |  |  | 22
| 1 | 3 | 227 | 265 | 258 | 283 | 282.69 | 251 | 279 | 313.3794 | 282 | 265.8 | 272.9642 | 292.21 | 323.2152 | 261 | 268.5 | 297.7 | 269.4766 |  |  |  |  | 16
| 1 | 4 | 223 | 269 | 240 | 283 | 277.6 | 299 | 282 | 312.4944 | 265 | 267 | 269.564 | 290.77 | 316.638 | 265 | 271.5 | 296.4 | 260.416 |  |  |  |  | 275.1158
| 1 | 5 | 220 | 266 | 229 | 277 | 279.85 | 276 | 278 | 311.3244 | 268 | 269.8 | 280.6604 | 292.26 | 303.8868 | 254 | 273.5 | 302.6 | 264.2156 |  |  |  |  | 41
| 1 | 6 | 235 | 279 | 234 | 283 | 276.82 | 265 | 283 | 312.3744 | 261 | 263.9 | 274.8899 | 292.92 | 314.6724 | 260 | 280.5 | 295.6 | 283.644
| Mean |  | 226.5 | 270.1667 | 246.5 | 281 | 277.9067 | 269.8333 | 280.6667 | 312.6711 | 269 | 267.45 | 274.0188 | 291.7967 | 315.2835 | 258.3333 | 273.5833 | 294.8 | 267.9501
| Median |  | 227 | 268 | 247 | 282 | 277.605 | 269 | 281 | 312.6919 | 266.5 | 267.75 | 273.2047 | 291.915 | 316.3923 | 258.5 | 272.5 | 296.3 | 265.6214
| Std Dev. |  | 5.0498 | 5.6006 | 14.1386 | 2.5298 | 3.2677 | 17.3022 | 1.9664 | 0.80996 | 8.6948 | 2.333 | 3.6932 | 0.82065 | 6.312 | 4.5461 | 4.7793 | 7.5453 | 8.3124
| Rel.Std.Dev. |  | 2.229% | 2.073% | 5.736% | 0.9003% | 1.176% | 6.412% | 0.7006% | 0.259% | 3.232% | 0.8723% | 1.348% | 0.2812% | 2.002% | 1.76% | 1.747% | 2.559% | 3.102%
| PDM3 |  | -17.67% | -1.799% | -10.4% | 2.139% | 1.014% | -1.92% | 2.018% | 13.65% | -2.223% | -2.786% | -0.3987% | 6.063% | 14.6% | -6.1% | -0.557% | 7.155% | -2.605%
| Z-Score (Absolute) |  | 4.48 | 0.33 | 2.58 | 0.71 | 0.41 | 0.36 | 0.67 | 3.72 | 0.44 | 0.58 | 0.04 | 1.73 | 3.97 | 1.45 | 0 | 2.02 | 0.54
| Analytical results for CaO in OREAS 702 (Certified Value 22.21 wt.%) | [22 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | PF*MS | BF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  |  |  | 2
3
| 1 | 1 | 21.41 | 22 | 22.86 | 23.13 | 22.8258 | 21.41 | 21.0228 | 22.44 | 23.255 | 21.7355 | 22.42 | 22.177 | 22.4099 | 22.92 | 21.27 | 21.69 | 21.97 | 21.1966 |  |  |  | 1
| 1 | 2 | 21.41 | 21.6 | 22.61 | 23.27 | 22.5673 | 21.69 | 21.093 | 22.39 | 23.297 | 21.3557 | 22.5 | 22.317 | 22.2951 | 22.63 | 22.39 | 21.55 | 23.09 | 21.3166 |  |  |  | #N/A
| 1 | 3 | 21.97 | 21.9 | 22.72 | 22.44 | 23.1576 | 21.55 | 21.07 | 22.55 | 23.577 | 21.0315 | 22.57 | 22.499 | 22.3374 | 23.02 | 23.23 | 21.69 | 21.41 | 21.3543 |  |  |  | 16
| 1 | 4 | 21.97 | 22.2 | 22.89 | 23.43 | 23.726 | 21.55 | 21.0946 | 22.78 | 23.003 | 21.6733 | 22.77 | 21.953 | 22.2875 | 22.84 | 23.37 | 21.83 | 22.95 | 21.3703 |  |  |  | 22.2132
| 1 | 5 | 21.27 | 21.9 | 22.88 | 22.87 | 22.879 | 21.27 | 21.0647 | 22.9 | 22.863 | 21.6375 | 22.6 | 22.023 | 22.3541 | 23.01 | 23.09 | 21.83 | 22.11 | 21.1035 |  |  |  | 42
| 1 | 6 | 21.83 | 21.5 | 22.38 | 23.04 | 22.9378 | 21.55 | 21.0259 | 22.25 | 22.947 | 21.0041 | 22.54 | 22.065 | 22.4705 | 22.72 | 22.95 | 21.83 | 22.25 | 21.7342
| Mean |  | 21.6433 | 21.85 | 22.7233 | 23.03 | 23.0156 | 21.5033 | 21.0618 | 22.5517 | 23.157 | 21.4063 | 22.5667 | 22.1723 | 22.3591 | 22.8567 | 22.7167 | 21.7367 | 22.2967 | 21.3459
| Median |  | 21.62 | 21.9 | 22.79 | 23.085 | 22.9084 | 21.55 | 21.0674 | 22.495 | 23.129 | 21.4966 | 22.555 | 22.121 | 22.3457 | 22.88 | 23.02 | 21.76 | 22.18 | 21.3355
| Std Dev. |  | 0.31513 | 0.25884 | 0.20086 | 0.34681 | 0.3966 | 0.14459 | 0.031408 | 0.2462 | 0.26827 | 0.32799 | 0.11759 | 0.20474 | 0.070295 | 0.15782 | 0.78533 | 0.11431 | 0.63026 | 0.21604
| Rel.Std.Dev. |  | 1.456% | 1.185% | 0.884% | 1.506% | 1.723% | 0.6724% | 0.1491% | 1.092% | 1.159% | 1.532% | 0.5211% | 0.9234% | 0.3144% | 0.6905% | 3.457% | 0.5259% | 2.827% | 1.012%
| PDM3 |  | -2.566% | -1.635% | 2.296% | 3.677% | 3.612% | -3.196% | -5.183% | 1.524% | 4.249% | -3.633% | 1.591% | -0.1841% | 0.6566% | 2.897% | 2.266% | -2.145% | 0.3756% | -3.904%
| Z-Score (Absolute) |  | 0.82 | 0.57 | 0.48 | 0.85 | 0.83 | 0.99 | 1.52 | 0.27 | 1 | 1.11 | 0.29 | 0.18 | 0.04 | 0.64 | 0.47 | 0.71 | 0.04 | 1.18
| Analytical results for Cd in OREAS 702 (Certified Value < 10 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 02 | 10 | 14 | 15 | 17 | 18 | 22 | 23 | 25 | 26 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*OES | [11 empty] | 1
1
| 1 | 1 | < 0.8 | < 10 | 1 | < 10 | < 1 | 0.39 | < 2 | < 0.2 | < 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 0.8 | < 10 | < 1 | < 10 | < 1 | 0.36 | < 2 | 0.3 | < 1 | < 10 | [11 empty] | 1
| 1 | 3 | < 0.8 | < 10 | < 1 | < 10 | < 1 | 0.39 | < 2 | < 0.2 | < 1 | < 10 | [11 empty] | 16
| 1 | 4 | < 0.8 | < 10 | < 1 | < 10 | < 1 | 0.38 | < 2 | < 0.2 | < 1 | < 10 | [11 empty] | < 10
| 1 | 5 | < 0.8 | < 10 | < 1 | < 10 | < 1 | 0.38 | < 2 | < 0.2 | < 1 | < 10 | [11 empty] | 43
| 1 | 6 | < 0.8 | < 10 | < 1 | < 10 | < 1 | 0.38 | < 2 | < 0.2 | < 1 | < 10
| Mean |  |  |  | 1 |  |  | 0.38 |  | 0.3
| Median |  |  |  | 1 |  |  | 0.38 |  | 0.3
| Std Dev. | [6 empty] | 0.010954
| Rel.Std.Dev. | [6 empty] | 2.883%
PDM3
| Z-Score (Absolute) |  | 0.34 | 10.55 | 0.1 | 10.55 | 0.1 | 0.38 | 1.08 | 0.97 | 0.1 | 10.55
| Analytical results for Ce in OREAS 702 (Certified Value 56 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | 0
1
| 1 | 1 | 56.5 | 55.4 | 59.9 | 56.5 | 57.1 | 53.5 | 58.5 | 53 | 54.4 | 67.7136 | 53.3 | 63 | 52.5 | 56.1264 | 47.084 | 56.17 | 57.6138 | 52 | 55 | 58.3 | 53.5385 | 1
| 1 | 2 | 53.7 | 56.3 | 62.3 | 58.1 | 57.5 | 53.8 | 59.3 | 54 | 54.6 | 67.797 | 54.1 | 63 | 54.6 | 58.2815 | 47.2967 | 57.5 | 55.728 | 53 | 53 | 56.9 | 53.8411 | 2
| 1 | 3 | 53.6 | 52.6 | 58 | 56.5 | 56.1 | 54.3 | 54.8 | 54 | 55.2 | 67.6892 | 53.5 | 60 | 51.7 | 57.3776 | 47.112 | 56.69 | 56.3292 | 55 | 54 | 61.9 | 56.8291 | 16
| 1 | 4 | 52.7 | 56.3 | 58.8 | 55.7 | 55.6 | 53.5 | 57.8 | 56 | 54.7 | 67.5002 | 53.9 | 58 | 51.6 | 56.1447 | 48.1261 | 56.95 | 55.3699 | 53 | 53 | 59.2 | 53.2644 | 55.7091
| 1 | 5 | 54.5 | 54.4 | 60.6 | 59.5 | 55.7 | 53.2 | 55 | 53 | 55.7 | 67.4739 | 54.2 | 61 | 54.6 | 55.7365 | 48.6892 | 56.79 | 56.1421 | 50 | 53 | 58.7 | 53.3865 | 44
| 1 | 6 | 55 | 54.5 | 54.9 | 59.9 | 54 | 53.8 | 56.1 | 53 | 55.6 | 67.5862 | 53 | 60 | 52 | 57.4321 | 48.4813 | 56.4 | 58.4752 | 52 | 53 | 56.4 | 58.3236
| Mean |  | 54.3333 | 54.9167 | 59.0833 | 57.7 | 56 | 53.6833 | 56.9167 | 53.8333 | 55.0333 | 67.6267 | 53.6667 | 60.8333 | 52.8333 | 56.8498 | 47.7982 | 56.75 | 56.6097 | 52.5 | 53.5 | 58.5667 | 54.8639
| Median |  | 54.1 | 54.95 | 59.35 | 57.3 | 55.9 | 53.65 | 56.95 | 53.5 | 54.95 | 67.6377 | 53.7 | 60.5 | 52.25 | 56.7612 | 47.7114 | 56.74 | 56.2357 | 52.5 | 53 | 58.5 | 53.6898
| Std Dev. |  | 1.3246 | 1.4049 | 2.5325 | 1.739 | 1.2426 | 0.37639 | 1.8862 | 1.169 | 0.5465 | 0.12764 | 0.4761 | 1.9408 | 1.4038 | 0.99271 | 0.72117 | 0.46187 | 1.1917 | 1.6432 | 0.83666 | 1.9531 | 2.1622
| Rel.Std.Dev. |  | 2.438% | 2.558% | 4.286% | 3.014% | 2.219% | 0.7011% | 3.314% | 2.172% | 0.993% | 0.1887% | 0.8871% | 3.19% | 2.657% | 1.746% | 1.509% | 0.8139% | 2.105% | 3.13% | 1.564% | 3.335% | 3.941%
| PDM3 |  | -2.47% | -1.422% | 6.057% | 3.574% | 0.5221% | -3.636% | 2.168% | -3.367% | -1.213% | 21.39% | -3.666% | 9.198% | -5.162% | 2.048% | -14.2% | 1.868% | 1.617% | -5.761% | -3.965% | 5.129% | -1.517%
| Z-Score (Absolute) |  | 0.27 | 0.05 | 1.59 | 1.05 | 0.38 | 0.53 | 0.74 | 0.47 | 0 | 4.95 | 0.54 | 2.28 | 0.86 | 0.71 | 2.84 | 0.67 | 0.62 | 1 | 0.6 | 1.39 | 0.07
| Analytical results for Co in OREAS 702 (Certified Value 6.74 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 6.1 | 6 | 6.1 | 7 | 7.8941 | 6.6 | 7.58 | < 10 | 6.2702 | 6 | 7.0143 | 4.79 | 6 | < 50 | 7 | 7.0751 | [5 empty] | 1
| 1 | 2 | 6.1 | 6 | 5.9 | 7 | 7.8748 | 6.5 | 7.92 | < 10 | 6.098 | 5.9 | 6.6702 | 3.16 | 6 | < 50 | 7 | 6.632 | [5 empty] | #N/A
| 1 | 3 | 6 | 6 | 6.7 | 7 | 7.9108 | 6.7 | 7.33 | < 10 | 6.0378 | 5.7 | 7.2023 | 1.63 | 6 | < 50 | 7 | 6.801 | [5 empty] | 16
| 1 | 4 | 6.2 | 6 | 6.4 | 7 | 7.8713 | 7.5 | 7.78 | < 10 | 6.0402 | 5.6 | 6.9754 | 2.08 | 6 | < 50 | 7 | 6.5825 | [5 empty] | 6.7415
| 1 | 5 | 6.2 | 6 | 6.6 | 7 | 7.8984 | 6.5 | 7.69 | < 10 | 6.1516 | 5.7 | 7.034 | 4.65 | 6 | < 50 | 8 | 6.6585 | [5 empty] | 45
| 1 | 6 | 6.1 | 6 | 6.1 | 7 | 7.8538 | 6.7 | 7.22 | < 10 | 7.1823 | 5.9 | 7.0098 | 4.88 | 6 | < 50 | 7 | 7.9806
| Mean |  | 6.1167 | 6 | 6.3 | 7 | 7.8839 | 6.75 | 7.5867 |  | 6.2967 | 5.8 | 6.9843 | 3.5317 | 6 |  | 7.1667 | 6.9549
| Median |  | 6.1 | 6 | 6.25 | 7 | 7.8845 | 6.65 | 7.635 |  | 6.1248 | 5.8 | 7.0121 | 3.905 | 6 |  | 7 | 6.7297
| Std Dev. |  | 0.075277 | 0 | 0.31623 | 0 | 0.020912 | 0.37815 | 0.26815 |  | 0.44236 | 0.15492 | 0.17338 | 1.4501 | 0 |  | 0.40825 | 0.53312
| Rel.Std.Dev. |  | 1.231% | 0% | 5.019% | 0% | 0.2653% | 5.602% | 3.535% |  | 7.025% | 2.671% | 2.482% | 41.06% | 0% |  | 5.696% | 7.665%
| PDM3 |  | -9.268% | -11% | -6.548% | 3.835% | 16.95% | 0.1267% | 12.54% |  | -6.598% | -13.97% | 3.603% | -47.61% | -11% |  | 6.307% | 3.167%
| Z-Score (Absolute) |  | 0.46 | NA | 0.26 | NA | 1.55 | 0.26 | 1.21 | 1.74 | 0.26 | 0.83 | 0.52 | 3.41 | NA | 21.04 | NA | 0.49
Results from laboratories 6, 10, 23 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 702 (Certified Value 29.5 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 17 | 18 | 23 | 24 | 26 | [5 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | [5 empty] | 1
1
| 1 | 1 | 28 | 31 | 34 | 30 | 31 | 30 | 25 | 30 | < 50 | 21 | 34 | 66.4368 | < 70 | 30 | < 50 | 354.2 | [5 empty] | 1
| 1 | 2 | 29 | 29 | 32 | 30 | 31 | 30 | 30 | 203 | < 50 | 34 | 34 | 55.07 | < 70 | 27 | < 50 | 353.2 | [5 empty] | #N/A
| 1 | 3 | 29 | 28 | 31 | 31 | 31 | 30 | 23 | 34 | < 50 | 21 | 35 | 62.2496 | < 70 | 25 | < 50 | 348.14 | [5 empty] | 16
| 1 | 4 | 28 | 29 | 31 | 30 | 31 | 30 | 23 | 35 | < 50 | 34 | 36 | 53.5332 | < 70 | 26 | < 50 | 351.84 | [5 empty] | 29.4662
| 1 | 5 | 30 | 28 | 29 | 31 | 30 | 30 | 26 | 30 | < 50 | 21 | 34 | 59.9604 | < 70 | 27 | < 50 | 353.33 | [5 empty] | 46
| 1 | 6 | 28 | 28 | 28 | 31 | 30 | 30 | 27 | 39 | < 50 | 27 | 30 | 61.767 | < 70 | 20 | < 50 | 356.91
| Mean |  | 28.6667 | 28.8333 | 30.8333 | 30.5 | 30.6667 | 30 | 25.6667 | 61.8333 |  | 26.3333 | 33.8333 | 59.8362 |  | 25.8333 |  | 352.9367
| Median |  | 28.5 | 28.5 | 31 | 30.5 | 31 | 30 | 25.5 | 34.5 |  | 24 | 34 | 60.8637 |  | 26.5 |  | 353.265
| Std Dev. |  | 0.8165 | 1.169 | 2.137 | 0.54772 | 0.5164 | 0 | 2.6583 | 69.2399 |  | 6.377 | 2.0412 | 4.8079 |  | 3.3116 |  | 2.892
| Rel.Std.Dev. |  | 2.848% | 4.054% | 6.931% | 1.796% | 1.684% | 0% | 10.36% | 112% |  | 24.22% | 6.033% | 8.035% |  | 12.82% |  | 0.8194%
| PDM3 |  | -2.713% | -2.148% | 4.64% | 3.508% | 4.074% | 1.812% | -12.89% | 109.8% |  | -10.63% | 14.82% | 103.1% |  | -12.33% |  | 1098%
| Z-Score (Absolute) |  | 0.17 | 0.14 | 0.2 | 0.14 | 0.17 | NA | 0.69 | 5.54 | 0.8 | 0.59 | 0.72 | 5.19 | 0.78 | 0.66 | 0.8 | NA
Results from laboratory 1.26 were removed due to their systematic bias. Results from laboratory 6 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 702 (Certified Value 8.53 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | 26 |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 8.33 | 8.77 | 9.19 | 8.91 | 8.59 | 8.6 | 7.7 | 9 | 8.8 | 24 | 8.2 | 8.9579 | 7.12 | 8.1492 | 8 | 10 | 8.6 | 8.2872 |  |  |  | 1
| 1 | 2 | 8.17 | 8.62 | 9.01 | 9.01 | 9.11 | 8.7 | 7.8 | 9.16 | 8.9 | 25 | 8.2 | 8.7118 | 6.93 | 7.7819 | 8.3 | 9 | 8.6 | 8.4662 |  |  |  | #N/A
| 1 | 3 | 8.57 | 8.49 | 9.23 | 8.86 | 8.72 | 8.6 | 7.4 | 9.1 | 8.9 | 22 | 8.3 | 9.2583 | 7.12 | 7.6134 | 8.3 | 9 | 8.5 | 8.3815 |  |  |  | 16
| 1 | 4 | 8.22 | 8.65 | 8.96 | 8.84 | 8.74 | 8.6 | 7.2 | 9.03 | 9.1 | 20 | 8.2 | 8.7984 | 7 | 7.4137 | 8.3 | 9 | 8.6 | 8.2971 |  |  |  | 8.5304
| 1 | 5 | 8.64 | 8.38 | 9.44 | 8.96 | 8.39 | 8.6 | 7.1 | 9.08 | 9 | 17 | 8.8 | 8.9315 | 7.13 | 7.7263 | 8 | 9 | 8.6 | 8.4709 |  |  |  | 47
| 1 | 6 | 8.44 | 8.41 | 9.05 | 8.85 | 8.27 | 8.6 | 7.1 | 9.06 | 9.2 | 17 | 8 | 9.5233 | 7.1 | 7.8103 | 8.2 | 9 | 8.5 | 8.8074
| Mean |  | 8.395 | 8.5533 | 9.1467 | 8.905 | 8.6367 | 8.6167 | 7.3833 | 9.0717 | 8.9833 | 20.8333 | 8.2833 | 9.0302 | 7.0667 | 7.7491 | 8.1833 | 9.1667 | 8.5667 | 8.4517
| Median |  | 8.385 | 8.555 | 9.12 | 8.885 | 8.655 | 8.6 | 7.3 | 9.07 | 8.95 | 21 | 8.2 | 8.9447 | 7.11 | 7.7541 | 8.25 | 9 | 8.6 | 8.4238
| Std Dev. |  | 0.18876 | 0.15188 | 0.17761 | 0.068337 | 0.29636 | 0.040825 | 0.30605 | 0.056006 | 0.1472 | 3.4303 | 0.27142 | 0.30505 | 0.082381 | 0.24347 | 0.1472 | 0.40825 | 0.05164 | 0.19133
| Rel.Std.Dev. |  | 2.248% | 1.776% | 1.942% | 0.7674% | 3.431% | 0.4738% | 4.145% | 0.6174% | 1.639% | 16.47% | 3.277% | 3.378% | 1.166% | 3.142% | 1.799% | 4.454% | 0.6028% | 2.264%
| PDM3 |  | -1.587% | 0.2688% | 7.224% | 4.391% | 1.246% | 1.011% | -13.45% | 6.345% | 5.31% | 144.2% | -2.896% | 5.859% | -17.16% | -9.159% | -4.069% | 7.459% | 0.4251% | -0.9224%
| Z-Score (Absolute) |  | 0.31 | 0.01 | 1.1 | 0.64 | 0.14 | 0.11 | 2.2 | 0.96 | 0.79 | NA | 0.52 | 0.88 | 2.79 | 1.52 | 0.7 | NA | 0.01 | 0.2
Results from laboratories 14 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 702 (Certified Value 362 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 06 | 07 | 09 | 10 | 11 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [5 empty] | 0
0
| 1 | 1 | 330 | 340 | 351 | 376 | 350 | 370.7716 | 372 | 327 | 370.5502 | 391.1 | 391.4646 | 326.3 | 393 | 400 | 374 | 60.57 | [5 empty] | 1
| 1 | 2 | 350 | 340 | 347 | 361 | 350 | 371.0076 | 380 | 314 | 363.9542 | 389.1 | 386.1306 | 324.73 | 392 | 350 | 376 | 45.36 | [5 empty] | 19
| 1 | 3 | 350 | 340 | 339 | 369 | 350 | 370.3352 | 387 | 308 | 379.8953 | 395.7 | 384.4848 | 324.2 | 360 | 350 | 368 | 37.74 | [5 empty] | 16
| 1 | 4 | 350 | 330 | 354 | 398 | 350 | 370.6532 | 381 | 341 | 356.5783 | 390.3 | 378.8839 | 320.1 | 369 | 350 | 381 | 42.88 | [5 empty] | 362.1484
| 1 | 5 | 350 | 340 | 347 | 367 | 350 | 370.017 | 383 | 319 | 370.9552 | 388 | 385.1211 | 325.45 | 390 | 350 | 378 | 41.53 | [5 empty] | 48
| 1 | 6 | 360 | 350 | 349 | 370 | 350 | 371.3254 | 382 | 385 | 384.7241 | 395.8 | 388.6933 | 322.03 | 387 | 350 | 380 | 44.12
| Mean |  | 348.3333 | 340 | 347.8333 | 373.5 | 350 | 370.685 | 380.8333 | 332.3333 | 371.1095 | 391.6667 | 385.7964 | 323.8017 | 381.8333 | 358.3333 | 376.1667 | 45.3667
| Median |  | 350 | 340 | 348 | 369.5 | 350 | 370.7124 | 381.5 | 323 | 370.7527 | 390.7 | 385.6258 | 324.465 | 388.5 | 350 | 377 | 43.5
| Std Dev. |  | 9.8319 | 6.3246 | 5.0761 | 12.9422 | 0 | 0.46718 | 4.9565 | 28.2253 | 10.2493 | 3.3339 | 4.2566 | 2.316 | 13.8768 | 20.4124 | 4.7504 | 7.8964
| Rel.Std.Dev. |  | 2.823% | 1.86% | 1.459% | 3.465% | 0% | 0.126% | 1.301% | 8.493% | 2.762% | 0.8512% | 1.103% | 0.7153% | 3.634% | 5.696% | 1.263% | 17.41%
| PDM3 |  | -3.815% | -6.116% | -3.953% | 3.135% | -3.355% | 2.357% | 5.159% | -8.233% | 2.474% | 8.151% | 6.53% | -10.59% | 5.436% | -1.053% | 3.871% | -87.47%
| Z-Score (Absolute) |  | 0.66 | 1 | 0.68 | 0.37 | 0.59 | 0.25 | 0.67 | 1.31 | 0.27 | 1.11 | 0.87 | 1.66 | 0.71 | 0.25 | 0.48 | 13.04
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Dy in OREAS 702 (Certified Value 4.46 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 4.45 | 4.27 | 5 | 4.9 | 4.91 | 4.2 | 4.1 | 5 | 4.2 | 4.1298 | 4.35 | 5.01 | 4.73 | 5.9206 | 3.32 | 4.6424 | 4.1 | 5 | 4.54 | 4.7614 |  | 1
| 1 | 2 | 4.4 | 4.04 | 5 | 4.72 | 4.43 | 4.3 | 4 | 4 | 4.3 | 4.1399 | 4.3 | 4.84 | 4.56 | 5.1356 | 3.3 | 4.4844 | 4.4 | 4.5 | 4.48 | 4.6966 |  | #N/A
| 1 | 3 | 4.24 | 4.1 | 4.6 | 4.5 | 4.4 | 4.5 | 4 | 4 | 4.4 | 4.1079 | 4.3 | 4.84 | 4.26 | 5.2912 | 3.32 | 4.6985 | 4.4 | 4 | 4.36 | 4.904 |  | 16
| 1 | 4 | 4.18 | 4.28 | 4.62 | 4.3 | 4.42 | 4.3 | 5.1 | 5 | 4.4 | 4.1565 | 4.32 | 5.12 | 4.45 | 4.9412 | 3.24 | 4.3773 | 4.2 | 5 | 4.56 | 4.5887 |  | 4.4564
| 1 | 5 | 4.32 | 4.24 | 5.02 | 4.54 | 4.56 | 4.2 | 4.5 | 4 | 4.4 | 4.1185 | 4.36 | 4.41 | 4.25 | 5.1957 | 3.25 | 4.5943 | 3.8 | 4.5 | 4.64 | 4.1321 |  | 49
| 1 | 6 | 4.25 | 4.24 | 4.9 | 4.89 | 4.89 | 4.5 | 4.5 | 5 | 4.6 | 4.1005 | 4.35 | 4.36 | 3.81 | 4.7379 | 3.31 | 4.8876 | 4.3 | 5.5 | 4.45 | 4.4751
| Mean |  | 4.3067 | 4.195 | 4.8567 | 4.6417 | 4.6017 | 4.3333 | 4.3667 | 4.5 | 4.3833 | 4.1255 | 4.33 | 4.7633 | 4.3433 | 5.2037 | 3.29 | 4.6141 | 4.2 | 4.75 | 4.505 | 4.593
| Median |  | 4.285 | 4.24 | 4.95 | 4.63 | 4.495 | 4.3 | 4.3 | 4.5 | 4.4 | 4.1241 | 4.335 | 4.84 | 4.355 | 5.1657 | 3.305 | 4.6183 | 4.25 | 4.75 | 4.51 | 4.6427
| Std Dev. |  | 0.10309 | 0.09995 | 0.19572 | 0.23735 | 0.23794 | 0.13663 | 0.4274 | 0.54772 | 0.13292 | 0.020829 | 0.026833 | 0.3122 | 0.31873 | 0.40303 | 0.035777 | 0.17665 | 0.22804 | 0.5244 | 0.097108 | 0.26913
| Rel.Std.Dev. |  | 2.394% | 2.383% | 4.03% | 5.114% | 5.171% | 3.153% | 9.788% | 12.17% | 3.032% | 0.5049% | 0.6197% | 6.554% | 7.338% | 7.745% | 1.087% | 3.828% | 5.429% | 11.04% | 2.156% | 5.86%
| PDM3 |  | -3.361% | -5.866% | 8.981% | 4.157% | 3.259% | -2.762% | -2.014% | 0.9778% | -1.64% | -7.425% | -2.837% | 6.887% | -2.538% | 16.77% | -26.17% | 3.537% | -5.754% | 6.588% | 1.09% | 3.065%
| Z-Score (Absolute) |  | 0.25 | 0.61 | 1.52 | 0.83 | 0.7 | 0.16 | 0.05 | NA | 0 | 0.83 | 0.17 | 1.22 | 0.13 | 2.64 | 3.52 | 0.74 | 0.59 | 1.18 | 0.39 | 0.67
Results from laboratory 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 702 (Certified Value 2.26 ppm) | [22 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 2.24 | 2.34 | 2.5 | 2.3 | 2.37 | 2.1 | 2.9 | 2 | 2.2 | 1.9622 | 2.34 | 3.03 | 2.27 | 2.5809 | 1.29 | 2.4397 | 2.1 | 2.4 | 2.44 | 2.4041 |  | 1
| 1 | 2 | 2.27 | 2.13 | 2.39 | 2.38 | 2.33 | 2.2 | 2.5 | 2 | 2.2 | 2.0125 | 2.35 | 2.8 | 2.21 | 2.248 | 1.26 | 2.4247 | 2.2 | 2.3 | 2.4 | 2.1067 |  | #N/A
| 1 | 3 | 2.29 | 2.22 | 2.57 | 2.28 | 2.44 | 2.3 | 2 | 2 | 2.1 | 1.9766 | 2.34 | 2.5 | 2.34 | 2.306 | 1.3 | 2.3103 | 2.2 | 2.4 | 2.36 | 2.3145 |  | 16
| 1 | 4 | 2.15 | 2.24 | 2.46 | 2.12 | 2.34 | 2.3 | 2.1 | 2 | 2.2 | 1.9992 | 2.39 | 2.6 | 2.38 | 2.2025 | 1.3 | 2.2191 | 2.1 | 2.3 | 2.4 | 2.4332 |  | 2.2622
| 1 | 5 | 2.22 | 2.3 | 2.69 | 2.35 | 2.52 | 2.1 | 1.9 | 2 | 2.1 | 1.9784 | 2.32 | 2.6 | 2.31 | 2.2175 | 1.33 | 2.3644 | 2 | 2 | 2.27 | 2.2863 |  | 50
| 1 | 6 | 2.18 | 2.28 | 2.62 | 2.34 | 2.14 | 2.3 | 1.6 | 2 | 2.3 | 1.9388 | 2.26 | 2.65 | 2.05 | 2.0423 | 1.31 | 2.395 | 2.1 | 2 | 2.38 | 2.5207
| Mean |  | 2.225 | 2.2517 | 2.5383 | 2.295 | 2.3567 | 2.2167 | 2.1667 | 2 | 2.1833 | 1.978 | 2.3333 | 2.6967 | 2.26 | 2.2662 | 1.2983 | 2.3589 | 2.1167 | 2.2333 | 2.375 | 2.3442
| Median |  | 2.23 | 2.26 | 2.535 | 2.32 | 2.355 | 2.25 | 2.05 | 2 | 2.2 | 1.9775 | 2.34 | 2.625 | 2.29 | 2.2328 | 1.3 | 2.3797 | 2.1 | 2.3 | 2.39 | 2.3593
| Std Dev. |  | 0.053198 | 0.073326 | 0.1098 | 0.092898 | 0.12785 | 0.098319 | 0.46332 | 0 | 0.075277 | 0.02619 | 0.04274 | 0.19044 | 0.11832 | 0.17749 | 0.023166 | 0.082594 | 0.075277 | 0.18619 | 0.057879 | 0.14373
| Rel.Std.Dev. |  | 2.391% | 3.257% | 4.326% | 4.048% | 5.425% | 4.435% | 21.38% | 0% | 3.448% | 1.324% | 1.832% | 7.062% | 5.235% | 7.832% | 1.784% | 3.501% | 3.556% | 8.337% | 2.437% | 6.131%
| PDM3 |  | -1.644% | -0.4653% | 12.21% | 1.45% | 4.176% | -2.012% | -4.223% | -11.59% | -3.486% | -12.56% | 3.145% | 19.21% | -0.09693% | 0.1771% | -42.61% | 4.273% | -6.433% | -1.276% | 4.987% | 3.627%
| Z-Score (Absolute) |  | 0.28 | 0.07 | 2.23 | 0.28 | 0.77 | 0.35 | 0.75 | NA | 0.61 | 2.26 | 0.59 | 3.5 | 0 | 0.05 | 7.7 | 0.79 | 1.15 | 0.21 | 0.92 | 0.67
Results from laboratory 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 702 (Certified Value 1.01 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 0.96 | 1.04 | 1 | 1.07 | 1.15 | 1.01 | 0.9 | < 2 | 1 | 1.3146 | 0.96 | 1.69 | 0.87 | 1.244 | 1.09 | 1.0915 | 1 | 1.3 | 1.08 | 0.93992 |  | 1
| 1 | 2 | 1.01 | 1.08 | 1.14 | 1.03 | 1.04 | 1.01 | 0.9 | < 2 | 1 | 1.2874 | 0.92 | 1.75 | 0.97 | 1.01 | 1.11 | 0.99347 | 1.1 | 1.1 | 1.11 | 1.0408 |  | #N/A
| 1 | 3 | 0.96 | 1 | 1 | 1 | 0.91 | 1.02 | 0.9 | < 2 | 1 | 1.3723 | 1 | 1.25 | 0.86 | 1.0212 | 1.09 | 1.0299 | 1 | 1 | 1.08 | 1.0789 |  | 16
| 1 | 4 | 0.92 | 1.02 | 0.96 | 1.04 | 1.2 | 1 | 1 | < 2 | 1 | 1.317 | 0.97 | 1.46 | 0.85 | 0.9666 | 1.08 | 1.0053 | 1 | 1 | 1.1 | 0.98821 |  | 1.0105
| 1 | 5 | 1.04 | 1.08 | 1 | 1.02 | 0.88 | 0.96 | 1 | < 2 | 1 | 1.2946 | 0.96 | 1.14 | 0.99 | 1.0245 | 1.09 | 1.0188 | 0.9 | 1.2 | 1.06 | 0.9469 |  | 51
| 1 | 6 | 0.97 | 0.99 | 0.92 | 1.12 | 1.06 | 0.98 | 1 | < 2 | 1 | 1.2759 | 0.99 | 1.46 | 0.91 | 0.8821 | 1.1 | 1.0332 | 0.9 | 1 | 1.06 | 1.11
| Mean |  | 0.97667 | 1.035 | 1.0033 | 1.0467 | 1.04 | 0.99667 | 0.95 |  | 1 | 1.3103 | 0.96667 | 1.4583 | 0.90833 | 1.0247 | 1.0933 | 1.0287 | 0.98333 | 1.1 | 1.0817 | 1.0175
| Median |  | 0.965 | 1.03 | 1 | 1.035 | 1.05 | 1.005 | 0.95 |  | 1 | 1.3046 | 0.965 | 1.46 | 0.89 | 1.0156 | 1.09 | 1.0244 | 1 | 1.05 | 1.08 | 1.0145
| Std Dev. |  | 0.042269 | 0.038859 | 0.074207 | 0.04274 | 0.12696 | 0.022509 | 0.054772 |  | 0 | 0.034225 | 0.028048 | 0.23811 | 0.05947 | 0.12003 | 0.010328 | 0.034205 | 0.075277 | 0.12649 | 0.020412 | 0.070327
| Rel.Std.Dev. |  | 4.328% | 3.754% | 7.396% | 4.083% | 12.21% | 2.258% | 5.766% |  | 0% | 2.612% | 2.901% | 16.33% | 6.547% | 11.71% | 0.9446% | 3.325% | 7.655% | 11.5% | 1.887% | 6.912%
| PDM3 |  | -3.352% | 2.42% | -0.7135% | 3.575% | 2.915% | -1.373% | -5.991% |  | -1.043% | 29.66% | -4.342% | 44.31% | -10.11% | 1.404% | 8.193% | 1.796% | -2.693% | 8.852% | 7.038% | 0.6844%
| Z-Score (Absolute) |  | 0.78 | 0.17 | 0.35 | 0.36 | 0.25 | 0.46 | 1.22 | 0.4 | NA | 4.65 | 0.95 | 7.06 | 1.9 | 0 | 1.12 | 0.06 | 0.67 | 1.23 | 0.93 | 0.12
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe2O3 in OREAS 702 (Certified Value 3.38 wt.%) | [22 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | PF*MS | BF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  |  |  | 2
3
| 1 | 1 | 3.446 | 3.36 | 3.46 | 3.46 | 2.9703 | 3.474 | 3.3324 | 3.38 | 3.403 | 3.2805 | 3.4497 | 3.59 | 3.1638 | 3.42 | 3.474 | 3.288 | 3.36 | 3.3111 |  |  |  | 1
| 1 | 2 | 3.403 | 3.33 | 3.4 | 3.47 | 2.9323 | 3.46 | 3.3252 | 3.34 | 3.388 | 3.2182 | 3.3627 | 3.6 | 3.1691 | 3.39 | 3.36 | 3.317 | 3.603 | 3.2553 |  |  |  | #N/A
| 1 | 3 | 3.503 | 3.35 | 3.44 | 3.34 | 2.9952 | 3.374 | 3.3034 | 3.43 | 3.488 | 3.1712 | 3.4527 | 3.43 | 3.1567 | 3.43 | 3.26 | 3.303 | 3.317 | 3.2748 |  |  |  | 16
| 1 | 4 | 3.517 | 3.4 | 3.48 | 3.5 | 3.0747 | 3.417 | 3.3232 | 3.48 | 3.331 | 3.2275 | 3.4347 | 3.4 | 3.1081 | 3.35 | 3.274 | 3.317 | 3.517 | 3.3663 |  |  |  | 3.3839
| 1 | 5 | 3.417 | 3.35 | 3.45 | 3.42 | 2.9569 | 3.388 | 3.3062 | 3.34 | 3.317 | 3.2546 | 3.5023 | 3.43 | 3.1112 | 3.34 | 3.488 | 3.317 | 3.531 | 3.2796 |  |  |  | 52
| 1 | 6 | 3.488 | 3.3 | 3.4 | 3.44 | 2.9585 | 3.388 | 3.3179 | 3.39 | 3.317 | 3.1629 | 3.4903 | 3.37 | 3.0894 | 3.3 | 3.403 | 3.331 | 3.374 | 3.3399
| Mean |  | 3.4623 | 3.3483 | 3.4383 | 3.4383 | 2.9813 | 3.4168 | 3.3181 | 3.3933 | 3.374 | 3.2192 | 3.4487 | 3.47 | 3.1331 | 3.3717 | 3.3765 | 3.3122 | 3.4503 | 3.3045
| Median |  | 3.467 | 3.35 | 3.445 | 3.45 | 2.9644 | 3.4025 | 3.3206 | 3.385 | 3.3595 | 3.2229 | 3.4512 | 3.43 | 3.134 | 3.37 | 3.3815 | 3.317 | 3.4455 | 3.2954
| Std Dev. |  | 0.047209 | 0.033116 | 0.032506 | 0.055287 | 0.05011 | 0.041542 | 0.011322 | 0.054283 | 0.066903 | 0.045916 | 0.049422 | 0.099398 | 0.034071 | 0.050365 | 0.096945 | 0.014784 | 0.11491 | 0.042537
| Rel.Std.Dev. |  | 1.363% | 0.989% | 0.9454% | 1.608% | 1.681% | 1.216% | 0.3412% | 1.6% | 1.983% | 1.426% | 1.433% | 2.865% | 1.087% | 1.494% | 2.871% | 0.4464% | 3.33% | 1.287%
| PDM3 |  | 2.317% | -1.052% | 1.608% | 1.608% | -11.9% | 0.9726% | -1.947% | 0.2782% | -0.2932% | -4.869% | 1.915% | 2.544% | -7.413% | -0.3621% | -0.2193% | -2.12% | 1.963% | -2.347%
| Z-Score (Absolute) |  | 0.93 | 0.29 | 0.67 | 0.67 | 4.21 | 0.45 | 0.61 | 0.19 | 0.01 | 1.67 | 0.78 | 1.01 | 2.59 | 0.04 | 0.01 | 0.67 | 0.8 | 0.76
| Analytical results for Ga in OREAS 702 (Certified Value 14.2 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 25 | 26 |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*OES |  |  |  | 1
2
| 1 | 1 | 14 | 14.8 | 13.7 | 14.6 | 14 | 14 | 14.6 | 14 | 15.8692 | 13.4 | 15.5 | 13 | 13.5622 | 14.19 | 14.8706 | 13 | 15 | < 100 |  |  |  | 1
| 1 | 2 | 14.2 | 14.7 | 14.4 | 14.6 | 13.9 | 14 | 13.3 | 14 | 15.84 | 13.7 | 15.8 | 13 | 13.5923 | 14.85 | 14.4571 | 13 | 15 | < 100 |  |  |  | #N/A
| 1 | 3 | 14 | 14.5 | 13.9 | 14.6 | 13.6 | 14 | 14.1 | 14 | 15.8821 | 13.6 | 16.5 | 13 | 13.9117 | 14.26 | 14.673 | 13 | 15 | < 100 |  |  |  | 16
| 1 | 4 | 13.2 | 14.8 | 14 | 14.3 | 14 | 14 | 13.6 | 14 | 15.9048 | 13.3 | 15.3 | 14 | 13.7579 | 14.64 | 13.7582 | 13 | 15 | < 100 |  |  |  | 14.2344
| 1 | 5 | 14.6 | 14.6 | 14.7 | 14.6 | 13.6 | 14 | 14 | 13 | 15.84 | 13.1 | 15.3 | 14 | 13.7476 | 14.37 | 14.4599 | 13 | 15 | < 100 |  |  |  | 53
| 1 | 6 | 14.5 | 14.7 | 14.3 | 14.8 | 13.4 | 14 | 13.5 | 14 | 15.9172 | 13.2 | 16.4 | 14 | 13.6766 | 14.07 | 14.7759 | 13 | 15 | < 100
| Mean |  | 14.0833 | 14.6833 | 14.1667 | 14.5833 | 13.75 | 14 | 13.85 | 13.8333 | 15.8755 | 13.3833 | 15.8 | 13.5 | 13.7081 | 14.3967 | 14.4991 | 13 | 15
| Median |  | 14.1 | 14.7 | 14.15 | 14.6 | 13.75 | 14 | 13.8 | 14 | 15.8756 | 13.35 | 15.65 | 13.5 | 13.7121 | 14.315 | 14.5664 | 13 | 15
| Std Dev. |  | 0.49967 | 0.1169 | 0.36697 | 0.16021 | 0.251 | 0 | 0.47645 | 0.40825 | 0.032255 | 0.23166 | 0.53666 | 0.54772 | 0.12741 | 0.29446 | 0.39925 | 0 | 0
| Rel.Std.Dev. |  | 3.548% | 0.7962% | 2.59% | 1.099% | 1.825% | 0% | 3.44% | 2.951% | 0.2032% | 1.731% | 3.397% | 4.057% | 0.9295% | 2.045% | 2.754% | 0% | 0%
| PDM3 |  | -1.062% | 3.154% | -0.4761% | 2.451% | -3.403% | -1.647% | -2.701% | -2.818% | 11.53% | -5.979% | 11% | -5.16% | -3.698% | 1.14% | 1.859% | -8.672% | 5.378%
| Z-Score (Absolute) |  | 0.39 | 0.35 | 0.28 | 0.23 | 0.8 | NA | 0.67 | NA | 1.82 | 1.25 | 1.73 | NA | 0.85 | 0 | 0.13 | NA | NA | 43.92
Results from laboratories 6, 10, 15, 23 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 702 (Certified Value 5.07 ppm) | [22 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 5.11 | 5.42 | 5.03 | 5.36 | 5.63 | 4.8 | 5.21 | 6 | 5.2 | 5.8516 | 5.1 | 5.7 | 4.73 | 6.9592 | 4.68 | 5.4977 | 4.4 | 5 | 6 | 5.1587 |  | 1
| 1 | 2 | 4.87 | 5.12 | 5.15 | 5.33 | 5.5 | 4.7 | 5.56 | 5 | 5.2 | 5.8992 | 5.08 | 5.9 | 4.77 | 5.8659 | 4.71 | 5.3237 | 4.5 | 5 | 5 | 4.8504 |  | #N/A
| 1 | 3 | 4.91 | 5.04 | 5.27 | 5.27 | 5.23 | 4.9 | 4.89 | 6 | 5.3 | 5.923 | 5.11 | 5.3 | 4.49 | 6.0915 | 4.67 | 5.468 | 4.7 | 5 | 5 | 5.0455 |  | 16
| 1 | 4 | 4.8 | 5.53 | 5.03 | 4.98 | 5.26 | 4.8 | 4.76 | 6 | 5.3 | 5.8768 | 5.12 | 5.2 | 5.03 | 5.5855 | 4.72 | 5.2759 | 4.6 | 5 | 6 | 5.1513 |  | 5.0726
| 1 | 5 | 4.97 | 5.23 | 5.36 | 5.27 | 5.35 | 4.7 | 5.07 | 5 | 5.4 | 5.8657 | 5.2 | 4.9 | 5.08 | 5.9955 | 4.68 | 5.308 | 4.3 | 5 | 7 | 4.6775 |  | 54
| 1 | 6 | 5.04 | 5.33 | 5.19 | 5.63 | 5.28 | 4.9 | 5.41 | 5 | 5.3 | 5.9297 | 5.09 | 4.8 | 4.76 | 5.4067 | 4.64 | 5.6237 | 4.6 | 5 | 5 | 5.141
| Mean |  | 4.95 | 5.2783 | 5.1717 | 5.3067 | 5.375 | 4.8 | 5.15 | 5.5 | 5.2833 | 5.891 | 5.1167 | 5.3 | 4.81 | 5.984 | 4.6833 | 5.4162 | 4.5167 | 5 | 5.6667 | 5.0041
| Median |  | 4.94 | 5.28 | 5.17 | 5.3 | 5.315 | 4.8 | 5.14 | 5.5 | 5.3 | 5.888 | 5.105 | 5.25 | 4.765 | 5.9307 | 4.68 | 5.3958 | 4.55 | 5 | 5.5 | 5.0933
| Std Dev. |  | 0.11367 | 0.18454 | 0.13121 | 0.20849 | 0.15783 | 0.089443 | 0.30509 | 0.54772 | 0.075277 | 0.03156 | 0.043205 | 0.43359 | 0.21643 | 0.54198 | 0.028752 | 0.13587 | 0.1472 | 0 | 0.8165 | 0.19818
| Rel.Std.Dev. |  | 2.296% | 3.496% | 2.537% | 3.929% | 2.936% | 1.863% | 5.924% | 9.959% | 1.425% | 0.5357% | 0.8444% | 8.181% | 4.499% | 9.057% | 0.6139% | 2.509% | 3.259% | 0% | 14.41% | 3.96%
| PDM3 |  | -2.417% | 4.055% | 1.953% | 4.614% | 5.961% | -5.374% | 1.525% | 8.425% | 4.154% | 16.13% | 0.8683% | 4.482% | -5.177% | 17.97% | -7.674% | 6.773% | -10.96% | -1.432% | 11.71% | -1.351%
| Z-Score (Absolute) |  | 0.77 | 0.43 | 0.04 | 0.53 | 0.78 | 1.31 | 0.04 | NA | 0.44 | 2.65 | 0.16 | 0.5 | 1.27 | 2.99 | 1.73 | 0.93 | 2.34 | 0.58 | NA | 0.57
Results from laboratories 9 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 702 (Indicative Value 0.97 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 06 | 07 | 10 | 14 | 15 | 17 | 18 | 22 | 23 | 25 | [10 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 0.5 | < 1 | 1.3 | 1 | 3.5 | 2 | 1.1163 | < 0.05 | < 10 | < 1 | < 2 | [10 empty] | 1
| 1 | 2 | 0.6 | 1 | 1.1 | 1 | 3.2 | 1 | 1.1286 | < 0.05 | < 10 | < 1 | < 2 | [10 empty] | 14
| 1 | 3 | 1 | < 1 | 1.1 | < 1 | 2.7 | 2 | 1.1212 | < 0.05 | < 10 | < 1 | < 2 | [10 empty] | 16
| 1 | 4 | 1 | < 1 | 1.1 | 1 | 2.9 | 1 | 1.1241 | < 0.05 | < 10 | 1 | < 2 | [10 empty] | 0.96871
| 1 | 5 | 1.2 | < 1 | 1.3 | 1 | 2.5 | 1 | 1.1637 | < 0.05 | < 10 | 1 | < 2 | [10 empty] | 20
| 1 | 6 | 1.3 | < 1 | 1.2 | 1 | 2.4 | 1 | 1.1232 | < 0.05 | < 10 | 1 | < 2
| Mean |  | 1.02 | 1 | 1.1833 | 1 | 2.8667 | 1.3333 | 1.1295 |  |  | 1
| Median |  | 1 | 1 | 1.15 | 1 | 2.8 | 1 | 1.1236 |  |  | 1
| Std Dev. |  | 0.26833 |  | 0.098319 | 0 | 0.42269 | 0.5164 | 0.01722 |  |  | 0
| Rel.Std.Dev. |  | 26.31% |  | 8.309% | 0% | 14.74% | 38.73% | 1.525% |  |  | 0%
| PDM3 |  | 5.294% | 3.23% | 22.16% | 3.23% | 195.9% | 37.64% | 16.6% |  |  | 3.23%
| Z-Score (Absolute) |  | 0.29 | 1.12 | 0.49 | 0.22 | 5.03 | 0.9 | 0.35 | 2.63 | 10.79 | 0.67 | 0
| Analytical results for Hf in OREAS 702 (Certified Value 4.52 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | [6 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 4.21 | 4.05 | 5 | 5.02 | 4.67 | 4.2 | 4.4 | 4.6 | 20 | 3 | 3.155 | 5 | 11.9577 | 4 | 11 | [6 empty] | 1
| 1 | 2 | 4.74 | 4.28 | 5.08 | 4.52 | 5 | 4.3 | 4.3 | 4.3 | 17 | 4 | 3.145 | 4.9 | 12.0147 | 4 | 6 | [6 empty] | #N/A
| 1 | 3 | 4.43 | 4.2 | 4.94 | 4.8 | 4.76 | 4.3 | 4.7 | 4.4 | 21 | 3 | 3.0792 | 4.8 | 12.0049 | 5 | 5 | [6 empty] | 16
| 1 | 4 | 4.17 | 4.35 | 4.75 | 4.68 | 4.42 | 4.1 | 4.2 | 4.5 | 16 | 3 | 3.0802 | 4.8 | 11.6737 | 5 | 6 | [6 empty] | 4.5206
| 1 | 5 | 4.65 | 4.04 | 4.8 | 4.84 | 4.37 | 4.2 | 4.1 | 4.5 | 15 | 3 | 3.1241 | 4.7 | 11.9178 | 4 | 6 | [6 empty] | 55
| 1 | 6 | 4.53 | 4.24 | 4.63 | 4.67 | 4.37 | 4.3 | 3.9 | 4.3 | 12 | 8 | 3.3323 | 5.1 | 11.9351 | 4 | 5
| Mean |  | 4.455 | 4.1933 | 4.8667 | 4.755 | 4.5983 | 4.2333 | 4.2667 | 4.4333 | 16.8333 | 4 | 3.1526 | 4.8833 | 11.9173 | 4.3333 | 6.5
| Median |  | 4.48 | 4.22 | 4.87 | 4.74 | 4.545 | 4.25 | 4.25 | 4.45 | 16.5 | 3 | 3.1346 | 4.85 | 11.9464 | 4 | 6
| Std Dev. |  | 0.23098 | 0.12517 | 0.16896 | 0.17178 | 0.25639 | 0.08165 | 0.27325 | 0.12111 | 3.3116 | 2 | 0.093598 | 0.1472 | 0.12526 | 0.5164 | 2.2583
| Rel.Std.Dev. |  | 5.185% | 2.985% | 3.472% | 3.613% | 5.576% | 1.929% | 6.404% | 2.732% | 19.67% | 50% | 2.969% | 3.014% | 1.051% | 11.92% | 34.74%
| PDM3 |  | -1.45% | -7.239% | 7.656% | 5.186% | 1.721% | -6.354% | -5.616% | -1.929% | 272.4% | -11.52% | -30.26% | 8.025% | 163.6% | -4.142% | 43.79%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.06 | 0.77 | 0.37 | 0.57 | 0.49 | 0.06 | NA | NA | 3.36 | 1.1 | 19.23 | NA | NA
Results from laboratories 14, 15, 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 702 (Certified Value 0.84 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 0.82 | 0.81 | 0.87 | 0.88 | 0.78 | 0.8 | 1 | < 2 | 0.8 | 0.77517 | 0.81 | 1.3 | 0.76 | 1.1069 | 0.52 | 0.92476 | 0.74 | 1 | 0.88 | 0.82521 |  | 1
| 1 | 2 | 0.82 | 0.8 | 0.94 | 0.88 | 0.9 | 0.8 | 1 | < 2 | 0.8 | 0.76053 | 0.82 | 0.96 | 0.86 | 0.9652 | 0.52 | 0.86065 | 0.85 | 0.9 | 0.84 | 0.81518 |  | #N/A
| 1 | 3 | 0.82 | 0.76 | 0.96 | 0.82 | 0.78 | 0.8 | 0.7 | < 2 | 0.9 | 0.79973 | 0.81 | 1.07 | 0.73 | 0.9946 | 0.52 | 0.85261 | 0.85 | 0.8 | 0.87 | 0.84671 |  | 16
| 1 | 4 | 0.81 | 0.81 | 0.94 | 0.8 | 0.82 | 0.8 | 0.9 | < 2 | 0.9 | 0.78441 | 0.82 | 0.84 | 0.8 | 0.9573 | 0.52 | 0.86199 | 0.76 | 0.9 | 0.88 | 0.80504 |  | 0.84143
| 1 | 5 | 0.84 | 0.84 | 0.98 | 0.91 | 0.87 | 0.8 | 0.8 | < 2 | 0.9 | 0.78121 | 0.83 | 0.76 | 0.77 | 0.9802 | 0.52 | 0.89345 | 0.68 | 0.8 | 0.85 | 0.80597 |  | 56
| 1 | 6 | 0.84 | 0.75 | 0.92 | 0.91 | 0.86 | 0.8 | 0.6 | < 2 | 0.9 | 0.77157 | 0.83 | 0.65 | 0.69 | 0.9091 | 0.52 | 0.89767 | 0.78 | 0.9 | 0.86 | 0.80419
| Mean |  | 0.825 | 0.795 | 0.935 | 0.86667 | 0.835 | 0.8 | 0.83333 |  | 0.86667 | 0.77877 | 0.82 | 0.93 | 0.76833 | 0.98555 | 0.52 | 0.88185 | 0.77667 | 0.88333 | 0.86333 | 0.81705
| Median |  | 0.82 | 0.805 | 0.94 | 0.88 | 0.84 | 0.8 | 0.85 |  | 0.9 | 0.77819 | 0.82 | 0.9 | 0.765 | 0.9727 | 0.52 | 0.87772 | 0.77 | 0.9 | 0.865 | 0.81058
| Std Dev. |  | 0.012247 | 0.033912 | 0.037815 | 0.046332 | 0.049699 | 1.2162e-16 | 0.1633 |  | 0.05164 | 0.013226 | 0.0089443 | 0.23358 | 0.058452 | 0.066171 | 0 | 0.028017 | 0.065929 | 0.075277 | 0.01633 | 0.016626
| Rel.Std.Dev. |  | 1.485% | 4.266% | 4.044% | 5.346% | 5.952% | 1.52e-14% | 19.6% |  | 5.958% | 1.698% | 1.091% | 25.12% | 7.608% | 6.714% | 0% | 3.177% | 8.489% | 8.522% | 1.891% | 2.035%
| PDM3 |  | -1.952% | -5.517% | 11.12% | 3% | -0.7636% | -4.923% | -0.9617% |  | 3% | -7.446% | -2.546% | 10.53% | -8.687% | 17.13% | -38.2% | 4.805% | -7.696% | 4.981% | 2.604% | -2.897%
| Z-Score (Absolute) |  | 0.07 | 0.46 | 1.37 | 0.48 | 0.07 | NA | NA | 2.22 | NA | 0.67 | 0.13 | 1.31 | 0.81 | 2.04 | 4.06 | 0.68 | 0.7 | NA | 0.44 | 0.17
Results from laboratories 6, 7, 10 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 702 (Indicative Value 0.14 ppm) | [22 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 06 | 07 | 10 | 14 | 17 | 18 | 23 | 25 | [12 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 0.3 | < 0.2 | 0.09 | < 0.1 | 0.31 | < 10 | 0.067 | < 0.2 | < 0.3 | [12 empty] | 1
| 1 | 2 | < 0.3 | < 0.2 | 0.12 | < 0.1 | 0.32 | < 10 | 0.071 | < 0.2 | < 0.3 | [12 empty] | 11
| 1 | 3 | < 0.3 | < 0.2 | 0.09 | 0.1 | 0.32 | < 10 | 0.061 | < 0.2 | < 0.3 | [12 empty] | 16
| 1 | 4 | < 0.3 | < 0.2 | 0.09 | 0.1 | 0.29 | < 10 | 0.063 | < 0.2 | < 0.3 | [12 empty] | 0.13921
| 1 | 5 | < 0.3 | < 0.2 | 0.14 | 0.1 | 0.29 | < 10 | 0.066 | < 0.2 | < 0.3 | [12 empty] | 17
| 1 | 6 | < 0.3 | < 0.2 | 0.13 | 0.1 | 0.26 | < 10 | 0.063 | < 0.2 | < 0.3
| Mean |  |  |  | 0.11 | 0.1 | 0.29833 |  | 0.065167
| Median |  |  |  | 0.105 | 0.1 | 0.3 |  | 0.0645
| Std Dev. |  |  |  | 0.022804 | 0 | 0.023166 |  | 0.0036009
| Rel.Std.Dev. |  |  |  | 20.73% | 0% | 7.765% |  | 5.526%
| PDM3 |  |  |  | -20.98% | -28.17% | 114.3% |  | -53.19%
| Z-Score (Absolute) |  | 0.67 | 0.17 | 0 | 0.45 | 3.17 | 82.43 | 0.76 | 0.17 | 0.67
| Analytical results for K2O in OREAS 702 (Certified Value 2.32 wt.%) | [22 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | wt.%
|  |  | PF*MS | BF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  |  |  |  | 2
3
| 1 | 1 | 2.313 | 2.22 | 2.3 | 2.32 | 2.0396 | 2.361 | 2.35 | 2.939 | 2.1382 | 2.36 | 2.37 | 2.3124 | 2.31 | 2.433 | 2.29 | 2.277 | 2.0413 |  |  |  |  | 1
| 1 | 2 | 2.337 | 2.19 | 2.27 | 2.33 | 2.0184 | 2.373 | 2.32 | 2.915 | 2.1498 | 2.355 | 2.42 | 2.4351 | 2.36 | 2.433 | 2.29 | 2.373 | 2.05 |  |  |  |  | #N/A
| 1 | 3 | 2.337 | 2.21 | 2.29 | 2.27 | 2.1247 | 2.325 | 2.34 | 2.939 | 2.1373 | 2.397 | 2.39 | 2.4669 | 2.39 | 2.229 | 2.29 | 2.241 | 2.0202 |  |  |  |  | 16
| 1 | 4 | 2.373 | 2.24 | 2.3 | 2.34 | 2.2059 | 2.325 | 2.29 | 2.807 | 2.1419 | 2.348 | 2.37 | 2.4548 | 2.31 | 2.301 | 2.29 | 2.289 | 2.0092 |  |  |  |  | 2.3156
| 1 | 5 | 2.289 | 2.2 | 2.32 | 2.28 | 2.0929 | 2.325 | 2.29 | 2.807 | 2.1915 | 2.36 | 2.36 | 2.5116 | 2.29 | 2.433 | 2.29 | 2.265 | 2.0276 |  |  |  |  | 57
| 1 | 6 | 2.289 | 2.16 | 2.33 | 2.32 | 2.0897 | 2.313 | 2.31 | 2.807 | 2.1371 | 2.458 | 2.39 | 2.4364 | 2.32 | 2.397 | 2.29 | 2.289 | 2.0317
| Mean |  | 2.323 | 2.2033 | 2.3017 | 2.31 | 2.0952 | 2.337 | 2.3167 | 2.869 | 2.1493 | 2.3797 | 2.3833 | 2.4362 | 2.33 | 2.371 | 2.29 | 2.289 | 2.03
| Median |  | 2.325 | 2.205 | 2.3 | 2.32 | 2.0913 | 2.325 | 2.315 | 2.861 | 2.14 | 2.36 | 2.38 | 2.4456 | 2.315 | 2.415 | 2.29 | 2.283 | 2.0297
| Std Dev. |  | 0.03257 | 0.027325 | 0.02137 | 0.028284 | 0.06653 | 0.024 | 0.025033 | 0.068481 | 0.021239 | 0.042003 | 0.021602 | 0.066806 | 0.037417 | 0.086339 | 4.8648e-16 | 0.0449 | 0.014613
| Rel.Std.Dev. |  | 1.402% | 1.24% | 0.9284% | 1.224% | 3.175% | 1.027% | 1.081% | 2.387% | 0.9882% | 1.765% | 0.9064% | 2.742% | 1.606% | 3.641% | 2.124e-14% | 1.962% | 0.7198%
| PDM3 |  | 0.3193% | -4.849% | -0.602% | -0.2421% | -9.518% | 0.9239% | 0.04575% | 23.9% | -7.182% | 2.766% | 2.925% | 5.208% | 0.6216% | 2.392% | -1.106% | -1.149% | -12.33%
| Z-Score (Absolute) |  | 0.08 | 1.41 | 0.19 | 0.08 | 2.75 | 0.25 | 0 | 6.85 | 2.07 | 0.78 | 0.83 | 1.48 | 0.17 | 0.67 | 0.35 | 0.35 | 3.55
| Analytical results for La in OREAS 702 (Certified Value 27.8 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 1
2
| 1 | 1 | 29 | 27.6 | 29.9 | 28.5 | 28.1 | 26.9 | 26.7 | 33 | 27.8 | 32.2053 | 27.1 | 28.77 | 25.8 | 29.4123 | 33.2 | 28.794 | 27 | 29 | 28.99 | 27.9408 |  | 1
| 1 | 2 | 26.9 | 27.8 | 31.4 | 29 | 28.9 | 27.4 | 29.8 | 34 | 28.2 | 32.2266 | 27.8 | 29.5 | 27.6 | 26.136 | 33 | 27.8897 | 26 | 27 | 28.61 | 27.7909 |  | #N/A
| 1 | 3 | 27.3 | 26.2 | 29.5 | 28.1 | 27.7 | 27.8 | 28.2 | 33 | 27.7 | 32.2013 | 27.5 | 29.31 | 26.6 | 27.1274 | 33.3 | 28.0544 | 27 | 28 | 28.29 | 29.4973 |  | 16
| 1 | 4 | 26.9 | 28 | 29.7 | 28.2 | 28.2 | 27.7 | 27.1 | 35 | 27.9 | 32.3582 | 27.7 | 27.47 | 25.2 | 28.6572 | 32.3 | 27.2553 | 26 | 27 | 29.46 | 27.4386 |  | 27.8141
| 1 | 5 | 27.8 | 27 | 30.5 | 29.5 | 27.8 | 26.6 | 25.8 | 33 | 28 | 32.2655 | 27.7 | 27.98 | 26.5 | 27.5444 | 32.7 | 28.214 | 26 | 27 | 29.23 | 28.0517 |  | 58
| 1 | 6 | 27.7 | 26.6 | 27.9 | 29.2 | 27.1 | 27.2 | 25.3 | 33 | 28.4 | 32.2572 | 27 | 28.03 | 25.3 | 28.6569 | 32.4 | 28.9071 | 26 | 27 | 28.22 | 30.407
| Mean |  | 27.6 | 27.2 | 29.8167 | 28.75 | 27.9667 | 27.2667 | 27.15 | 33.5 | 28 | 32.2524 | 27.4667 | 28.51 | 26.1667 | 27.9224 | 32.8167 | 28.1857 | 26.3333 | 27.5 | 28.8 | 28.521
| Median |  | 27.5 | 27.3 | 29.8 | 28.75 | 27.95 | 27.3 | 26.9 | 33 | 27.95 | 32.2419 | 27.6 | 28.4 | 26.15 | 28.1007 | 32.85 | 28.1342 | 26 | 27 | 28.8 | 27.9962
| Std Dev. |  | 0.78486 | 0.71554 | 1.1635 | 0.56833 | 0.59889 | 0.46332 | 1.6477 | 0.83666 | 0.26077 | 0.058125 | 0.33862 | 0.81001 | 0.9136 | 1.2054 | 0.41673 | 0.61019 | 0.5164 | 0.83666 | 0.50778 | 1.1637
| Rel.Std.Dev. |  | 2.844% | 2.631% | 3.902% | 1.977% | 2.141% | 1.699% | 6.069% | 2.497% | 0.9313% | 0.1802% | 1.233% | 2.841% | 3.491% | 4.317% | 1.27% | 2.165% | 1.961% | 3.042% | 1.763% | 4.08%
| PDM3 |  | -0.7696% | -2.208% | 7.2% | 3.365% | 0.5486% | -1.968% | -2.388% | 20.44% | 0.6685% | 15.96% | -1.249% | 2.502% | -5.923% | 0.3894% | 17.99% | 1.336% | -5.324% | -1.129% | 3.545% | 2.542%
| Z-Score (Absolute) |  | 0.35 | 0.71 | 1.67 | 0.7 | 0.02 | 0.65 | 0.76 | 5.02 | 0.02 | 3.88 | 0.47 | 0.48 | 1.65 | 0.06 | 4.39 | 0.18 | 1.5 | 0.44 | 0.74 | 0.49
| Analytical results for Li in OREAS 702 (Certified Value 37.9 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 09 | 10 | 11 | 14 | 15 | 17 | 18 | 22 | 23 | 24 | 25 | 26 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [8 empty] | 1
2
| 1 | 1 | 39 | 42 | 37 | 39.4353 | 43.44 | 41 | 60 | 37.7 | 33.47 | 42 | 40 | 36 | 33.58 | [8 empty] | 1
| 1 | 2 | 39 | 44 | 37 | 39.5936 | 40.35 | 44 | 40 | 37.5 | 31.64 | 42 | 40 | 34 | 34.4 | [8 empty] | #N/A
| 1 | 3 | 38 | 41 | 37 | 39.0956 | 39.75 | 40 | 40 | 38.3 | 33.79 | 38 | 40 | 33 | 33.67 | [8 empty] | 16
| 1 | 4 | 39 | 43 | 37 | 39.3121 | 39.13 | 41 | 40 | 38 | 32.53 | 39 | 40 | 33 | 34.1 | [8 empty] | 37.9487
| 1 | 5 | 39 | 41 | 37 | 38.9806 | 39.14 | 44 | 40 | 38.1 | 31.62 | 41 | 40 | 32 | 33.92 | [8 empty] | 59
| 1 | 6 | 40 | 40 | 37 | 38.9366 | 38.27 | 42 | 40 | 38 | 31.38 | 41 | 40 | 33 | 34.48
| Mean |  | 39 | 41.8333 | 37 | 39.2257 | 40.0133 | 42 | 43.3333 | 37.9333 | 32.405 | 40.5 | 40 | 33.5 | 34.025
| Median |  | 39 | 41.5 | 37 | 39.2039 | 39.445 | 41.5 | 40 | 38 | 32.085 | 41 | 40 | 33 | 34.01
| Std Dev. |  | 0.63246 | 1.472 | 0 | 0.26356 | 1.817 | 1.6733 | 8.165 | 0.28752 | 1.0315 | 1.6432 | 0 | 1.3784 | 0.37104
| Rel.Std.Dev. |  | 1.622% | 3.519% | 0% | 0.6719% | 4.541% | 3.984% | 18.84% | 0.758% | 3.183% | 4.057% | 0% | 4.115% | 1.09%
| PDM3 |  | 2.77% | 10.24% | -2.5% | 3.365% | 5.441% | 10.68% | 14.19% | -0.04048% | -14.61% | 6.723% | 5.405% | -11.72% | -10.34%
| Z-Score (Absolute) |  | 0 | 0.96 | 0.67 | 0.08 | 0.34 | 1.01 | NA | 0.36 | 2.22 | 0.51 | NA | 1.85 | 1.68
Results from laboratories 17 and 24 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 702 (Certified Value 0.29 ppm) | [22 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 0.31 | 0.31 | 0.24 | 0.32 | 0.27 | 0.29 | 0.3 | < 2 | 0.28 | 0.25313 | 0.28 | 0.6 | 0.25 | 0.3862 | 0.13 | 0.31722 | 0.3 | < 0.5 | 0.31 | < 0.5 |  | 1
| 1 | 2 | 0.31 | 0.27 | 0.31 | 0.29 | 0.26 | 0.29 | 0.3 | < 2 | 0.28 | 0.25903 | 0.27 | 0.49 | 0.31 | 0.3491 | 0.13 | 0.3644 | 0.4 | < 0.5 | 0.32 | < 0.5 |  | #N/A
| 1 | 3 | 0.28 | 0.3 | 0.32 | 0.3 | 0.3 | 0.31 | 0.3 | < 2 | 0.28 | 0.24333 | 0.29 | 0.44 | 0.21 | 0.3531 | 0.13 | 0.35716 | 0.3 | < 0.5 | 0.3 | < 0.5 |  | 16
| 1 | 4 | 0.28 | 0.29 | 0.28 | 0.26 | 0.32 | 0.3 | 0.3 | < 2 | 0.29 | 0.26913 | 0.27 | 0.4 | 0.34 | 0.342 | 0.13 | 0.33747 | 0.3 | < 0.5 | 0.31 | < 0.5 |  | 0.28664
| 1 | 5 | 0.28 | 0.29 | 0.29 | 0.3 | 0.3 | 0.27 | 0.3 | < 2 | 0.28 | 0.25143 | 0.29 | 0.38 | 0.22 | 0.3506 | 0.13 | 0.33769 | 0.2 | < 0.5 | 0.3 | < 0.5 |  | 60
| 1 | 6 | 0.29 | 0.32 | 0.29 | 0.27 | 0.28 | 0.29 | 0.3 | < 2 | 0.29 | 0.25233 | 0.29 | 0.34 | 0.28 | 0.3194 | 0.13 | 0.33782 | 0.3 | < 0.5 | 0.3 | < 0.5
| Mean |  | 0.29167 | 0.29667 | 0.28833 | 0.29 | 0.28833 | 0.29167 | 0.3 |  | 0.28333 | 0.25473 | 0.28167 | 0.44167 | 0.26833 | 0.35007 | 0.13 | 0.34196 | 0.3 |  | 0.30667
| Median |  | 0.285 | 0.295 | 0.29 | 0.295 | 0.29 | 0.29 | 0.3 |  | 0.28 | 0.25273 | 0.285 | 0.42 | 0.265 | 0.34985 | 0.13 | 0.33775 | 0.3 |  | 0.305
| Std Dev. |  | 0.01472 | 0.017512 | 0.027869 | 0.021909 | 0.022286 | 0.013292 | 0 |  | 0.005164 | 0.0086605 | 0.0098319 | 0.093041 | 0.051153 | 0.021548 | 0 | 0.016748 | 0.063246 |  | 0.008165
| Rel.Std.Dev. |  | 5.047% | 5.903% | 9.665% | 7.555% | 7.729% | 4.557% | 0% |  | 1.823% | 3.4% | 3.491% | 21.07% | 19.06% | 6.155% | 0% | 4.898% | 21.08% |  | 2.662%
| PDM3 |  | 1.753% | 3.497% | 0.5899% | 1.171% | 0.5899% | 1.753% | 4.66% |  | -1.154% | -11.13% | -1.736% | 54.08% | -6.387% | 22.13% | -54.65% | 19.3% | 4.66% |  | 6.986%
| Z-Score (Absolute) |  | 0.09 | 0.26 | 0.03 | 0.03 | 0.03 | 0.09 | NA | 25.01 | 0.21 | 1.21 | 0.26 | 5.37 | 0.73 | 2.14 | 5.6 | 1.86 | NA | 1.38 | 0.62 | 1.38
Results from laboratories 7 and 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for MgO in OREAS 702 (Certified Value 1.01 wt.%) | [22 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | PF*MS | BF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  |  |  | 2
3
| 1 | 1 | 0.962 | 1.02 | 1.03 | 1.04 | 1.0057 | 0.978 | 1.0097 | 1.05 | 0.879 | 0.99475 | 1.066 | 0.995 | 0.96132 | 1.09 | 1.111 | 1.012 | 0.978 | 0.94262 |  |  |  | 1
| 1 | 2 | 0.945 | 1.02 | 1.02 | 1.03 | 1.0011 | 1.012 | 1.0038 | 1.03 | 0.945 | 0.96756 | 1.061 | 0.978 | 0.95539 | 1.06 | 1.1259 | 0.995 | 1.028 | 0.93045 |  |  |  | #N/A
| 1 | 3 | 0.978 | 1.01 | 1.02 | 1 | 1.0236 | 0.995 | 1.0121 | 1.04 | 0.912 | 0.98447 | 1.074 | 0.995 | 0.95964 | 1.1 | 1.0331 | 0.978 | 0.978 | 0.92945 |  |  |  | 16
| 1 | 4 | 0.978 | 1.03 | 1.04 | 1.05 | 1.0619 | 1.012 | 1.0114 | 1.02 | 0.945 | 0.98779 | 1.061 | 0.995 | 0.94827 | 1.05 | 1.0463 | 0.995 | 1.012 | 0.95358 |  |  |  | 1.0051
| 1 | 5 | 0.945 | 1.02 | 1.03 | 1.02 | 1.0123 | 0.995 | 1.011 | 1.02 | 0.912 | 0.98497 | 1.066 | 0.995 | 0.94589 | 1.02 | 1.1127 | 1.012 | 0.978 | 0.93419 |  |  |  | 61
| 1 | 6 | 0.978 | 1 | 1.01 | 1.03 | 1.0173 | 0.995 | 1.0013 | 1.03 | 0.962 | 0.95131 | 1.137 | 0.995 | 0.95162 | 1.03 | 1.1044 | 1.012 | 0.995 | 0.95849
| Mean |  | 0.96433 | 1.0167 | 1.025 | 1.0283 | 1.0203 | 0.99783 | 1.0082 | 1.0317 | 0.92583 | 0.97847 | 1.0775 | 0.99217 | 0.95369 | 1.0583 | 1.0889 | 1.0007 | 0.99483 | 0.94146
| Median |  | 0.97 | 1.02 | 1.025 | 1.03 | 1.0148 | 0.995 | 1.0104 | 1.03 | 0.9285 | 0.98472 | 1.066 | 0.995 | 0.95351 | 1.055 | 1.1077 | 1.0035 | 0.9865 | 0.9384
| Std Dev. |  | 0.016207 | 0.010328 | 0.010488 | 0.017224 | 0.0219 | 0.012797 | 0.004523 | 0.01169 | 0.030394 | 0.016043 | 0.029535 | 0.0069402 | 0.0061782 | 0.031885 | 0.038968 | 0.01388 | 0.021189 | 0.012301
| Rel.Std.Dev. |  | 1.681% | 1.016% | 1.023% | 1.675% | 2.146% | 1.282% | 0.4486% | 1.133% | 3.283% | 1.64% | 2.741% | 0.6995% | 0.6478% | 3.013% | 3.579% | 1.387% | 2.13% | 1.307%
| PDM3 |  | -4.058% | 1.148% | 1.978% | 2.309% | 1.512% | -0.7253% | 0.3096% | 2.641% | -7.889% | -2.651% | 7.201% | -1.289% | -5.117% | 5.294% | 8.335% | -0.4434% | -1.024% | -6.333%
| Z-Score (Absolute) |  | 1.08 | 0.33 | 0.56 | 0.65 | 0.43 | 0.18 | 0.11 | 0.74 | 2.12 | 0.7 | 1.98 | 0.33 | 1.37 | 1.46 | 2.29 | 0.11 | 0.25 | 1.7
| Analytical results for MnO in OREAS 702 (Certified Value 0.04 wt.%) | [22 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | PF*MS | BF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  |  |  | 3
3
| 1 | 1 | 0.0387 | 0.04 | 0.04 | 0.04 | 0.04132 | 0.0413 | 0.03793 | 0.04 | 0.04571 | 0.04145 | 0.040954 | 0.04 | 0.042684 | 0.04 | 0.04455 | 0.04 | 0.04132 | 0.040331 |  |  |  | 1
| 1 | 2 | 0.0387 | 0.04 | 0.04 | 0.04 | 0.04339 | 0.0413 | 0.037842 | 0.04 | 0.04403 | 0.04093 | 0.040069 | 0.05 | 0.043101 | 0.04 | 0.04545 | 0.04 | 0.04158 | 0.039584 |  |  |  | #N/A
| 1 | 3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04313 | 0.0413 | 0.038019 | 0.04 | 0.04377 | 0.03964 | 0.040301 | 0.04 | 0.043427 | 0.04 | 0.04235 | 0.04 | 0.04145 | 0.040047 |  |  |  | 16
| 1 | 4 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04403 | 0.04 | 0.038333 | 0.04 | 0.04274 | 0.04145 | 0.039797 | 0.04 | 0.042274 | 0.04 | 0.04196 | 0.04 | 0.04196 | 0.040613 |  |  |  | 0.040068
| 1 | 5 | 0.0387 | 0.04 | 0.04 | 0.04 | 0.04145 | 0.04 | 0.038129 | 0.04 | 0.04313 | 0.04145 | 0.040174 | 0.04 | 0.043483 | 0.04 | 0.04442 | 0.04 | 0.04313 | 0.039545 |  |  |  | 62
| 1 | 6 | 0.0387 | 0.04 | 0.04 | 0.04 | 0.04287 | 0.0413 | 0.038578 | 0.04 | 0.04429 | 0.03964 | 0.040558 | 0.04 | 0.042976 | 0.04 | 0.04403 | 0.04 | 0.04106 | 0.041035
| Mean |  | 0.039133 | 0.04 | 0.04 | 0.04 | 0.042698 | 0.040867 | 0.038139 | 0.04 | 0.043945 | 0.04076 | 0.040309 | 0.041667 | 0.042991 | 0.04 | 0.043793 | 0.04 | 0.04175 | 0.040193
| Median |  | 0.0387 | 0.04 | 0.04 | 0.04 | 0.043 | 0.0413 | 0.038074 | 0.04 | 0.0439 | 0.04119 | 0.040238 | 0.04 | 0.043038 | 0.04 | 0.044225 | 0.04 | 0.041515 | 0.040189
| Std Dev. |  | 0.00067132 | 0 | 0 | 0 | 0.0010886 | 0.00067132 | 0.00027447 | 0 | 0.0010381 | 0.00089062 | 0.00040396 | 0.0040825 | 0.0004589 | 0 | 0.0013572 | 0 | 0.0007387 | 0.00058598
| Rel.Std.Dev. |  | 1.715% | 0% | 0% | 0% | 2.55% | 1.643% | 0.7197% | 0% | 2.362% | 2.185% | 1.002% | 9.798% | 1.067% | 0% | 3.099% | 0% | 1.769% | 1.458%
| PDM3 |  | -2.334% | -0.1709% | -0.1709% | -0.1709% | 6.563% | 1.992% | -4.816% | -0.1709% | 9.675% | 1.726% | 0.5998% | 3.989% | 7.293% | -0.1709% | 9.296% | -0.1709% | 4.197% | 0.31%
| Z-Score (Absolute) |  | 1.27 | 0.29 | 0.29 | 0.29 | 2.88 | 0.75 | 2.48 | 0.29 | 4.34 | 0.6 | 0.07 | 1.66 | 3.22 | 0.29 | 4.17 | 0.26 | 1.76 | 0.07
| Analytical results for Mo in OREAS 702 (Certified Value 588 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  | 0
0
| 1 | 1 | 667 | 636 | 587 | 634 | 561 | 614 | 584 | 584 | 568.94 | 616 | 512 | 505.9538 | 555.93 | 571.2828 | 635.7267 | 548 | 609 | 600 | 582.6272 |  |  | 1
| 1 | 2 | 650 | 633 | 592 | 629 | 553 | 593 | 587 | 594 | 566.71 | 613 | 508 | 501.3679 | 561.59 | 561.0366 | 619.3284 | 538 | 614 | 650 | 576.9826 |  |  | 21
| 1 | 3 | 666 | 633 | 593 | 616 | 560 | 611 | 600 | 592 | 566.96 | 635 | 517 | 505.7886 | 561.74 | 564.5545 | 656.2986 | 551 | 620 | 593 | 593.6896 |  |  | 16
| 1 | 4 | 656 | 638 | 582 | 635 | 559 | 605 | 611 | 597 | 565.87 | 595 | 514 | 499.4309 | 558.64 | 550.3948 | 627.0693 | 552 | 634 | 594 | 579.2999 |  |  | 587.5244
| 1 | 5 | 684 | 636 | 574 | 625 | 554 | 602 | 596 | 599 | 567.68 | 598 | 515 | 502.4825 | 560.5 | 567.6769 | 602.7468 | 543 | 614 | 606 | 585.9892 |  |  | 63
| 1 | 6 | 649 | 643 | 566 | 614 | 574 | 601 | 591 | 597 | 568.44 | 587 | 492 | 500.7866 | 553.12 | 580.4644 | 604.89 | 549 | 637 | 593 | 620.5873
| Mean |  | 662 | 636.5 | 582.3333 | 625.5 | 560.1667 | 604.3333 | 594.8333 | 593.8333 | 567.4333 | 607.3333 | 509.6667 | 502.635 | 558.5867 | 565.9017 | 624.3433 | 546.8333 | 621.3333 | 606 | 589.8626
| Median |  | 661 | 636 | 584.5 | 627 | 559.5 | 603.5 | 593.5 | 595.5 | 567.32 | 605.5 | 513 | 501.9252 | 559.57 | 566.1157 | 623.1988 | 548.5 | 617 | 597 | 584.3082
| Std Dev. |  | 13.2212 | 3.7283 | 10.6333 | 8.9163 | 7.5211 | 7.5277 | 9.8268 | 5.4191 | 1.1433 | 17.4662 | 9.1797 | 2.6934 | 3.4517 | 10.1008 | 20.1402 | 5.3448 | 11.5528 | 22.154 | 16.146
| Rel.Std.Dev. |  | 1.997% | 0.5857% | 1.826% | 1.425% | 1.343% | 1.246% | 1.652% | 0.9126% | 0.2015% | 2.876% | 1.801% | 0.5359% | 0.6179% | 1.785% | 3.226% | 0.9774% | 1.859% | 3.656% | 2.737%
| PDM3 |  | 12.68% | 8.336% | -0.8835% | 6.464% | -4.656% | 2.861% | 1.244% | 1.074% | -3.42% | 3.372% | -13.25% | -14.45% | -4.925% | -3.68% | 6.267% | -6.926% | 5.754% | 3.145% | 0.398%
| Z-Score (Absolute) |  | 1.65 | 1.03 | 0.28 | 0.76 | 0.81 | 0.25 | 0.02 | 0 | 0.64 | 0.33 | 2.03 | 2.2 | 0.85 | 0.67 | 0.74 | 1.13 | 0.66 | 0.29 | 0.1
| Analytical results for Na2O in OREAS 702 (Certified Value 1.77 wt.%) | [22 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 05 | 06 | 07 | 13 | 18 | [16 empty] | wt.%
|  |  | BF*OES | BF*OES/MS | BF*OES | BF*OES/MS | PF*OES/MS | [16 empty] | 2
3
| 1 | 1 | 1.78 | 1.82 | 1.8 | 1.74 | 1.766 | [16 empty] | 1
| 1 | 2 | 1.75 | 1.79 | 1.8 | 1.73 | 1.698 | [16 empty] | #N/A
| 1 | 3 | 1.76 | 1.8 | 1.75 | 1.73 | 1.82 | [16 empty] | 16
| 1 | 4 | 1.8 | 1.81 | 1.82 | 1.7 | 1.698 | [16 empty] | 1.7683
| 1 | 5 | 1.76 | 1.83 | 1.77 | 1.7 | 1.739 | [16 empty] | 64
| 1 | 6 | 1.74 | 1.83 | 1.8 | 1.71 | 1.806
| Mean |  | 1.765 | 1.8133 | 1.79 | 1.7183 | 1.7545
| Median |  | 1.76 | 1.815 | 1.8 | 1.72 | 1.7525
| Std Dev. |  | 0.021679 | 0.01633 | 0.025298 | 0.017224 | 0.052344
| Rel.Std.Dev. |  | 1.228% | 0.9005% | 1.413% | 1.002% | 2.983%
| PDM3 |  | -0.1845% | 2.549% | 1.229% | -2.824% | -0.7783%
| Z-Score (Absolute) |  | 0 | 1.3 | 0.67 | 1.26 | 0.28
| Analytical results for Nb in OREAS 702 (Certified Value 8.78 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 17 | 18 | 19 | 23 | 25 | 26 | [5 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | BF*OES/MS | PF*OES | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
3
| 1 | 1 | 9.05 | 8.8 | 8.78 | 8.81 | 8.96 | 10 | 10 | 10 | 9.2 | < 10 | 8.7809 | 3 | 8.5628 | 12 | 12 | 9.4534 | [5 empty] | 1
| 1 | 2 | 8.41 | 8.6 | 8.8 | 8.84 | 8.34 | 10 | 9.3 | 11 | 8.7 | < 10 | 8.6045 | 3.2 | 8.2624 | 9 | 12 | 9.3436 | [5 empty] | 1
| 1 | 3 | 8.93 | 8.6 | 8.87 | 8.98 | 8 | 9 | 9.7 | 11 | 8.8 | < 10 | 8.8655 | 3 | 8.5432 | 8 | 14 | 8.92 | [5 empty] | 16
| 1 | 4 | 8.78 | 8.6 | 8.72 | 8.58 | 8.26 | 9 | 8.8 | 9 | 9.3 | < 10 | 8.422 | 2.9 | 7.6992 | 8 | 12 | 8.4379 | [5 empty] | 8.7778
| 1 | 5 | 8.94 | 9.3 | 8.84 | 9.02 | 8.16 | 8 | 8.8 | 9 | 8.6 | < 10 | 8.3738 | 2.9 | 7.5548 | 8 | 13 | 9.7392 | [5 empty] | 65
| 1 | 6 | 8.77 | 9.4 | 8.9 | 9.01 | 7.91 | 10 | 9.3 | 10 | 8.4 | < 10 | 8.6937 | 3.1 | 8.3751 | 8 | 12 | 13.2737
| Mean |  | 8.8133 | 8.8833 | 8.8183 | 8.8733 | 8.2717 | 9.3333 | 9.3167 | 10 | 8.8333 |  | 8.6234 | 3.0167 | 8.1663 | 8.8333 | 12.5 | 9.8613
| Median |  | 8.855 | 8.7 | 8.82 | 8.91 | 8.21 | 9.5 | 9.3 | 10 | 8.75 |  | 8.6491 | 3 | 8.3188 | 8 | 12 | 9.3985
| Std Dev. |  | 0.2242 | 0.37103 | 0.065243 | 0.16896 | 0.37301 | 0.8165 | 0.47924 | 0.89443 | 0.35024 |  | 0.19574 | 0.1169 | 0.43456 | 1.6021 | 0.83666 | 1.7324
| Rel.Std.Dev. |  | 2.544% | 4.177% | 0.7399% | 1.904% | 4.509% | 8.748% | 5.144% | 8.944% | 3.965% |  | 2.27% | 3.875% | 5.321% | 18.14% | 6.693% | 17.57%
| PDM3 |  | 0.4042% | 1.202% | 0.4612% | 1.088% | -5.767% | 6.328% | 6.138% | 13.92% | 0.6321% |  | -1.76% | -65.63% | -6.967% | 0.6321% | 42.4% | 12.34%
| Z-Score (Absolute) |  | 0 | 0.13 | 0 | 0.11 | 1.06 | NA | 0.97 | NA | 0.03 | 7.42 | 0.37 | 11.28 | 1.26 | NA | NA | 2.03
Results from laboratories 6, 9, 23 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 702 (Certified Value 25.7 ppm) | [22 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 1
2
| 1 | 1 | 26.6 | 24.7 | 28.5 | 25.7 | 27.1 | 25.4 | 26.9 | 28 | 25.2 | 29.0787 | 24 | 23.2 | 24 | 34.0375 | 26.2 | 25.8157 | 24 | 26 | 25.8 | 24.2987 |  | 1
| 1 | 2 | 25.1 | 25.5 | 29 | 27.3 | 27.7 | 25.3 | 26.9 | 26 | 25.4 | 29.4901 | 24.4 | 20.7 | 25.4 | 30.9775 | 26.7 | 25.4519 | 24 | 24 | 25.4 | 25.8374 |  | #N/A
| 1 | 3 | 24.8 | 24 | 28 | 25.7 | 26 | 25.3 | 26.2 | 27 | 25.2 | 29.5557 | 23.9 | 20.4 | 24.7 | 31.0292 | 26.4 | 25.5936 | 26 | 24 | 27.2 | 26.6392 |  | 16
| 1 | 4 | 24.5 | 26.1 | 27.1 | 24.9 | 26 | 25.5 | 26.7 | 28 | 25.9 | 29.4706 | 24.2 | 20.6 | 26.1 | 27.5404 | 26.6 | 25.5614 | 22 | 24 | 26.8 | 24.0994 |  | 25.6724
| 1 | 5 | 25.3 | 24.5 | 28 | 27.2 | 25.8 | 24.8 | 26.5 | 26 | 25.9 | 29.5717 | 24.2 | 18.4 | 26.3 | 31.6566 | 26.4 | 25.8375 | 23 | 24 | 27.7 | 24.3618 |  | 66
| 1 | 6 | 26 | 24.1 | 24.8 | 28.8 | 26.5 | 25 | 26.4 | 28 | 25.5 | 29.4156 | 23.7 | 21.4 | 24.3 | 25.8935 | 26.2 | 26.122 | 23 | 24 | 25.1 | 26.9689
| Mean |  | 25.3833 | 24.8167 | 27.5667 | 26.6 | 26.5167 | 25.2167 | 26.6 | 27.1667 | 25.5167 | 29.4304 | 24.0667 | 20.7833 | 25.1333 | 30.1891 | 26.4167 | 25.7303 | 23.6667 | 24.3333 | 26.3333 | 25.3676
| Median |  | 25.2 | 24.6 | 28 | 26.45 | 26.25 | 25.3 | 26.6 | 27.5 | 25.45 | 29.4803 | 24.1 | 20.65 | 25.05 | 31.0033 | 26.4 | 25.7046 | 23.5 | 24 | 26.3 | 25.0996
| Std Dev. |  | 0.78337 | 0.82563 | 1.4949 | 1.4283 | 0.74677 | 0.26394 | 0.28284 | 0.98319 | 0.31885 | 0.18153 | 0.25033 | 1.5549 | 0.95219 | 2.9583 | 0.20412 | 0.24372 | 1.3663 | 0.8165 | 1.0501 | 1.2778
| Rel.Std.Dev. |  | 3.086% | 3.327% | 5.423% | 5.369% | 2.816% | 1.047% | 1.063% | 3.619% | 1.25% | 0.6168% | 1.04% | 7.481% | 3.789% | 9.799% | 0.7727% | 0.9472% | 5.773% | 3.355% | 3.988% | 5.037%
| PDM3 |  | -1.126% | -3.333% | 7.378% | 3.613% | 3.289% | -1.775% | 3.613% | 5.82% | -0.6067% | 14.64% | -6.255% | -19.04% | -2.1% | 17.59% | 2.899% | 0.2256% | -7.813% | -5.216% | 2.574% | -1.187%
| Z-Score (Absolute) |  | 0.17 | 0.58 | 1.4 | 0.7 | 0.64 | 0.29 | 0.7 | 1.11 | 0.08 | 2.75 | 1.12 | 3.49 | 0.35 | 3.29 | 0.57 | 0.08 | 1.41 | 0.93 | 0.51 | 0.18
| Analytical results for Ni in OREAS 702 (Indicative Value 18.8 ppm) | [22 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 22 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [6 empty] | 1
1
| 1 | 1 | 30 | < 20 | < 20 | 22 | < 20 | < 20 | 18 | 13 | 22.555 | 10 | 18.87 | < 20 | < 100 | 51 | 29.82 | [6 empty] | 1
| 1 | 2 | 20 | < 20 | < 20 | 42 | < 20 | < 20 | 14 | 16 | 16.3859 | 10.4 | 12.62 | < 20 | < 100 | < 50 | 24.62 | [6 empty] | 12
| 1 | 3 | 20 | < 20 | < 20 | 43 | < 20 | < 20 | 20 | 45 | 20.1533 | 11.1 | 15.91 | < 20 | < 100 | < 50 | < 20 | [6 empty] | 16
| 1 | 4 | 20 | < 20 | < 20 | 49 | 20 | < 20 | 18 | 22 | 17.6011 | 10.1 | 15.02 | < 20 | < 100 | 53 | < 20 | [6 empty] | 18.7573
| 1 | 5 | 20 | < 20 | < 20 | 29 | 20 | < 20 | 22 | 15 | 22.504 | 10.7 | 17.7 | < 20 | < 100 | 51 | 26.28 | [6 empty] | 18
| 1 | 6 | 20 | < 20 | < 20 | 47 | < 20 | < 20 | 17 | 25 | 17.9369 | 10.8 | 16.74 | < 20 | < 100 | < 50 | 22.65
| Mean |  | 21.6667 |  |  | 38.6667 | 20 |  | 18.1667 | 22.6667 | 19.5227 | 10.5167 | 16.1433 |  |  | 51.6667 | 25.8425
| Median |  | 20 |  |  | 42.5 | 20 |  | 18 | 19 | 19.0451 | 10.55 | 16.325 |  |  | 51 | 25.45
| Std Dev. |  | 4.0825 |  |  | 10.7455 | 0 |  | 2.7142 | 11.8434 | 2.6278 | 0.42622 | 2.1883 |  |  | 1.1547 | 3.0386
| Rel.Std.Dev. |  | 18.84% |  |  | 27.79% | 0% |  | 14.94% | 52.25% | 13.46% | 4.053% | 13.56% |  |  | 2.235% | 11.76%
| PDM3 |  | 15.51% |  |  | 106.1% | 6.625% |  | -3.149% | 20.84% | 4.08% | -43.93% | -13.94% |  |  | 175.4% | 37.77%
| Z-Score (Absolute) |  | 0.31 | 0.72 | 0.72 | 1.81 | 0.43 | 0.72 | 0 | 0.4 | 0.12 | 0.67 | 0.18 | 0.72 | 2.81 | 1.78 | 0.21
| Analytical results for P2O5 in OREAS 702 (Certified Value 0.133 wt.%) | [22 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | [7 empty] | wt.%
|  |  | BF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | [7 empty] | 3
3
| 1 | 1 | 0.14 | 0.13 | 0.137 | 0.12833 | 0.137 | 0.13 | 0.12902 | 0.08089 | 0.1008 | 0.13658 | 0.1404 | 0.13 | 0.13956 | 0.137 | [7 empty] | 1
| 1 | 2 | 0.13 | 0.13 | 0.137 | 0.12627 | 0.137 | 0.13 | 0.12948 | 0.07677 | 0.1421 | 0.13498 | 0.13673 | 0.12 | 0.13956 | 0.137 | [7 empty] | #N/A
| 1 | 3 | 0.13 | 0.13 | 0.137 | 0.12925 | 0.137 | 0.13 | 0.12627 | 0.08112 | 0.1696 | 0.13612 | 0.13721 | 0.13 | 0.13131 | 0.137 | [7 empty] | 16
| 1 | 4 | 0.14 | 0.14 | 0.137 | 0.13268 | 0.137 | 0.13 | 0.12512 | 0.08158 | 0.149 | 0.13543 | 0.13459 | 0.13 | 0.12787 | 0.16 | [7 empty] | 0.13294
| 1 | 5 | 0.14 | 0.13 | 0.137 | 0.12673 | 0.137 | 0.12 | 0.11848 | 0.08135 | 0.1581 | 0.13658 | 0.13446 | 0.13 | 0.13223 | 0.16 | [7 empty] | 67
| 1 | 6 | 0.13 | 0.12 | 0.137 | 0.12581 | 0.137 | 0.13 | 0.1281 | 0.07746 | 0.1558 | 0.13658 | 0.13678 | 0.13 | 0.13429 | 0.137
| Mean |  | 0.135 | 0.13 | 0.137 | 0.12818 | 0.137 | 0.12833 | 0.12608 | 0.079862 | 0.1459 | 0.13604 | 0.13669 | 0.12833 | 0.13414 | 0.14467
| Median |  | 0.135 | 0.13 | 0.137 | 0.12753 | 0.137 | 0.13 | 0.12718 | 0.081005 | 0.1524 | 0.13635 | 0.13675 | 0.13 | 0.13326 | 0.137
| Std Dev. |  | 0.0054772 | 0.0063246 | 0 | 0.0025607 | 0 | 0.0040825 | 0.0040744 | 0.0021511 | 0.023944 | 0.00068946 | 0.0021641 | 0.0040825 | 0.004685 | 0.011877
| Rel.Std.Dev. |  | 4.057% | 4.865% | 0% | 1.998% | 0% | 3.181% | 3.232% | 2.693% | 16.41% | 0.5068% | 1.583% | 3.181% | 3.493% | 8.21%
| PDM3 |  | 1.55% | -2.212% | 3.054% | -3.582% | 3.054% | -3.465% | -5.162% | -39.93% | 9.749% | 2.336% | 2.823% | -3.465% | 0.9001% | 8.821%
| Z-Score (Absolute) |  | 0.05 | 0.57 | 0.37 | 0.8 | 0.37 | 0.78 | 1.06 | 6.83 | 1.41 | 0.18 | 0.27 | 0.78 | 0.05 | 1.32
| Analytical results for Pb in OREAS 702 (Certified Value 17.6 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 06 | 07 | 10 | 14 | 15 | 17 | 18 | 22 | 23 | 25 | 26 | [9 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*OES | [9 empty] | 1
1
| 1 | 1 | 17.2 | 13 | 19 | < 20 | 16 | < 20 | 19.5831 | 14.7 | 24.83 | 17 | < 100 | < 100 | [9 empty] | 1
| 1 | 2 | 17 | 15 | 18 | < 20 | 22 | < 20 | 19.3458 | 15 | 20.87 | 16 | < 100 | < 100 | [9 empty] | 23
| 1 | 3 | 16.6 | 14 | 18 | < 20 | 22 | < 20 | 19.3067 | 14.8 | 23.35 | 16 | < 100 | < 100 | [9 empty] | 16
| 1 | 4 | 17.8 | 13 | 15 | < 20 | 17 | < 20 | 19.0357 | 14.9 | 19.37 | 16 | < 100 | < 100 | [9 empty] | 17.5514
| 1 | 5 | 16.8 | 14 | 21 | < 20 | 16 | < 20 | 19.711 | 15.2 | 21.69 | 16 | < 100 | < 100 | [9 empty] | 68
| 1 | 6 | 17.6 | 13 | 18 | < 20 | 19 | < 20 | 19.9043 | 14.4 | 23.47 | 16 | < 100 | < 100
| Mean |  | 17.1667 | 13.6667 | 18.1667 |  | 18.6667 |  | 19.4811 | 14.8333 | 22.2633 | 16.1667
| Median |  | 17.1 | 13.5 | 18 |  | 18 |  | 19.4644 | 14.85 | 22.52 | 16
| Std Dev. |  | 0.46332 | 0.8165 | 1.9408 |  | 2.8048 |  | 0.31287 | 0.27325 | 1.9933 | 0.40825
| Rel.Std.Dev. |  | 2.699% | 5.974% | 10.68% |  | 15.03% |  | 1.606% | 1.842% | 8.953% | 2.525%
| PDM3 |  | -2.192% | -22.13% | 3.506% |  | 6.354% |  | 10.99% | -15.49% | 26.85% | -7.89%
| Z-Score (Absolute) |  | 0.1 | 0.79 | 0.1 | 1.51 | 0.2 | 1.51 | 0.36 | 0.56 | 0.91 | 0.3 | 6.38 | 6.38
| Analytical results for Pr in OREAS 702 (Certified Value 6.71 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 6.92 | 6.43 | 7.11 | 6.75 | 6.86 | 6.36 | 6.6 | 7 | 6.36 | 7.2075 | 6.45 | 27 | 6.04 | 6.203 | 6.92 | 7.2538 | 6.2 | 7.5 | 7.01 | 6.3965 |  | 1
| 1 | 2 | 6.55 | 6.39 | 7.52 | 6.9 | 6.8 | 6.45 | 7 | 7 | 6.4 | 7.1786 | 6.55 | 25.6 | 6.57 | 6.812 | 7.02 | 6.963 | 6.3 | 7 | 6.89 | 6.1848 |  | #N/A
| 1 | 3 | 6.59 | 6.31 | 7.02 | 6.86 | 6.6 | 6.62 | 6.5 | 7 | 6.49 | 7.1728 | 6.55 | 26.7 | 6.04 | 6.6764 | 6.99 | 7.0922 | 6.9 | 7 | 7.15 | 6.7735 |  | 16
| 1 | 4 | 6.32 | 6.53 | 7.16 | 6.61 | 7.06 | 6.57 | 7 | 7 | 6.52 | 7.1652 | 6.56 | 25.5 | 6.22 | 6.1111 | 7.05 | 6.8762 | 6.3 | 7 | 7.14 | 6.2705 |  | 6.714
| 1 | 5 | 6.6 | 6.38 | 7.19 | 7.31 | 6.87 | 6.44 | 6.5 | 7 | 6.42 | 7.1589 | 6.52 | 26.1 | 6.44 | 6.9896 | 7 | 6.9721 | 5.9 | 6.5 | 7.2 | 6.5175 |  | 69
| 1 | 6 | 6.68 | 6.38 | 6.62 | 7.05 | 6.66 | 6.55 | 6.4 | 7 | 6.76 | 7.2062 | 6.41 | 27.4 | 6.28 | 5.9503 | 6.96 | 7.2341 | 6.1 | 7 | 6.77 | 7.3442
| Mean |  | 6.61 | 6.4033 | 7.1033 | 6.9133 | 6.8083 | 6.4983 | 6.6667 | 7 | 6.4917 | 7.1815 | 6.5067 | 26.3833 | 6.265 | 6.4571 | 6.99 | 7.0652 | 6.2833 | 7 | 7.0267 | 6.5812
| Median |  | 6.595 | 6.385 | 7.135 | 6.88 | 6.83 | 6.5 | 6.55 | 7 | 6.455 | 7.1757 | 6.535 | 26.4 | 6.25 | 6.4397 | 6.995 | 7.0321 | 6.25 | 7 | 7.075 | 6.457
| Std Dev. |  | 0.19453 | 0.073121 | 0.29152 | 0.24402 | 0.16449 | 0.097451 | 0.26583 | 0 | 0.144 | 0.020733 | 0.062183 | 0.77309 | 0.21314 | 0.42397 | 0.045607 | 0.15469 | 0.33714 | 0.31623 | 0.16908 | 0.42685
| Rel.Std.Dev. |  | 2.943% | 1.142% | 4.104% | 3.53% | 2.416% | 1.5% | 3.987% | 0% | 2.218% | 0.2887% | 0.9557% | 2.93% | 3.402% | 6.566% | 0.6525% | 2.19% | 5.366% | 4.518% | 2.406% | 6.486%
| PDM3 |  | -1.549% | -4.627% | 5.799% | 2.969% | 1.405% | -3.212% | -0.7047% | 4.26% | -3.311% | 6.964% | -3.088% | 293% | -6.687% | -3.827% | 4.111% | 5.232% | -6.414% | 4.26% | 4.657% | -1.978%
| Z-Score (Absolute) |  | 0.15 | 0.67 | 1.12 | 0.63 | 0.36 | 0.43 | 0 | NA | 0.45 | 1.32 | 0.41 | 50.49 | 1.03 | 0.54 | 0.83 | 1.02 | 0.98 | 0.85 | 0.92 | 0.22
| Analytical results for Rb in OREAS 702 (Certified Value 120 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  | 0
0
| 1 | 1 | 122.5 | 121 | 115.5 | 126.5 | 123.5 | 120 | 110.1 | 121 | 133.1919 | 117.2 | 144 | 118 | 112.7554 | 119.9 | 122.1923 | 109 | 125 | 119 | 111.2724 |  |  | 1
| 1 | 2 | 121 | 121 | 118 | 125 | 124 | 121 | 109.8 | 120 | 132.3823 | 117 | 144 | 121 | 111.7657 | 120.1 | 122.1181 | 110 | 120 | 120 | 109.3297 |  |  | 2
| 1 | 3 | 124.5 | 118.5 | 120 | 129 | 120.5 | 120 | 114.2 | 120 | 133.127 | 120.4 | 148 | 119 | 113.2883 | 121.2 | 123.5133 | 110 | 120 | 118 | 114.8524 |  |  | 16
| 1 | 4 | 120 | 120 | 117 | 128 | 128 | 119 | 109.9 | 119 | 133.4551 | 119.7 | 143 | 118 | 111.4326 | 119.5 | 119.6919 | 111 | 120 | 119 | 110.0357 |  |  | 119.5442
| 1 | 5 | 124.5 | 121.5 | 121 | 130 | 121 | 120 | 111.4 | 118 | 132.9126 | 118.3 | 145 | 116 | 112.6141 | 119.5 | 123.4583 | 108 | 120 | 119 | 112.0065 |  |  | 70
| 1 | 6 | 121 | 125.5 | 117.5 | 127.5 | 118 | 119 | 109.9 | 119 | 132.9448 | 119.6 | 143 | 115 | 111.4985 | 120.6 | 122.168 | 109 | 115 | 120 | 112.9829
| Mean |  | 122.25 | 121.25 | 118.1667 | 127.6667 | 122.5 | 119.8333 | 110.8833 | 119.5 | 133.0023 | 118.7 | 144.5 | 117.8333 | 112.2258 | 120.1333 | 122.1903 | 109.5 | 120 | 119.1667 | 111.7466
| Median |  | 121.75 | 121 | 117.75 | 127.75 | 122.25 | 120 | 110 | 119.5 | 133.0359 | 118.95 | 144 | 118 | 112.1899 | 120 | 122.1802 | 109.5 | 120 | 119 | 111.6394
| Std Dev. |  | 1.917 | 2.3399 | 2.0166 | 1.7795 | 3.4641 | 0.75277 | 1.7314 | 1.0488 | 0.3612 | 1.4142 | 1.8708 | 2.137 | 0.7655 | 0.66533 | 1.3861 | 1.0488 | 3.1623 | 0.75277 | 2.0117
| Rel.Std.Dev. |  | 1.568% | 1.93% | 1.707% | 1.394% | 2.828% | 0.6282% | 1.561% | 0.8777% | 0.2716% | 1.191% | 1.295% | 1.814% | 0.6821% | 0.5538% | 1.134% | 0.9578% | 2.635% | 0.6317% | 1.8%
| PDM3 |  | 2.263% | 1.427% | -1.152% | 6.795% | 2.473% | 0.2419% | -7.245% | -0.03696% | 11.26% | -0.7062% | 20.88% | -1.431% | -6.122% | 0.4928% | 2.214% | -8.402% | 0.3813% | -0.3158% | -6.523%
| Z-Score (Absolute) |  | 0.69 | 0.41 | 0.48 | 2.24 | 0.76 | 0 | 2.56 | 0.1 | 3.77 | 0.32 | 7.06 | 0.57 | 2.18 | 0.09 | 0.67 | 2.96 | 0.05 | 0.19 | 2.31
| Analytical results for Re in OREAS 702 (Indicative Value 0.037 ppm) | [22 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 10 | 14 | 18 | 25 | [16 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | [16 empty] | 3
3
| 1 | 1 | 0.02 | < 0.1 | < 1 | 0.041 | 0.05 | [16 empty] | 1
| 1 | 2 | 0.02 | < 0.1 | < 1 | 0.042 | 0.04 | [16 empty] | 13
| 1 | 3 | 0.01 | < 0.1 | < 1 | 0.043 | 0.06 | [16 empty] | 16
| 1 | 4 | 0.02 | < 0.1 | < 1 | 0.038 | 0.05 | [16 empty] | 0.036556
| 1 | 5 | 0.02 | < 0.1 | < 1 | 0.038 | 0.07 | [16 empty] | 19
| 1 | 6 | 0.02 | < 0.1 | < 1 | 0.036 | 0.04
| Mean |  | 0.018333 |  |  | 0.039667 | 0.051667
| Median |  | 0.02 |  |  | 0.0395 | 0.05
| Std Dev. |  | 0.0040825 |  |  | 0.0027325 | 0.01169
| Rel.Std.Dev. |  | 22.27% |  |  | 6.889% | 22.63%
| PDM3 |  | -49.85% |  |  | 8.511% | 41.34%
| Z-Score (Absolute) |  | 2.07 | 0 | 29.37 | 0.67 | 0.11
| Analytical results for S in OREAS 702 (Certified Value 0.346 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 07 | 09 | 10 | 11 | 14 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | [9 empty] | wt.%
|  |  | BF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [9 empty] | 3
3
| 1 | 1 | 0.33 | 0.3322 | 0.38 | 0.28859 | 0.36 | 0.22 | 0.39379 | 0.35 | 0.358 | 0.34 | 0.35 | 0.44446 | [9 empty] | 1
| 1 | 2 | 0.34 | 0.3282 | 0.39 | 0.2884 | 0.35 | 0.23 | 0.36407 | 0.36 | 0.371 | 0.33 | 0.43 | 0.46317 | [9 empty] | #N/A
| 1 | 3 | 0.32 | 0.3328 | 0.38 | 0.28826 | 0.35 | 0.23 | 0.41302 | 0.33 | 0.339 | 0.32 | 0.37 | 0.48717 | [9 empty] | 16
| 1 | 4 | 0.33 | 0.3883 | 0.39 | 0.29401 | 0.34 | 0.22 | 0.34495 | 0.34 | 0.343 | 0.34 | 0.39 | 0.47372 | [9 empty] | 0.34622
| 1 | 5 | 0.33 | 0.3278 | 0.38 | 0.289 | 0.35 | 0.23 | 0.34348 | 0.35 | 0.358 | 0.33 | 0.37 | 0.4513 | [9 empty] | 71
| 1 | 6 | 0.32 | 0.3341 | 0.38 | 0.28859 | 0.33 | 0.23 | 0.37047 | 0.35 | 0.356 | 0.34 | 0.36 | 0.46946
| Mean |  | 0.32833 | 0.34057 | 0.38333 | 0.28947 | 0.34667 | 0.22667 | 0.37163 | 0.34667 | 0.35417 | 0.33333 | 0.37833 | 0.46488
| Median |  | 0.33 | 0.3325 | 0.38 | 0.28859 | 0.35 | 0.23 | 0.36727 | 0.35 | 0.357 | 0.335 | 0.37 | 0.46632
| Std Dev. |  | 0.0075277 | 0.023522 | 0.005164 | 0.0022359 | 0.010328 | 0.005164 | 0.027448 | 0.010328 | 0.011583 | 0.008165 | 0.028577 | 0.015493
| Rel.Std.Dev. |  | 2.293% | 6.907% | 1.347% | 0.7724% | 2.979% | 2.278% | 7.386% | 2.979% | 3.271% | 2.449% | 7.553% | 3.333%
| PDM3 |  | -5.165% | -1.632% | 10.72% | -16.39% | 0.1298% | -34.53% | 7.34% | 0.1298% | 2.296% | -3.721% | 9.276% | 34.27%
| Z-Score (Absolute) |  | 0.57 | 0.19 | 1.14 | 1.78 | 0 | 3.74 | 0.78 | 0 | 0.23 | 0.42 | 0.99 | 3.68
| Analytical results for Sb in OREAS 702 (Indicative Value 0.41 ppm) | [22 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 06 | 07 | 10 | 14 | 15 | 17 | 18 | 22 | 23 | 25 | [10 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | 0.6 | < 0.5 | < 0.2 | 0.9 | 9.4 | < 0.5 | < 1 | 0.34 | < 32 | < 1 | < 4 | [10 empty] | 1
| 1 | 2 | 0.4 | < 0.5 | 0.4 | 0.8 | 9.13 | < 0.5 | < 1 | 0.37 | < 32 | < 1 | < 4 | [10 empty] | 3
| 1 | 3 | 0.5 | < 0.5 | 0.2 | 0.7 | 9.43 | < 0.5 | < 1 | 0.35 | < 32 | < 1 | < 4 | [10 empty] | 16
| 1 | 4 | 0.3 | < 0.5 | < 0.2 | 0.7 | 8.45 | 0.5 | < 1 | 0.36 | < 32 | < 1 | < 4 | [10 empty] | 0.40867
| 1 | 5 | 0.6 | < 0.5 | 0.2 | 0.8 | 9.26 | < 0.5 | < 1 | 0.37 | < 32 | < 1 | < 4 | [10 empty] | 20
| 1 | 6 | 0.4 | < 0.5 | 0.2 | 0.7 | 8.55 | < 0.5 | < 1 | 0.37 | < 32 | < 1 | < 4
| Mean |  | 0.46667 |  | 0.25 | 0.76667 | 9.0367 | 0.5 |  | 0.36
| Median |  | 0.45 |  | 0.2 | 0.75 | 9.195 | 0.5 |  | 0.365
| Std Dev. |  | 0.12111 |  | 0.1 | 0.08165 | 0.43043 |  |  | 0.012649
| Rel.Std.Dev. |  | 25.95% |  | 40% | 10.65% | 4.763% |  |  | 3.514%
| PDM3 |  | 14.19% |  | -38.83% | 87.6% | 2111% | 22.35% |  | -11.91%
| Z-Score (Absolute) |  | 0.09 | 0.67 | 0.81 | 0.72 | 23.03 | 0.56 | 0 | 0.38 | 41.81 | 0 | 4.05
| Analytical results for Sc in OREAS 702 (Certified Value 5.81 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 26 |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 6.5 | 6.4 | 6.7 | 6 | 5.2 | 7 | < 10 | 5 | < 20 | 7 | < 5 | 5 | 5.955 | 5.4828 | 5.8 | 6.0863 | 7 | < 30 |  |  |  | 1
| 1 | 2 | 6.3 | 6.6 | 8.2 | 6.7 | 6.5 | 7 | < 10 | 5 | < 20 | 6 | < 5 | 5 | 6.0979 | 5.5836 | 5.7 | 5.9591 | 7 | < 30 |  |  |  | 4
| 1 | 3 | 6.6 | 6 | 5.9 | 5.6 | 5.2 | 7 | < 10 | 5 | < 20 | 6 | 17 | 5 | 6.0189 | 5.6092 | 5.7 | 5.8436 | 7 | < 30 |  |  |  | 16
| 1 | 4 | 6 | 6.4 | 4.9 | 5.6 | 6.6 | 7 | < 10 | 5 | < 20 | 6 | < 5 | 5 | 5.8763 | 4.9844 | 5.7 | 5.9249 | 6 | < 30 |  |  |  | 5.8088
| 1 | 5 | 7.3 | 6.8 | 5.1 | 5.7 | 7 | 7 | < 10 | 5 | < 20 | 6 | 13 | 6 | 5.9455 | 5.8717 | 5.8 | 5.872 | 7 | < 30 |  |  |  | 72
| 1 | 6 | 6 | 5.9 | 5.3 | 5.7 | 5.2 | 7 | < 10 | 5 | < 20 | 6 | < 5 | 5 | 5.8354 | 5.6112 | 5.8 | 5.784 | 7 | < 30
| Mean |  | 6.45 | 6.35 | 6.0167 | 5.8833 | 5.95 | 7 |  | 5 |  | 6.1667 | 15 | 5.1667 | 5.9548 | 5.5238 | 5.75 | 5.9117 | 6.8333
| Median |  | 6.4 | 6.4 | 5.6 | 5.7 | 5.85 | 7 |  | 5 |  | 6 | 15 | 5 | 5.9503 | 5.5964 | 5.75 | 5.8985 | 7
| Std Dev. |  | 0.48477 | 0.34496 | 1.2529 | 0.42622 | 0.83845 | 0 |  | 0 |  | 0.40825 | 2.8284 | 0.40825 | 0.094965 | 0.294 | 0.054772 | 0.1053 | 0.40825
| Rel.Std.Dev. |  | 7.516% | 5.433% | 20.82% | 7.245% | 14.09% | 0% |  | 0% |  | 6.62% | 18.86% | 7.902% | 1.595% | 5.322% | 0.9526% | 1.781% | 5.974%
| PDM3 |  | 11.04% | 9.317% | 3.579% | 1.283% | 2.431% | 20.51% |  | -13.92% |  | 6.161% | 158.2% | -11.05% | 2.514% | -4.906% | -1.012% | 1.771% | 17.64%
| Z-Score (Absolute) |  | 1.12 | 0.89 | 0.14 | 0.16 | 0.01 | NA | 2.14 | NA | 9.1 | NA | NA | NA | 0.01 | 0.96 | 0.45 | 0.09 | NA | 20.34
Results from laboratories 6, 9, 13, 14, 15 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 702 (Indicative Value < 10 ppm) | [22 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 02 | 10 | 17 | 18 | 23 | 25 | [15 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [15 empty] | 1
1
| 1 | 1 | < 3 | < 20 | < 10 | < 1 | < 50 | < 10 | [15 empty] | 1
| 1 | 2 | < 3 | < 20 | < 10 | < 1 | < 50 | < 10 | [15 empty] | 11
| 1 | 3 | < 3 | < 20 | < 10 | < 1 | < 50 | < 10 | [15 empty] | 16
| 1 | 4 | < 3 | < 20 | < 10 | < 1 | < 50 | < 10 | [15 empty] | < 10
| 1 | 5 | 3 | < 20 | < 10 | < 1 | < 50 | < 10 | [15 empty] | 17
| 1 | 6 | 3 | < 20 | < 10 | < 1 | < 50 | < 10
| Mean |  | 3
| Median |  | 3
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.54 | 0.9 | 0 | 0.81 | 3.6 | 0
| Analytical results for SiO2 in OREAS 702 (Certified Value 42.01 wt.%) | [22 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 05 | 06 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | [7 empty] | wt.%
|  |  | BF*OES | BF*OES/MS | BF*OES | PF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [7 empty] | 2
3
| 1 | 1 | 39.9 | 42.52 | 42.39 | 42.14 | 41.95 | 41.93 | 42.8697 | 41.11 | 44.0146 | 41.66 | 41.29 | 40.688 | 41.29 | 42.2031 | [7 empty] | 1
| 1 | 2 | 39.2 | 41.88 | 42.28 | 41.29 | 42.28 | 41.93 | 42.0313 | 41.02 | 44.476 | 41.64 | 43.43 | 40.88 | 43.21 | 42.7174 | [7 empty] | #N/A
| 1 | 3 | 39.5 | 42.06 | 40.77 | 41.5 | 41.9 | 43.85 | 42.849 | 40.88 | 43.6151 | 40.98 | 44.71 | 40.944 | 40.64 | 42.8624 | [7 empty] | 16
| 1 | 4 | 40.3 | 42.3 | 42.58 | 41.5 | 41.75 | 41.5 | 42.4764 | 40.72 | 42.0129 | 41.61 | 44.71 | 41.672 | 42.57 | 43.0782 | [7 empty] | 42.0143
| 1 | 5 | 39.5 | 42.6 | 41.54 | 41.5 | 41.85 | 42.36 | 42.5799 | 40.44 | 43.8582 | 41.2 | 44.71 | 41.436 | 41.07 | 41.9708 | [7 empty] | 73
| 1 | 6 | 38.8 | 42.57 | 42.04 | 41.93 | 42.29 | 41.71 | 43.4493 | 40.78 | 43.101 | 41.96 | 44.28 | 41.372 | 41.5 | 42.8996
| Mean |  | 39.5333 | 42.3217 | 41.9333 | 41.6433 | 42.0033 | 42.2133 | 42.7093 | 40.825 | 43.513 | 41.5083 | 43.855 | 41.1653 | 41.7133 | 42.6219
| Median |  | 39.5 | 42.41 | 42.16 | 41.5 | 41.925 | 41.93 | 42.7144 | 40.83 | 43.7366 | 41.625 | 44.495 | 41.158 | 41.395 | 42.7899
| Std Dev. |  | 0.52409 | 0.29748 | 0.67325 | 0.32104 | 0.22801 | 0.85123 | 0.47401 | 0.23814 | 0.86355 | 0.35465 | 1.3516 | 0.38234 | 0.97617 | 0.43626
| Rel.Std.Dev. |  | 1.326% | 0.7029% | 1.606% | 0.7709% | 0.5428% | 2.016% | 1.11% | 0.5833% | 1.985% | 0.8544% | 3.082% | 0.9288% | 2.34% | 1.024%
| PDM3 |  | -5.905% | 0.7315% | -0.1928% | -0.8831% | -0.02622% | 0.4736% | 1.654% | -2.831% | 3.567% | -1.204% | 4.381% | -2.021% | -0.7165% | 1.446%
| Z-Score (Absolute) |  | 2.95 | 0.43 | 0.04 | 0.39 | 0.04 | 0.3 | 0.9 | 1.38 | 1.87 | 0.56 | 2.28 | 0.97 | 0.31 | 0.79
| Analytical results for Sm in OREAS 702 (Certified Value 5.6 ppm) | [22 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  | 2
3
| 1 | 1 | 5.76 | 5.67 | 6.33 | 5.42 | 6.81 | 5.5 | 5.1 | 6 | 5.6944 | 5.06 | 40.9 | 5.4 | 7.5851 | 5.65 | 5.9369 | 4.9 | 6 | 5.57 | 5.6633 |  |  | 1
| 1 | 2 | 5.49 | 6.06 | 5.86 | 5.62 | 5.95 | 5.4 | 5.2 | 6 | 5.4907 | 5.23 | 34.5 | 5.5 | 6.1439 | 5.72 | 5.6465 | 4.9 | 5.5 | 5.59 | 5.0668 |  |  | #N/A
| 1 | 3 | 5.37 | 6.26 | 6.04 | 5.12 | 6.35 | 5.6 | 5.3 | 6 | 5.5144 | 5.28 | 51 | 5.4 | 6.3222 | 5.69 | 5.825 | 5.2 | 6 | 5.48 | 5.7949 |  |  | 16
| 1 | 4 | 5.38 | 6.05 | 5.81 | 5.14 | 6.14 | 5.6 | 5.4 | 6 | 5.4682 | 5.25 | 37.3 | 4.5 | 5.7503 | 5.65 | 5.8149 | 5 | 5.5 | 5.57 | 5.4687 |  |  | 5.5985
| 1 | 5 | 5.6 | 5.74 | 6.17 | 5.54 | 5.77 | 5.3 | 5.2 | 6 | 5.5073 | 5.36 | 47.4 | 5.5 | 6.3146 | 5.67 | 5.8277 | 4.8 | 5.5 | 5.84 | 4.8879 |  |  | 74
| 1 | 6 | 5.6 | 5.82 | 5.95 | 6.34 | 5.29 | 5.4 | 5.5 | 6 | 5.4878 | 5.12 | 46.5 | 5 | 5.4544 | 5.6 | 6.262 | 5 | 5.5 | 5.55 | 5.9144
| Mean |  | 5.5333 | 5.9333 | 6.0267 | 5.53 | 6.0517 | 5.4667 | 5.2833 | 6 | 5.5271 | 5.2167 | 42.9333 | 5.2167 | 6.2617 | 5.6633 | 5.8855 | 4.9667 | 5.6667 | 5.6 | 5.466
| Median |  | 5.545 | 5.935 | 5.995 | 5.48 | 6.045 | 5.45 | 5.25 | 6 | 5.499 | 5.24 | 43.7 | 5.4 | 6.2293 | 5.66 | 5.8263 | 4.95 | 5.5 | 5.57 | 5.566
| Std Dev. |  | 0.14989 | 0.22624 | 0.19664 | 0.44645 | 0.51801 | 0.12111 | 0.1472 | 0 | 0.083535 | 0.1093 | 6.399 | 0.39707 | 0.7328 | 0.040825 | 0.20665 | 0.13663 | 0.2582 | 0.12361 | 0.41031
| Rel.Std.Dev. |  | 2.709% | 3.813% | 3.263% | 8.073% | 8.56% | 2.215% | 2.786% | 0% | 1.511% | 2.095% | 14.9% | 7.612% | 11.7% | 0.7209% | 3.511% | 2.751% | 4.556% | 2.207% | 7.507%
| PDM3 |  | -1.164% | 5.981% | 7.648% | -1.223% | 8.095% | -2.354% | -5.629% | 7.172% | -1.274% | -6.82% | 666.9% | -6.82% | 11.85% | 1.159% | 5.127% | -11.29% | 1.218% | 0.02729% | -2.367%
| Z-Score (Absolute) |  | 0.07 | 0.82 | 1.03 | 0.08 | 1.09 | 0.22 | 0.63 | NA | 0.09 | 0.78 | 83.69 | 0.78 | 1.56 | 0.22 | 0.71 | 1.34 | 0.22 | 0.07 | 0.23
Results from laboratory 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 702 (Certified Value 4.36 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 25 | 26 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS |  |  |  |  | 2
3
| 1 | 1 | 4.3 | 4.8 | 4.1 | 4.9 | 4.2 | 4 | 5 | < 20 | 4 | 12 | < 5 | < 10 | 4.4 | 4.4927 | < 5 | < 5 | < 10 |  |  |  |  | 1
| 1 | 2 | 4.3 | 4.8 | 4.5 | 4.7 | 3.8 | 4 | 4 | < 20 | 4 | 14 | < 5 | < 10 | 4.3 | 4.4963 | < 5 | < 5 | < 10 |  |  |  |  | #N/A
| 1 | 3 | 4.6 | 4.7 | 4.8 | 4.6 | 4.1 | 4 | 4 | < 20 | 4 | 14 | < 5 | < 10 | 4.5 | 4.1766 | < 5 | < 5 | < 10 |  |  |  |  | 16
| 1 | 4 | 3.9 | 4.9 | 4.5 | 4.7 | 4.5 | 4 | 5 | < 20 | 4 | 14 | < 5 | < 10 | 4.3 | 4.1512 | < 5 | < 5 | < 10 |  |  |  |  | 4.3564
| 1 | 5 | 4.4 | 4.4 | 4.6 | 4.9 | 4.2 | 4 | 4 | < 20 | 4 | 14 | < 5 | < 10 | 4.3 | 4.2497 | < 5 | < 5 | < 10 |  |  |  |  | 75
| 1 | 6 | 4 | 4.3 | 4.1 | 5 | 4.1 | 4 | 4 | < 20 | 5 | 15 | < 5 | < 10 | 4.3 | 4.4026 | < 5 | < 5 | < 10
| Mean |  | 4.25 | 4.65 | 4.4333 | 4.8 | 4.15 | 4 | 4.3333 |  | 4.1667 | 13.8333 |  |  | 4.35 | 4.3282
| Median |  | 4.3 | 4.75 | 4.5 | 4.8 | 4.15 | 4 | 4 |  | 4 | 14 |  |  | 4.3 | 4.3262
| Std Dev. |  | 0.25884 | 0.2429 | 0.28048 | 0.15492 | 0.22583 | 0 | 0.5164 |  | 0.40825 | 0.98319 |  |  | 0.083666 | 0.1558
| Rel.Std.Dev. |  | 6.09% | 5.224% | 6.327% | 3.227% | 5.442% | 0% | 11.92% |  | 9.798% | 7.107% |  |  | 1.923% | 3.6%
| PDM3 |  | -2.443% | 6.739% | 1.766% | 10.18% | -4.738% | -8.181% | -0.5297% |  | -4.355% | 217.5% |  |  | -0.1471% | -0.6478%
| Z-Score (Absolute) |  | 0.16 | 0.54 | 0.16 | 0.81 | 0.34 | NA | 0.01 | 9.95 | NA | NA | 3.24 | 1.16 | 0.01 | 0.02 | 3.24 | 3.24 | 1.16
Results from laboratories 6, 13 and 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 702 (Certified Value 195 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | 0
0
| 1 | 1 | 194 | 196 | 211 | 204 | 206 | 194 | 197.7 | 176 | 200 | 202.6332 | 197 | 200 | 181 | 192.0214 | 194.9 | 191.8543 | 170.2 | 206 | 190 | 193 | 173.12 | 1
| 1 | 2 | 192 | 192 | 213 | 206 | 207 | 193 | 198.3 | 178 | 190 | 203.5099 | 196.1 | 196 | 181 | 190.4206 | 192 | 193.5351 | 170.74 | 209 | 190 | 192 | 173.67 | #N/A
| 1 | 3 | 197.5 | 189.5 | 211 | 207 | 202 | 193 | 195.6 | 181 | 190 | 202.8535 | 202.3 | 213 | 176 | 191.3987 | 197.2 | 195.1127 | 169.62 | 192 | 190 | 193 | 174.34 | 16
| 1 | 4 | 192 | 193 | 211 | 197.5 | 208 | 196 | 193.9 | 184 | 190 | 202.2612 | 204 | 208 | 182 | 191.9233 | 192.1 | 196.6789 | 165.22 | 193 | 190 | 192 | 174.3 | 194.7788
| 1 | 5 | 202 | 187.5 | 214 | 208 | 204 | 195 | 196 | 179 | 190 | 203.9514 | 202.2 | 197 | 184 | 192.0313 | 194 | 194.9166 | 167.94 | 205 | 190 | 190 | 171.72 | 76
| 1 | 6 | 193.5 | 188 | 214 | 206 | 199 | 192 | 193.5 | 180 | 200 | 202.9959 | 202.9 | 254 | 177 | 189.7564 | 196.6 | 199.4425 | 168.15 | 204 | 190 | 192 | 177.41
| Mean |  | 195.1667 | 191 | 212.3333 | 204.75 | 204.3333 | 193.8333 | 195.8333 | 179.6667 | 193.3333 | 203.0342 | 200.75 | 211.3333 | 180.1667 | 191.2586 | 194.4667 | 195.2567 | 168.645 | 201.5 | 190 | 192 | 174.0933
| Median |  | 193.75 | 190.75 | 212 | 206 | 205 | 193.5 | 195.8 | 179.5 | 190 | 202.9247 | 202.25 | 204 | 181 | 191.661 | 194.45 | 195.0147 | 168.885 | 204.5 | 190 | 192 | 173.985
| Std Dev. |  | 3.9073 | 3.2711 | 1.5055 | 3.7914 | 3.3862 | 1.472 | 1.9408 | 2.7325 | 5.164 | 0.60992 | 3.3279 | 21.9241 | 3.0605 | 0.95908 | 2.1961 | 2.6161 | 2.0115 | 7.1764 | 0 | 1.0954 | 1.8901
| Rel.Std.Dev. |  | 2.002% | 1.713% | 0.709% | 1.852% | 1.657% | 0.7594% | 0.991% | 1.521% | 2.671% | 0.3004% | 1.658% | 10.37% | 1.699% | 0.5015% | 1.129% | 1.34% | 1.193% | 3.561% | 0% | 0.5705% | 1.086%
| PDM3 |  | 0.1991% | -1.94% | 9.013% | 5.119% | 4.905% | -0.4854% | 0.5414% | -7.759% | -0.7421% | 4.238% | 3.066% | 8.499% | -7.502% | -1.807% | -0.1603% | 0.2454% | -13.42% | 3.451% | -2.453% | -1.427% | -10.62%
| Z-Score (Absolute) |  | 0.08 | 0.37 | 1.92 | 1.1 | 1.06 | 0.07 | 0.15 | 1.59 | 0.12 | 0.92 | 0.67 | 1.81 | 1.53 | 0.34 | 0 | 0.08 | 2.77 | 0.75 | 0.48 | 0.26 | 2.19
| Analytical results for Ta in OREAS 702 (Certified Value 0.78 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 25 | 26 |  |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS |  |  |  | 2
2
| 1 | 1 | 0.7 | 0.8 | 1 | 0.7 | 0.7 | 0.8 | 1.1 | < 10 | 0.5 | 1.4 | 5.44 | < 10 | < 5 | 0.11 | 0.84496 | 0.7 | < 3 | < 1 |  |  |  | 1
| 1 | 2 | 0.8 | 0.74 | 0.9 | 0.7 | 0.6 | 0.8 | 0.9 | < 10 | 0.5 | 1.1 | 6.86 | < 10 | < 5 | 0.1 | 0.87269 | 0.8 | < 3 | < 1 |  |  |  | #N/A
| 1 | 3 | 0.8 | 0.77 | 0.9 | 0.7 | 0.7 | 0.8 | 1.3 | < 10 | 0.6 | 1 | 5.5 | < 10 | < 5 | 0.1 | 0.87359 | 0.7 | < 3 | < 1 |  |  |  | 16
| 1 | 4 | 0.7 | 0.82 | 0.9 | 0.6 | 0.7 | 0.7 | 0.8 | < 10 | 0.6 | 0.9 | 5.34 | < 10 | < 5 | 0.12 | 0.79331 | 0.7 | < 3 | < 1 |  |  |  | 0.78373
| 1 | 5 | 0.8 | 0.8 | 0.9 | 0.7 | 0.9 | 0.7 | 1.3 | < 10 | 0.5 | 0.9 | 4.65 | < 10 | < 5 | 0.11 | 0.78647 | 0.7 | < 3 | < 1 |  |  |  | 77
| 1 | 6 | 0.8 | 0.77 | 0.9 | 0.7 | 0.6 | 0.8 | 0.6 | < 10 | 0.5 | 0.8 | 5.64 | < 10 | < 5 | 0.11 | 0.81499 | 0.6 | < 3 | 1.3041
| Mean |  | 0.76667 | 0.78333 | 0.91667 | 0.68333 | 0.7 | 0.76667 | 1 |  | 0.53333 | 1.0167 | 5.5717 |  |  | 0.10833 | 0.831 | 0.7 |  | 1.3041
| Median |  | 0.8 | 0.785 | 0.9 | 0.7 | 0.7 | 0.8 | 1 |  | 0.5 | 0.95 | 5.47 |  |  | 0.11 | 0.82998 | 0.7 |  | 1.3041
| Std Dev. |  | 0.05164 | 0.028752 | 0.040825 | 0.040825 | 0.10954 | 0.05164 | 0.28284 |  | 0.05164 | 0.2137 | 0.71973 |  |  | 0.0075277 | 0.038487 | 0.063246
| Rel.Std.Dev. |  | 6.736% | 3.67% | 4.454% | 5.974% | 15.65% | 6.736% | 28.28% |  | 9.682% | 21.02% | 12.92% |  |  | 6.949% | 4.631% | 9.035%
| PDM3 |  | -2.177% | -0.05031% | 16.96% | -12.81% | -10.68% | -2.177% | 27.6% |  | -31.95% | 29.72% | 610.9% |  |  | -86.18% | 6.032% | -10.68% |  | 66.4%
| Z-Score (Absolute) |  | 0.15 | 0.09 | 0.4 | 0.46 | 0.39 | 0.15 | 0.71 | 15.45 | 1.01 | 0.77 | 17.56 | 15.45 | 6.24 | 2.58 | 0.09 | 0.39 | 2.55 | 0.64
| Analytical results for Tb in OREAS 702 (Certified Value 0.78 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 0.73 | 0.76 | 0.79 | 0.85 | 0.75 | 0.7 | 0.8 | < 3 | 0.78 | 0.77015 | 0.74 | 6.75 | 0.79 | 0.9594 | 0.63 | 0.85163 | 0.7 | 0.8 | 0.86 | 0.79183 |  | 1
| 1 | 2 | 0.77 | 0.74 | 0.77 | 0.79 | 0.73 | 0.7 | 0.7 | < 3 | 0.81 | 0.78223 | 0.76 | 6.2 | 0.75 | 0.9387 | 0.64 | 0.84821 | 0.8 | 0.8 | 0.82 | 0.7783 |  | #N/A
| 1 | 3 | 0.75 | 0.7 | 0.76 | 0.81 | 0.75 | 0.8 | 0.7 | < 3 | 0.79 | 0.79783 | 0.74 | 6.51 | 0.68 | 0.9332 | 0.64 | 0.81421 | 0.8 | 0.7 | 0.82 | 0.75815 |  | 16
| 1 | 4 | 0.73 | 0.78 | 0.81 | 0.75 | 0.78 | 0.7 | 0.8 | < 3 | 0.8 | 0.79443 | 0.77 | 4.85 | 0.76 | 0.8652 | 0.63 | 0.76902 | 0.7 | 0.8 | 0.84 | 0.72345 |  | 0.78335
| 1 | 5 | 0.75 | 0.75 | 0.77 | 0.81 | 0.7 | 0.7 | 0.8 | < 3 | 0.78 | 0.80399 | 0.76 | 4.72 | 0.75 | 0.8938 | 0.64 | 0.82888 | 0.6 | 0.7 | 0.85 | 0.72466 |  | 78
| 1 | 6 | 0.75 | 0.75 | 0.76 | 0.85 | 0.71 | 0.7 | 0.8 | < 3 | 0.8 | 0.79787 | 0.75 | 6.3 | 0.67 | 0.7784 | 0.64 | 0.8319 | 0.7 | 0.8 | 0.8 | 0.77335
| Mean |  | 0.74667 | 0.74667 | 0.77667 | 0.81 | 0.73667 | 0.71667 | 0.76667 |  | 0.79333 | 0.79108 | 0.75333 | 5.8883 | 0.73333 | 0.89478 | 0.63667 | 0.82397 | 0.71667 | 0.76667 | 0.83167 | 0.75829
| Median |  | 0.75 | 0.75 | 0.77 | 0.81 | 0.74 | 0.7 | 0.8 |  | 0.795 | 0.79613 | 0.755 | 6.25 | 0.75 | 0.9135 | 0.64 | 0.83039 | 0.7 | 0.8 | 0.83 | 0.76575
| Std Dev. |  | 0.015055 | 0.026583 | 0.019664 | 0.037947 | 0.029439 | 0.040825 | 0.05164 |  | 0.012111 | 0.012538 | 0.012111 | 0.87621 | 0.04761 | 0.066334 | 0.005164 | 0.030181 | 0.075277 | 0.05164 | 0.022286 | 0.028626
| Rel.Std.Dev. |  | 2.016% | 3.56% | 2.532% | 4.685% | 3.996% | 5.696% | 6.736% |  | 1.527% | 1.585% | 1.608% | 14.88% | 6.492% | 7.413% | 0.8111% | 3.663% | 10.5% | 6.736% | 2.68% | 3.775%
| PDM3 |  | -4.683% | -4.683% | -0.853% | 3.402% | -5.959% | -8.512% | -2.13% |  | 1.275% | 0.9869% | -3.832% | 651.7% | -6.385% | 14.23% | -18.73% | 5.186% | -8.512% | -2.13% | 6.168% | -3.199%
| Z-Score (Absolute) |  | 0.65 | 0.65 | 0.13 | 0.46 | 0.82 | NA | NA | 12.49 | 0.17 | 0.13 | 0.53 | 89.05 | 0.88 | 1.93 | 2.57 | 0.7 | NA | NA | 0.83 | 0.45
Results from laboratories 6, 7, 23 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 702 (Indicative Value 1.11 ppm) | [22 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 02 | 10 | 14 | 17 | 18 | 23 | 25 | [14 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [14 empty] | 2
2
| 1 | 1 | 1.2 | 1 | < 5 | < 1 | 1.05 | 1.1 | < 5 | [14 empty] | 1
| 1 | 2 | 1.5 | 1 | < 5 | < 1 | 0.99 | 1.2 | < 5 | [14 empty] | 12
| 1 | 3 | 1.4 | 1 | < 5 | < 1 | 1.02 | 1.1 | < 5 | [14 empty] | 16
| 1 | 4 | 1.3 | 1 | < 5 | < 1 | 1.05 | 2.5 | < 5 | [14 empty] | 1.1096
| 1 | 5 | 1.1 | 1 | < 5 | < 1 | 0.99 | 1.2 | < 5 | [14 empty] | 18
| 1 | 6 | 1.3 | 1 | < 5 | < 1 | 1.01 | 1 | < 5
| Mean |  | 1.3 | 1 |  |  | 1.0183 | 1.35
| Median |  | 1.3 | 1 |  |  | 1.015 | 1.15
| Std Dev. |  | 0.14142 | 0 |  |  | 0.027142 | 0.56833
| Rel.Std.Dev. |  | 10.88% | 0% |  |  | 2.665% | 42.1%
| PDM3 |  | 17.16% | -9.876% |  |  | -8.224% | 21.67%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.7 | 1.8 | 0.63 | 0.11 | 2.7
| Analytical results for Th in OREAS 702 (Certified Value 9.88 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 10.3 | 9.9 | 10.45 | 10.15 | 10.5 | 9.4 | 9.6 | < 50 | 10.2 | 9.3 | 16.04 | 9.9 | 10.6617 | 9.3924 | 9.98 | 10.4753 | 9.2 | 10.5 | 11 | 9.0123 |  | 1
| 1 | 2 | 9.84 | 10 | 10.95 | 10.45 | 10.55 | 9.4 | 10.3 | < 50 | 10 | 9.7 | 13.32 | 9.8 | 10.5897 | 9.3823 | 10.06 | 10.1894 | 9.5 | 10.5 | 10 | 9.3674 |  | 1
| 1 | 3 | 9.76 | 9.8 | 10.1 | 10.1 | 10.1 | 9.6 | 9.5 | < 50 | 10 | 9.3 | 11.81 | 9.1 | 10.7198 | 9.5025 | 10.18 | 10.2841 | 9.6 | 11 | 11 | 9.7289 |  | 16
| 1 | 4 | 9.57 | 10 | 10.4 | 9.96 | 10 | 9.8 | 10.1 | < 50 | 10.6 | 9.5 | 10.74 | 8.9 | 9.2895 | 9.6741 | 10.06 | 9.9334 | 9.1 | 10 | 10 | 9.2073 |  | 9.8764
| 1 | 5 | 9.84 | 9.4 | 10.8 | 10.25 | 9.79 | 9.3 | 9.2 | < 50 | 10.2 | 9.5 | 10.17 | 9.6 | 11.516 | 9.7677 | 10.1 | 10.3928 | 8.7 | 9.5 | 11 | 9.1592 |  | 79
| 1 | 6 | 10.05 | 9.8 | 9.43 | 10.65 | 9.58 | 9.6 | 10 | < 50 | 10.7 | 9.2 | 10.4 | 9.1 | 9.7828 | 9.6497 | 10.12 | 10.5405 | 9 | 10 | 10 | 10.5132
| Mean |  | 9.8933 | 9.8167 | 10.355 | 10.26 | 10.0867 | 9.5167 | 9.7833 |  | 10.2833 | 9.4167 | 12.08 | 9.4 | 10.4266 | 9.5615 | 10.0833 | 10.3026 | 9.1833 | 10.25 | 10.5 | 9.498
| Median |  | 9.84 | 9.85 | 10.425 | 10.2 | 10.05 | 9.5 | 9.8 |  | 10.2 | 9.4 | 11.275 | 9.35 | 10.6257 | 9.5761 | 10.08 | 10.3384 | 9.15 | 10.25 | 10.5 | 9.2873
| Std Dev. |  | 0.25216 | 0.22286 | 0.54475 | 0.2514 | 0.38417 | 0.18348 | 0.41673 |  | 0.29944 | 0.18348 | 2.2609 | 0.41952 | 0.78257 | 0.15949 | 0.067429 | 0.22089 | 0.33116 | 0.5244 | 0.54772 | 0.55429
| Rel.Std.Dev. |  | 2.549% | 2.27% | 5.261% | 2.45% | 3.809% | 1.928% | 4.26% |  | 2.912% | 1.949% | 18.72% | 4.463% | 7.506% | 1.668% | 0.6687% | 2.144% | 3.606% | 5.116% | 5.216% | 5.836%
| PDM3 |  | 0.1713% | -0.605% | 4.846% | 3.884% | 2.129% | -3.643% | -0.9425% |  | 4.12% | -4.655% | 22.31% | -4.824% | 5.57% | -3.189% | 2.095% | 4.315% | -7.018% | 3.783% | 6.314% | -3.831%
| Z-Score (Absolute) |  | 0.43 | 0.6 | 0.61 | 0.4 | 0.01 | 1.27 | 0.67 | 33.53 | 0.45 | 1.5 | 4.49 | 1.54 | 0.77 | 1.17 | 0 | 0.49 | 2.02 | 0.37 | NA | 1.32
Results from laboratory 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 702 (Certified Value 0.237 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | 3
3
| 1 | 1 | 0.29 | 0.231 | 0.26 | 0.26 | 0.234 | 0.2398 | 0.2403 | 0.2423 | 0.25 | 0.23047 | 0.24 | 0.24 | 0.2285 | 0.2308 | 0.24 | 0.24017 | 0.228 | 0.256 | 0.23 | 0.232 | 0.22755 | 1
| 1 | 2 | 0.28 | 0.232 | 0.27 | 0.25 | 0.228 | 0.235 | 0.2383 | 0.2421 | 0.25 | 0.2299 | 0.24 | 0.22 | 0.2256 | 0.229 | 0.246 | 0.24041 | 0.222 | 0.261 | 0.23 | 0.245 | 0.22701 | #N/A
| 1 | 3 | 0.29 | 0.236 | 0.27 | 0.25 | 0.228 | 0.2404 | 0.2445 | 0.2437 | 0.25 | 0.22981 | 0.24 | 0.25 | 0.2181 | 0.2332 | 0.24 | 0.24155 | 0.228 | 0.241 | 0.23 | 0.232 | 0.22892 | 16
| 1 | 4 | 0.29 | 0.237 | 0.27 | 0.25 | 0.234 | 0.2392 | 0.2385 | 0.2475 | 0.24 | 0.22988 | 0.24 | 0.24 | 0.2248 | 0.2308 | 0.24 | 0.23909 | 0.234 | 0.241 | 0.24 | 0.239 | 0.22875 | 0.23741
| 1 | 5 | 0.3 | 0.231 | 0.27 | 0.26 | 0.228 | 0.2374 | 0.2413 | 0.2404 | 0.25 | 0.22897 | 0.24 | 0.24 | 0.2279 | 0.2338 | 0.24 | 0.24222 | 0.228 | 0.255 | 0.24 | 0.237 | 0.22502 | 80
| 1 | 6 | 0.29 | 0.234 | 0.27 | 0.26 | 0.228 | 0.2356 | 0.2375 | 0.2401 | 0.25 | 0.23014 | 0.24 | 0.24 | 0.2199 | 0.235 | 0.24 | 0.23801 | 0.228 | 0.252 | 0.24 | 0.235 | 0.22946
| Mean |  | 0.29 | 0.2335 | 0.26833 | 0.255 | 0.23 | 0.2379 | 0.24007 | 0.24268 | 0.24833 | 0.22986 | 0.24 | 0.23833 | 0.22413 | 0.2321 | 0.241 | 0.24024 | 0.228 | 0.251 | 0.235 | 0.23667 | 0.22778
| Median |  | 0.29 | 0.233 | 0.27 | 0.255 | 0.228 | 0.2383 | 0.2394 | 0.2422 | 0.25 | 0.22989 | 0.24 | 0.24 | 0.2252 | 0.232 | 0.24 | 0.24029 | 0.228 | 0.2535 | 0.235 | 0.236 | 0.22815
| Std Dev. |  | 0.0063246 | 0.0025884 | 0.0040825 | 0.0054772 | 0.0030984 | 0.0022583 | 0.0025843 | 0.0027059 | 0.0040825 | 0.0004988 | 0 | 0.0098319 | 0.0042467 | 0.0022583 | 0.0024495 | 0.0015457 | 0.0037947 | 0.0082704 | 0.0054772 | 0.0049261 | 0.001631
| Rel.Std.Dev. |  | 2.181% | 1.109% | 1.521% | 2.148% | 1.347% | 0.9493% | 1.076% | 1.115% | 1.644% | 0.217% | 0% | 4.125% | 1.895% | 0.973% | 1.016% | 0.6434% | 1.664% | 3.295% | 2.331% | 2.081% | 0.716%
| PDM3 |  | 22.15% | -1.647% | 13.03% | 7.409% | -3.121% | 0.2063% | 1.119% | 2.221% | 4.601% | -3.18% | 1.091% | 0.3889% | -5.592% | -2.237% | 1.512% | 1.193% | -3.964% | 5.724% | -1.015% | -0.3132% | -4.055%
| Z-Score (Absolute) |  | 5.59 | 0.52 | 3.25 | 1.8 | 0.92 | 0.05 | 0.19 | 0.47 | 1.08 | 0.92 | 0.16 | 0 | 1.54 | 0.67 | 0.27 | 0.21 | 1.14 | 1.37 | 0.36 | 0.18 | 1.14
| Analytical results for Tl in OREAS 702 (Certified Value 0.7 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 06 | 10 | 14 | 15 | 17 | 18 | 23 | 25 | [12 empty] | ppm
|  |  | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | 0.7 | 0.6 | 0.8 | 0.91 | 0.7 | < 1 | 0.45 | < 1 | 0.7 | [12 empty] | 1
| 1 | 2 | 0.71 | 0.6 | 0.8 | 0.73 | 0.8 | < 1 | 0.42 | < 1 | 0.69 | [12 empty] | 21
| 1 | 3 | 0.64 | 0.6 | 0.8 | 0.68 | 0.7 | < 1 | 0.45 | < 1 | 0.7 | [12 empty] | 16
| 1 | 4 | 0.68 | 0.6 | 0.8 | 0.65 | 0.8 | < 1 | 0.43 | < 1 | 0.71 | [12 empty] | 0.70067
| 1 | 5 | 0.7 | 0.6 | 0.8 | 0.61 | 0.8 | < 1 | 0.46 | < 1 | 0.69 | [12 empty] | 81
| 1 | 6 | 0.73 | 0.6 | 0.8 | 0.65 | 0.7 | < 1 | 0.47 | < 1 | 0.69
| Mean |  | 0.69333 | 0.6 | 0.8 | 0.705 | 0.75 |  | 0.44667 |  | 0.69667
| Median |  | 0.7 | 0.6 | 0.8 | 0.665 | 0.75 |  | 0.45 |  | 0.695
| Std Dev. |  | 0.030768 | 0 | 1.2162e-16 | 0.10803 | 0.054772 |  | 0.018619 |  | 0.008165
| Rel.Std.Dev. |  | 4.438% | 0% | 1.52e-14% | 15.32% | 7.303% |  | 4.168% |  | 1.172%
| PDM3 |  | -1.047% | -14.37% | 14.18% | 0.6185% | 7.041% |  | -36.25% |  | -0.5709%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.77 | 0.08 | 0.41 | 1.4 | 1.78 | 1.4 | 0.02
| Analytical results for Tm in OREAS 702 (Certified Value 0.31 ppm) | [22 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 0.32 | 0.29 | 0.29 | 0.34 | 0.3 | 0.29 | 0.4 | < 2 | 0.3 | 0.25164 | 0.29 | 16.04 | 0.28 | 0.3576 | 0.16 | 0.34612 | 0.28 | 0.3 | 0.34 | < 0.5 |  | 1
| 1 | 2 | 0.31 | 0.29 | 0.31 | 0.33 | 0.32 | 0.3 | 0.4 | < 2 | 0.32 | 0.24954 | 0.31 | 13.32 | 0.35 | 0.3177 | 0.17 | 0.33196 | 0.39 | 0.3 | 0.34 | < 0.5 |  | #N/A
| 1 | 3 | 0.3 | 0.29 | 0.31 | 0.31 | 0.26 | 0.32 | 0.3 | < 2 | 0.33 | 0.27894 | 0.31 | 11.81 | 0.28 | 0.3326 | 0.17 | 0.3505 | 0.35 | 0.3 | 0.32 | < 0.5 |  | 16
| 1 | 4 | 0.31 | 0.29 | 0.31 | 0.29 | 0.3 | 0.31 | 0.3 | < 2 | 0.31 | 0.257 | 0.31 | 10.74 | 0.31 | 0.3004 | 0.17 | 0.31909 | 0.31 | 0.3 | 0.35 | < 0.5 |  | 0.31169
| 1 | 5 | 0.32 | 0.32 | 0.3 | 0.32 | 0.3 | 0.29 | 0.3 | < 2 | 0.3 | 0.25847 | 0.33 | 10.17 | 0.29 | 0.3098 | 0.16 | 0.32361 | 0.29 | 0.3 | 0.33 | < 0.5 |  | 82
| 1 | 6 | 0.32 | 0.3 | 0.31 | 0.31 | 0.29 | 0.32 | 0.4 | < 2 | 0.32 | 0.25776 | 0.31 | 10.4 | 0.29 | 0.2756 | 0.17 | 0.32886 | 0.3 | 0.3 | 0.33 | < 0.5
| Mean |  | 0.31333 | 0.29667 | 0.305 | 0.31667 | 0.295 | 0.305 | 0.35 |  | 0.31333 | 0.25889 | 0.31 | 12.08 | 0.3 | 0.31562 | 0.16667 | 0.33336 | 0.32 | 0.3 | 0.335
| Median |  | 0.315 | 0.29 | 0.31 | 0.315 | 0.3 | 0.305 | 0.35 |  | 0.315 | 0.25738 | 0.31 | 11.275 | 0.29 | 0.31375 | 0.17 | 0.33041 | 0.305 | 0.3 | 0.335
| Std Dev. |  | 0.008165 | 0.012111 | 0.0083666 | 0.017512 | 0.019748 | 0.013784 | 0.054772 |  | 0.012111 | 0.01046 | 0.012649 | 2.2609 | 0.026833 | 0.028007 | 0.005164 | 0.012468 | 0.041952 | 0 | 0.010488
| Rel.Std.Dev. |  | 2.606% | 4.082% | 2.743% | 5.53% | 6.694% | 4.519% | 15.65% |  | 3.865% | 4.04% | 4.08% | 18.72% | 8.944% | 8.874% | 3.098% | 3.74% | 13.11% | 0% | 3.131%
| PDM3 |  | 0.5272% | -4.82% | -2.146% | 1.597% | -5.355% | -2.146% | 12.29% |  | 0.5272% | -16.94% | -0.5423% | 3776% | -3.751% | 1.26% | -46.53% | 6.951% | 2.666% | -3.751% | 7.478%
| Z-Score (Absolute) |  | 0.08 | 0.76 | 0.34 | 0.25 | 0.84 | 0.34 | NA | 34.81 | 0.08 | 2.67 | 0.08 | 595.16 | 0.59 | 0.2 | 7.33 | 1.1 | 0.42 | NA | 1.18 | 3.12
Results from laboratories 7 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 702 (Certified Value 4.2 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  | 2
2
| 1 | 1 | 4.21 | 4.1 | 4.21 | 4.1 | 3.94 | 4.6 | < 50 | 4.2 | 3.5325 | 3.8 | 5 | 3.48 | 4.2175 | 4.4848 | 3.7 | 4.7463 | 4 | 4.5 | 6 | 4.4703 |  | 1
| 1 | 2 | 3.95 | 4.5 | 4.03 | 4.06 | 5.13 | 4.7 | < 50 | 4.2 | 3.4851 | 4.2 | 5.1 | 4.1 | 3.8984 | 4.1892 | 4 | 4.2712 | 3 | 4.5 | 8 | 4.2039 |  | #N/A
| 1 | 3 | 3.84 | 3.9 | 4.45 | 4.28 | 4.76 | 4.7 | < 50 | 4 | 3.5141 | 3.8 | 5 | 4.14 | 4.0084 | 4.4555 | 4 | 4.6625 | 4 | 4.5 | 7 | 4.0501 |  | 16
| 1 | 4 | 4.22 | 4.1 | 4.12 | 4.08 | 4.41 | 4.3 | < 50 | 4.3 | 3.5478 | 4.1 | 4.7 | 4.08 | 3.7388 | 4.0943 | 3.9 | 4.644 | 4 | 4.5 | 6 | 4.5266 |  | 4.1971
| 1 | 5 | 4.32 | 3.7 | 4.33 | 4.02 | 5.36 | 3.9 | < 50 | 4.2 | 3.5395 | 3.6 | 4.1 | 3.55 | 4.0564 | 4.5723 | 3.7 | 4.5397 | 4 | 5 | 9 | 5.4293 |  | 83
| 1 | 6 | 5.74 | 5.3 | 4.86 | 5.58 | 3.58 | 4 | < 50 | 4.2 | 3.5178 | 4.2 | 4 | 3.24 | 3.4247 | 4.5493 | 3.7 | 4.192 | 4 | 4.5 | 6 | 4.5427
| Mean |  | 4.38 | 4.2667 | 4.3333 | 4.3533 | 4.53 | 4.3667 |  | 4.1833 | 3.5228 | 3.95 | 4.65 | 3.765 | 3.8907 | 4.3909 | 3.8333 | 4.5093 | 3.8333 | 4.5833 | 7 | 4.5371
| Median |  | 4.215 | 4.1 | 4.27 | 4.09 | 4.585 | 4.45 |  | 4.2 | 3.5252 | 3.95 | 4.85 | 3.815 | 3.9534 | 4.4701 | 3.8 | 4.5919 | 4 | 4.5 | 6.5 | 4.4984
| Std Dev. |  | 0.69054 | 0.57155 | 0.2979 | 0.60764 | 0.68801 | 0.3559 |  | 0.098319 | 0.022444 | 0.251 | 0.48477 | 0.38863 | 0.27871 | 0.19981 | 0.15055 | 0.22631 | 0.40825 | 0.20412 | 1.2649 | 0.47933
| Rel.Std.Dev. |  | 15.77% | 13.4% | 6.875% | 13.96% | 15.19% | 8.15% |  | 2.35% | 0.6371% | 6.354% | 10.43% | 10.32% | 7.163% | 4.55% | 3.928% | 5.019% | 10.65% | 4.454% | 18.07% | 10.56%
| PDM3 |  | 4.359% | 1.658% | 3.247% | 3.723% | 7.933% | 4.041% |  | -0.3271% | -16.07% | -5.887% | 10.79% | -10.29% | -7.299% | 4.619% | -8.666% | 7.439% | -8.666% | 9.203% | 66.78% | 8.103%
| Z-Score (Absolute) |  | 0.08 | 0.36 | 0.1 | 0.03 | 0.65 | 0.03 | 78.67 | 0.67 | 3.19 | 1.56 | 1.11 | 2.27 | 1.79 | 0.12 | 2.01 | 0.57 | NA | 0.85 | NA | 0.68
| Analytical results for V in OREAS 702 (Certified Value 41.6 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | 1
2
| 1 | 1 | 44 | 42 | 44 | 47 | 44 | 41 | 44 | 39 | 40 | 38.13 | 42 | 40.57 | 40 | 38.2357 | 39 | 41.0551 | 32.49 | 41 | < 100 | 46 | 39.6812 | 1
| 1 | 2 | 44 | 41 | 46 | 46 | 44 | 40 | 42 | 39 | 40 | 37.95 | 39 | 39.86 | 41 | 38.1496 | 38 | 41.2453 | 32.43 | 42 | < 100 | 46 | 38.4294 | #N/A
| 1 | 3 | 45 | 40 | 46 | 46 | 43 | 42 | 42 | 39 | 40 | 38.2 | 43 | 44.47 | 40 | 38.7828 | 39 | 41.0876 | 34.01 | 38 | < 100 | 46 | 39.1591 | 16
| 1 | 4 | 44 | 43 | 46 | 47 | 45 | 42 | 41 | 46 | 40 | 37.7 | 47 | 40.09 | 42 | 38.7451 | 38 | 39.5358 | 32.42 | 39 | < 100 | 47 | 38.7187 | 41.6191
| 1 | 5 | 46 | 42 | 45 | 47 | 43 | 41 | 43 | 39 | 40 | 37.87 | 43 | 41.77 | 40 | 38.7216 | 39 | 40.4712 | 31.03 | 41 | < 100 | 47 | 37.4147 | 84
| 1 | 6 | 44 | 39 | 47 | 45 | 43 | 39 | 42 | 39 | 40 | 37.76 | 41 | 40.51 | 40 | 38.6352 | 39 | 42.1215 | 31.52 | 39 | < 100 | 47 | 42.5108
| Mean |  | 44.5 | 41.1667 | 45.6667 | 46.3333 | 43.6667 | 40.8333 | 42.3333 | 40.1667 | 40 | 37.935 | 42.5 | 41.2117 | 40.5 | 38.545 | 38.6667 | 40.9194 | 32.3167 | 40 |  | 46.5 | 39.319
| Median |  | 44 | 41.5 | 46 | 46.5 | 43.5 | 41 | 42 | 39 | 40 | 37.91 | 42.5 | 40.54 | 40 | 38.6784 | 39 | 41.0713 | 32.425 | 40 |  | 46.5 | 38.9389
| Std Dev. |  | 0.83666 | 1.472 | 1.0328 | 0.8165 | 0.8165 | 1.169 | 1.0328 | 2.8577 | 0 | 0.19927 | 2.6646 | 1.7273 | 0.83666 | 0.27855 | 0.5164 | 0.86177 | 1.0209 | 1.5492 |  | 0.54772 | 1.7391
| Rel.Std.Dev. |  | 1.88% | 3.576% | 2.262% | 1.762% | 1.87% | 2.863% | 2.44% | 7.115% | 0% | 0.5253% | 6.27% | 4.191% | 2.066% | 0.7227% | 1.336% | 2.106% | 3.159% | 3.873% |  | 1.178% | 4.423%
| PDM3 |  | 6.922% | -1.087% | 9.725% | 11.33% | 4.92% | -1.888% | 1.716% | -3.49% | -3.89% | -8.852% | 2.117% | -0.979% | -2.689% | -7.386% | -7.094% | -1.681% | -22.35% | -3.89% |  | 11.73% | -5.527%
| Z-Score (Absolute) |  | 1.51 | 0.1 | 2 | 2.28 | 1.16 | 0.04 | 0.6 | 0.32 | 0.39 | 1.26 | 0.67 | 0.12 | 0.18 | 1 | 0.95 | 0 | 3.62 | 0.39 | 3.83 | 2.35 | 0.67
| Analytical results for WO3 in OREAS 702 (Certified Value 0.286 wt.%) | [22 empty] | Certified
WO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | 3
3
| 1 | 1 | 0.29634 | 0.2774 | 0.27994 | 0.2888 | 0.2913 | 0.26985 | 0.31462 | 0.30125 | 0.26481 | 0.28168 | 0.27508 | 0.29318 | 0.27722 | 0.29412 | 0.28635 | 0.28862 | 0.30631 | 0.25346 | 0.2825 | 0.27755 | 0.27393 | 1
| 1 | 2 | 0.29634 | 0.2837 | 0.28625 | 0.2888 | 0.2862 | 0.26355 | 0.2986 | 0.27982 | 0.27364 | 0.28224 | 0.27753 | 0.29596 | 0.28306 | 0.29536 | 0.28916 | 0.28633 | 0.30818 | 0.25346 | 0.285 | 0.26582 | 0.26807 | #N/A
| 1 | 3 | 0.2976 | 0.2724 | 0.27994 | 0.2875 | 0.285 | 0.26985 | 0.30731 | 0.28082 | 0.26733 | 0.2846 | 0.27563 | 0.29684 | 0.28449 | 0.29734 | 0.28915 | 0.28565 | 0.30757 | 0.23328 | 0.2812 | 0.28776 | 0.28299 | 16
| 1 | 4 | 0.28246 | 0.2787 | 0.28499 | 0.2774 | 0.2913 | 0.26733 | 0.30125 | 0.28448 | 0.27112 | 0.28276 | 0.2766 | 0.29394 | 0.28808 | 0.29347 | 0.29303 | 0.28304 | 0.30898 | 0.24463 | 0.2812 | 0.29205 | 0.27713 | 0.28601
| 1 | 5 | 0.3039 | 0.2762 | 0.28372 | 0.29 | 0.29 | 0.26103 | 0.30365 | 0.27692 | 0.27742 | 0.28738 | 0.27694 | 0.28877 | 0.28993 | 0.29697 | 0.28437 | 0.28762 | 0.31623 | 0.24337 | 0.2862 | 0.29331 | 0.27317 | 85
| 1 | 6 | 0.29507 | 0.2724 | 0.28625 | 0.2926 | 0.2812 | 0.2812 | 0.29886 | 0.2865 | 0.27238 | 0.2848 | 0.27772 | 0.28826 | 0.28818 | 0.29671 | 0.29066 | 0.28772 | 0.30608 | 0.23959 | 0.2875 | 0.26431 | 0.29482
| Mean |  | 0.29528 | 0.2768 | 0.28352 | 0.28752 | 0.2875 | 0.2688 | 0.30405 | 0.28497 | 0.27112 | 0.28391 | 0.27658 | 0.29283 | 0.28516 | 0.29566 | 0.28879 | 0.28649 | 0.30889 | 0.24463 | 0.28393 | 0.28013 | 0.27835
| Median |  | 0.29634 | 0.2768 | 0.28436 | 0.2888 | 0.2881 | 0.26859 | 0.30245 | 0.28265 | 0.27175 | 0.28368 | 0.27677 | 0.29356 | 0.28629 | 0.29603 | 0.28915 | 0.28697 | 0.30788 | 0.244 | 0.28375 | 0.28265 | 0.27553
| Std Dev. |  | 0.007019 | 0.0042582 | 0.0029237 | 0.0052469 | 0.0040684 | 0.0070151 | 0.0061154 | 0.0086737 | 0.0045124 | 0.0021181 | 0.0010502 | 0.0035944 | 0.0046524 | 0.0016031 | 0.0030746 | 0.001999 | 0.0037593 | 0.0078962 | 0.002683 | 0.012928 | 0.0094482
| Rel.Std.Dev. |  | 2.377% | 1.538% | 1.031% | 1.825% | 1.415% | 2.61% | 2.011% | 3.044% | 1.664% | 0.746% | 0.3797% | 1.228% | 1.632% | 0.5422% | 1.065% | 0.6977% | 1.217% | 3.228% | 0.945% | 4.615% | 3.394%
| PDM3 |  | 3.242% | -3.221% | -0.8735% | 0.5257% | 0.5198% | -6.018% | 6.306% | -0.3665% | -5.208% | -0.7355% | -3.297% | 2.382% | -0.2987% | 3.373% | 0.9693% | 0.1683% | 7.999% | -14.47% | -0.7272% | -2.056% | -2.678%
| Z-Score (Absolute) |  | 1.05 | 0.83 | 0.15 | 0.26 | 0.26 | 1.65 | 1.95 | 0 | 1.41 | 0.11 | 0.85 | 0.8 | 0.02 | 1.09 | 0.39 | 0.16 | 2.44 | 4.11 | 0.11 | 0.49 | 0.67
| Analytical results for Y in OREAS 702 (Certified Value 22.8 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | 1
2
| 1 | 1 | 24 | 23.6 | 24.2 | 24.7 | 24.2 | 21 | 22.1 | 24 | 21.6 | 15.0229 | 21.2 | 23.34 | 21.54 | 29.9304 | 12.7 | 23.8519 | 19.6 | 21.2 | 26 | 23 | 22.4791 | 1
| 1 | 2 | 24 | 23 | 24.5 | 23.9 | 25.1 | 21 | 22.5 | 24 | 21.5 | 14.9743 | 21.4 | 22.82 | 22.21 | 25.9396 | 12.6 | 24.4749 | 19.79 | 21.4 | 24 | 23 | 21.7572 | #N/A
| 1 | 3 | 23.5 | 23.4 | 24.8 | 23.1 | 25.1 | 22 | 21.2 | 25 | 21.9 | 14.9902 | 21.6 | 22.92 | 20.85 | 25.871 | 12.8 | 24.8131 | 19.22 | 21.3 | 24 | 23 | 22.7176 | 16
| 1 | 4 | 23.5 | 22.4 | 24.5 | 21.8 | 24.4 | 22 | 21.4 | 27 | 21.6 | 15.0206 | 21.5 | 23.1 | 21.52 | 24.1741 | 12.7 | 23.4547 | 19.16 | 21 | 24 | 24 | 21.8828 | 22.8092
| 1 | 5 | 24 | 22.8 | 23.7 | 23.7 | 24.2 | 20 | 22.1 | 25 | 21.7 | 15.0413 | 22 | 22.12 | 21.41 | 24.6554 | 12.7 | 24.594 | 18.93 | 20 | 24 | 25 | 21.319 | 86
| 1 | 6 | 23.7 | 24 | 23.9 | 23.8 | 22.8 | 21 | 20.9 | 25 | 22.6 | 14.9606 | 21.6 | 22.82 | 20.36 | 21.8487 | 12.7 | 25.0945 | 18.49 | 21.7 | 23 | 24 | 21.3289
| Mean |  | 23.7833 | 23.2 | 24.2667 | 23.5 | 24.3 | 21.1667 | 21.7 | 25 | 21.8167 | 15.0017 | 21.55 | 22.8533 | 21.315 | 25.4032 | 12.7 | 24.3805 | 19.1983 | 21.1 | 24.1667 | 23.6667 | 21.9141
| Median |  | 23.85 | 23.2 | 24.35 | 23.75 | 24.3 | 21 | 21.75 | 25 | 21.65 | 15.0054 | 21.55 | 22.87 | 21.465 | 25.2632 | 12.7 | 24.5344 | 19.19 | 21.25 | 24 | 23.5 | 21.82
| Std Dev. |  | 0.24833 | 0.57966 | 0.41312 | 0.97775 | 0.8438 | 0.75277 | 0.6229 | 1.0954 | 0.40702 | 0.031463 | 0.26646 | 0.4104 | 0.63733 | 2.6723 | 0.063246 | 0.61451 | 0.46628 | 0.58652 | 0.98319 | 0.8165 | 0.58086
| Rel.Std.Dev. |  | 1.044% | 2.499% | 1.702% | 4.161% | 3.472% | 3.556% | 2.87% | 4.382% | 1.866% | 0.2097% | 1.236% | 1.796% | 2.99% | 10.52% | 0.498% | 2.52% | 2.429% | 2.78% | 4.068% | 3.45% | 2.651%
| PDM3 |  | 4.271% | 1.713% | 6.39% | 3.028% | 6.536% | -7.201% | -4.863% | 9.605% | -4.352% | -34.23% | -5.521% | 0.1933% | -6.551% | 11.37% | -44.32% | 6.889% | -15.83% | -7.494% | 5.951% | 3.759% | -3.924%
| Z-Score (Absolute) |  | 0.44 | 0.17 | 0.67 | 0.31 | 0.69 | 0.8 | 0.55 | 1.02 | 0.49 | 3.75 | 0.62 | 0 | 0.73 | 1.22 | 4.84 | 0.73 | 1.74 | 0.84 | 0.63 | 0.39 | 0.45
| Analytical results for Yb in OREAS 702 (Certified Value 1.99 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 16 | 18 | 19 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  | 2
3
| 1 | 1 | 2.02 | 1.94 | 2.04 | 2.17 | 2.19 | 1.9 | 2.42 | 2 | 2 | 1.98 | 23 | 1.8 | 2.1795 | 0.98 | 2.1739 | 1.9 | 2 | 2.21 | 1.9246 |  |  | 1
| 1 | 2 | 2.02 | 1.9 | 2.08 | 2.27 | 2.19 | 1.9 | 2.2 | 2 | 2 | 1.96 | 24 | 2 | 1.9582 | 1.01 | 2.1823 | 2 | 2 | 2.09 | 1.6435 |  |  | #N/A
| 1 | 3 | 1.95 | 1.76 | 2.04 | 2.01 | 1.91 | 2.1 | 1.77 | 2 | 2 | 1.97 | 24.4 | 2.1 | 1.9746 | 0.99 | 2.1072 | 1.9 | 1.5 | 2.03 | 1.9303 |  |  | 16
| 1 | 4 | 1.93 | 2.11 | 2 | 1.9 | 2.29 | 2 | 1.33 | 2 | 1.9 | 2.04 | 23.8 | 2 | 1.8703 | 0.98 | 2.1434 | 1.9 | 1.5 | 2.2 | 2.1384 |  |  | 1.9907
| 1 | 5 | 2 | 1.78 | 2 | 2.09 | 1.98 | 1.9 | 2.01 | < 2 | 2 | 2.03 | 23.4 | 2.1 | 1.8665 | 0.99 | 2.2443 | 1.7 | 2 | 1.96 | 1.9576 |  |  | 87
| 1 | 6 | 1.94 | 1.84 | 2.05 | 2.08 | 1.8 | 2.1 | 1.51 | 2 | 2.1 | 1.91 | 23.7 | 1.9 | 1.7226 | 1.02 | 2.2457 | 1.9 | 2 | 2.06 | 2.1281
| Mean |  | 1.9767 | 1.8883 | 2.035 | 2.0867 | 2.06 | 1.9833 | 1.8733 | 2 | 2 | 1.9817 | 23.7167 | 1.9833 | 1.9286 | 0.995 | 2.1828 | 1.8833 | 1.8333 | 2.0917 | 1.9538
| Median |  | 1.975 | 1.87 | 2.04 | 2.085 | 2.085 | 1.95 | 1.89 | 2 | 2 | 1.975 | 23.75 | 2 | 1.9143 | 0.99 | 2.1781 | 1.9 | 2 | 2.075 | 1.9439
| Std Dev. |  | 0.041312 | 0.12844 | 0.030822 | 0.12754 | 0.19142 | 0.098319 | 0.41524 | 0 | 0.063246 | 0.047924 | 0.48339 | 0.1169 | 0.15208 | 0.016432 | 0.054908 | 0.098319 | 0.2582 | 0.09786 | 0.18009
| Rel.Std.Dev. |  | 2.09% | 6.802% | 1.515% | 6.112% | 9.292% | 4.957% | 22.17% | 0% | 3.162% | 2.418% | 2.038% | 5.894% | 7.885% | 1.651% | 2.515% | 5.22% | 14.08% | 4.679% | 9.218%
| PDM3 |  | -0.7028% | -5.14% | 2.228% | 4.823% | 3.483% | -0.3679% | -5.894% | 0.4693% | 0.4693% | -0.4516% | 1091% | -0.3679% | -3.117% | -50.02% | 9.653% | -5.391% | -7.903% | 5.074% | -1.854%
| Z-Score (Absolute) |  | 0.05 | 0.74 | 0.41 | 0.82 | 0.61 | 0.01 | 0.86 | NA | 0.14 | 0.01 | 170.74 | 0.01 | 0.42 | 7.76 | 1.57 | 0.78 | 1.17 | 0.86 | 0.23
Results from laboratory 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 702 (Certified Value 59 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 06 | 07 | 09 | 10 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | [6 empty] | 0
0
| 1 | 1 | 60 | 50 | 61 | 58 | 60 | 59 | 54 | 73.2536 | 57 | 49.9463 | 55.16 | 64 | < 100 | 52 | 62.04 | [6 empty] | 1
| 1 | 2 | 70 | 50 | 65 | 56 | 60 | 61 | 55 | 74.2838 | 56 | 50.4496 | 54.79 | 65 | < 100 | 60 | 66.08 | [6 empty] | 22
| 1 | 3 | 70 | 50 | 63 | 59 | 60 | 59 | 59 | 71.9548 | 56 | 49.8268 | 54.07 | 60 | < 100 | 53 | 61.74 | [6 empty] | 16
| 1 | 4 | 70 | 50 | 59 | 86 | 60 | 60 | 58 | 70.7978 | 56 | 51.9696 | 56.98 | 60 | < 100 | 61 | 58.71 | [6 empty] | 59.2041
| 1 | 5 | 70 | 50 | 65 | 58 | 60 | 58 | 56 | 70.3439 | 56 | 49.6642 | 53.66 | 65 | < 100 | 50 | 63.33 | [6 empty] | 88
| 1 | 6 | 70 | 50 | 66 | 55 | 60 | 59 | 62 | 70.2433 | 57 | 48.8825 | 55.46 | 62 | < 100 | 54 | 63.31
| Mean |  | 68.3333 | 50 | 63.1667 | 62 | 60 | 59.3333 | 57.3333 | 71.8129 | 56.3333 | 50.1232 | 55.02 | 62.6667 |  | 55 | 62.535
| Median |  | 70 | 50 | 64 | 58 | 60 | 59 | 57 | 71.3763 | 56 | 49.8865 | 54.975 | 63 |  | 53.5 | 62.675
| Std Dev. |  | 4.0825 | 0 | 2.7142 | 11.8491 | 0 | 1.0328 | 2.9439 | 1.6645 | 0.5164 | 1.0378 | 1.1713 | 2.3381 |  | 4.4721 | 2.421
| Rel.Std.Dev. |  | 5.974% | 0% | 4.297% | 19.11% | 0% | 1.741% | 5.135% | 2.318% | 0.9167% | 2.07% | 2.129% | 3.731% |  | 8.131% | 3.872%
| PDM3 |  | 15.42% | -15.55% | 6.693% | 4.722% | 1.344% | 0.2182% | -3.16% | 21.3% | -4.849% | -15.34% | -7.067% | 5.848% |  | -7.101% | 5.626%
| Z-Score (Absolute) |  | 1.58 | 1.64 | 0.67 | 0.47 | 0.12 | 0 | 0.35 | 2.2 | 0.53 | 1.62 | 0.76 | 0.59 | 1.64 | 0.76 | 0.56
| Analytical results for Zr in OREAS 702 (Certified Value 167 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 15 | 17 | 18 | 19 | 23 | 24 | [6 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | BF*OES/MS | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | [6 empty] | 0
0
| 1 | 1 | 161 | 169 | 177 | 168 | 180 | 164 | 149 | 165 | 167 | 141.8 | 106.0108 | 202.9 | 173.1603 | 152 | 140 | [6 empty] | 1
| 1 | 2 | 172 | 173 | 172 | 167 | 179 | 164 | 147 | 165 | 163.1 | 136.2 | 107.5527 | 199.7 | 174.9283 | 158 | 155 | [6 empty] | 4
| 1 | 3 | 164 | 165 | 178 | 177 | 172 | 167 | 156 | 165 | 167 | 111.2 | 107.1361 | 198.2 | 179.1942 | 183 | 130 | [6 empty] | 16
| 1 | 4 | 162 | 174 | 173 | 168 | 172 | 159 | 156 | 171 | 166.1 | 127.8 | 105.7141 | 204.1 | 173.7788 | 203 | 135 | [6 empty] | 167.3616
| 1 | 5 | 169 | 163 | 175 | 174 | 168 | 166 | 149 | 166 | 169.9 | 137.9 | 105.4577 | 203 | 174.8379 | 174 | 130 | [6 empty] | 89
| 1 | 6 | 171 | 164 | 162 | 168 | 167 | 169 | 158 | 168 | 161 | 135.4 | 112.0511 | 208.2 | 175.4639 | 160 | 135
| Mean |  | 166.5 | 168 | 172.8333 | 170.3333 | 173 | 164.8333 | 152.5 | 166.6667 | 165.6833 | 131.7167 | 107.3204 | 202.6833 | 175.2272 | 171.6667 | 137.5
| Median |  | 166.5 | 167 | 174 | 168 | 172 | 165 | 152.5 | 165.5 | 166.55 | 135.8 | 106.5735 | 202.95 | 174.8831 | 167 | 135
| Std Dev. |  | 4.7645 | 4.7329 | 5.7764 | 4.1312 | 5.4406 | 3.4303 | 4.6797 | 2.4221 | 3.1632 | 11.044 | 2.459 | 3.5097 | 2.1168 | 19.1067 | 9.3541
| Rel.Std.Dev. |  | 2.862% | 2.817% | 3.342% | 2.425% | 3.145% | 2.081% | 3.069% | 1.453% | 1.909% | 8.385% | 2.291% | 1.732% | 1.208% | 11.13% | 6.803%
| PDM3 |  | -0.5148% | 0.3815% | 3.269% | 1.776% | 3.369% | -1.511% | -8.88% | -0.4152% | -1.003% | -21.3% | -35.88% | 21.11% | 4.7% | 2.572% | -17.84%
| Z-Score (Absolute) |  | 0.02 | 0.15 | 0.67 | 0.4 | 0.69 | 0.2 | 1.55 | 0 | 0.11 | 3.82 | 6.49 | 3.94 | 0.94 | 0.55 | 3.19
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C423: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C442: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C460: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C478: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C496: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C515: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C534: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C552: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C570: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C606: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C661: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 702 (Certified Value 0.178 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.16 | 0.12 | 0.16 | 0.17 | 0.2 | 0.18 | < 2 | 0.19 | 0.1 | < 0.5 | < 0.5 | 0.19 | 0.18 | < 5 | < 1 | 0.12 | 0.17325 | 0.1 | 0.2 | 0.3 | 0.16 |  |  | 1
| 1 | 2 | 0.16 | 0.13 | 0.16 | 0.18 | 0.18 | 0.18 | < 2 | 0.2 | 0.1 | < 0.5 | < 0.5 | 0.21 | 0.18 | < 5 | < 1 | 0.13 | 0.17111 | 0.2 | < 0.2 | 0.2 | 0.18 |  |  | 11
| 1 | 3 | 0.18 | 0.12 | 0.16 | 0.16 | 0.28 | 0.14 | < 2 | 0.19 | 0.1 | < 0.5 | < 0.5 | 0.21 | 0.19 | < 5 | < 1 | 0.13 | 0.17869 | 0.2 | < 0.2 | 0.3 | 0.18 |  |  | 16
| 1 | 4 | 0.18 | 0.12 | 0.16 | 0.15 | 0.18 | 0.19 | < 2 | 0.19 | 0.1 | < 0.5 | < 0.5 | 0.18 | 0.17 | < 5 | < 1 | 0.12 | 0.15617 | 0.2 | < 0.2 | 0.2 | 0.17 |  |  | 0.17836
| 1 | 5 | 0.17 | 0.1 | 0.16 | 0.16 | 0.19 | 0.19 | < 2 | 0.19 | 0.1 | < 0.5 | < 0.5 | 0.21 | 0.19 | < 5 | < 1 | 0.15 | 0.16885 | 0.2 | < 0.2 | 0.2 | 0.2 |  |  | 91
| 1 | 6 | 0.18 | 0.12 | 0.18 | 0.16 | 0.18 | 0.22 | < 2 | 0.19 | 0.1 | < 0.5 | < 0.5 | 0.2 | 0.16 | < 5 | < 1 | 0.12 | 0.16739 | 0.2 | < 0.2 | 0.1 | 0.17
| Mean |  | 0.17167 | 0.11833 | 0.16333 | 0.16333 | 0.20167 | 0.18333 |  | 0.19167 | 0.1 |  |  | 0.2 | 0.17833 |  |  | 0.12833 | 0.16924 | 0.18333 | 0.2 | 0.21667 | 0.17667
| Median |  | 0.175 | 0.12 | 0.16 | 0.16 | 0.185 | 0.185 |  | 0.19 | 0.1 |  |  | 0.205 | 0.18 |  |  | 0.125 | 0.16998 | 0.2 | 0.2 | 0.2 | 0.175
| Std Dev. |  | 0.0098319 | 0.0098319 | 0.008165 | 0.010328 | 0.0392 | 0.02582 |  | 0.0040825 | 1.5202e-17 |  |  | 0.012649 | 0.01169 |  |  | 0.01169 | 0.0075273 | 0.040825 |  | 0.075277 | 0.013663
| Rel.Std.Dev. |  | 5.727% | 8.309% | 4.999% | 6.323% | 19.44% | 14.08% |  | 2.13% | 1.52e-14% |  |  | 6.325% | 6.555% |  |  | 9.109% | 4.448% | 22.27% |  | 34.74% | 7.734%
| PDM3 |  | -3.751% | -33.65% | -8.424% | -8.424% | 13.07% | 2.79% |  | 7.462% | -43.93% |  |  | 12.13% | -0.01348% |  |  | -28.05% | -5.112% | 2.79% | 12.13% | 21.48% | -0.9479%
| Z-Score (Absolute) |  | 0.39 | 2.19 | 0.67 | 0.67 | 0.62 | 0 | 27.53 | 0.28 | NA | 2.25 | 2.25 | 0.56 | 0.17 | 78.11 | 10.68 | 1.85 | 0.48 | NA | NA | NA | 0.22
Results from laboratories 10, 21, 23 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 702 (Certified Value 5.1 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.2 | 5.11 | 5.03 | 5.09 | 5.21 | 4.98 | 5.23 | 4.7842 | 4.95 | 5.1587 | 4.7771 | 5.1 | 5.1 | 5.45 | 5.4186 | 5.26 | 5.0669 | 5.47 | 5.0158 | 5.15 | 5.062 | 5.08 |  | 1
| 1 | 2 | 5.14 | 5.07 | 4.96 | 5.12 | 4.94 | 4.99 | 5.27 | 4.7966 | 4.95 | 5.1679 | 4.6505 | 5.24 | 4.98 | 5.49 | 5.4555 | 5.29 | 5.1216 | 5.53 | 4.8859 | 5.24 | 5.121 | 5.14 |  | #N/A
| 1 | 3 | 5.4 | 5.01 | 4.95 | 4.92 | 4.89 | 5.07 | 5.09 | 4.8528 | 5.03 | 5.1582 | 4.6765 | 5.23 | 4.92 | 6 | 5.4096 | 5.27 | 5.1778 | 5.5 | 5.1206 | 5.17 | 5.05 | 5.24 |  | 16
| 1 | 4 | 5.16 | 5.06 | 5.04 | 4.99 | 4.94 | 5.15 | 5.18 | 4.7763 | 4.97 | 5.1421 | 4.8553 | 5.21 | 4.96 | 5.73 | 5.2875 | 5.27 | 5.1699 | 5.5 | 5.1715 | 5.2 | 5.021 | 5.15 |  | 5.1021
| 1 | 5 | 4.94 | 5.09 | 5.02 | 5.08 | 5.01 | 5.11 | 5.18 | 4.6988 | 4.91 | 5.157 | 4.8545 | 5.18 | 5 | 5.39 | 5.5675 | 5.25 | 5.1303 | 5.54 | 4.9941 | 5.2 | 5.06 | 5.18 |  | 92
| 1 | 6 | 5.15 | 5.03 | 4.93 | 5.22 | 5.12 | 5.13 | 5.21 | 4.72 | 4.86 | 5.1439 | 4.6677 | 5.21 | 4.94 | 5.19 | 5.5227 | 5.25 | 5.1459 | 5.54 | 5.2631 | 5.18 | 5.087 | 5.15
| Mean |  | 5.165 | 5.0617 | 4.9883 | 5.07 | 5.0183 | 5.0717 | 5.1933 | 4.7715 | 4.945 | 5.1547 | 4.7469 | 5.195 | 4.9833 | 5.5417 | 5.4436 | 5.265 | 5.1354 | 5.5133 | 5.0752 | 5.19 | 5.0668 | 5.1567
| Median |  | 5.155 | 5.065 | 4.99 | 5.085 | 4.975 | 5.09 | 5.195 | 4.7803 | 4.95 | 5.1576 | 4.7268 | 5.21 | 4.97 | 5.47 | 5.437 | 5.265 | 5.1381 | 5.515 | 5.0682 | 5.19 | 5.061 | 5.15
| Std Dev. |  | 0.14694 | 0.037103 | 0.047081 | 0.10431 | 0.12319 | 0.072226 | 0.060882 | 0.055428 | 0.057184 | 0.0098466 | 0.094623 | 0.050892 | 0.06377 | 0.284 | 0.09788 | 0.015166 | 0.040018 | 0.028048 | 0.13607 | 0.030984 | 0.034044 | 0.052409
| Rel.Std.Dev. |  | 2.845% | 0.733% | 0.9438% | 2.057% | 2.455% | 1.424% | 1.172% | 1.162% | 1.156% | 0.191% | 1.993% | 0.9796% | 1.28% | 5.125% | 1.798% | 0.288% | 0.7793% | 0.5087% | 2.681% | 0.597% | 0.6719% | 1.016%
| PDM3 |  | 1.233% | -0.7921% | -2.229% | -0.6288% | -1.641% | -0.5961% | 1.789% | -6.48% | -3.079% | 1.031% | -6.961% | 1.821% | -2.327% | 8.616% | 6.693% | 3.193% | 0.6527% | 8.06% | -0.5275% | 1.723% | -0.6909% | 1.07%
| Z-Score (Absolute) |  | 0.46 | 0.34 | 0.9 | 0.27 | 0.67 | 0.26 | 0.68 | 2.57 | 1.24 | 0.38 | 2.76 | 0.69 | 0.94 | 3.36 | 2.61 | 1.23 | 0.23 | 3.14 | 0.23 | 0.65 | 0.3 | 0.4
| Analytical results for As in OREAS 702 (Certified Value 3.54 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 21 | 25 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 3.8 | 2.4 | 2.8 | 2.7 | 3.3 | 3.6 | < 20 | 3.3 | 3 | < 5 | < 5 | 4.23 | < 10 | 4.6 | 3.4996 | 3.8 | 3.3 | [6 empty] | 1
| 1 | 2 | 3.1 | 2.2 | 3.3 | 2.9 | 2.3 | 3.4 | < 20 | 2 | 3 | < 5 | 5 | 4.03 | < 10 | 5 | 3.7968 | 4.8 | 3.6 | [6 empty] | 12
| 1 | 3 | 3.7 | 2.6 | 2.9 | 3 | 2.7 | 3.6 | < 20 | 3.2 | 3 | < 5 | < 5 | 3.94 | < 10 | 5.1 | 3.6102 | 3.5 | 3.9 | [6 empty] | 16
| 1 | 4 | 3.7 | 3.3 | 3.2 | 3 | 2.7 | 3.6 | < 20 | 3 | 3 | < 5 | < 5 | 4.1 | < 10 | 4.8 | 3.6832 | 4.9 | 3.9 | [6 empty] | 3.5448
| 1 | 5 | 3.5 | 2.5 | 3.8 | 3.2 | 3.5 | 3.2 | < 20 | 3.7 | 3 | < 5 | < 5 | 4.12 | < 10 | 5.4 | 3.7757 | 4.7 | 3.5 | [6 empty] | 93
| 1 | 6 | 3.5 | 2.3 | 3.7 | 2.8 | 3.1 | 3.6 | < 20 | 2.8 | 3 | < 5 | < 5 | 4.12 | < 10 | 5.2 | 3.5196 | 4.8 | 3.5
| Mean |  | 3.55 | 2.55 | 3.2833 | 2.9333 | 2.9333 | 3.5 |  | 3 | 3 |  | 5 | 4.09 |  | 5.0167 | 3.6475 | 4.4167 | 3.6167
| Median |  | 3.6 | 2.45 | 3.25 | 2.95 | 2.9 | 3.6 |  | 3.1 | 3 |  | 5 | 4.11 |  | 5.05 | 3.6467 | 4.75 | 3.55
| Std Dev. |  | 0.251 | 0.3937 | 0.40702 | 0.17512 | 0.44572 | 0.16733 |  | 0.57619 | 0 |  |  | 0.09757 |  | 0.28577 | 0.12621 | 0.6047 | 0.24014
| Rel.Std.Dev. |  | 7.07% | 15.44% | 12.4% | 5.97% | 15.2% | 4.781% |  | 19.21% | 0% |  |  | 2.386% |  | 5.696% | 3.46% | 13.69% | 6.64%
| PDM3 |  | 0.1469% | -28.06% | -7.376% | -17.25% | -17.25% | -1.264% |  | -15.37% | -15.37% |  | 41.05% | 15.38% |  | 41.52% | 2.898% | 24.6% | 2.028%
| Z-Score (Absolute) |  | 0 | 1.09 | 0.29 | 0.67 | 0.67 | 0.05 | 7.05 | 0.6 | NA | 1.15 | NA | 0.59 | 1.59 | 1.6 | 0.11 | 0.95 | 0.07
Results from laboratories 10 and 13 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 702 (Indicative Value 22.1 ppm) | [24 empty] | Indicative
B | Round | Replicate | Lab | [22 empty] | 1
| No | No | 14 | [22 empty] | ppm
|  |  | 4A*OES/MS | [22 empty] | 1
1
| 1 | 1 | 22.18 | [22 empty] | 1
| 1 | 2 | 21.91 | [22 empty] | 13
| 1 | 3 | 22.25 | [22 empty] | 16
| 1 | 4 | 22.06 | [22 empty] | 22.0533
| 1 | 5 | 22.31 | [22 empty] | 22
| 1 | 6 | 21.61
| Mean |  | 22.0533
| Median |  | 22.12
| Std Dev. |  | 0.2599
| Rel.Std.Dev. |  | 1.178%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 702 (Certified Value 657 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 670 | 660 | 660 | 630 | 680 | 654 | 663 | 617 | 679 | 656.5897 | 612.8503 | 648 | 670 | 629 | 609.6456 | 652 | 640.138 | 654 | 680.5403 | 638 | 640 |  |  | 1
| 1 | 2 | 660 | 650 | 650 | 640 | 640 | 654 | 668 | 620 | 673 | 656.0757 | 598.23 | 691 | 640 | 655 | 625.9444 | 648 | 646.8554 | 658 | 672.0204 | 658 | 646 |  |  | #N/A
| 1 | 3 | 690 | 650 | 650 | 630 | 640 | 664 | 646 | 618 | 692 | 655.5287 | 598.94 | 698 | 654 | 695 | 627.6815 | 648 | 654.2066 | 653 | 680.8985 | 666 | 662 |  |  | 16
| 1 | 4 | 650 | 650 | 670 | 630 | 640 | 673 | 660 | 616 | 688 | 658.1766 | 582.28 | 666 | 665 | 660 | 619.0347 | 646 | 651.7555 | 652 | 678.8823 | 648 | 652 |  |  | 656.8759
| 1 | 5 | 630 | 650 | 670 | 650 | 650 | 671 | 655 | 609 | 692 | 656.4896 | 590.19 | 689 | 654 | 622 | 619.9695 | 653 | 649.5105 | 668 | 676.056 | 671 | 650 |  |  | 94
| 1 | 6 | 670 | 650 | 650 | 650 | 670 | 644 | 656 | 608 | 699 | 657.2927 | 614.5867 | 673 | 683 | 606 | 614.2597 | 654 | 647.5741 | 659 | 660.5484 | 666 | 648
| Mean |  | 661.6667 | 651.6667 | 658.3333 | 638.3333 | 653.3333 | 660 | 658 | 614.6667 | 687.1667 | 656.6922 | 599.5128 | 677.5 | 661 | 644.5 | 619.4226 | 650.1667 | 648.34 | 657.3333 | 674.8243 | 657.8333 | 649.6667
| Median |  | 665 | 650 | 655 | 635 | 645 | 659 | 658 | 616.5 | 690 | 656.5397 | 598.585 | 681 | 659.5 | 642 | 619.5021 | 650 | 648.5423 | 656 | 677.4692 | 662 | 649
| Std Dev. |  | 20.4124 | 4.0825 | 9.8319 | 9.8319 | 17.5119 | 11.2606 | 7.5631 | 4.9666 | 9.5376 | 0.93268 | 12.5775 | 18.7483 | 14.9666 | 32.0297 | 6.8328 | 3.2506 | 4.848 | 5.9217 | 7.7358 | 12.6241 | 7.3121
| Rel.Std.Dev. |  | 3.085% | 0.6265% | 1.493% | 1.54% | 2.68% | 1.706% | 1.149% | 0.808% | 1.388% | 0.142% | 2.098% | 2.767% | 2.264% | 4.97% | 1.103% | 0.5% | 0.7478% | 0.9009% | 1.146% | 1.919% | 1.126%
| PDM3 |  | 0.7293% | -0.793% | 0.2219% | -2.823% | -0.5393% | 0.4756% | 0.1711% | -6.426% | 4.611% | -0.02796% | -8.733% | 3.14% | 0.6278% | -1.884% | -5.702% | -1.021% | -1.299% | 0.06964% | 2.732% | 0.1458% | -1.097%
| Z-Score (Absolute) |  | 0.51 | 0.52 | 0.17 | 1.9 | 0.35 | 0.34 | 0.14 | 4.34 | 3.15 | 0 | 5.91 | 2.15 | 0.45 | 1.26 | 3.85 | 0.67 | 0.86 | 0.07 | 1.87 | 0.12 | 0.73
| Analytical results for Be in OREAS 702 (Certified Value 5.24 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.28 | 5.27 | 5.36 | 5.08 | 5.3 | 6 | 5.7 | 4.64 | 5.6 | 5.1261 | 4.6207 | 6 | 5.38 | 6.1 | 5.0343 | 5.5 | 4.7834 | 6 | 5.455 | 5.22 | 5.23 |  |  | 1
| 1 | 2 | 5.18 | 5.2 | 5.3 | 5.22 | 5.09 | 5.19 | 5.8 | 4.69 | 5.6 | 5.1855 | 4.6897 | 6 | 5.33 | 6.1 | 4.9805 | 5.45 | 4.8269 | 6 | 5.8219 | 5.34 | 5.24 |  |  | #N/A
| 1 | 3 | 5.41 | 5.17 | 5.23 | 5.07 | 5.01 | 5.4 | 5.7 | 4.71 | 5.5 | 5.207 | 4.744 | 6 | 5.12 | 6.3 | 4.5916 | 5.43 | 4.8865 | 6 | 5.4013 | 5.43 | 5.21 |  |  | 16
| 1 | 4 | 5.17 | 5.14 | 5.42 | 5 | 5.05 | 5.68 | 5.4 | 4.67 | 5.4 | 5.2023 | 4.62 | 6 | 5.2 | 5.9 | 4.3816 | 5.48 | 4.8467 | 6 | 5.5032 | 5.3 | 5.12 |  |  | 5.2428
| 1 | 5 | 4.96 | 5.17 | 5.38 | 5.15 | 5.09 | 5.38 | 5.2 | 4.62 | 5.6 | 5.2669 | 4.6717 | 6 | 5.24 | 5.7 | 4.3889 | 5.45 | 4.8356 | 6 | 5.5804 | 5.49 | 5.27 |  |  | 95
| 1 | 6 | 5.28 | 5.17 | 5.32 | 5.16 | 5.25 | 6.11 | 5.7 | 4.63 | 5.5 | 5.2471 | 4.6383 | 6 | 5.25 | 5.8 | 4.6065 | 5.41 | 4.8727 | 6 | 5.695 | 5.48 | 5.2
| Mean |  | 5.2133 | 5.1867 | 5.335 | 5.1133 | 5.1317 | 5.6267 | 5.5833 | 4.66 | 5.5333 | 5.2058 | 4.6641 | 6 | 5.2533 | 5.9833 | 4.6639 | 5.4533 | 4.842 | 6 | 5.5761 | 5.3767 | 5.2117
| Median |  | 5.23 | 5.17 | 5.34 | 5.115 | 5.09 | 5.54 | 5.7 | 4.655 | 5.55 | 5.2046 | 4.655 | 6 | 5.245 | 6 | 4.599 | 5.45 | 4.8412 | 6 | 5.5418 | 5.385 | 5.22
| Std Dev. |  | 0.15148 | 0.045019 | 0.066858 | 0.078401 | 0.116 | 0.36844 | 0.23166 | 0.035777 | 0.08165 | 0.049405 | 0.048138 | 0 | 0.092448 | 0.22286 | 0.2833 | 0.03266 | 0.036482 | 0 | 0.15818 | 0.10783 | 0.051153
| Rel.Std.Dev. |  | 2.906% | 0.868% | 1.253% | 1.533% | 2.261% | 6.548% | 4.149% | 0.7677% | 1.476% | 0.949% | 1.032% | 0% | 1.76% | 3.725% | 6.074% | 0.5989% | 0.7534% | 0% | 2.837% | 2.005% | 0.9815%
| PDM3 |  | -0.5624% | -1.071% | 1.758% | -2.47% | -2.12% | 7.321% | 6.495% | -11.12% | 5.541% | -0.7058% | -11.04% | 14.44% | 0.2006% | 14.12% | -11.04% | 4.015% | -7.646% | 14.44% | 6.358% | 2.553% | -0.5942%
| Z-Score (Absolute) |  | 0 | 0.07 | 0.34 | 0.28 | 0.23 | 1.16 | 1.04 | 1.55 | 0.9 | 0.02 | 1.54 | NA | 0.11 | 2.16 | 1.54 | 0.67 | 1.04 | NA | 1.02 | 0.46 | 0
Results from laboratories 13 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 702 (Certified Value 277 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 282 | 292 | 301 | 288 | 298 | 270 | 290 | 256.65 | 264 | 336.0321 | 268.7819 | 296.7 | 230 | 264.97 | 233.0595 | 266.1 | 271.3564 | 291.2 | 324.9923 | 281 | 267.5 | 272.55 |  | 1
| 1 | 2 | 278 | 284 | 297 | 295 | 286 | 266 | 286 | 275.27 | 259 | 335.9221 | 269.8971 | 307.6 | 232 | 263.03 | 262.0007 | 268.16 | 280.0918 | 295.5 | 326.9952 | 286 | 266.4 | 278 |  | 22
| 1 | 3 | 291 | 285 | 294 | 281 | 282 | 270 | 285 | 265.14 | 259 | 336.5511 | 269.3026 | 304.8 | 231 | 273 | 274.7543 | 268.78 | 280.9202 | 293.5 | 324.6519 | 293 | 264.9 | 281.05 |  | 16
| 1 | 4 | 278 | 281 | 299 | 282 | 281 | 269 | 286 | 258.52 | 253 | 336.9621 | 257.0117 | 294 | 232 | 257.25 | 280.4785 | 270.19 | 279.3758 | 291.2 | 329.0653 | 283 | 252.6 | 277.19 |  | 277.4113
| 1 | 5 | 270 | 284 | 297 | 296 | 288 | 263 | 280 | 254.08 | 256 | 336.9251 | 267.2917 | 305.5 | 231 | 255.39 | 240.2843 | 270.81 | 278.7356 | 290.8 | 324.2094 | 298 | 253.5 | 275.79 |  | 96
| 1 | 6 | 282 | 279 | 293 | 296 | 296 | 251 | 288 | 253.84 | 264 | 337.0681 | 260.846 | 304.2 | 231 | 257.9 | 281.7233 | 271.68 | 277.7572 | 290.4 | 334.3871 | 293 | 254.3 | 279.92
| Mean |  | 280.1667 | 284.1667 | 296.8333 | 289.6667 | 288.5 | 264.8333 | 285.8333 | 260.5833 | 259.1667 | 336.5768 | 265.5218 | 302.1333 | 231.1667 | 261.9233 | 262.0501 | 269.2867 | 278.0395 | 292.1 | 327.3835 | 289 | 259.8667 | 277.4167
| Median |  | 280 | 284 | 297 | 291.5 | 287 | 267.5 | 286 | 257.585 | 259 | 336.7381 | 268.0368 | 304.5 | 231 | 260.465 | 268.3775 | 269.485 | 279.0557 | 291.2 | 325.9937 | 289.5 | 259.6 | 277.595
| Std Dev. |  | 6.8823 | 4.446 | 2.9944 | 7.0048 | 7.0922 | 7.3052 | 3.3714 | 8.294 | 4.3551 | 0.49742 | 5.3195 | 5.4456 | 0.75277 | 6.5408 | 20.9869 | 2.0271 | 3.4503 | 1.986 | 3.8803 | 6.6633 | 7.0797 | 3.0394
| Rel.Std.Dev. |  | 2.457% | 1.565% | 1.009% | 2.418% | 2.458% | 2.758% | 1.18% | 3.183% | 1.68% | 0.1478% | 2.003% | 1.802% | 0.3256% | 2.497% | 8.009% | 0.7528% | 1.241% | 0.6799% | 1.185% | 2.306% | 2.724% | 1.096%
| PDM3 |  | 0.9932% | 2.435% | 7.001% | 4.418% | 3.997% | -4.534% | 3.036% | -6.066% | -6.577% | 21.33% | -4.286% | 8.912% | -16.67% | -5.583% | -5.537% | -2.929% | 0.2264% | 5.295% | 18.01% | 4.177% | -6.324% | 0.001918%
| Z-Score (Absolute) |  | 0.05 | 0.25 | 0.86 | 0.51 | 0.46 | 0.69 | 0.33 | 0.9 | 0.97 | 2.78 | 0.66 | 1.12 | 2.32 | 0.83 | 0.83 | 0.48 | 0.05 | 0.63 | 2.34 | 0.48 | 0.93 | 0.08
| Analytical results for Ca in OREAS 702 (Certified Value 15.63 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 16.15 | 16.05 | 15.95 | 15.5 | 16.75 | 14.9 | 16.24 | 15.9169 | 15.2 | 14.8691 | 14.9757 | 16.38 | 14.92 | 14.58 | 15.76 | 16.19 | 14.5594 | 15.91 | 15.2386 | 15.7 | 14.88 |  |  | 1
| 1 | 2 | 16.05 | 16 | 15.7 | 15.85 | 15.95 | 15 | 16.38 | 16.2931 | 15.3 | 14.8429 | 14.8608 | 16.79 | 14.86 | > 15 | 15.91 | 16.17 | 14.7392 | 15.94 | 14.4178 | 16 | 15.28 |  |  | #N/A
| 1 | 3 | 16.85 | 15.85 | 15.6 | 15.4 | 15.75 | 15 | 15.83 | 16.3157 | 15.5 | 14.8914 | 14.3193 | 16.76 | 14.71 | > 15 | 15.93 | 16.26 | 14.8338 | 15.99 | 15.362 | 15.8 | 15.47 |  |  | 16
| 1 | 4 | 15.95 | 15.85 | 16.2 | 15.45 | 15.75 | 15.2 | 16.16 | 16.3211 | 15.3 | 14.868 | 14.7579 | 16.66 | 14.84 | > 15 | 15.53 | 16.07 | 14.8087 | 16.02 | 15.7703 | 15.9 | 15.13 |  |  | 15.6337
| 1 | 5 | 15.25 | 15.8 | 16.05 | 15.8 | 16.1 | 15.2 | 16.19 | 16.0615 | 15.1 | 14.8531 | 15.4273 | 16.58 | 14.88 | 14.57 | 16.21 | 16.25 | 14.8166 | 16 | 15.2008 | 15.9 | 15.17 |  |  | 97
| 1 | 6 | 16.2 | 15.85 | 15.6 | 15.9 | 16.55 | 15.7 | 16.32 | 15.9562 | 15.4 | 14.8498 | 14.0464 | 16.69 | 14.8 | 14.47 | 15.9 | 16.23 | 14.7399 | 16.04 | 15.9763 | 15.9 | 15.23
| Mean |  | 16.075 | 15.9 | 15.85 | 15.65 | 16.1417 | 15.1667 | 16.1867 | 16.1441 | 15.3 | 14.8624 | 14.7312 | 16.6433 | 14.835 | 14.54 | 15.8733 | 16.195 | 14.7496 | 15.9833 | 15.3276 | 15.8667 | 15.1933
| Median |  | 16.1 | 15.85 | 15.825 | 15.65 | 16.025 | 15.1 | 16.215 | 16.1773 | 15.3 | 14.8606 | 14.8093 | 16.675 | 14.85 | 14.57 | 15.905 | 16.21 | 14.7743 | 15.995 | 15.3003 | 15.9 | 15.2
| Std Dev. |  | 0.51357 | 0.1 | 0.25298 | 0.22361 | 0.42002 | 0.28752 | 0.19284 | 0.188 | 0.14142 | 0.017588 | 0.49001 | 0.14895 | 0.073144 | 0.060828 | 0.22331 | 0.070356 | 0.10143 | 0.049261 | 0.54216 | 0.10328 | 0.19398
| Rel.Std.Dev. |  | 3.195% | 0.6289% | 1.596% | 1.429% | 2.602% | 1.896% | 1.191% | 1.165% | 0.9243% | 0.1183% | 3.326% | 0.895% | 0.493% | 0.4183% | 1.407% | 0.4344% | 0.6877% | 0.3082% | 3.537% | 0.6509% | 1.277%
| PDM3 |  | 2.822% | 1.703% | 1.383% | 0.104% | 3.249% | -2.988% | 3.537% | 3.264% | -2.135% | -4.934% | -5.773% | 6.458% | -5.109% | -6.996% | 1.532% | 3.59% | -5.655% | 2.236% | -1.958% | 1.49% | -2.817%
| Z-Score (Absolute) |  | 0.44 | 0.1 | 0 | 0.39 | 0.57 | 1.34 | 0.66 | 0.57 | 1.07 | 1.93 | 2.19 | 1.55 | 1.98 | 2.56 | 0.05 | 0.67 | 2.15 | 0.26 | 1.02 | 0.03 | 1.28
| Analytical results for Cd in OREAS 702 (Indicative Value 0.21 ppm) | [24 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 0.11 | 0.03 | 0.17 | 0.07 | 0.25 | 0.33 | < 5 | 0.22 | 0.14 | 0.6294 | < 0.5 | < 0.5 | 0.26 | < 5 | 1.087 | 0.36 | < 0.5 | 0.3 | < 0.3 | < 2 | 1 | 0.21 |  | 1
| 1 | 2 | 0.07 | < 0.02 | 0.17 | 0.07 | 0.24 | 0.26 | < 5 | 0.26 | 0.15 | 0.62 | < 0.5 | < 0.5 | 0.28 | < 5 | 1.0903 | 0.38 | < 0.5 | 0.3 | < 0.3 | < 2 | 1 | 0.21 |  | 1
| 1 | 3 | 0.11 | < 0.02 | 0.16 | 0.08 | 0.22 | 0.31 | < 5 | 0.23 | 0.16 | 0.62 | < 0.5 | < 0.5 | 0.21 | < 5 | 1.0922 | 0.39 | < 0.5 | 0.3 | < 0.3 | < 2 | 1 | 0.22 |  | 16
| 1 | 4 | 0.1 | 0.04 | 0.13 | 0.1 | 0.23 | 0.33 | < 5 | 0.23 | 0.14 | 0.6 | < 0.5 | < 0.5 | 0.27 | < 5 | 1.0782 | 0.35 | < 0.5 | 0.4 | < 0.3 | < 2 | 1 | 0.21 |  | 0.20553
| 1 | 5 | 0.09 | < 0.02 | 0.2 | 0.05 | 0.23 | 0.24 | < 5 | 0.23 | 0.15 | 0.63 | < 0.5 | < 0.5 | 0.21 | < 5 | 1.0939 | 0.38 | < 0.5 | 0.3 | < 0.3 | < 2 | 1 | 0.23 |  | 23
| 1 | 6 | 0.12 | < 0.02 | 0.18 | 0.1 | 0.23 | 0.29 | < 5 | 0.23 | 0.16 | 0.64 | < 0.5 | < 0.5 | 0.3 | < 5 | 1.0829 | 0.39 | < 0.5 | 0.4 | < 0.3 | < 2 | 1 | 0.22
| Mean |  | 0.1 | 0.035 | 0.16833 | 0.078333 | 0.23333 | 0.29333 |  | 0.23333 | 0.15 | 0.62323 |  |  | 0.255 |  | 1.0874 | 0.375 |  | 0.33333 |  |  | 1 | 0.21667
| Median |  | 0.105 | 0.035 | 0.17 | 0.075 | 0.23 | 0.3 |  | 0.23 | 0.15 | 0.6247 |  |  | 0.265 |  | 1.0886 | 0.38 |  | 0.3 |  |  | 1 | 0.215
| Std Dev. |  | 0.017889 | 0.0070711 | 0.023166 | 0.019408 | 0.010328 | 0.037238 |  | 0.013663 | 0.0089443 | 0.013606 |  |  | 0.037283 |  | 0.005961 | 0.016432 |  | 0.05164 |  |  | 0 | 0.008165
| Rel.Std.Dev. |  | 17.89% | 20.2% | 13.76% | 24.78% | 4.426% | 12.69% |  | 5.855% | 5.963% | 2.183% |  |  | 14.62% |  | 0.5482% | 4.382% |  | 15.49% |  |  | 0% | 3.768%
| PDM3 |  | -51.34% | -82.97% | -18.1% | -61.89% | 13.53% | 42.72% |  | 13.53% | -27.02% | 203.2% |  |  | 24.07% |  | 429.1% | 82.46% |  | 62.18% |  |  | 386.6% | 5.42%
| Z-Score (Absolute) |  | 1.01 | 1.56 | 0.55 | 1.16 | 0.11 | 0.29 | 15.17 | 0.11 | 0.67 | 2.52 | 0 | 0 | 0.03 | 15.17 | 5.65 | 0.84 | 0 | 0.56 | 0.67 | 5.06 | 5.06 | 0.22
| Analytical results for Ce in OREAS 702 (Certified Value 59 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | 25 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 61.3 | 57.4 | 62.1 | 53.1 | 59.5 | 58.3 | 65.3 | 25.7 | 65.8604 | 57 | 61 | 64.19 | 46.51 | 56.81 | 56.4003 | 51 | 54.066 | 56.08 | [5 empty] | 1
| 1 | 2 | 57.1 | 57.4 | 60.8 | 56.5 | 56.8 | 57.8 | 65.1 | 25.8 | 65.6086 | 62 | 59 | 61.63 | 45.81 | 58.72 | 59.1665 | 49 | 56.329 | 58.86 | [5 empty] | 2
| 1 | 3 | 61.2 | 59.1 | 60.4 | 59 | 56.7 | 57.2 | 64.3 | 27 | 65.5542 | 59 | 59 | 64.13 | 45.17 | 57.37 | 57.9276 | 49 | 55.089 | 58.51 | [5 empty] | 16
| 1 | 4 | 62.8 | 57.9 | 60.7 | 54.5 | 59.4 | 56.6 | 62 | 27.8 | 65.8924 | 57 | 60 | 60.9 | 46.18 | 58.73 | 57.9924 | 48 | 53.957 | 58.56 | [5 empty] | 59.2081
| 1 | 5 | 57.7 | 58.9 | 62 | 55.9 | 56.7 | 54.8 | 61.3 | 26.7 | 65.6023 | 60 | 58 | 63.11 | 45.67 | 58.4 | 59.2903 | 49 | 53.232 | 57.05 | [5 empty] | 98
| 1 | 6 | 60 | 57.9 | 59.4 | 56.7 | 60.6 | 56 | 62.3 | 25.2 | 65.4282 | 60 | 61 | 56.51 | 46.14 | 57.81 | 58.3445 | 45 | 55.868 | 56.25
| Mean |  | 60.0167 | 58.1 | 60.9 | 55.95 | 58.2833 | 56.7833 | 63.3833 | 26.3667 | 65.6577 | 59.1667 | 59.6667 | 61.745 | 45.9133 | 57.9733 | 58.1869 | 48.5 | 54.7568 | 57.5517
| Median |  | 60.6 | 57.9 | 60.75 | 56.2 | 58.1 | 56.9 | 63.3 | 26.25 | 65.6055 | 59.5 | 59.5 | 62.37 | 45.975 | 58.105 | 58.1685 | 49 | 54.5775 | 57.78
| Std Dev. |  | 2.2212 | 0.73485 | 1.0198 | 2.0197 | 1.7498 | 1.2719 | 1.7256 | 0.96885 | 0.18167 | 1.9408 | 1.2111 | 2.8866 | 0.46924 | 0.78181 | 1.0487 | 1.9748 | 1.2049 | 1.2457
| Rel.Std.Dev. |  | 3.701% | 1.265% | 1.675% | 3.61% | 3.002% | 2.24% | 2.722% | 3.675% | 0.2767% | 3.28% | 2.03% | 4.675% | 1.022% | 1.349% | 1.802% | 4.072% | 2.2% | 2.164%
| PDM3 |  | 1.366% | -1.872% | 2.858% | -5.503% | -1.562% | -4.095% | 7.052% | -55.47% | 10.89% | -0.06998% | 0.7745% | 4.285% | -22.45% | -2.085% | -1.725% | -18.09% | -7.518% | -2.798%
| Z-Score (Absolute) |  | 0.62 | 0.01 | 0.91 | 0.73 | 0.05 | 0.45 | 1.74 | 10.54 | 2.49 | 0.34 | 0.51 | 1.19 | 4.06 | 0.06 | 0.01 | 3.2 | 1.12 | 0.2
| Analytical results for Co in OREAS 702 (Certified Value 6.47 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.2 | 6.8 | 7.3 | 6.2 | 7.2 | 6.6 | < 5 | 5.8 | 6 | 7.711 | 8.1661 | 7 | 6.45 | 6.8 | 6.7 | 5.845 | 6.4 | 6.5863 | 6.4 | 3.4026 | 5.5 | 6.2 |  | 1
| 1 | 2 | 6 | 6.9 | 7.2 | 6.4 | 6.8 | 6.4 | < 5 | 5.8 | 7 | 7.7587 | 7.994 | 7 | 6.36 | 7 | 5.6 | 5.8722 | 6.6 | 6.6369 | 6.4 | 3.5772 | 6.9 | 6.2 |  | #N/A
| 1 | 3 | 6.5 | 6.9 | 7.2 | 6.8 | 6.6 | 6.5 | < 5 | 5.7 | 6 | 7.7166 | 8.075 | 7 | 6.34 | 7.1 | 5.74 | 5.6576 | 6.5 | 6.5947 | 6.3 | 4.1874 | 6.5 | 6.4 |  | 16
| 1 | 4 | 6.3 | 6.8 | 7.1 | 6.5 | 6.9 | 6.7 | < 5 | 5.7 | 6 | 7.7446 | 8.12 | 7 | 6.74 | 6.6 | 6.64 | 5.4521 | 6.6 | 6.7074 | 6.3 | 4.0945 | 6.7 | 6.3 |  | 6.4747
| 1 | 5 | 6.2 | 6.9 | 7.1 | 6.9 | 6.9 | 6.4 | < 5 | 5.8 | 7 | 7.7786 | 8.09 | 7 | 6.57 | 7 | 5.95 | 5.5028 | 6.6 | 6.2524 | 6.3 | 3.5328 | 7.1 | 6.3 |  | 99
| 1 | 6 | 6.4 | 6.8 | 6.9 | 6.6 | 6.9 | 6.5 | 5 | 5.7 | 7 | 7.7896 | 8.075 | 7 | 7.06 | 7.2 | 5.72 | 5.8736 | 6.5 | 6.423 | 6.5 | 3.3982 | 7 | 6.3
| Mean |  | 6.2667 | 6.85 | 7.1333 | 6.5667 | 6.8833 | 6.5167 | 5 | 5.75 | 6.5 | 7.7499 | 8.0867 | 7 | 6.5867 | 6.95 | 6.0583 | 5.7006 | 6.5333 | 6.5334 | 6.3667 | 3.6988 | 6.6167 | 6.2833
| Median |  | 6.25 | 6.85 | 7.15 | 6.55 | 6.9 | 6.5 | 5 | 5.75 | 6.5 | 7.7516 | 8.0825 | 7 | 6.51 | 7 | 5.845 | 5.7513 | 6.55 | 6.5905 | 6.35 | 3.555 | 6.8 | 6.3
| Std Dev. |  | 0.17512 | 0.054772 | 0.13663 | 0.2582 | 0.19408 | 0.1169 |  | 0.054772 | 0.54772 | 0.032026 | 0.057083 | 0 | 0.27522 | 0.21679 | 0.48738 | 0.19129 | 0.08165 | 0.16654 | 0.08165 | 0.35093 | 0.58793 | 0.075277
| Rel.Std.Dev. |  | 2.794% | 0.7996% | 1.915% | 3.932% | 2.82% | 1.794% |  | 0.9526% | 8.427% | 0.4133% | 0.7059% | 0% | 4.178% | 3.119% | 8.045% | 3.356% | 1.25% | 2.549% | 1.282% | 9.488% | 8.886% | 1.198%
| PDM3 |  | -3.213% | 5.796% | 10.17% | 1.42% | 6.311% | 0.6477% | -22.78% | -11.19% | 0.3903% | 19.69% | 24.9% | 8.113% | 1.729% | 7.34% | -6.431% | -11.96% | 0.9051% | 0.9068% | -1.669% | -42.87% | 2.192% | -2.956%
| Z-Score (Absolute) |  | 0.57 | 0.67 | 1.28 | 0.07 | 0.75 | 0.04 | NA | 1.67 | NA | 2.59 | 3.31 | NA | 0.11 | 0.89 | 1.01 | 1.77 | 0 | 0 | 0.36 | 6.04 | 0.18 | 0.53
Results from laboratories 7, 10 and 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 702 (Certified Value 21.4 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 20 | 20 | 19 | 21 | 20 | 18 | < 20 | 22.8 | 21 | 25.5829 | 21 | 23.7 | 12 | 32.56 | 27.2197 | 26 | 19.6296 | 23 | 35.1428 | 21 | 21 |  |  | 1
| 1 | 2 | 19 | 19 | 18 | 21 | 19 | 19 | < 20 | 23 | 19 | 25.9612 | 21 | 24.18 | 14 | 28.41 | 25.7845 | 26 | 19.5803 | 23 | 32.2209 | 28 | 20 |  |  | #N/A
| 1 | 3 | 20 | 20 | 19 | 21 | 21 | 19 | < 20 | 23 | 20 | 25.07 | 21 | 24.12 | 15 | 30.28 | 21.7602 | 26 | 18.8701 | 23 | 32.118 | 27 | 21 |  |  | 16
| 1 | 4 | 19 | 20 | 19 | 20 | 20 | 18 | < 20 | 22.7 | 20 | 24.49 | 22 | 23.83 | 15 | 37.11 | 21.3724 | 26 | 18.8757 | 23 | 37.2634 | 27 | 19 |  |  | 21.3867
| 1 | 5 | 19 | 20 | 20 | 21 | 19 | 19 | < 20 | 22.5 | 22 | 25.0883 | 20 | 23.54 | 14 | 31.29 | 32.3516 | 26 | 18.2382 | 23 | 32.3855 | 25 | 19 |  |  | 100
| 1 | 6 | 23 | 20 | 21 | 21 | 20 | 23 | < 20 | 22.7 | 22 | 26.0967 | 22 | 23.51 | 22 | 31.07 | 35.4064 | 27 | 18.4979 | 23 | 36.9586 | 20 | 19
| Mean |  | 20 | 19.8333 | 19.3333 | 20.8333 | 19.8333 | 19.3333 |  | 22.7833 | 20.6667 | 25.3815 | 21.1667 | 23.8133 | 15.3333 | 31.7867 | 27.3158 | 26.1667 | 18.9486 | 23 | 34.3482 | 24.6667 | 19.8333
| Median |  | 19.5 | 20 | 19 | 21 | 20 | 19 |  | 22.75 | 20.5 | 25.3356 | 21 | 23.765 | 14.5 | 31.18 | 26.5021 | 26 | 18.8729 | 23 | 33.7642 | 26 | 19.5
| Std Dev. |  | 1.5492 | 0.40825 | 1.0328 | 0.40825 | 0.75277 | 1.8619 |  | 0.19408 | 1.2111 | 0.61088 | 0.75277 | 0.28577 | 3.4448 | 2.9451 | 5.6455 | 0.40825 | 0.56251 | 0 | 2.4206 | 3.3862 | 0.98319
| Rel.Std.Dev. |  | 7.746% | 2.058% | 5.342% | 1.96% | 3.795% | 9.631% |  | 0.8518% | 5.86% | 2.407% | 3.556% | 1.2% | 22.47% | 9.265% | 20.67% | 1.56% | 2.969% | 0% | 7.047% | 13.73% | 4.957%
| PDM3 |  | -6.484% | -7.263% | -9.601% | -2.588% | -7.263% | -9.601% |  | 6.53% | -3.367% | 18.68% | -1.029% | 11.35% | -28.3% | 48.63% | 27.72% | 22.35% | -11.4% | 7.543% | 60.61% | 15.34% | -7.263%
| Z-Score (Absolute) |  | 0.29 | 0.35 | 0.52 | 0 | 0.35 | 0.52 | 3.75 | 0.67 | 0.06 | 1.57 | 0.12 | 1.03 | 1.9 | 3.79 | 2.24 | 1.84 | 0.65 | 0.75 | 4.67 | 1.33 | 0.35
| Analytical results for Cs in OREAS 702 (Certified Value 8.77 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 8.56 | 9.18 | 9.71 | 8.7 | 9.03 | 9.34 | 9.6 | 8.7 | 8.7 | 8.02 | 8 | 6.0212 | 7.91 | 8.57 | 7.4 | 8.656 | 8.89 | [6 empty] | 1
| 1 | 2 | 8.38 | 9.19 | 9.47 | 9.19 | 8.47 | 8.62 | 9.3 | 9.3 | 8.5 | 8.13 | 8 | 6.1299 | 7.94 | 8.6621 | 7.8 | 9.129 | 9.11 | [6 empty] | #N/A
| 1 | 3 | 8.89 | 9.21 | 9.43 | 9.57 | 8.63 | 8.67 | 9.2 | 9.2 | 8.8 | 8.05 | 8 | 6.2122 | 8.07 | 8.686 | 7.8 | 8.88 | 9.24 | [6 empty] | 16
| 1 | 4 | 8.95 | 9.34 | 9.61 | 9.22 | 8.71 | 8.76 | 9.4 | 9 | 8.8 | 8.44 | 8 | 5.8098 | 7.97 | 8.5327 | 7.6 | 8.748 | 9.18 | [6 empty] | 8.7725
| 1 | 5 | 8.5 | 9.27 | 9.44 | 9.55 | 8.71 | 8.66 | 9.2 | 8.9 | 8.8 | 8.21 | 8 | 6.794 | 7.94 | 8.6916 | 7.9 | 8.163 | 9.24 | [6 empty] | 101
| 1 | 6 | 8.81 | 9.08 | 9.12 | 9.31 | 8.9 | 8.52 | 9.1 | 8.6 | 8.7 | 8.61 | 8 | 6.9665 | 8.09 | 8.5825 | 8 | 9.074 | 9.11
| Mean |  | 8.6817 | 9.2117 | 9.4633 | 9.2567 | 8.7417 | 8.7617 | 9.3 | 8.95 | 8.7167 | 8.2433 | 8 | 6.3223 | 7.9867 | 8.6208 | 7.75 | 8.775 | 9.1283
| Median |  | 8.685 | 9.2 | 9.455 | 9.265 | 8.71 | 8.665 | 9.25 | 8.95 | 8.75 | 8.17 | 8 | 6.171 | 7.955 | 8.6223 | 7.8 | 8.814 | 9.145
| Std Dev. |  | 0.23267 | 0.087958 | 0.20096 | 0.31709 | 0.19823 | 0.29383 | 0.17889 | 0.27386 | 0.1169 | 0.23424 | 0 | 0.45608 | 0.075011 | 0.067508 | 0.21679 | 0.35076 | 0.13045
| Rel.Std.Dev. |  | 2.68% | 0.9549% | 2.124% | 3.426% | 2.268% | 3.354% | 1.923% | 3.06% | 1.341% | 2.842% | 0% | 7.214% | 0.9392% | 0.7831% | 2.797% | 3.997% | 1.429%
| PDM3 |  | -1.035% | 5.006% | 7.875% | 5.519% | -0.3515% | -0.1235% | 6.013% | 2.023% | -0.6364% | -6.032% | -8.806% | -27.93% | -8.958% | -1.729% | -11.66% | 0.02852% | 4.056%
| Z-Score (Absolute) |  | 0.11 | 0.74 | 1.15 | 0.81 | 0.02 | 0.02 | 0.88 | 0.32 | 0.06 | 0.82 | NA | 3.92 | 1.23 | 0.21 | 1.61 | 0.04 | 0.61
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 702 (Certified Value 358 ppm) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 372 | 366 | 358 | 360 | 368 | 345 | 358 | 321.6 | 340 | 371.1071 | 347.336 | 379.8 | 351 | 365.2 | 303.81 | 316.0946 | 381.9 | 348.3658 | 358.4 | 359.9995 | 339 | 357.2 |  | 1
| 1 | 2 | 363 | 359 | 355 | 370 | 352 | 338 | 362 | 322.5 | 341 | 369.8915 | 338.4192 | 398.2 | 351 | 377.9 | 299.48 | 323.2621 | 383.8 | 353.4947 | 356.1 | 361.7581 | 347 | 360.3 |  | 19
| 1 | 3 | 385 | 354 | 348 | 361 | 349 | 340 | 352 | 321.6 | 348 | 369.2228 | 343.9556 | 391.9 | 351 | 405.7 | 308.41 | 308.709 | 381.6 | 353.1294 | 355.4 | 360.0756 | 355 | 369.6 |  | 16
| 1 | 4 | 369 | 357 | 359 | 360 | 354 | 343 | 353 | 322.2 | 342 | 370.4626 | 343.1383 | 383.2 | 354 | 379.4 | 306.7 | 297.4067 | 383.9 | 352.433 | 354.7 | 364.6417 | 355 | 359.8 |  | 357.6129
| 1 | 5 | 352 | 355 | 358 | 371 | 357 | 336 | 351 | 317.9 | 343 | 369.3439 | 334.03 | 380.8 | 354 | 369.1 | 309.9 | 291.186 | 381 | 344.8773 | 360.6 | 358.1312 | 354 | 362.4 |  | 102
| 1 | 6 | 369 | 353 | 350 | 375 | 364 | 333 | 356 | 315.5 | 344 | 369.1152 | 348.9 | 385.4 | 353 | 358.2 | 310.32 | 307.6317 | 380.6 | 346.2788 | 362.3 | 360.9456 | 354 | 362.8
| Mean |  | 368.3333 | 357.3333 | 354.6667 | 366.1667 | 357.3333 | 339.1667 | 355.3333 | 320.2167 | 343 | 369.8572 | 342.6298 | 386.55 | 352.3333 | 375.9167 | 306.4367 | 307.3817 | 382.1333 | 349.7632 | 357.9167 | 360.9253 | 350.6667 | 362.0167
| Median |  | 369 | 356 | 356.5 | 365.5 | 355.5 | 339 | 354.5 | 321.6 | 342.5 | 369.6177 | 343.5469 | 384.3 | 352 | 373.5 | 307.555 | 308.1704 | 381.75 | 350.3994 | 357.25 | 360.5106 | 354 | 361.35
| Std Dev. |  | 10.8382 | 4.761 | 4.6332 | 6.6156 | 7.3121 | 4.446 | 4.1793 | 2.8492 | 2.8284 | 0.79329 | 5.5707 | 7.1486 | 1.5055 | 16.6031 | 4.1553 | 11.762 | 1.4052 | 3.751 | 3.0538 | 2.1859 | 6.4704 | 4.2268
| Rel.Std.Dev. |  | 2.942% | 1.332% | 1.306% | 1.807% | 2.046% | 1.311% | 1.176% | 0.8898% | 0.8246% | 0.2145% | 1.626% | 1.849% | 0.4273% | 4.417% | 1.356% | 3.827% | 0.3677% | 1.072% | 0.8532% | 0.6056% | 1.845% | 1.168%
| PDM3 |  | 2.998% | -0.07819% | -0.8239% | 2.392% | -0.07819% | -5.158% | -0.6375% | -10.46% | -4.086% | 3.424% | -4.19% | 8.092% | -1.476% | 5.118% | -14.31% | -14.05% | 6.857% | -2.195% | 0.08493% | 0.9262% | -1.942% | 1.231%
| Z-Score (Absolute) |  | 0.74 | 0.06 | 0.1 | 0.61 | 0.06 | 1.06 | 0.06 | 2.23 | 0.82 | 0.84 | 0.85 | 1.87 | 0.25 | 1.21 | 3.08 | 3.02 | 1.59 | 0.41 | 0.1 | 0.28 | 0.35 | 0.35
| Analytical results for Dy in OREAS 702 (Certified Value 3.18 ppm) | [24 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 3.29 | 3.3 | 2.91 | 3.4864 | 3.13 | 3.31 | 2.7884 | 2.762 | [15 empty] | 1
| 1 | 2 | 3.3 | 3.4 | 3.11 | 3.5298 | 3.1 | 3.21 | 3.0242 | 2.562 | [15 empty] | #N/A
| 1 | 3 | 3.23 | 3.3 | 3.04 | 3.5222 | 3.09 | 3.22 | 3.2583 | 2.564 | [15 empty] | 16
| 1 | 4 | 3.39 | 3.1 | 2.99 | 3.5552 | 3.02 | 3.24 | 2.898 | 2.521 | [15 empty] | 3.1816
| 1 | 5 | 3.3 | 3.2 | 2.92 | 3.5065 | 3.04 | 3.28 | 2.9823 | 2.304 | [15 empty] | 103
| 1 | 6 | 3.04 | 3 | 2.86 | 3.4557 | 3.13 | 3.24 | 2.9304 | 2.559
| Mean |  | 3.2583 | 3.2167 | 2.9717 | 3.5093 | 3.085 | 3.25 | 2.9803 | 2.5453
| Median |  | 3.295 | 3.25 | 2.955 | 3.5144 | 3.095 | 3.24 | 2.9563 | 2.5605
| Std Dev. |  | 0.11856 | 0.1472 | 0.092826 | 0.034931 | 0.045935 | 0.037947 | 0.1583 | 0.14601
| Rel.Std.Dev. |  | 3.639% | 4.576% | 3.124% | 0.9954% | 1.489% | 1.168% | 5.311% | 5.736%
| PDM3 |  | 2.412% | 1.102% | -6.599% | 10.3% | -3.036% | 2.15% | -6.328% | -20%
| Z-Score (Absolute) |  | 0.52 | 0.32 | 0.87 | 1.74 | 0.32 | 0.48 | 0.83 | 2.94
| Analytical results for Er in OREAS 702 (Certified Value 1.26 ppm) | [24 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 1.27 | 1.3 | 1.23 | 1.4753 | 1.31 | 1.3 | 1.2343 | 1.219 | [15 empty] | 1
| 1 | 2 | 1.38 | 1.3 | 1.3 | 1.4171 | 1.27 | 1.31 | 1.2501 | 1.171 | [15 empty] | #N/A
| 1 | 3 | 1.26 | 1.2 | 1.27 | 1.4112 | 1.21 | 1.29 | 1.2273 | 1.119 | [15 empty] | 16
| 1 | 4 | 1.34 | 1.3 | 1.28 | 1.4389 | 1.33 | 1.35 | 1.2204 | 1.183 | [15 empty] | 1.2626
| 1 | 5 | 1.33 | 1.3 | 1.24 | 1.5457 | 1.29 | 1.31 | 1.2402 | 1.024 | [15 empty] | 104
| 1 | 6 | 1.3 | 1.2 | 1.2 | 1.4764 | 1.29 | 1.27 | 1.2355 | 1.218
| Mean |  | 1.3133 | 1.2667 | 1.2533 | 1.4607 | 1.2833 | 1.305 | 1.2346 | 1.1557
| Median |  | 1.315 | 1.3 | 1.255 | 1.4571 | 1.29 | 1.305 | 1.2349 | 1.177
| Std Dev. |  | 0.045461 | 0.05164 | 0.036697 | 0.050021 | 0.041312 | 0.026646 | 0.010267 | 0.074242
| Rel.Std.Dev. |  | 3.461% | 4.077% | 2.928% | 3.424% | 3.219% | 2.042% | 0.8316% | 6.424%
| PDM3 |  | 4.017% | 0.321% | -0.735% | 15.69% | 1.641% | 3.357% | -2.217% | -8.47%
| Z-Score (Absolute) |  | 0.76 | 0.16 | 0.43 | 3.67 | 0.16 | 0.59 | 0.8 | 2.36
| Analytical results for Eu in OREAS 702 (Certified Value 1.06 ppm) | [24 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 1.13 | 1.1 | 1.03 | 1.2541 | 1.15 | 1.01 | 0.92305 | 1.439 | [15 empty] | 1
| 1 | 2 | 0.98 | 1.2 | 1.1 | 1.3107 | 1.03 | 1.02 | 0.96027 | 1.36 | [15 empty] | #N/A
| 1 | 3 | 1.1 | 1.1 | 1.08 | 1.2243 | 1.07 | 1.02 | 0.9983 | 1.308 | [15 empty] | 16
| 1 | 4 | 1.04 | 1.1 | 1.06 | 1.2666 | 1.15 | 1.03 | 0.9208 | 1.31 | [15 empty] | 1.055
| 1 | 5 | 1.07 | 1.1 | 1.04 | 1.315 | 1.15 | 1.05 | 0.96054 | 1.204 | [15 empty] | 105
| 1 | 6 | 1.02 | 1.1 | 1.02 | 1.315 | 1.17 | 1.02 | 0.97828 | 1.342
| Mean |  | 1.0567 | 1.1167 | 1.055 | 1.2809 | 1.12 | 1.025 | 0.95687 | 1.3272
| Median |  | 1.055 | 1.1 | 1.05 | 1.2886 | 1.15 | 1.02 | 0.96041 | 1.326
| Std Dev. |  | 0.05465 | 0.040825 | 0.030822 | 0.038328 | 0.056214 | 0.013784 | 0.030479 | 0.076989
| Rel.Std.Dev. |  | 5.172% | 3.656% | 2.922% | 2.992% | 5.019% | 1.345% | 3.185% | 5.801%
| PDM3 |  | 0.1547% | 5.842% | -0.003264% | 21.41% | 6.158% | -2.847% | -9.304% | 25.79%
| Z-Score (Absolute) |  | 0.43 | 0.43 | 0.45 | 2.76 | 0.47 | 0.88 | 1.84 | 3.41
| Analytical results for Fe in OREAS 702 (Certified Value 2.34 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.4 | 2.36 | 2.31 | 2.3 | 2.43 | 2.34 | 2.34 | 2.2572 | 2.24 | 2.3829 | 2.2946 | 2.38 | 2.32 | 2.52 | 2.3419 | 2.5806 | 2.57 | 2.3503 | 2.36 | 2.2357 | 2.4 | 2.36 |  | 1
| 1 | 2 | 2.36 | 2.33 | 2.29 | 2.36 | 2.31 | 2.35 | 2.34 | 2.2671 | 2.24 | 2.3589 | 2.2985 | 2.42 | 2.28 | 2.59 | 2.3341 | 2.585 | 2.52 | 2.365 | 2.38 | 2.0693 | 2.43 | 2.38 |  | #N/A
| 1 | 3 | 2.47 | 2.34 | 2.28 | 2.3 | 2.3 | 2.39 | 2.28 | 2.2589 | 2.29 | 2.34 | 2.2839 | 2.42 | 2.24 | 2.8 | 2.3754 | 2.5916 | 2.57 | 2.3945 | 2.37 | 2.2846 | 2.4 | 2.44 |  | 16
| 1 | 4 | 2.35 | 2.32 | 2.34 | 2.29 | 2.31 | 2.42 | 2.33 | 2.2472 | 2.26 | 2.3581 | 2.1798 | 2.37 | 2.26 | 2.62 | 2.3576 | 2.5399 | 2.52 | 2.3564 | 2.37 | 2.3425 | 2.43 | 2.39 |  | 2.3355
| 1 | 5 | 2.24 | 2.32 | 2.33 | 2.36 | 2.33 | 2.4 | 2.31 | 2.232 | 2.21 | 2.3338 | 2.3506 | 2.37 | 2.28 | 2.6 | 2.4087 | 2.6686 | 2.56 | 2.3713 | 2.38 | 2.245 | 2.43 | 2.41 |  | 106
| 1 | 6 | 2.39 | 2.31 | 2.28 | 2.39 | 2.39 | 2.37 | 2.33 | 2.235 | 2.21 | 2.3477 | 2.1711 | 2.4 | 2.25 | 2.47 | 2.4022 | 2.7467 | 2.53 | 2.3658 | 2.39 | 2.3724 | 2.41 | 2.39
| Mean |  | 2.3683 | 2.33 | 2.305 | 2.3333 | 2.345 | 2.3783 | 2.3217 | 2.2496 | 2.2417 | 2.3536 | 2.2631 | 2.3933 | 2.2717 | 2.6 | 2.37 | 2.6187 | 2.545 | 2.3672 | 2.375 | 2.2582 | 2.4167 | 2.395
| Median |  | 2.375 | 2.325 | 2.3 | 2.33 | 2.32 | 2.38 | 2.33 | 2.2522 | 2.24 | 2.3529 | 2.2893 | 2.39 | 2.27 | 2.595 | 2.3665 | 2.5883 | 2.545 | 2.3654 | 2.375 | 2.2648 | 2.42 | 2.39
| Std Dev. |  | 0.075741 | 0.017889 | 0.025884 | 0.041793 | 0.05282 | 0.030605 | 0.023166 | 0.013994 | 0.030605 | 0.017432 | 0.071771 | 0.023381 | 0.028577 | 0.11296 | 0.030963 | 0.075373 | 0.02429 | 0.015326 | 0.010488 | 0.10697 | 0.015055 | 0.027386
| Rel.Std.Dev. |  | 3.198% | 0.7677% | 1.123% | 1.791% | 2.252% | 1.287% | 0.9978% | 0.6221% | 1.365% | 0.7407% | 3.171% | 0.9769% | 1.258% | 4.345% | 1.306% | 2.878% | 0.9544% | 0.6474% | 0.4416% | 4.737% | 0.623% | 1.143%
| PDM3 |  | 1.406% | -0.2354% | -1.306% | -0.09264% | 0.4069% | 1.834% | -0.5922% | -3.679% | -4.018% | 0.7737% | -3.1% | 2.476% | -2.733% | 11.33% | 1.476% | 12.13% | 8.97% | 1.358% | 1.691% | -3.308% | 3.475% | 2.548%
| Z-Score (Absolute) |  | 0.15 | 0.56 | 1.02 | 0.5 | 0.28 | 0.33 | 0.71 | 2.04 | 2.18 | 0.13 | 1.79 | 0.61 | 1.63 | 4.41 | 0.18 | 4.75 | 3.4 | 0.13 | 0.27 | 1.88 | 1.03 | 0.64
| Analytical results for Ga in OREAS 702 (Certified Value 14 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 24 | 25 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 13.5 | 13.05 | 14.35 | 12.95 | 14.55 | 14.6 | 14.9 | 11.31 | 13.7 | 15.5742 | 14.91 | 14.8 | 13.501 | 13.159 | 14.37 | 13.8611 | 14 | 13.4277 | 16 | 13.53 |  |  |  | 1
| 1 | 2 | 13.3 | 13.25 | 14.3 | 13.1 | 13.25 | 13.8 | 15.6 | 11.31 | 13.5 | 15.4854 | 14.75 | 14.8 | 12.692 | 13.2652 | 14.72 | 14.1454 | 14 | 12.6709 | 17 | 13.65 |  |  |  | #N/A
| 1 | 3 | 14.25 | 13.2 | 13.95 | 13.8 | 13.5 | 14.6 | 14.3 | 11.07 | 13.9 | 15.5074 | 14.5 | 15.5 | 11.334 | 12.8592 | 15.35 | 14.2318 | 14 | 13.5468 | 16 | 13.86 |  |  |  | 16
| 1 | 4 | 13.85 | 13.25 | 14.35 | 13.1 | 13.8 | 14.2 | 14.1 | 11.19 | 13.5 | 15.4645 | 14.44 | 14.2 | 12.979 | 12.2971 | 14.41 | 14.5622 | 14 | 14.5686 | 16 | 13.59 |  |  |  | 13.9892
| 1 | 5 | 13.6 | 13.2 | 14.35 | 13.55 | 13.55 | 14.6 | 14.3 | 11.44 | 13.9 | 15.4866 | 14.55 | 14.9 | 13.484 | 12.3002 | 15.44 | 14.0361 | 14 | 13.5478 | 16 | 13.57 |  |  |  | 107
| 1 | 6 | 13.55 | 13.1 | 13.8 | 13.5 | 14.15 | 14.2 | 14.2 | 11 | 13.8 | 15.5222 | 14.41 | 14.3 | 12.81 | 12.8658 | 14.88 | 13.8292 | 14 | 14.3455 | 16 | 13.65
| Mean |  | 13.675 | 13.175 | 14.1833 | 13.3333 | 13.8 | 14.3333 | 14.5667 | 11.22 | 13.7167 | 15.5067 | 14.5933 | 14.75 | 12.8 | 12.7911 | 14.8617 | 14.111 | 14 | 13.6845 | 16.1667 | 13.6417
| Median |  | 13.575 | 13.2 | 14.325 | 13.3 | 13.675 | 14.4 | 14.3 | 11.25 | 13.75 | 15.497 | 14.525 | 14.8 | 12.8945 | 12.8625 | 14.8 | 14.0908 | 14 | 13.5473 | 16 | 13.62
| Std Dev. |  | 0.33279 | 0.082158 | 0.24427 | 0.33116 | 0.47749 | 0.3266 | 0.5785 | 0.16517 | 0.18348 | 0.038553 | 0.19623 | 0.46797 | 0.794 | 0.41359 | 0.45587 | 0.27087 | 0 | 0.6855 | 0.40825 | 0.11669
| Rel.Std.Dev. |  | 2.434% | 0.6236% | 1.722% | 2.484% | 3.46% | 2.279% | 3.971% | 1.472% | 1.338% | 0.2486% | 1.345% | 3.173% | 6.203% | 3.233% | 3.067% | 1.92% | 0% | 5.009% | 2.525% | 0.8554%
| PDM3 |  | -2.246% | -5.82% | 1.388% | -4.688% | -1.353% | 2.46% | 4.128% | -19.8% | -1.948% | 10.85% | 4.319% | 5.438% | -8.501% | -8.565% | 6.237% | 0.8704% | 0.07716% | -2.178% | 15.57% | -2.484%
| Z-Score (Absolute) |  | 0.1 | 0.68 | 0.49 | 0.49 | 0.05 | 0.67 | 0.94 | 2.96 | 0.05 | 2.04 | 0.97 | 1.15 | 1.12 | 1.13 | 1.28 | 0.41 | NA | 0.09 | NA | 0.14
Results from laboratories 21 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 702 (Certified Value 4.52 ppm) | [24 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 4.95 | 5 | 3.72 | 5.5866 | 4.85 | 4.7 | 4.1109 | 4.697 | [15 empty] | 1
| 1 | 2 | 4.49 | 5.3 | 3.91 | 5.6402 | 4.49 | 4.74 | 4.2963 | 4.494 | [15 empty] | #N/A
| 1 | 3 | 5.24 | 4.9 | 3.98 | 5.6008 | 4.65 | 4.67 | 4.343 | 4.357 | [15 empty] | 16
| 1 | 4 | 4.76 | 4.8 | 3.97 | 5.6509 | 4.48 | 4.69 | 4.2496 | 4.342 | [15 empty] | 4.5163
| 1 | 5 | 4.56 | 4.8 | 3.82 | 5.5808 | 4.73 | 4.73 | 4.4467 | 4.31 | [15 empty] | 108
| 1 | 6 | 4.66 | 4.6 | 3.67 | 5.6826 | 4.85 | 4.61 | 4.2801 | 4.439
| Mean |  | 4.7767 | 4.9 | 3.845 | 5.6236 | 4.675 | 4.69 | 4.2878 | 4.4398
| Median |  | 4.71 | 4.85 | 3.865 | 5.6205 | 4.69 | 4.695 | 4.2882 | 4.398
| Std Dev. |  | 0.27847 | 0.23664 | 0.13035 | 0.040557 | 0.16562 | 0.046904 | 0.11065 | 0.14308
| Rel.Std.Dev. |  | 5.83% | 4.829% | 3.39% | 0.7212% | 3.543% | 1% | 2.58% | 3.223%
| PDM3 |  | 5.765% | 8.495% | -14.86% | 24.52% | 3.513% | 3.846% | -5.061% | -1.694%
| Z-Score (Absolute) |  | 0.28 | 0.64 | 2.45 | 2.76 | 0.02 | 0.02 | 1.16 | 0.71
| Analytical results for Ge in OREAS 702 (Indicative Value 0.098 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 14 | 15 | 17 | 18 | 19 | 21 | 25 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.05 | < 0.05 | 0.05 | < 0.05 | 0.13 | 0.17 | < 0.5 | 0.03 | 0.9 | 0.09 | < 0.1 | 0.27726 | 0.07 | < 0.2 | 3 | < 0.05 | [7 empty] | 1
| 1 | 2 | 0.05 | < 0.05 | 0.05 | < 0.05 | 0.11 | 0.13 | < 0.5 | 0.02 | 0.9 | 0.11 | < 0.1 | 0.23986 | 0.07 | < 0.2 | 3 | < 0.05 | [7 empty] | 14
| 1 | 3 | 0.05 | < 0.05 | < 0.05 | 0.05 | 0.08 | 0.28 | < 0.5 | 0.02 | 0.9 | 0.11 | < 0.1 | 0.21346 | 0.07 | < 0.2 | 2.9 | < 0.05 | [7 empty] | 16
| 1 | 4 | 0.06 | < 0.05 | < 0.05 | 0.05 | 0.1 | 0.27 | < 0.5 | 0.02 | 1 | 0.1 | < 0.1 | 0.18842 | 0.08 | < 0.2 | 2.8 | < 0.05 | [7 empty] | 0.098024
| 1 | 5 | < 0.05 | < 0.05 | < 0.05 | 0.05 | 0.09 | 0.18 | < 0.5 | 0.02 | 0.9 | 0.11 | < 0.1 | 0.26242 | 0.08 | < 0.2 | 3 | < 0.05 | [7 empty] | 22
| 1 | 6 | 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.06 | 0.14 | 0.7 | 0.03 | 0.9 | 0.1 | < 0.1 | 0.24989 | 0.08 | < 0.2 | 3 | < 0.05
| Mean |  | 0.052 |  | 0.05 | 0.05 | 0.095 | 0.195 | 0.7 | 0.023333 | 0.91667 | 0.10333 |  | 0.23855 | 0.075 |  | 2.95
| Median |  | 0.05 |  | 0.05 | 0.05 | 0.095 | 0.175 | 0.7 | 0.02 | 0.9 | 0.105 |  | 0.24487 | 0.075 |  | 3
| Std Dev. |  | 0.0044721 |  | 0 | 8.4984e-18 | 0.02429 | 0.06473 |  | 0.005164 | 0.040825 | 0.008165 |  | 0.032682 | 0.0054772 |  | 0.083666
| Rel.Std.Dev. |  | 8.6% |  | 0% | 1.7e-14% | 25.57% | 33.19% |  | 22.13% | 4.454% | 7.902% |  | 13.7% | 7.303% |  | 2.836%
| PDM3 |  | -46.95% |  | -48.99% | -48.99% | -3.085% | 98.93% | 614.1% | -76.2% | 835.1% | 5.416% |  | 143.4% | -23.49% |  | 2909%
| Z-Score (Absolute) |  | 0.45 | 0.72 | 0.62 | 0.57 | 0.12 | 1.33 | 2.9 | 0.74 | 10.04 | 0.22 | 0.42 | 1.85 | 0.12 | 0.18 | 34.6 | 0.72
| Analytical results for Hf in OREAS 702 (Certified Value 1.26 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 25 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.3 | 1.2 | 1.4 | 1.4 | 1.3 | 1.3 | 1.3 | 1.09 | 1.1 | 1.3 | 1.38 | 1.48 | 1.3751 | 1.3 | 13.3881 | 1 | 1.3 | [6 empty] | 1
| 1 | 2 | 1.2 | 1.2 | 1.6 | 1.3 | 1.3 | 1.2 | 1.1 | 1.17 | 1.1 | 1.1 | 1.32 | 1.52 | 1.449 | 1.2 | 13.9015 | 1 | 1.3 | [6 empty] | #N/A
| 1 | 3 | 1.4 | 1.2 | 1.4 | 1.4 | 1.3 | 1.3 | 0.9 | 1.16 | 1.1 | 1.3 | 1.42 | 1.45 | 1.5788 | 1.4 | 13.768 | 0.9 | 1.3 | [6 empty] | 16
| 1 | 4 | 1.2 | 1.3 | 1.4 | 1.4 | 1.2 | 1.1 | 1.1 | 1.14 | 1.1 | 1.2 | 1.33 | 1.52 | 1.4543 | 1.2 | 13.4924 | 1 | 1.4 | [6 empty] | 1.2612
| 1 | 5 | 1.2 | 1.2 | 1.3 | 1.5 | 1.3 | 1.3 | 0.9 | 1.08 | 1 | 1.1 | 1.36 | 1.48 | 1.4624 | 1.2 | 13.7948 | 1 | 1.4 | [6 empty] | 109
| 1 | 6 | 1.3 | 1.3 | 1.3 | 1.2 | 1.2 | 1.3 | 1.4 | 1.04 | 1.1 | 1.3 | 1.24 | 1.27 | 1.5101 | 1.4 | 13.7171 | 0.9 | 1.4
| Mean |  | 1.2667 | 1.2333 | 1.4 | 1.3667 | 1.2667 | 1.25 | 1.1167 | 1.1133 | 1.0833 | 1.2167 | 1.3417 | 1.4533 | 1.4716 | 1.2833 | 13.677 | 0.96667 | 1.35
| Median |  | 1.25 | 1.2 | 1.4 | 1.4 | 1.3 | 1.3 | 1.1 | 1.115 | 1.1 | 1.25 | 1.345 | 1.48 | 1.4583 | 1.25 | 13.7425 | 1 | 1.35
| Std Dev. |  | 0.08165 | 0.05164 | 0.10954 | 0.10328 | 0.05164 | 0.083666 | 0.20412 | 0.051251 | 0.040825 | 0.098319 | 0.061455 | 0.093737 | 0.068089 | 0.098319 | 0.19583 | 0.05164 | 0.054772
| Rel.Std.Dev. |  | 6.446% | 4.187% | 7.825% | 7.557% | 4.077% | 6.693% | 18.28% | 4.603% | 3.768% | 8.081% | 4.58% | 6.45% | 4.627% | 7.661% | 1.432% | 5.342% | 4.057%
| PDM3 |  | 0.4297% | -2.213% | 11% | 8.358% | 0.4297% | -0.8918% | -11.46% | -11.73% | -14.11% | -3.535% | 6.376% | 15.23% | 16.68% | 1.751% | 984.4% | -23.36% | 7.037%
| Z-Score (Absolute) |  | 0 | 0.22 | 0.9 | 0.67 | 0 | 0.11 | 1.01 | 1.03 | 1.24 | 0.34 | 0.51 | 1.26 | 1.38 | 0.11 | 83.68 | 2.02 | 0.56
| Analytical results for Hg in OREAS 702 (Indicative Value 1.02 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [21 empty] | 1
| No | No | 09 | 19 | [21 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [21 empty] | 2
2
| 1 | 1 | 0.96 | < 1 | [21 empty] | 1
| 1 | 2 | 1.08 | < 1 | [21 empty] | 17
| 1 | 3 | 1.08 | < 1 | [21 empty] | 16
| 1 | 4 | 1 | < 1 | [21 empty] | 1.015
| 1 | 5 | 1 | < 1 | [21 empty] | 23
| 1 | 6 | 0.97 | < 1
| Mean |  | 1.015
| Median |  | 1
| Std Dev. |  | 0.05282
| Rel.Std.Dev. |  | 5.204%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 702 (Certified Value 0.5 ppm) | [24 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.59 | 0.6 | 0.47 | 0.57647 | 0.51 | 0.53 | 0.48243 | 0.434 | [15 empty] | 1
| 1 | 2 | 0.57 | 0.6 | 0.48 | 0.56784 | 0.48 | 0.53 | 0.50132 | 0.401 | [15 empty] | #N/A
| 1 | 3 | 0.57 | 0.5 | 0.48 | 0.57006 | 0.5 | 0.51 | 0.49466 | 0.407 | [15 empty] | 16
| 1 | 4 | 0.57 | 0.6 | 0.47 | 0.58428 | 0.5 | 0.53 | 0.49159 | 0.407 | [15 empty] | 0.50436
| 1 | 5 | 0.53 | 0.5 | 0.46 | 0.55383 | 0.47 | 0.53 | 0.50561 | 0.393 | [15 empty] | 110
| 1 | 6 | 0.53 | 0.5 | 0.45 | 0.59106 | 0.52 | 0.54 | 0.48903 | 0.413
| Mean |  | 0.56 | 0.55 | 0.46833 | 0.57392 | 0.49667 | 0.52833 | 0.49411 | 0.40917
| Median |  | 0.57 | 0.55 | 0.47 | 0.57326 | 0.5 | 0.53 | 0.49313 | 0.407
| Std Dev. |  | 0.024495 | 0.054772 | 0.01169 | 0.013139 | 0.018619 | 0.0098319 | 0.0084038 | 0.01392
| Rel.Std.Dev. |  | 4.374% | 9.959% | 2.496% | 2.289% | 3.749% | 1.861% | 1.701% | 3.402%
| PDM3 |  | 11.03% | 9.049% | -7.143% | 13.79% | -1.526% | 4.753% | -2.033% | -18.87%
| Z-Score (Absolute) |  | 1.35 | NA | 0.6 | 1.65 | 0 | 0.67 | 0.05 | 1.86
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 702 (Certified Value 0.097 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 14 | 15 | 17 | 18 | 21 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.095 | 0.09 | 0.105 | 0.087 | 0.101 | 0.11 | 0.1 | 0.108 | 0.08 | 0.1 | 0.1 | < 10 | 0.072 | 0.11 | < 0.1 | 0.094 | [7 empty] | 1
| 1 | 2 | 0.092 | 0.09 | 0.098 | 0.088 | 0.09 | 0.1 | < 0.1 | 0.111 | 0.07 | 0.103 | 0.1 | < 10 | 0.072 | 0.1 | 0.1 | 0.094 | [7 empty] | 11
| 1 | 3 | 0.095 | 0.096 | 0.101 | 0.095 | 0.087 | 0.14 | 0.1 | 0.115 | 0.09 | 0.101 | 0.1 | < 10 | 0.069 | 0.1 | 0.1 | 0.097 | [7 empty] | 16
| 1 | 4 | 0.099 | 0.097 | 0.104 | 0.096 | 0.099 | 0.09 | 0.1 | 0.106 | 0.09 | 0.103 | 0.1 | < 10 | 0.066 | 0.09 | < 0.1 | 0.094 | [7 empty] | 0.097278
| 1 | 5 | 0.091 | 0.096 | 0.102 | 0.094 | 0.095 | 0.12 | < 0.1 | 0.103 | 0.09 | 0.101 | 0.1 | < 10 | 0.066 | 0.11 | < 0.1 | 0.095 | [7 empty] | 111
| 1 | 6 | 0.093 | 0.09 | 0.098 | 0.09 | 0.095 | 0.09 | 0.1 | 0.101 | 0.09 | 0.102 | 0.09 | < 10 | 0.064 | 0.11 | < 0.1 | 0.095
| Mean |  | 0.094167 | 0.093167 | 0.10133 | 0.091667 | 0.0945 | 0.10833 | 0.1 | 0.10733 | 0.085 | 0.10167 | 0.098333 |  | 0.068167 | 0.10333 | 0.1 | 0.094833
| Median |  | 0.094 | 0.093 | 0.1015 | 0.092 | 0.095 | 0.105 | 0.1 | 0.107 | 0.09 | 0.1015 | 0.1 |  | 0.0675 | 0.105 | 0.1 | 0.0945
| Std Dev. |  | 0.0028577 | 0.0034881 | 0.0029439 | 0.0038297 | 0.005282 | 0.019408 | 0 | 0.005164 | 0.0083666 | 0.0012111 | 0.0040825 |  | 0.0033714 | 0.008165 | 0 | 0.001169
| Rel.Std.Dev. |  | 3.035% | 3.744% | 2.905% | 4.178% | 5.589% | 17.91% | 0% | 4.811% | 9.843% | 1.191% | 4.152% |  | 4.946% | 7.902% | 0% | 1.233%
| PDM3 |  | -3.198% | -4.226% | 4.169% | -5.768% | -2.856% | 11.36% | 2.798% | 10.34% | -12.62% | 4.512% | 1.085% |  | -29.93% | 6.225% | 2.798% | -2.513%
| Z-Score (Absolute) |  | 0.33 | 0.46 | 0.64 | 0.66 | 0.28 | 1.58 | NA | 1.45 | 1.56 | 0.69 | 0.24 | 661.28 | 3.83 | 0.91 | NA | 0.24
Results from laboratories 7 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 702 (Certified Value 1.91 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.91 | 1.91 | 1.86 | 1.85 | 1.96 | 1.87 | 1.9593 | 1.88 | 2.0587 | 1.8738 | 1.91 | 1.92 | 2.03 | 1.8987 | 1.854 | 2 | 1.9126 | 1.9 | 1.9487 | 1.94 | 1.94 |  |  | 1
| 1 | 2 | 1.88 | 1.88 | 1.86 | 1.89 | 1.87 | 1.87 | 1.9864 | 1.87 | 2.0322 | 1.8783 | 1.95 | 1.88 | 2.12 | 1.8853 | 1.869 | 1.97 | 1.9289 | 1.91 | 1.8975 | 1.94 | 1.96 |  |  | #N/A
| 1 | 3 | 1.98 | 1.86 | 1.83 | 1.84 | 1.85 | 1.89 | 2.0002 | 1.9 | 2.0594 | 1.8332 | 1.95 | 1.86 | 2.28 | 1.9226 | 1.848 | 2 | 1.9494 | 1.9 | 1.9948 | 1.96 | 2 |  |  | 16
| 1 | 4 | 1.89 | 1.87 | 1.87 | 1.85 | 1.88 | 1.92 | 1.9961 | 1.88 | 2.0605 | 1.7299 | 1.95 | 1.88 | 2.12 | 1.9382 | 1.804 | 1.97 | 1.9446 | 1.89 | 2.0246 | 1.86 | 1.97 |  |  | 1.9104
| 1 | 5 | 1.84 | 1.87 | 1.88 | 1.89 | 1.9 | 1.9 | 1.9557 | 1.88 | 2.0805 | 1.8582 | 1.94 | 1.88 | 2.04 | 1.9434 | 1.912 | 1.99 | 1.9363 | 1.91 | 1.9557 | 1.91 | 1.98 |  |  | 112
| 1 | 6 | 1.9 | 1.84 | 1.82 | 1.92 | 1.91 | 1.9 | 1.9629 | 1.88 | 2.0735 | 1.7118 | 1.93 | 1.87 | 2.01 | 1.9533 | 1.894 | 1.98 | 1.9444 | 1.89 | 2.0591 | 1.88 | 1.97
| Mean |  | 1.9 | 1.8717 | 1.8533 | 1.8733 | 1.895 | 1.8917 | 1.9768 | 1.8817 | 2.0608 | 1.8142 | 1.9383 | 1.8817 | 2.1 | 1.9236 | 1.8635 | 1.985 | 1.936 | 1.9 | 1.9801 | 1.915 | 1.97
| Median |  | 1.895 | 1.87 | 1.86 | 1.87 | 1.89 | 1.895 | 1.9747 | 1.88 | 2.0599 | 1.8457 | 1.945 | 1.88 | 2.08 | 1.9304 | 1.8615 | 1.985 | 1.9404 | 1.9 | 1.9753 | 1.925 | 1.97
| Std Dev. |  | 0.046043 | 0.023166 | 0.023381 | 0.031411 | 0.038341 | 0.019408 | 0.019783 | 0.0098319 | 0.016559 | 0.074204 | 0.016021 | 0.020412 | 0.0998 | 0.026739 | 0.037904 | 0.013784 | 0.013572 | 0.0089443 | 0.058037 | 0.038859 | 0.02
| Rel.Std.Dev. |  | 2.423% | 1.238% | 1.262% | 1.677% | 2.023% | 1.026% | 1.001% | 0.5225% | 0.8035% | 4.09% | 0.8265% | 1.085% | 4.752% | 1.39% | 2.034% | 0.6944% | 0.701% | 0.4708% | 2.931% | 2.029% | 1.015%
| PDM3 |  | -0.5441% | -2.027% | -2.987% | -1.94% | -0.8058% | -0.9803% | 3.474% | -1.504% | 7.873% | -5.035% | 1.462% | -1.504% | 9.925% | 0.6907% | -2.455% | 3.905% | 1.343% | -0.5441% | 3.647% | 0.2411% | 3.12%
| Z-Score (Absolute) |  | 0 | 0.53 | 0.87 | 0.5 | 0.09 | 0.16 | 1.44 | 0.34 | 3.01 | 1.61 | 0.72 | 0.34 | 3.74 | 0.44 | 0.68 | 1.59 | 0.67 | 0 | 1.5 | 0.28 | 1.31
| Analytical results for La in OREAS 702 (Certified Value 28.8 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 30.8 | 28.6 | 30.9 | 26.9 | 28.6 | 28.2 | 32.7 | 17.8 | 31.4813 | 25.8347 | 29.1 | 29.6 | 32.3 | 25.19 | 29 | 27.5607 | 35 | 26.122 | 27.4 |  |  |  |  | 1
| 1 | 2 | 28.7 | 28.7 | 30.2 | 27.6 | 27.4 | 27.5 | 33 | 17.8 | 31.5021 | 26.1126 | 31.7 | 28.6 | 31.6 | 25.05 | 29.3 | 28.362 | 34 | 27.747 | 29.1 |  |  |  |  | #N/A
| 1 | 3 | 31.4 | 29.4 | 30.6 | 29.6 | 27.3 | 28 | 32 | 18.5 | 31.4773 | 26.3661 | 30.6 | 28.6 | 32.4 | 24.65 | 29 | 27.9652 | 34 | 26.658 | 28.8 |  |  |  |  | 16
| 1 | 4 | 31.2 | 28.8 | 30.6 | 27.4 | 28.3 | 27.6 | 30.9 | 19 | 31.4231 | 25.54 | 29.7 | 29.3 | 30.5 | 24.53 | 29 | 28.1368 | 34 | 25.731 | 28.4 |  |  |  |  | 28.7942
| 1 | 5 | 29.3 | 29.3 | 31.2 | 27.7 | 27.5 | 27.2 | 30.5 | 18.1 | 31.3065 | 25.8 | 30.3 | 28.4 | 31.4 | 24.35 | 29 | 28.1929 | 35 | 23.599 | 28.2 |  |  |  |  | 113
| 1 | 6 | 30.6 | 28.9 | 29.9 | 28.3 | 29.2 | 27.8 | 31.2 | 17.7 | 31.4034 | 26.1817 | 30.3 | 29.8 | 29.2 | 25.5 | 28.7 | 27.6411 | 35 | 26.999 | 27.8
| Mean |  | 30.3333 | 28.95 | 30.5667 | 27.9167 | 28.05 | 27.7167 | 31.7167 | 18.15 | 31.4323 | 25.9725 | 30.2833 | 29.05 | 31.2333 | 24.8783 | 29 | 27.9765 | 34.5 | 26.1427 | 28.2833
| Median |  | 30.7 | 28.85 | 30.6 | 27.65 | 27.9 | 27.7 | 31.6 | 17.95 | 31.4502 | 25.9736 | 30.3 | 28.95 | 31.5 | 24.85 | 29 | 28.051 | 34.5 | 26.39 | 28.3
| Std Dev. |  | 1.0875 | 0.32711 | 0.46762 | 0.9411 | 0.77136 | 0.36009 | 1.0108 | 0.50892 | 0.072149 | 0.30137 | 0.87731 | 0.59245 | 1.2111 | 0.43947 | 0.18974 | 0.31832 | 0.54772 | 1.4292 | 0.62743
| Rel.Std.Dev. |  | 3.585% | 1.13% | 1.53% | 3.371% | 2.75% | 1.299% | 3.187% | 2.804% | 0.2295% | 1.16% | 2.897% | 2.039% | 3.877% | 1.766% | 0.6543% | 1.138% | 1.588% | 5.467% | 2.218%
| PDM3 |  | 5.345% | 0.5409% | 6.155% | -3.048% | -2.585% | -3.742% | 10.15% | -36.97% | 9.162% | -9.8% | 5.171% | 0.8882% | 8.471% | -13.6% | 0.7146% | -2.84% | 19.82% | -9.209% | -1.774%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.79 | 0.5 | 0.44 | 0.6 | 1.35 | 5.26 | 1.21 | 1.45 | 0.65 | 0.05 | 1.11 | 1.98 | 0.02 | 0.47 | 2.71 | 1.37 | 0.32
| Analytical results for Li in OREAS 702 (Certified Value 36.9 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 39.5 | 38.9 | 36.9 | 37.2 | 39.8 | 39.4 | 38 | 35 | 35 | 38.2313 | 38.2 | 34.9 | 34 | 16.56 | 30 | 36.6 | 34.7484 | 38.2 | 37 | 34.5 |  |  |  | 1
| 1 | 2 | 38.9 | 38.3 | 36.2 | 38.5 | 38 | 37.6 | 38 | 35 | 34.9 | 38.345 | 39.5 | 36.9 | 33 | 15.62 | 30 | 35.9 | 35.343 | 36.3 | 39 | 34.9 |  |  |  | #N/A
| 1 | 3 | 40.8 | 37.9 | 36 | 37.1 | 37.8 | 38.7 | 35 | 35 | 35 | 38.3806 | 39 | 36.9 | 37 | 16.03 | 30 | 36.6 | 35.361 | 36.3 | 39 | 35.8 |  |  |  | 16
| 1 | 4 | 38.8 | 37.6 | 37.7 | 36.8 | 37.4 | 38.8 | 37 | 34 | 35.1 | 38.3174 | 38.5 | 36.4 | 35 | 16.62 | 30 | 35.9 | 35.2793 | 37 | 38 | 34.3 |  |  |  | 36.9187
| 1 | 5 | 37.4 | 38 | 36.6 | 38.3 | 38.3 | 36.3 | 36 | 35 | 34.7 | 38.2677 | 37.8 | 35.7 | 34 | 16.27 | 30 | 36.3 | 35.2508 | 36.7 | 39 | 35.9 |  |  |  | 114
| 1 | 6 | 39 | 37.9 | 36.1 | 38.6 | 39.3 | 37.5 | 37 | 35 | 35.4 | 38.2245 | 38.7 | 36.9 | 35 | 17.04 | 30 | 36.1 | 35.4698 | 36 | 40 | 35.3
| Mean |  | 39.0667 | 38.1 | 36.5833 | 37.75 | 38.4333 | 38.05 | 36.8333 | 34.8333 | 35.0167 | 38.2944 | 38.6167 | 36.2833 | 34.6667 | 16.3567 | 30 | 36.2333 | 35.242 | 36.75 | 38.6667 | 35.1167
| Median |  | 38.95 | 37.95 | 36.4 | 37.75 | 38.15 | 38.15 | 37 | 35 | 35 | 38.2925 | 38.6 | 36.65 | 34.5 | 16.415 | 30 | 36.2 | 35.3111 | 36.5 | 39 | 35.1
| Std Dev. |  | 1.1021 | 0.45166 | 0.64317 | 0.80187 | 0.92664 | 1.1292 | 1.169 | 0.40825 | 0.23166 | 0.063399 | 0.59805 | 0.82563 | 1.3663 | 0.4969 | 0 | 0.32042 | 0.25349 | 0.79183 | 1.0328 | 0.66458
| Rel.Std.Dev. |  | 2.821% | 1.185% | 1.758% | 2.124% | 2.411% | 2.968% | 3.174% | 1.172% | 0.6616% | 0.1656% | 1.549% | 2.276% | 3.941% | 3.038% | 0% | 0.8843% | 0.7193% | 2.155% | 2.671% | 1.892%
| PDM3 |  | 5.818% | 3.2% | -0.9084% | 2.252% | 4.103% | 3.064% | -0.2312% | -5.649% | -5.152% | 3.726% | 4.599% | -1.721% | -6.1% | -55.7% | -18.74% | -1.856% | -4.541% | -0.4569% | 4.735% | -4.881%
| Z-Score (Absolute) |  | 1.02 | 0.61 | 0.04 | 0.46 | 0.75 | 0.59 | 0.07 | 0.78 | 0.7 | 0.69 | 0.83 | 0.16 | 0.85 | 8.62 | 2.83 | 0.18 | 0.6 | 0.04 | 0.85 | 0.66
| Analytical results for Lu in OREAS 702 (Certified Value 0.14 ppm) | [24 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 15 | 18 | 19 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 0.13 | 0.14 | 0.17 | 0.14622 | 0.14 | 0.16 | 0.14 | 0.12657 | 0.128 | [14 empty] | 1
| 1 | 2 | 0.14 | 0.16 | 0.19 | 0.15024 | 0.14 | 0.22 | 0.14 | 0.12982 | 0.149 | [14 empty] | #N/A
| 1 | 3 | 0.12 | 0.16 | 0.19 | 0.1451 | 0.13 | 0.17 | 0.14 | 0.12899 | 0.128 | [14 empty] | 16
| 1 | 4 | 0.14 | 0.12 | 0.2 | 0.14482 | 0.14 | 0.16 | 0.15 | 0.13053 | 0.124 | [14 empty] | 0.14038
| 1 | 5 | 0.12 | 0.14 | 0.18 | 0.16277 | 0.14 | 0.16 | 0.15 | 0.12691 | 0.12 | [14 empty] | 115
| 1 | 6 | 0.15 | 0.14 | 0.18 | 0.1464 | 0.14 | 0.14 | 0.14 | 0.13166 | 0.121
| Mean |  | 0.13333 | 0.14333 | 0.185 | 0.14926 | 0.13833 | 0.16833 | 0.14333 | 0.12908 | 0.12833
| Median |  | 0.135 | 0.14 | 0.185 | 0.14631 | 0.14 | 0.16 | 0.14 | 0.1294 | 0.126
| Std Dev. |  | 0.012111 | 0.015055 | 0.010488 | 0.0068967 | 0.0040825 | 0.027142 | 0.005164 | 0.0020171 | 0.010671
| Rel.Std.Dev. |  | 9.083% | 10.5% | 5.669% | 4.621% | 2.951% | 16.12% | 3.603% | 1.563% | 8.315%
| PDM3 |  | -5.017% | 2.107% | 31.79% | 6.327% | -1.455% | 19.92% | 2.107% | -8.047% | -8.579%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.81 | 0.4 | 0.34 | 1.69 | 0 | 0.96 | 1.01
| Analytical results for Mg in OREAS 702 (Certified Value 0.587 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.57 | 0.56 | 0.57 | 0.55 | 0.6 | 0.58 | 0.6 | 0.5748 | 0.576 | 0.63715 | 0.55957 | 0.58 | 0.6 | 0.6 | 0.59616 | 0.588 | 0.59 | 0.57424 | 0.62 | 0.5943 | 0.632 | 0.55 |  | 1
| 1 | 2 | 0.56 | 0.56 | 0.56 | 0.57 | 0.57 | 0.59 | 0.6 | 0.5793 | 0.579 | 0.63696 | 0.55677 | 0.59 | 0.6 | 0.62 | 0.59196 | 0.593 | 0.58 | 0.5837 | 0.63 | 0.5863 | 0.634 | 0.56 |  | #N/A
| 1 | 3 | 0.59 | 0.55 | 0.56 | 0.55 | 0.56 | 0.59 | 0.57 | 0.5793 | 0.588 | 0.63364 | 0.55998 | 0.59 | 0.57 | 0.66 | 0.59923 | 0.588 | 0.58 | 0.58842 | 0.62 | 0.6045 | 0.634 | 0.57 |  | 16
| 1 | 4 | 0.56 | 0.55 | 0.57 | 0.55 | 0.56 | 0.6 | 0.59 | 0.5782 | 0.585 | 0.63295 | 0.5423 | 0.59 | 0.57 | 0.62 | 0.60172 | 0.573 | 0.57 | 0.58561 | 0.63 | 0.6106 | 0.605 | 0.56 |  | 0.58653
| 1 | 5 | 0.54 | 0.55 | 0.57 | 0.56 | 0.57 | 0.6 | 0.59 | 0.5674 | 0.576 | 0.63518 | 0.55211 | 0.59 | 0.6 | 0.59 | 0.60725 | 0.603 | 0.58 | 0.58356 | 0.62 | 0.5909 | 0.622 | 0.56 |  | 116
| 1 | 6 | 0.57 | 0.55 | 0.56 | 0.57 | 0.59 | 0.6 | 0.59 | 0.5686 | 0.588 | 0.63447 | 0.54288 | 0.59 | 0.6 | 0.6 | 0.61576 | 0.619 | 0.58 | 0.58187 | 0.62 | 0.6171 | 0.614 | 0.56
| Mean |  | 0.565 | 0.55333 | 0.565 | 0.55833 | 0.575 | 0.59333 | 0.59 | 0.5746 | 0.582 | 0.63506 | 0.55227 | 0.58833 | 0.59 | 0.615 | 0.60201 | 0.594 | 0.58 | 0.5829 | 0.62333 | 0.60062 | 0.6235 | 0.56
| Median |  | 0.565 | 0.55 | 0.565 | 0.555 | 0.57 | 0.595 | 0.59 | 0.5765 | 0.582 | 0.63482 | 0.55444 | 0.59 | 0.6 | 0.61 | 0.60048 | 0.5905 | 0.58 | 0.58363 | 0.62 | 0.5994 | 0.627 | 0.56
| Std Dev. |  | 0.016432 | 0.005164 | 0.0054772 | 0.0098319 | 0.016432 | 0.008165 | 0.010954 | 0.0053855 | 0.0056921 | 0.0017199 | 0.0080086 | 0.0040825 | 0.015492 | 0.0251 | 0.0084811 | 0.01562 | 0.0063246 | 0.0047928 | 0.005164 | 0.012048 | 0.012062 | 0.0063246
| Rel.Std.Dev. |  | 2.908% | 0.9332% | 0.9694% | 1.761% | 2.858% | 1.376% | 1.857% | 0.9373% | 0.978% | 0.2708% | 1.45% | 0.6939% | 2.626% | 4.081% | 1.409% | 2.63% | 1.09% | 0.8222% | 0.8284% | 2.006% | 1.935% | 1.129%
| PDM3 |  | -3.67% | -5.66% | -3.67% | -4.807% | -1.966% | 1.16% | 0.5919% | -2.034% | -0.7721% | 8.274% | -5.841% | 0.3077% | 0.5919% | 4.854% | 2.64% | 1.274% | -1.113% | -0.6184% | 6.275% | 2.402% | 6.303% | -4.523%
| Z-Score (Absolute) |  | 0.89 | 1.39 | 0.89 | 1.17 | 0.46 | 0.33 | 0.19 | 0.47 | 0.16 | 2.12 | 1.43 | 0.12 | 0.19 | 1.26 | 0.7 | 0.36 | 0.24 | 0.12 | 1.62 | 0.64 | 1.63 | 1.1
| Analytical results for Mn in OREAS 702 (Certified Value 0.031 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.0311 | 0.0305 | 0.0308 | 0.0301 | 0.0319 | 0.0304 | 0.0323 | 0.0292 | 0.0296 | 0.031134 | 0.031694 | 0.0342 | 0.0307 | 0.03 | 0.0296 | 0.030784 | 0.0325 | 0.03106 | 0.0313 | 0.0293 | 0.0305 | 0.0298 |  | 1
| 1 | 2 | 0.0303 | 0.0302 | 0.0306 | 0.0311 | 0.0305 | 0.0304 | 0.0327 | 0.0295 | 0.0296 | 0.031016 | 0.031306 | 0.0354 | 0.03 | 0.04 | 0.029518 | 0.03154 | 0.0323 | 0.031702 | 0.0314 | 0.0275 | 0.031 | 0.0301 |  | #N/A
| 1 | 3 | 0.032 | 0.0299 | 0.0303 | 0.0302 | 0.0301 | 0.0309 | 0.0315 | 0.0294 | 0.0302 | 0.030876 | 0.031135 | 0.0355 | 0.0299 | 0.04 | 0.030385 | 0.030305 | 0.0327 | 0.031885 | 0.0312 | 0.0298 | 0.0316 | 0.0307 |  | 16
| 1 | 4 | 0.0306 | 0.03 | 0.0307 | 0.0299 | 0.0306 | 0.0314 | 0.032 | 0.0292 | 0.0299 | 0.030914 | 0.030427 | 0.0348 | 0.0299 | 0.04 | 0.0301 | 0.02958 | 0.0321 | 0.0315 | 0.0313 | 0.0305 | 0.0307 | 0.0303 |  | 0.030685
| 1 | 5 | 0.0293 | 0.03 | 0.0308 | 0.0309 | 0.0308 | 0.0313 | 0.0318 | 0.0289 | 0.0298 | 0.030984 | 0.031916 | 0.0342 | 0.0302 | 0.03 | 0.030456 | 0.029402 | 0.0325 | 0.031387 | 0.0313 | 0.0294 | 0.0318 | 0.0305 |  | 117
| 1 | 6 | 0.0309 | 0.0297 | 0.0301 | 0.0311 | 0.0312 | 0.0299 | 0.0312 | 0.0291 | 0.0292 | 0.030939 | 0.029938 | 0.0345 | 0.03 | 0.03 | 0.030602 | 0.031541 | 0.0324 | 0.031464 | 0.0315 | 0.0308 | 0.0315 | 0.0301
| Mean |  | 0.0307 | 0.03005 | 0.03055 | 0.03055 | 0.03085 | 0.030717 | 0.031917 | 0.029217 | 0.029717 | 0.030977 | 0.031069 | 0.034767 | 0.030117 | 0.035 | 0.03011 | 0.030525 | 0.032417 | 0.0315 | 0.031333 | 0.02955 | 0.031183 | 0.03025
| Median |  | 0.03075 | 0.03 | 0.03065 | 0.03055 | 0.0307 | 0.03065 | 0.0319 | 0.0292 | 0.0297 | 0.030962 | 0.03122 | 0.03465 | 0.03 | 0.035 | 0.030242 | 0.030544 | 0.03245 | 0.031482 | 0.0313 | 0.0296 | 0.03125 | 0.0302
| Std Dev. |  | 0.00089666 | 0.00027386 | 0.0002881 | 0.00054314 | 0.00062849 | 0.00058452 | 0.00054191 | 0.0002137 | 0.00033714 | 9.1451e-05 | 0.00075619 | 0.00057504 | 0.00030605 | 0.0054772 | 0.00045784 | 0.00093092 | 0.00020412 | 0.00028158 | 0.00010328 | 0.0011675 | 0.00052694 | 0.00032094
| Rel.Std.Dev. |  | 2.921% | 0.9114% | 0.943% | 1.778% | 2.037% | 1.903% | 1.698% | 0.7314% | 1.135% | 0.2952% | 2.434% | 1.654% | 1.016% | 15.65% | 1.521% | 3.05% | 0.6297% | 0.8939% | 0.3296% | 3.951% | 1.69% | 1.061%
| PDM3 |  | 0.04751% | -2.071% | -0.4413% | -0.4413% | 0.5363% | 0.1018% | 4.012% | -4.786% | -3.157% | 0.951% | 1.251% | 13.3% | -1.853% | 14.06% | -1.875% | -0.5221% | 5.642% | 2.654% | 2.111% | -3.7% | 1.623% | -1.419%
| Z-Score (Absolute) |  | 0.01 | 0.75 | 0.18 | 0.18 | 0.16 | 0.01 | 1.37 | 1.69 | 1.12 | 0.3 | 0.41 | 4.6 | 0.67 | 4.86 | 0.68 | 0.21 | 1.94 | 0.9 | 0.71 | 1.31 | 0.54 | 0.52
| Analytical results for Mo in OREAS 702 (Certified Value 577 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 575 | 558 | 578 | 556 | 594 | 637 | 606 | 602.5 | 576 | 568.94 | 521.4276 | 594 | 536 | 624.538 | 556.58 | 515.453 | 567.9 | 594.8082 | 560 | 633.5799 | 557 | 591 | 595.21 | 1
| 1 | 2 | 565 | 552 | 568 | 566 | 571 | 619 | 592 | 607.7 | 584 | 566.71 | 533.7065 | 601 | 533 | 630.92 | 547.981 | 523.0645 | 575.94 | 595.5652 | 558 | 652.4831 | 571 | 580 | 605.2 | 21
| 1 | 3 | 581 | 549 | 561 | 544 | 560 | 624 | 593 | 608 | 585 | 566.96 | 510.786 | 599.3 | 526 | 657.98 | 564.302 | 521.3883 | 577.26 | 601.3651 | 551 | 637.0572 | 577 | 584 | 618.12 | 16
| 1 | 4 | 567 | 546 | 581 | 549 | 567 | 629 | 594 | 601.6 | 582 | 566.87 | 510.83 | 592.7 | 529 | 625.3916 | 558.342 | 506.6976 | 574.82 | 592.3491 | 553 | 641.2011 | 564 | 580 | 605.41 | 576.9731
| 1 | 5 | 541 | 552 | 578 | 565 | 569 | 608 | 599 | 602.9 | 583 | 567.68 | 516.805 | 594.7 | 537 | 630.2819 | 567.98 | 509.5629 | 583.84 | 595.2255 | 550 | 624.852 | 584 | 575 | 607.03 | 118
| 1 | 6 | 576 | 547 | 570 | 570 | 591 | 625 | 594 | 601.4 | 582 | 568.44 | 513.634 | 600.5 | 534 | 612.58 | 568.169 | 511.8354 | 579.17 | 583.3184 | 551 | 656.202 | 581 | 578 | 606.48
| Mean |  | 567.5 | 550.6667 | 572.6667 | 558.3333 | 575.3333 | 623.6667 | 596.3333 | 604.0167 | 582 | 567.6 | 517.8649 | 597.0333 | 532.5 | 630.2819 | 560.559 | 514.6669 | 576.4883 | 593.7719 | 553.8333 | 640.8959 | 572.3333 | 581.3333 | 606.2417
| Median |  | 571 | 550.5 | 574 | 560.5 | 570 | 624.5 | 594 | 602.7 | 582.5 | 567.32 | 515.2195 | 597 | 533.5 | 627.8368 | 561.322 | 513.6442 | 576.6 | 595.0168 | 552 | 639.1291 | 574 | 580 | 605.945
| Std Dev. |  | 14.2794 | 4.3665 | 7.6333 | 10.3666 | 13.8371 | 9.7502 | 5.3166 | 3.0222 | 3.1623 | 0.92111 | 8.7439 | 3.6418 | 4.2308 | 15.0855 | 7.8224 | 6.5405 | 5.2664 | 5.9222 | 4.1673 | 11.7839 | 10.3859 | 5.5737 | 7.2835
| Rel.Std.Dev. |  | 2.516% | 0.793% | 1.333% | 1.857% | 2.405% | 1.563% | 0.8916% | 0.5003% | 0.5433% | 0.1623% | 1.688% | 0.61% | 0.7945% | 2.393% | 1.395% | 1.271% | 0.9135% | 0.9974% | 0.7525% | 1.839% | 1.815% | 0.9588% | 1.201%
| PDM3 |  | -1.642% | -4.559% | -0.7464% | -3.231% | -0.2842% | 8.093% | 3.355% | 4.687% | 0.8713% | -1.625% | -10.24% | 3.477% | -7.708% | 9.239% | -2.845% | -10.8% | -0.08402% | 2.912% | -4.011% | 11.08% | -0.8042% | 0.7557% | 5.073%
| Z-Score (Absolute) |  | 0.25 | 0.79 | 0.09 | 0.55 | 0 | 1.55 | 0.67 | 0.92 | 0.21 | 0.25 | 1.85 | 0.7 | 1.38 | 1.76 | 0.47 | 1.95 | 0.04 | 0.59 | 0.69 | 2.11 | 0.1 | 0.19 | 0.99
| Analytical results for Na in OREAS 702 (Certified Value 1.32 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.35 | 1.34 | 1.31 | 1.32 | 1.37 | 1.3 | 1.35 | 1.247 | 1.32 | 1.4347 | 1.2545 | 1.4 | 1.29 | 1.43 | 1.3139 | 1.31 | 1.35 | 1.3167 | 1.31 | 1.254 | 1.37 | 1.3 | 1.31 | 1
| 1 | 2 | 1.33 | 1.33 | 1.29 | 1.32 | 1.31 | 1.3 | 1.36 | 1.2607 | 1.3 | 1.4431 | 1.2421 | 1.43 | 1.27 | 1.48 | 1.2987 | 1.316 | 1.3 | 1.3235 | 1.32 | 1.2412 | 1.39 | 1.315 | 1.33 | #N/A
| 1 | 3 | 1.4 | 1.32 | 1.3 | 1.28 | 1.29 | 1.32 | 1.31 | 1.2513 | 1.32 | 1.4235 | 1.2041 | 1.43 | 1.26 | 1.59 | 1.3243 | 1.3 | 1.36 | 1.3324 | 1.31 | 1.2969 | 1.39 | 1.305 | 1.35 | 16
| 1 | 4 | 1.33 | 1.33 | 1.32 | 1.3 | 1.31 | 1.33 | 1.33 | 1.2485 | 1.31 | 1.4349 | 1.2679 | 1.42 | 1.27 | 1.49 | 1.3256 | 1.268 | 1.29 | 1.3285 | 1.32 | 1.3199 | 1.33 | 1.3 | 1.33 | 1.3211
| 1 | 5 | 1.3 | 1.33 | 1.32 | 1.32 | 1.32 | 1.32 | 1.33 | 1.2342 | 1.32 | 1.429 | 1.2135 | 1.41 | 1.28 | 1.42 | 1.3401 | 1.351 | 1.31 | 1.3194 | 1.33 | 1.2612 | 1.35 | 1.325 | 1.34 | 119
| 1 | 6 | 1.34 | 1.31 | 1.29 | 1.35 | 1.34 | 1.3 | 1.34 | 1.2296 | 1.3 | 1.4238 | 1.202 | 1.42 | 1.26 | 1.4 | 1.3438 | 1.348 | 1.35 | 1.3213 | 1.32 | 1.3167 | 1.33 | 1.305 | 1.33
| Mean |  | 1.3417 | 1.3267 | 1.305 | 1.315 | 1.3233 | 1.3117 | 1.3367 | 1.2452 | 1.3117 | 1.4315 | 1.2307 | 1.4183 | 1.2717 | 1.4683 | 1.3244 | 1.3155 | 1.3267 | 1.3236 | 1.3183 | 1.2816 | 1.36 | 1.3083 | 1.3317
| Median |  | 1.335 | 1.33 | 1.305 | 1.32 | 1.315 | 1.31 | 1.335 | 1.2477 | 1.315 | 1.4319 | 1.2278 | 1.42 | 1.27 | 1.455 | 1.325 | 1.313 | 1.33 | 1.3224 | 1.32 | 1.279 | 1.36 | 1.305 | 1.33
| Std Dev. |  | 0.033116 | 0.010328 | 0.013784 | 0.023452 | 0.028048 | 0.013292 | 0.017512 | 0.011454 | 0.0098319 | 0.0075567 | 0.027943 | 0.01169 | 0.01169 | 0.069113 | 0.016704 | 0.031124 | 0.030111 | 0.0058861 | 0.0075277 | 0.033891 | 0.027568 | 0.0098319 | 0.013292
| Rel.Std.Dev. |  | 2.468% | 0.7785% | 1.056% | 1.783% | 2.119% | 1.013% | 1.31% | 0.9198% | 0.7496% | 0.5279% | 2.271% | 0.8242% | 0.9193% | 4.707% | 1.261% | 2.366% | 2.27% | 0.4447% | 0.571% | 2.644% | 2.027% | 0.7515% | 0.9981%
| PDM3 |  | 1.559% | 0.4233% | -1.217% | -0.4598% | 0.171% | -0.7121% | 1.18% | -5.742% | -0.7121% | 8.361% | -6.843% | 7.362% | -3.74% | 11.15% | 0.2506% | -0.422% | 0.4233% | 0.1942% | -0.2075% | -2.984% | 2.947% | -0.9644% | 0.8018%
| Z-Score (Absolute) |  | 0.93 | 0.17 | 0.93 | 0.42 | 0 | 0.59 | 0.67 | 3.95 | 0.59 | 5.47 | 4.69 | 4.8 | 2.61 | 7.33 | 0.05 | 0.4 | 0.17 | 0.02 | 0.25 | 2.11 | 1.85 | 0.76 | 0.42
| Analytical results for Nb in OREAS 702 (Certified Value 8.68 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 24 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 8.4 | 8.6 | 9.6 | 8.2 | 9.1 | 9.4 | 9 | 7.94 | 8.3 | 4.8241 | 8.3 | 4.23 | 2.9 | < 1 | 4.2 | 8.6328 | 5.6 | 10 | 8.3 |  |  |  |  | 1
| 1 | 2 | 8 | 8.7 | 9.5 | 8.5 | 8.5 | 9.2 | 8.9 | 8.47 | 8.4 | 5.1547 | 9 | 4.37 | 2.7 | < 1 | 4 | 8.6887 | 5.7 | 10 | 8.4 |  |  |  |  | 1
| 1 | 3 | 8.8 | 8.6 | 9.4 | 9 | 8.6 | 9.4 | 8.7 | 8.36 | 8.5 | 4.9896 | 8.4 | 4.45 | 4.6 | < 1 | 4.1 | 8.7058 | 5.5 | 10 | 8.6 |  |  |  |  | 16
| 1 | 4 | 8.7 | 8.7 | 9.2 | 8.6 | 8.8 | 9.5 | 8.9 | 8.13 | 8.6 | 5.0242 | 8.4 | 4.4 | 1.9 | < 1 | 3.7 | 8.605 | 5.5 | 9 | 8.4 |  |  |  |  | 8.6839
| 1 | 5 | 8.3 | 8.6 | 9.4 | 8.8 | 8.7 | 9.2 | 8.7 | 8.1 | 8.4 | 4.8996 | 8.4 | 4.24 | 4.5 | 1.2375 | 4 | 8.5747 | 5.3 | 10 | 8.4 |  |  |  |  | 120
| 1 | 6 | 8.7 | 8.5 | 8.9 | 8.8 | 8.9 | 9.2 | 9 | 7.9 | 8.5 | 5.1145 | 8.6 | 4.23 | 2.3 | < 1 | 3.9 | 8.6358 | 4.6 | 9 | 8.4
| Mean |  | 8.4833 | 8.6167 | 9.3333 | 8.65 | 8.7667 | 9.3167 | 8.8667 | 8.15 | 8.45 | 5.0011 | 8.5167 | 4.32 | 3.15 | 1.2375 | 3.9833 | 8.6405 | 5.3667 | 9.6667 | 8.4167
| Median |  | 8.55 | 8.6 | 9.4 | 8.7 | 8.75 | 9.3 | 8.9 | 8.115 | 8.45 | 5.0069 | 8.4 | 4.305 | 2.8 | 1.2375 | 4 | 8.6343 | 5.5 | 10 | 8.4
| Std Dev. |  | 0.30605 | 0.075277 | 0.25033 | 0.28107 | 0.21602 | 0.13292 | 0.13663 | 0.22627 | 0.10488 | 0.1255 | 0.25626 | 0.098387 | 1.138 |  | 0.17224 | 0.049534 | 0.39833 | 0.5164 | 0.098319
| Rel.Std.Dev. |  | 3.608% | 0.8736% | 2.682% | 3.249% | 2.464% | 1.427% | 1.541% | 2.776% | 1.241% | 2.509% | 3.009% | 2.277% | 36.13% |  | 4.324% | 0.5733% | 7.422% | 5.342% | 1.168%
| PDM3 |  | -2.31% | -0.7746% | 7.478% | -0.3907% | 0.9528% | 7.286% | 2.104% | -6.148% | -2.694% | -42.41% | -1.926% | -50.25% | -63.73% | -85.75% | -54.13% | -0.5004% | -38.2% | 11.32% | -3.078%
| Z-Score (Absolute) |  | 0.03 | 0.28 | 1.63 | 0.34 | 0.56 | 1.6 | 0.75 | 0.6 | 0.03 | 6.52 | 0.09 | 7.8 | 10 | 14.76 | 8.44 | 0.33 | 5.83 | NA | 0.09
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 702 (Certified Value 24.8 ppm) | [24 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 1
1
| 1 | 1 | 25.1 | 30.4 | 18.56 | 28.1868 | 27 | 25.5 | 22.986 | 22.396 | [15 empty] | 1
| 1 | 2 | 24.9 | 29.7 | 19.04 | 28.3729 | 26.1 | 26.4 | 24.0458 | 21.386 | [15 empty] | #N/A
| 1 | 3 | 25.8 | 28.3 | 19.38 | 28.4062 | 26.2 | 26 | 24.2475 | 21.005 | [15 empty] | 16
| 1 | 4 | 24.7 | 28.2 | 19.67 | 28.1566 | 26.7 | 26.6 | 23.9211 | 20.854 | [15 empty] | 24.8154
| 1 | 5 | 24.6 | 27.3 | 18.65 | 28.1245 | 26.5 | 26.6 | 24.4452 | 20.4788 | [15 empty] | 121
| 1 | 6 | 25.5 | 27.2 | 17.8 | 28.1417 | 26.3 | 25.8 | 24.3004 | 21.183
| Mean |  | 25.1 | 28.5167 | 18.85 | 28.2314 | 26.4667 | 26.15 | 23.991 | 21.2171
| Median |  | 25 | 28.25 | 18.845 | 28.1717 | 26.4 | 26.2 | 24.1466 | 21.094
| Std Dev. |  | 0.46904 | 1.2891 | 0.66573 | 0.1246 | 0.33862 | 0.45497 | 0.52644 | 0.65443
| Rel.Std.Dev. |  | 1.869% | 4.52% | 3.532% | 0.4413% | 1.279% | 1.74% | 2.194% | 3.084%
| PDM3 |  | 1.147% | 14.92% | -24.04% | 13.77% | 6.654% | 5.378% | -3.322% | -14.5%
| Z-Score (Absolute) |  | 0.17 | 0.92 | 2.15 | 0.83 | 0.27 | 0.17 | 0.52 | 1.4
| Analytical results for Ni in OREAS 702 (Certified Value 11.3 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 10.3 | 9.6 | 14.8 | 10.1 | 12.2 | 11.2 | 8 | 11 | 11 | 10.4802 | 11.7599 | 11.9 | 11.7 | 7.8 | 11.99 | 16.2189 | 13.3 | 20.2396 | 11.3 | 12.4514 | < 20 | 11.5 |  | 1
| 1 | 2 | 10 | 9.4 | 11.8 | 10.3 | 11.4 | 10.9 | 8 | 11.1 | 10 | 10.3891 | 12.0739 | 12.4 | 11.5 | 7.8 | 11.62 | 16.4503 | 14.6 | 20.3621 | 11.5 | 10.9065 | < 20 | 11.4 |  | 12
| 1 | 3 | 11 | 9.9 | 11.8 | 11.2 | 11.4 | 11.4 | 8 | 11.1 | 10 | 10.3815 | 11.824 | 12.5 | 11.3 | 7.5 | 11.9 | 15.7105 | 15.7 | 19.6465 | 12.5 | 12.5507 | < 20 | 11.7 |  | 16
| 1 | 4 | 10.6 | 9.5 | 11.6 | 10.6 | 11.2 | 11.4 | 8 | 11.1 | 10 | 10.4629 | 12.1 | 12.1 | 11.6 | 8.1 | 10.99 | 15.739 | 15.3 | 19.6919 | 10.8 | 12.8844 | < 20 | 11.4 |  | 11.2795
| 1 | 5 | 10.4 | 9.9 | 11.6 | 11.2 | 11.3 | 11.2 | 9 | 10.8 | 10 | 10.4329 | 11.77 | 12.1 | 11.3 | 7.7 | 10.58 | 15.2542 | 15.5 | 19.5801 | 11.2 | 12.2079 | < 20 | 11.5 |  | 122
| 1 | 6 | 10.8 | 10 | 11.3 | 11.5 | 11.6 | 11.2 | 9 | 11.1 | 10 | 10.5171 | 11.8217 | 11.9 | 11.7 | 10.8 | 12.97 | 16.695 | 14.5 | 19.3104 | 11.1 | 12.1744 | < 20 | 11.4
| Mean |  | 10.5167 | 9.7167 | 12.15 | 10.8167 | 11.5167 | 11.2167 | 8.3333 | 11.0333 | 10.1667 | 10.4439 | 11.8916 | 12.15 | 11.5167 | 8.2833 | 11.675 | 16.0113 | 14.8167 | 19.8051 | 11.4 | 12.1959 |  | 11.4833
| Median |  | 10.5 | 9.75 | 11.7 | 10.9 | 11.4 | 11.2 | 8 | 11.1 | 10 | 10.4479 | 11.8229 | 12.1 | 11.55 | 7.8 | 11.76 | 15.979 | 14.95 | 19.6692 | 11.25 | 12.3297 |  | 11.45
| Std Dev. |  | 0.36009 | 0.24833 | 1.3111 | 0.56362 | 0.36009 | 0.18348 | 0.5164 | 0.12111 | 0.40825 | 0.053023 | 0.15379 | 0.251 | 0.18348 | 1.2481 | 0.83644 | 0.53684 | 0.88638 | 0.40803 | 0.58652 | 0.6824 |  | 0.1169
| Rel.Std.Dev. |  | 3.424% | 2.556% | 10.79% | 5.211% | 3.127% | 1.636% | 6.197% | 1.098% | 4.016% | 0.5077% | 1.293% | 2.066% | 1.593% | 15.07% | 7.164% | 3.353% | 5.982% | 2.06% | 5.145% | 5.595% |  | 1.018%
| PDM3 |  | -6.763% | -13.86% | 7.717% | -4.104% | 2.102% | -0.5574% | -26.12% | -2.183% | -9.866% | -7.408% | 5.426% | 7.717% | 2.102% | -26.56% | 3.506% | 41.95% | 31.36% | 75.58% | 1.068% | 8.124% |  | 1.807%
| Z-Score (Absolute) |  | 0.99 | 1.8 | 0.66 | 0.69 | 0.02 | 0.29 | NA | 0.47 | NA | 1.07 | 0.4 | 0.66 | 0.02 | 3.25 | 0.18 | 4.56 | 3.35 | 8.4 | 0.1 | 0.7 | 1.52 | 0.02
Results from laboratories 7 and 10 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 702 (Certified Value 0.058 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.058 | 0.058 | 0.058 | 0.057 | 0.058 | 0.056 | 0.06 | 0.0546 | 0.056 | 0.054902 | 0.054979 | 0.06 | 0.0581 | 0.0567 | 0.045637 | 0.057 | 0.0591 | 0.059351 | 0.057 | 0.0576 | 0.0613 | 0.0582 |  | 1
| 1 | 2 | 0.057 | 0.059 | 0.057 | 0.058 | 0.056 | 0.057 | 0.06 | 0.0554 | 0.057 | 0.055085 | 0.053895 | 0.06 | 0.0597 | 0.0552 | 0.047839 | 0.058 | 0.0586 | 0.060395 | 0.058 | 0.0546 | 0.0628 | 0.0588 |  | #N/A
| 1 | 3 | 0.06 | 0.057 | 0.056 | 0.055 | 0.055 | 0.057 | 0.06 | 0.0557 | 0.057 | 0.055024 | 0.05767 | 0.06 | 0.0587 | 0.0611 | 0.049701 | 0.057 | 0.0596 | 0.060512 | 0.058 | 0.0592 | 0.0637 | 0.0602 |  | 16
| 1 | 4 | 0.057 | 0.057 | 0.058 | 0.056 | 0.056 | 0.058 | 0.06 | 0.055 | 0.057 | 0.055133 | 0.055325 | 0.06 | 0.0582 | 0.0554 | 0.049719 | 0.056 | 0.0584 | 0.059382 | 0.058 | 0.06 | 0.062 | 0.0588 |  | 0.05756
| 1 | 5 | 0.055 | 0.056 | 0.058 | 0.057 | 0.057 | 0.057 | 0.06 | 0.0549 | 0.057 | 0.054984 | 0.058307 | 0.06 | 0.0572 | 0.0553 | 0.049058 | 0.059 | 0.0592 | 0.060033 | 0.058 | 0.0584 | 0.0643 | 0.0593 |  | 123
| 1 | 6 | 0.058 | 0.057 | 0.056 | 0.058 | 0.058 | 0.055 | 0.06 | 0.0541 | 0.057 | 0.055138 | 0.053435 | 0.06 | 0.0585 | 0.0563 | 0.048181 | 0.057 | 0.0592 | 0.059477 | 0.058 | 0.0612 | 0.0635 | 0.0587
| Mean |  | 0.0575 | 0.057333 | 0.057167 | 0.056833 | 0.056667 | 0.056667 | 0.06 | 0.05495 | 0.056833 | 0.055044 | 0.055602 | 0.06 | 0.0584 | 0.056667 | 0.048356 | 0.057333 | 0.059017 | 0.059858 | 0.057833 | 0.0585 | 0.062933 | 0.059
| Median |  | 0.0575 | 0.057 | 0.0575 | 0.057 | 0.0565 | 0.057 | 0.06 | 0.05495 | 0.057 | 0.055055 | 0.055152 | 0.06 | 0.05835 | 0.05585 | 0.04862 | 0.057 | 0.05915 | 0.059755 | 0.058 | 0.0588 | 0.06315 | 0.0588
| Std Dev. |  | 0.0016432 | 0.0010328 | 0.00098319 | 0.001169 | 0.0012111 | 0.0010328 | 0 | 0.00056833 | 0.00040825 | 9.2388e-05 | 0.0019833 | 0 | 0.00081976 | 0.0022545 | 0.0015393 | 0.0010328 | 0.00044008 | 0.00052435 | 0.00040825 | 0.002283 | 0.0011255 | 0.00068411
| Rel.Std.Dev. |  | 2.858% | 1.801% | 1.72% | 2.057% | 2.137% | 1.823% | 0% | 1.034% | 0.7183% | 0.1678% | 3.567% | 0% | 1.404% | 3.978% | 3.183% | 1.801% | 0.7457% | 0.876% | 0.7059% | 3.903% | 1.788% | 1.16%
| PDM3 |  | -0.1046% | -0.3942% | -0.6837% | -1.263% | -1.552% | -1.552% | 4.239% | -4.535% | -1.263% | -4.371% | -3.402% | 4.239% | 1.459% | -1.552% | -15.99% | -0.3942% | 2.53% | 3.992% | 0.4745% | 1.633% | 9.335% | 2.501%
| Z-Score (Absolute) |  | 0.1 | 0 | 0.1 | 0.3 | 0.4 | 0.4 | 1.61 | 1.44 | 0.3 | 1.38 | 1.05 | 1.61 | 0.64 | 0.4 | 5.42 | 0 | 1.02 | 1.52 | 0.3 | 0.7 | 3.38 | 1.01
| Analytical results for Pb in OREAS 702 (Certified Value 17.3 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 17.2 | 16.8 | 18.4 | 15.7 | 17.2 | 17.6 | 21 | 10 | 16 | 21.5528 | 18.1 | 17.8 | 18.3 | 34.25 | 17.5448 | 15 | 13.0276 | 15.5 | 21.4889 | 15 | 17 | 17.4 |  | 1
| 1 | 2 | 15.2 | 16.4 | 18.2 | 16.6 | 16.4 | 18.4 | 33 | 9 | 16 | 20.9 | 18.7 | 16.9 | 18.3 | 37.4 | 19.0303 | 15.4 | 13.585 | 15.6 | 22.0002 | 18 | 17 | 17.5 |  | 23
| 1 | 3 | 16.9 | 16 | 17.5 | 17.2 | 16 | 17.1 | 29 | 10 | 16 | 20.21 | 19 | 17.7 | 18.9 | 35.43 | 20.4409 | 16 | 12.7293 | 14.1 | 23.1661 | 18 | 17 | 18 |  | 16
| 1 | 4 | 16.3 | 16.4 | 19 | 16.1 | 16.2 | 17.3 | 24 | 10 | 16 | 20.28 | 17.7 | 16.9 | 17.9 | 36.18 | 20.4119 | 16.1 | 12.4909 | 14.6 | 25.1219 | 18 | 16 | 17.3 |  | 17.2722
| 1 | 5 | 16 | 16.6 | 17.8 | 17.2 | 16.2 | 17 | < 20 | 9 | 16 | 20.2 | 18.7 | 17.2 | 17.9 | 34.93 | 17.3014 | 15.4 | 12.8251 | 14.9 | 20.8695 | 19 | 16 | 17.2 |  | 124
| 1 | 6 | 17.3 | 16.8 | 17.6 | 17 | 16.4 | 18.5 | 28 | 10 | 16 | 20.862 | 18.8 | 17.1 | 17.4 | 34.29 | 19.9313 | 16 | 12.5879 | 14.9 | 21.43 | 18 | 16 | 17.4
| Mean |  | 16.4833 | 16.5 | 18.0833 | 16.6333 | 16.4 | 17.65 | 27 | 9.6667 | 16 | 20.6675 | 18.5 | 17.2667 | 18.1167 | 35.4133 | 19.1101 | 15.65 | 12.8743 | 14.9333 | 22.3461 | 17.6667 | 16.5 | 17.4667
| Median |  | 16.6 | 16.5 | 18 | 16.8 | 16.3 | 17.45 | 28 | 10 | 16 | 20.571 | 18.7 | 17.15 | 18.1 | 35.18 | 19.4808 | 15.7 | 12.7772 | 14.9 | 21.7446 | 18 | 16.5 | 17.4
| Std Dev. |  | 0.8085 | 0.30332 | 0.56716 | 0.62183 | 0.41952 | 0.65345 | 4.6368 | 0.5164 | 0 | 0.53921 | 0.49396 | 0.39328 | 0.50761 | 1.2152 | 1.4049 | 0.44609 | 0.39523 | 0.56095 | 1.5653 | 1.3663 | 0.54772 | 0.28048
| Rel.Std.Dev. |  | 4.905% | 1.838% | 3.136% | 3.738% | 2.558% | 3.702% | 17.17% | 5.342% | 0% | 2.609% | 2.67% | 2.278% | 2.802% | 3.432% | 7.351% | 2.85% | 3.07% | 3.756% | 7.005% | 7.734% | 3.32% | 1.606%
| PDM3 |  | -4.567% | -4.471% | 4.696% | -3.699% | -5.05% | 2.187% | 56.32% | -44.03% | -7.366% | 19.66% | 7.108% | -0.03209% | 4.889% | 105% | 10.64% | -9.392% | -25.46% | -13.54% | 29.38% | 2.284% | -4.471% | 1.126%
| Z-Score (Absolute) |  | 0.57 | 0.56 | 0.46 | 0.47 | 0.62 | 0.18 | 4.37 | 4.94 | 0.88 | 2.12 | 0.73 | 0.06 | 0.48 | 11.59 | 1.12 | 1.1 | 2.88 | 1.56 | 3.2 | 0.19 | 0.56 | 0.06
| Analytical results for Pr in OREAS 702 (Certified Value 6.89 ppm) | [24 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 7.03 | 7.54 | 4.75 | 7.1169 | 6.75 | 6.96 | 6.2201 | 6.778 | [15 empty] | 1
| 1 | 2 | 6.97 | 7.68 | 5.03 | 7.186 | 6.62 | 7.1 | 6.5198 | 7.175 | [15 empty] | #N/A
| 1 | 3 | 6.71 | 7.49 | 5.12 | 7.1004 | 6.46 | 7 | 6.4749 | 6.86 | [15 empty] | 16
| 1 | 4 | 7.11 | 7.31 | 5.14 | 7.1067 | 6.82 | 7.12 | 6.3189 | 6.665 | [15 empty] | 6.892
| 1 | 5 | 6.45 | 7.15 | 4.92 | 7.0745 | 6.67 | 7.14 | 6.6131 | 6.05 | [15 empty] | 125
| 1 | 6 | 6.61 | 7.17 | 4.66 | 7.0866 | 6.88 | 7.05 | 6.4364 | 6.893
| Mean |  | 6.8133 | 7.39 | 4.9367 | 7.1119 | 6.7 | 7.0617 | 6.4305 | 6.7368
| Median |  | 6.84 | 7.4 | 4.975 | 7.1036 | 6.71 | 7.075 | 6.4557 | 6.819
| Std Dev. |  | 0.26212 | 0.21401 | 0.19765 | 0.039291 | 0.15113 | 0.07111 | 0.14147 | 0.37687
| Rel.Std.Dev. |  | 3.847% | 2.896% | 4.004% | 0.5525% | 2.256% | 1.007% | 2.2% | 5.594%
| PDM3 |  | -1.142% | 7.225% | -28.37% | 3.19% | -2.786% | 2.461% | -6.696% | -2.252%
| Z-Score (Absolute) |  | 0.08 | 1.33 | 3.98 | 0.73 | 0.16 | 0.62 | 0.75 | 0.08
| Analytical results for Rb in OREAS 702 (Certified Value 123 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 0
0
| 1 | 1 | 124 | 120 | 132 | 111 | 124.5 | 125 | 132 | 117.2 | 115 | 132.0563 | 122.4 | 6.7 | 123.7 | 118.9467 | 119.5 | 122.3394 | 130.4 | 100.536 | 120.9 |  |  |  |  | 1
| 1 | 2 | 123 | 121 | 131.5 | 116 | 119 | 122 | 129.1 | 125.7 | 116 | 131.2535 | 127.9 | 6.6 | 132.9 | 120.4093 | 122.3 | 120.9516 | 131.8 | 104.559 | 120.5 |  |  |  |  | 2
| 1 | 3 | 133.5 | 120.5 | 129.5 | 121 | 119 | 125 | 125.5 | 122.9 | 115 | 131.9918 | 127.1 | 6.5 | 128.8 | 118.4303 | 123 | 121.3623 | 131.2 | 101.583 | 122.8 |  |  |  |  | 16
| 1 | 4 | 130 | 121.5 | 131 | 116 | 121 | 125 | 129.5 | 120.7 | 115 | 132.3172 | 127.6 | 6.8 | 122.2 | 114.7131 | 123.2 | 125.2379 | 130.8 | 102.892 | 120.9 |  |  |  |  | 123.4713
| 1 | 5 | 124 | 121.5 | 131 | 122 | 122 | 120 | 125.1 | 119.3 | 114 | 131.7793 | 123.2 | 6.7 | 116.1 | 114.5227 | 123.4 | 121.9382 | 133.3 | 95.304 | 121.5 |  |  |  |  | 126
| 1 | 6 | 124.5 | 120 | 125.5 | 117.5 | 124 | 119 | 128.8 | 116 | 114 | 131.8112 | 128.1 | 6.9 | 122.3 | 119.0912 | 122.7 | 117.4173 | 132.1 | 107.331 | 121.1
| Mean |  | 126.5 | 120.75 | 130.0833 | 117.25 | 121.5833 | 122.6667 | 128.3333 | 120.3 | 114.8333 | 131.8682 | 126.05 | 6.7 | 124.3333 | 117.6856 | 122.35 | 121.5411 | 131.6 | 102.0342 | 121.2833
| Median |  | 124.25 | 120.75 | 131 | 116.75 | 121.5 | 123.5 | 128.95 | 120 | 115 | 131.9015 | 127.35 | 6.7 | 123 | 118.6885 | 122.85 | 121.6502 | 131.5 | 102.2375 | 121
| Std Dev. |  | 4.2426 | 0.6892 | 2.3962 | 3.9718 | 2.3752 | 2.7325 | 2.6113 | 3.6105 | 0.75277 | 0.35812 | 2.5524 | 0.14142 | 5.8346 | 2.4649 | 1.4488 | 2.5237 | 1.0412 | 4.0721 | 0.81097
| Rel.Std.Dev. |  | 3.354% | 0.5708% | 1.842% | 3.387% | 1.954% | 2.228% | 2.035% | 3.001% | 0.6555% | 0.2716% | 2.025% | 2.111% | 4.693% | 2.094% | 1.184% | 2.076% | 0.7911% | 3.991% | 0.6687%
| PDM3 |  | 2.453% | -2.204% | 5.355% | -5.039% | -1.529% | -0.6517% | 3.938% | -2.568% | -6.996% | 6.801% | 2.089% | -94.57% | 0.6982% | -4.686% | -0.9081% | -1.563% | 6.584% | -17.36% | -1.772%
| Z-Score (Absolute) |  | 0.77 | 0.13 | 1.32 | 0.67 | 0 | 0.17 | 1.05 | 0.2 | 1.05 | 1.6 | 0.7 | 17.88 | 0.43 | 0.61 | 0.12 | 0.01 | 1.56 | 3.04 | 0.05
| Analytical results for Re in OREAS 702 (Certified Value 0.035 ppm) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 14 | 18 | 21 | 24 | 25 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [9 empty] | 3
3
| 1 | 1 | 0.03 | 0.072 | 0.036 | 0.035 | 0.041 | 0.028 | 0.03 | 0.577 | < 0.05 | 0.013 | 0.216 | 0.049 | < 0.05 | 0.023 | [9 empty] | 1
| 1 | 2 | 0.032 | 0.075 | 0.04 | 0.038 | 0.039 | 0.03 | 0.04 | 0.625 | < 0.05 | 0.017 | 0.209 | 0.048 | < 0.05 | 0.023 | [9 empty] | 13
| 1 | 3 | 0.032 | 0.075 | 0.038 | 0.036 | 0.038 | 0.027 | 0.03 | 0.604 | < 0.05 | 0.022 | 0.218 | 0.044 | < 0.05 | 0.025 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.076 | 0.039 | 0.035 | 0.037 | 0.026 | 0.03 | 0.588 | < 0.05 | 0.02 | 0.219 | 0.049 | < 0.05 | 0.024 | [9 empty] | 0.035125
| 1 | 5 | 0.029 | 0.071 | 0.04 | 0.039 | 0.036 | 0.025 | 0.04 | 0.578 | < 0.05 | 0.028 | 0.216 | 0.046 | < 0.05 | 0.023 | [9 empty] | 127
| 1 | 6 | 0.035 | 0.073 | 0.035 | 0.038 | 0.038 | 0.044 | 0.05 | 0.561 | < 0.05 | 0.018 | 0.21 | 0.043 | < 0.05 | 0.023
| Mean |  | 0.031333 | 0.073667 | 0.038 | 0.036833 | 0.038167 | 0.03 | 0.036667 | 0.58883 |  | 0.019667 | 0.21467 | 0.0465 |  | 0.0235
| Median |  | 0.031 | 0.074 | 0.0385 | 0.037 | 0.038 | 0.0275 | 0.035 | 0.583 |  | 0.019 | 0.216 | 0.047 |  | 0.023
| Std Dev. |  | 0.0021602 | 0.0019664 | 0.0020976 | 0.0017224 | 0.0017224 | 0.0070711 | 0.008165 | 0.022675 |  | 0.0050859 | 0.0041793 | 0.0025884 |  | 0.00083666
| Rel.Std.Dev. |  | 6.894% | 2.669% | 5.52% | 4.676% | 4.513% | 23.57% | 22.27% | 3.851% |  | 25.86% | 1.947% | 5.567% |  | 3.56%
| PDM3 |  | -10.79% | 109.7% | 8.185% | 4.864% | 8.66% | -14.59% | 4.389% | 1576% |  | -44.01% | 511.2% | 32.38% |  | -33.1%
| Z-Score (Absolute) |  | 0.34 | 2.32 | 0.08 | 0.01 | 0.09 | 0.42 | 0.01 | 34.63 | 0.74 | 1.07 | 11.16 | 0.61 | 0.74 | 0.83
| Analytical results for S in OREAS 702 (Certified Value 0.349 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 24 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.36 | 0.38 | 0.35 | 0.35 | 0.37 | 0.34 | 0.35 | 0.3225 | 0.359 | 0.28407 | 0.43986 | 0.35 | 0.35 | 0.42 | 0.326 | 0.3 | 0.24 | 0.3741 | 0.35 | 0.3448 | 0.335 | 0.36 |  | 1
| 1 | 2 | 0.36 | 0.37 | 0.35 | 0.35 | 0.36 | 0.34 | 0.36 | 0.3266 | 0.363 | 0.28496 | 0.42044 | 0.37 | 0.35 | 0.44 | 0.348 | 0.3016 | 0.25 | 0.38047 | 0.36 | 0.347 | 0.315 | 0.37 |  | #N/A
| 1 | 3 | 0.37 | 0.37 | 0.35 | 0.34 | 0.35 | 0.35 | 0.35 | 0.3269 | 0.37 | 0.2858 | 0.41995 | 0.37 | 0.34 | 0.47 | 0.331 | 0.2944 | 0.25 | 0.38489 | 0.36 | 0.3429 | 0.315 | 0.38 |  | 16
| 1 | 4 | 0.36 | 0.36 | 0.35 | 0.34 | 0.36 | 0.35 | 0.35 | 0.3237 | 0.369 | 0.28531 | 0.42649 | 0.37 | 0.34 | 0.44 | 0.329 | 0.288 | 0.24 | 0.38525 | 0.36 | 0.339 | 0.32 | 0.37 |  | 0.34935
| 1 | 5 | 0.35 | 0.37 | 0.35 | 0.35 | 0.36 | 0.35 | 0.35 | 0.3201 | 0.36 | 0.28872 | 0.4341 | 0.37 | 0.35 | 0.42 | 0.335 | 0.3048 | 0.25 | 0.38035 | 0.36 | 0.3553 | 0.32 | 0.37 |  | 128
| 1 | 6 | 0.36 | 0.36 | 0.35 | 0.36 | 0.37 | 0.34 | 0.35 | 0.3213 | 0.358 | 0.28805 | 0.42491 | 0.37 | 0.34 | 0.42 | 0.325 | 0.2976 | 0.24 | 0.38122 | 0.36 | 0.3442 | 0.315 | 0.37
| Mean |  | 0.36 | 0.36833 | 0.35 | 0.34833 | 0.36167 | 0.345 | 0.35167 | 0.32352 | 0.36317 | 0.28615 | 0.42763 | 0.36667 | 0.345 | 0.435 | 0.33233 | 0.29773 | 0.245 | 0.38105 | 0.35833 | 0.34553 | 0.32 | 0.37
| Median |  | 0.36 | 0.37 | 0.35 | 0.35 | 0.36 | 0.345 | 0.35 | 0.3231 | 0.3615 | 0.28556 | 0.4257 | 0.37 | 0.345 | 0.43 | 0.33 | 0.2988 | 0.245 | 0.38084 | 0.36 | 0.3445 | 0.3175 | 0.37
| Std Dev. |  | 0.0063246 | 0.0075277 | 6.0809e-17 | 0.0075277 | 0.0075277 | 0.0054772 | 0.0040825 | 0.0027788 | 0.0051929 | 0.0018325 | 0.0078861 | 0.008165 | 0.0054772 | 0.019748 | 0.0084774 | 0.0059312 | 0.0054772 | 0.0040377 | 0.0040825 | 0.0054676 | 0.007746 | 0.0063246
| Rel.Std.Dev. |  | 1.757% | 2.044% | 1.737e-14% | 2.161% | 2.081% | 1.588% | 1.161% | 0.8589% | 1.43% | 0.6404% | 1.844% | 2.227% | 1.588% | 4.54% | 2.551% | 1.992% | 2.236% | 1.06% | 1.139% | 1.582% | 2.421% | 1.709%
| PDM3 |  | 3.048% | 5.433% | 0.1856% | -0.2915% | 3.525% | -1.246% | 0.6626% | -7.395% | 3.954% | -18.09% | 22.41% | 4.956% | -1.246% | 24.52% | -4.871% | -14.78% | -29.87% | 9.073% | 2.571% | -1.093% | -8.402% | 5.91%
| Z-Score (Absolute) |  | 0.37 | 0.71 | 0.03 | 0.1 | 0.44 | 0.24 | 0.03 | 1.11 | 0.5 | 2.62 | 3.11 | 0.64 | 0.24 | 3.41 | 0.75 | 2.15 | 4.28 | 1.22 | 0.3 | 0.21 | 1.25 | 0.78
| Analytical results for Sb in OREAS 702 (Certified Value 0.37 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.3 | 0.4 | 0.38 | 0.38 | 0.39 | 0.4 | < 20 | 0.45 | 0.4 | < 5 | < 0.5 | 11.35 | 0.36 | < 10 | < 1 | 0.28 | < 2.5 | 1 | < 0.8 | < 0.5 | < 0.5 |  |  | 1
| 1 | 2 | 0.29 | 0.35 | 0.39 | 0.35 | 0.36 | 0.3 | < 20 | 0.57 | 0.4 | < 5 | < 0.5 | 11.87 | 0.98 | < 10 | < 1 | 0.31 | < 2.5 | 0.7 | < 0.8 | < 0.5 | < 0.5 |  |  | 3
| 1 | 3 | 0.32 | 0.34 | 0.35 | 0.38 | 0.34 | 0.4 | < 20 | 0.45 | 0.4 | < 5 | < 0.5 | 11.28 | 0.6 | < 10 | < 1 | 0.33 | < 2.5 | 0.3 | < 0.8 | < 0.5 | < 0.5 |  |  | 16
| 1 | 4 | 0.32 | 0.36 | 0.4 | 0.34 | 0.36 | 0.4 | < 20 | 0.46 | 0.4 | < 5 | < 0.5 | 11.22 | 0.43 | < 10 | < 1 | 0.3 | < 2.5 | 0.5 | < 0.8 | < 0.5 | < 0.5 |  |  | 0.36513
| 1 | 5 | 0.29 | 0.35 | 0.39 | 0.35 | 0.35 | 0.4 | < 20 | 0.42 | 0.4 | < 5 | < 0.5 | 11.14 | 0.46 | < 10 | < 1 | 0.29 | < 2.5 | 0.6 | < 0.8 | < 0.5 | < 0.5 |  |  | 129
| 1 | 6 | 0.31 | 0.37 | 0.36 | 0.37 | 0.36 | 0.4 | < 20 | 0.45 | 0.4 | < 5 | < 0.5 | 11.82 | 0.39 | < 10 | < 1 | 0.28 | < 2.5 | 0.6 | < 0.8 | < 0.5 | < 0.5
| Mean |  | 0.305 | 0.36167 | 0.37833 | 0.36167 | 0.36 | 0.38333 |  | 0.46667 | 0.4 |  |  | 11.4467 | 0.53667 |  |  | 0.29833 |  | 0.61667
| Median |  | 0.305 | 0.355 | 0.385 | 0.36 | 0.36 | 0.4 |  | 0.45 | 0.4 |  |  | 11.315 | 0.445 |  |  | 0.295 |  | 0.6
| Std Dev. |  | 0.013784 | 0.02137 | 0.019408 | 0.017224 | 0.016733 | 0.040825 |  | 0.052409 | 6.0809e-17 |  |  | 0.31659 | 0.23261 |  |  | 0.019408 |  | 0.23166
| Rel.Std.Dev. |  | 4.519% | 5.909% | 5.13% | 4.762% | 4.648% | 10.65% |  | 11.23% | 1.52e-14% |  |  | 2.766% | 43.34% |  |  | 6.505% |  | 37.57%
| PDM3 |  | -16.47% | -0.9472% | 3.617% | -0.9472% | -1.404% | 4.987% |  | 27.81% | 9.552% |  |  | 3035% | 46.98% |  |  | -18.29% |  | 68.89%
| Z-Score (Absolute) |  | 0.45 | 0.15 | 0.06 | 0.15 | 0.16 | NA | 51.85 | 0.42 | NA | 11.39 | 0.75 | 59.65 | 0.8 | 24.87 | 0.6 | 0.49 | 4.64 | NA | 0.06 | 0.75 | 0.75
Results from laboratories 6, 10 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 702 (Certified Value 5.62 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 24 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.4 | 5.4 | 6.2 | 5.2 | 6.3 | 6 | 7 | 4.99 | 6 | 5.71 | 5.8776 | 6 | 5.94 | 7.3 | 5.3218 | 5.8 | 5.7882 | 7.1 | 6.156 | 6 | 5.2 |  |  | 1
| 1 | 2 | 5.3 | 5.5 | 6.1 | 5.3 | 5.9 | 6.1 | 7 | 5.06 | 6 | 5.6402 | 5.7847 | 6 | 5.8 | 6.8 | 4.9738 | 5.7 | 5.8314 | 7.1 | 5.793 | 6 | 5.2 |  |  | 4
| 1 | 3 | 5.7 | 5.5 | 6.1 | 5.7 | 5.9 | 5.8 | 7 | 5.07 | 6 | 5.5904 | 5.77 | 6 | 5.61 | 6.5 | 5.4878 | 5.8 | 5.8889 | 7.1 | 5.673 | 6 | 5.3 |  |  | 16
| 1 | 4 | 5.6 | 5.6 | 6.1 | 5.4 | 6 | 5.7 | 7 | 5.06 | 6 | 5.5405 | 5.6333 | 6 | 5.69 | 7.1 | 4.723 | 5.7 | 5.8164 | 7 | 5.641 | 5 | 5.2 |  |  | 5.6188
| 1 | 5 | 5.4 | 5.6 | 6.1 | 5.6 | 5.9 | 5.6 | 7 | 4.97 | 6 | 5.4807 | 5.7567 | 6 | 5.69 | 6.9 | 5.2828 | 5.8 | 5.8271 | 7.2 | 5.395 | 6 | 5.2 |  |  | 130
| 1 | 6 | 5.5 | 5.5 | 5.9 | 5.5 | 6.1 | 6.1 | 7 | 5.01 | 6 | 5.5106 | 5.74 | 6 | 5.62 | 7 | 5.4749 | 5.7 | 5.8091 | 7 | 5.86 | 5 | 5.2
| Mean |  | 5.4833 | 5.5167 | 6.0833 | 5.45 | 6.0167 | 5.8833 | 7 | 5.0267 | 6 | 5.5788 | 5.7604 | 6 | 5.725 | 6.9333 | 5.2107 | 5.75 | 5.8269 | 7.0833 | 5.753 | 5.6667 | 5.2167
| Median |  | 5.45 | 5.5 | 6.1 | 5.45 | 5.95 | 5.9 | 7 | 5.035 | 6 | 5.5655 | 5.7633 | 6 | 5.69 | 6.95 | 5.3023 | 5.75 | 5.8217 | 7.1 | 5.733 | 6 | 5.2
| Std Dev. |  | 0.1472 | 0.075277 | 0.098319 | 0.18708 | 0.16021 | 0.2137 | 0 | 0.042269 | 0 | 0.085833 | 0.078724 | 0 | 0.12534 | 0.27325 | 0.30267 | 0.054772 | 0.034023 | 0.075277 | 0.25399 | 0.5164 | 0.040825
| Rel.Std.Dev. |  | 2.684% | 1.365% | 1.616% | 3.433% | 2.663% | 3.632% | 0% | 0.8409% | 0% | 1.539% | 1.367% | 0% | 2.189% | 3.941% | 5.809% | 0.9526% | 0.5839% | 1.063% | 4.415% | 9.113% | 0.7826%
| PDM3 |  | -2.41% | -1.817% | 8.268% | -3.003% | 7.082% | 4.709% | 24.58% | -10.54% | 6.785% | -0.712% | 2.52% | 6.785% | 1.891% | 23.4% | -7.262% | 2.336% | 3.704% | 26.07% | 2.389% | 0.8527% | -7.156%
| Z-Score (Absolute) |  | 0.67 | 0.59 | 0.84 | 0.76 | 0.67 | 0.34 | NA | 1.83 | NA | 0.43 | 0.03 | NA | 0.06 | 2.99 | 1.36 | 0 | 0.19 | 3.37 | 0.01 | NA | 1.35
Results from laboratories 7, 10, 13 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 702 (Indicative Value 1.03 ppm) | [24 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 2 | < 1 | < 1 | < 1 | 1 | < 0.3 | < 5 | 0.47 | < 1 | < 5 | 1.8 | < 2 | < 10 | < 10 | < 1 | < 1.5 | 12.9171 | < 10 | < 1 |  |  |  |  | 1
| 1 | 2 | 1 | < 1 | < 1 | < 1 | 1 | < 0.3 | < 5 | 0.56 | < 1 | < 5 | 1.85 | < 2 | < 10 | < 10 | < 1 | < 1.5 | 9.6446 | < 10 | < 1 |  |  |  |  | 11
| 1 | 3 | 1 | 1 | < 1 | < 1 | 1 | < 0.3 | < 5 | 0.58 | < 1 | < 5 | 1.88 | < 2 | < 10 | < 10 | < 1 | < 1.5 | 9.8158 | < 10 | < 1 |  |  |  |  | 16
| 1 | 4 | 1 | 1 | < 1 | < 1 | 1 | < 0.3 | < 5 | 0.48 | < 1 | < 5 | 1.94 | < 2 | < 10 | < 10 | < 1 | < 1.5 | 10.6231 | < 10 | < 1 |  |  |  |  | 1.0333
| 1 | 5 | 2 | 1 | 1 | < 1 | 1 | < 0.3 | < 5 | 0.56 | < 1 | < 5 | 1.86 | < 2 | < 10 | < 10 | < 1 | < 1.5 | 11.3702 | < 10 | < 1 |  |  |  |  | 22
| 1 | 6 | 1 | 1 | 1 | < 1 | 1 | < 0.3 | < 5 | 0.54 | < 1 | < 5 | 1.68 | < 2 | < 10 | < 10 | < 1 | < 1.5 | 11.3148 | < 10 | < 1
| Mean |  | 1.3333 | 1 | 1 |  | 1 |  |  | 0.53167 |  |  | 1.835 | [5 empty] | 10.9476
| Median |  | 1 | 1 | 1 |  | 1 |  |  | 0.55 |  |  | 1.855 | [5 empty] | 10.9689
| Std Dev. |  | 0.5164 | 0 | 0 |  | 0 |  |  | 0.045789 |  |  | 0.088487 | [5 empty] | 1.2064
| Rel.Std.Dev. |  | 38.73% | 0% | 0% |  | 0% |  |  | 8.612% |  |  | 4.822% | [5 empty] | 11.02%
| PDM3 |  | 29.03% | -3.226% | -3.226% |  | -3.226% |  |  | -48.55% |  |  | 77.58% | [5 empty] | 959.4%
| Z-Score (Absolute) |  | 0.45 | 0.22 | 0.45 | 0.67 | 0 | 1.15 | 2.02 | 0.63 | 0.67 | 2.02 | 1.13 | 0 | 5.39 | 5.39 | 0.67 | 0.34 | 13.42 | 5.39 | 0.67
| Analytical results for Sm in OREAS 702 (Certified Value 5.17 ppm) | [24 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 5.09 | 5.9 | 4.04 | 6.0631 | 5.58 | 5.4 | 4.7738 | 4.607 | [15 empty] | 1
| 1 | 2 | 5.3 | 6.2 | 4.27 | 5.9725 | 5.4 | 5.47 | 5.1089 | 4.312 | [15 empty] | #N/A
| 1 | 3 | 5.58 | 5.8 | 4.38 | 6.0583 | 5.45 | 5.36 | 4.8067 | 4.342 | [15 empty] | 16
| 1 | 4 | 5.19 | 6.2 | 4.3 | 5.9504 | 5.36 | 5.51 | 4.8189 | 4.221 | [15 empty] | 5.1734
| 1 | 5 | 4.55 | 5.6 | 4.13 | 6.0222 | 5.23 | 5.5 | 5.084 | 4.04 | [15 empty] | 131
| 1 | 6 | 5.37 | 5.7 | 3.96 | 6.0864 | 5.48 | 5.48 | 4.9197 | 4.356
| Mean |  | 5.18 | 5.9 | 4.18 | 6.0255 | 5.4167 | 5.4533 | 4.9187 | 4.313
| Median |  | 5.245 | 5.85 | 4.2 | 6.0403 | 5.425 | 5.475 | 4.8693 | 4.327
| Std Dev. |  | 0.35088 | 0.25298 | 0.1631 | 0.054157 | 0.11843 | 0.059889 | 0.14628 | 0.18554
| Rel.Std.Dev. |  | 6.774% | 4.288% | 3.902% | 0.8988% | 2.186% | 1.098% | 2.974% | 4.302%
| PDM3 |  | 0.1277% | 14.05% | -19.2% | 16.47% | 4.702% | 5.411% | -4.924% | -16.63%
| Z-Score (Absolute) |  | 0.16 | 0.83 | 1.54 | 1 | 0.16 | 0.21 | 0.52 | 1.35
| Analytical results for Sn in OREAS 702 (Certified Value 4.1 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 4.1 | 4.1 | 4.7 | 3.9 | 4.2 | 4.3 | 4 | 4.04 | 4.2 | 4.0491 | 3.7 | 4.1 | 4.7 | < 20 | < 10 | 4 | 4.0858 | 4 | 3.0012 | 4 | 4.5 | 4.1 |  | 1
| 1 | 2 | 3.9 | 4.2 | 4.6 | 4 | 4 | 4 | 4 | 4.26 | 4.1 | 3.9505 | 3.9 | 4.2 | 4.8 | < 20 | < 10 | 3.9 | 4.09 | 4 | 2.8095 | 5 | 4.5 | 4.2 |  | #N/A
| 1 | 3 | 4.2 | 4.1 | 4.6 | 4.2 | 4.1 | 4.2 | 4 | 4.1 | 4.2 | 4.1224 | 3.8 | 4.3 | 4.9 | < 20 | < 10 | 4.2 | 3.9767 | 4 | 2.6392 | 4 | 4.5 | 4.3 |  | 16
| 1 | 4 | 4.2 | 4.2 | 4.5 | 4.1 | 4.1 | 4.2 | 4 | 4.15 | 4.1 | 3.77 | 3.8 | 4.3 | 4.5 | < 20 | < 10 | 4 | 4.2548 | 4.1 | 2.8804 | 4 | 4.5 | 4.2 |  | 4.1031
| 1 | 5 | 4.1 | 4.2 | 4.6 | 4.5 | 4.1 | 4.1 | 4 | 4 | 4.1 | 4.04 | 3.8 | 4.2 | 4.9 | < 20 | < 10 | 4 | 4.1452 | 4.1 | 2.448 | 4 | 4.5 | 4.2 |  | 132
| 1 | 6 | 4 | 4.1 | 4.3 | 4 | 4.4 | 4.2 | 5 | 3.88 | 4.2 | 3.97 | 3.7 | 4.2 | 4.6 | < 20 | < 10 | 4.1 | 4.1158 | 4 | 2.4879 | 4 | 4.5 | 4.2
| Mean |  | 4.0833 | 4.15 | 4.55 | 4.1167 | 4.15 | 4.1667 | 4.1667 | 4.0717 | 4.15 | 3.9837 | 3.7833 | 4.2167 | 4.7333 |  |  | 4.0333 | 4.1114 | 4.0333 | 2.711 | 4.1667 | 4.5 | 4.2
| Median |  | 4.1 | 4.15 | 4.6 | 4.05 | 4.1 | 4.2 | 4 | 4.07 | 4.15 | 4.005 | 3.8 | 4.2 | 4.75 |  |  | 4 | 4.1029 | 4 | 2.7244 | 4 | 4.5 | 4.2
| Std Dev. |  | 0.1169 | 0.054772 | 0.13784 | 0.2137 | 0.13784 | 0.10328 | 0.40825 | 0.1306 | 0.054772 | 0.12137 | 0.075277 | 0.075277 | 0.1633 |  |  | 0.10328 | 0.090505 | 0.05164 | 0.22213 | 0.40825 | 0 | 0.063246
| Rel.Std.Dev. |  | 2.863% | 1.32% | 3.029% | 5.191% | 3.321% | 2.479% | 9.798% | 3.208% | 1.32% | 3.047% | 1.99% | 1.785% | 3.45% |  |  | 2.561% | 2.201% | 1.28% | 8.194% | 9.798% | 0% | 1.506%
| PDM3 |  | -0.4813% | 1.143% | 10.89% | 0.3311% | 1.143% | 1.55% | 1.55% | -0.7657% | 1.143% | -2.91% | -7.793% | 2.768% | 15.36% |  |  | -1.7% | 0.2023% | -1.7% | -33.93% | 1.55% | 9.674% | 2.362%
| Z-Score (Absolute) |  | 0.46 | 0 | 2.77 | 0.23 | 0 | 0.12 | NA | 0.54 | 0 | 1.15 | 2.54 | 0.46 | 4.03 | 40.46 | 5.88 | 0.81 | 0.27 | 0.81 | 9.95 | NA | 2.42 | 0.35
Results from laboratories 7 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 702 (Certified Value 198 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 208 | 203 | 205 | 198 | 211 | 206 | 197 | 190 | 193 | 205.9278 | 182.1951 | 205 | 203 | 216.3 | 174.955 | 184.9645 | 198.1 | 204.2937 | 190 | 182.9063 | 181 | 200.9 |  | 1
| 1 | 2 | 205 | 200 | 202 | 206 | 197.5 | 203 | 198 | 193 | 193 | 202.5479 | 185.6632 | 210 | 206 | 222.2 | 177.23 | 189.7069 | 194.9 | 205.3233 | 192 | 179.8382 | 188 | 204.4 |  | #N/A
| 1 | 3 | 217 | 198.5 | 199.5 | 200 | 197.5 | 198 | 193 | 193 | 192 | 201.2464 | 180.014 | 209 | 202 | 237.2 | 178.312 | 189.0848 | 199.6 | 207.329 | 190 | 176.7457 | 189 | 208.2 |  | 16
| 1 | 4 | 205 | 199 | 207 | 201 | 197.5 | 209 | 196 | 192 | 194 | 203.8192 | 184.415 | 208 | 206 | 221.7 | 179.397 | 185.955 | 194.2 | 206.3096 | 191 | 187.5047 | 186 | 203.9 |  | 197.5478
| 1 | 5 | 198 | 199.5 | 205 | 206 | 202 | 193 | 195 | 190 | 194 | 198.9764 | 183.15 | 207 | 204 | 212.3 | 180.74 | 185.9495 | 196.8 | 205.3778 | 194 | 187.157 | 190 | 204.3 |  | 133
| 1 | 6 | 207 | 197.5 | 199.5 | 209 | 207 | 200 | 196 | 190 | 195 | 201.5188 | 180.1025 | 208 | 204 | 215.6 | 179.242 | 190.8544 | 198 | 203.1195 | 193 | 186.6374 | 190 | 206.8
| Mean |  | 206.6667 | 199.5833 | 203 | 203.3333 | 202.0833 | 201.5 | 195.8333 | 191.3333 | 193.5 | 202.3394 | 182.59 | 207.8333 | 204.1667 | 220.8833 | 178.3127 | 187.7525 | 196.9333 | 205.2922 | 191.6667 | 183.4649 | 187.3333 | 204.75
| Median |  | 206 | 199.25 | 203.5 | 203.5 | 199.75 | 201.5 | 196 | 191 | 193.5 | 202.0334 | 182.6725 | 208 | 204 | 219 | 178.777 | 187.5199 | 197.4 | 205.3506 | 191.5 | 184.7719 | 188.5 | 204.35
| Std Dev. |  | 6.1536 | 1.8819 | 3.1464 | 4.274 | 5.7742 | 5.7533 | 1.7224 | 1.5055 | 1.0488 | 2.3786 | 2.2833 | 1.7224 | 1.6021 | 8.8457 | 2.0188 | 2.4278 | 2.0607 | 1.4771 | 1.633 | 4.4414 | 3.4448 | 2.5288
| Rel.Std.Dev. |  | 2.978% | 0.9429% | 1.55% | 2.102% | 2.857% | 2.855% | 0.8795% | 0.7869% | 0.542% | 1.176% | 1.251% | 0.8287% | 0.7847% | 4.005% | 1.132% | 1.293% | 1.046% | 0.7195% | 0.852% | 2.421% | 1.839% | 1.235%
| PDM3 |  | 4.616% | 1.03% | 2.76% | 2.929% | 2.296% | 2.001% | -0.8679% | -3.146% | -2.049% | 2.426% | -7.572% | 5.207% | 3.351% | 11.81% | -9.737% | -4.958% | -0.311% | 3.92% | -2.977% | -7.129% | -5.171% | 3.646%
| Z-Score (Absolute) |  | 0.76 | 0.12 | 0.3 | 0.35 | 0.19 | 0.12 | 0.58 | 1.14 | 0.87 | 0.22 | 2.23 | 0.9 | 0.45 | 2.52 | 2.76 | 1.59 | 0.45 | 0.59 | 1.1 | 2.12 | 1.64 | 0.52
| Analytical results for Ta in OREAS 702 (Certified Value 0.72 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 24 | 25 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.68 | 0.77 | 0.83 | 0.77 | 0.77 | 0.82 | 0.8 | 0.59 | 0.65 | 0.7 | 0.56 | 0.33 | < 5 | 0.26 | 0.70875 | 0.5 | 1 | 0.77 | [5 empty] | 1
| 1 | 2 | 0.69 | 0.75 | 0.81 | 0.77 | 0.7 | 0.74 | 0.8 | 0.61 | 0.67 | 0.7 | 0.7 | 0.64 | < 5 | 0.27 | 0.72953 | 0.5 | 1 | 0.77 | [5 empty] | #N/A
| 1 | 3 | 0.72 | 0.77 | 0.8 | 0.8 | 0.74 | 0.76 | 0.8 | 0.61 | 0.63 | 0.7 | 0.6 | 0.57 | < 5 | 0.32 | 0.70321 | 0.6 | 1 | 0.79 | [5 empty] | 16
| 1 | 4 | 0.73 | 0.77 | 0.85 | 0.74 | 0.73 | 0.79 | 0.8 | 0.58 | 0.65 | 0.6 | 0.64 | 0.25 | < 5 | 0.29 | 0.75726 | 0.5 | 1 | 0.77 | [5 empty] | 0.72409
| 1 | 5 | 0.7 | 0.75 | 0.82 | 0.8 | 0.72 | 0.79 | 0.7 | 0.67 | 0.68 | 0.6 | 0.72 | 0.46 | < 5 | 0.25 | 0.70818 | 0.5 | 1 | 0.8 | [5 empty] | 134
| 1 | 6 | 0.74 | 0.74 | 0.77 | 0.76 | 0.74 | 0.84 | 0.7 | 0.59 | 0.65 | 0.7 | 0.61 | 0.23 | < 5 | 0.26 | 0.65879 | 0.4 | 1 | 0.77
| Mean |  | 0.71 | 0.75833 | 0.81333 | 0.77333 | 0.73333 | 0.79 | 0.76667 | 0.60833 | 0.655 | 0.66667 | 0.63833 | 0.41333 |  | 0.275 | 0.71095 | 0.5 | 1 | 0.77833
| Median |  | 0.71 | 0.76 | 0.815 | 0.77 | 0.735 | 0.79 | 0.8 | 0.6 | 0.65 | 0.7 | 0.625 | 0.395 |  | 0.265 | 0.70846 | 0.5 | 1 | 0.77
| Std Dev. |  | 0.023664 | 0.013292 | 0.027325 | 0.023381 | 0.023381 | 0.036878 | 0.05164 | 0.032506 | 0.017607 | 0.05164 | 0.061455 | 0.17049 |  | 0.025884 | 0.032507 | 0.063246 | 0 | 0.013292
| Rel.Std.Dev. |  | 3.333% | 1.753% | 3.36% | 3.023% | 3.188% | 4.668% | 6.736% | 5.344% | 2.688% | 7.746% | 9.627% | 41.25% |  | 9.412% | 4.572% | 12.65% | 0% | 1.708%
| PDM3 |  | -1.945% | 4.73% | 12.33% | 6.801% | 1.277% | 9.103% | 5.881% | -15.99% | -9.541% | -7.93% | -11.84% | -42.92% |  | -62.02% | -1.814% | -30.95% | 38.11% | 7.492%
| Z-Score (Absolute) |  | 0.12 | 0.36 | 0.91 | 0.51 | 0.11 | 0.68 | NA | 1.14 | 0.67 | NA | 0.84 | 3.08 | 17.76 | 4.47 | 0.11 | NA | NA | 0.56
Results from laboratories 7, 13, 21 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 702 (Certified Value 0.66 ppm) | [24 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 15 | 18 | 19 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 0.72 | 0.7 | 0.3 | 0.72573 | 0.6 | 0.69 | 0.69 | 0.56397 | 0.669 | [14 empty] | 1
| 1 | 2 | 0.7 | 0.69 | 0.4 | 0.7294 | 0.61 | 0.8 | 0.68 | 0.60789 | 0.626 | [14 empty] | #N/A
| 1 | 3 | 0.65 | 0.67 | 0.4 | 0.7148 | 0.62 | 0.72 | 0.67 | 0.59856 | 0.6 | [14 empty] | 16
| 1 | 4 | 0.65 | 0.65 | 0.4 | 0.73307 | 0.62 | 0.66 | 0.71 | 0.61364 | 0.626 | [14 empty] | 0.65995
| 1 | 5 | 0.63 | 0.62 | 0.3 | 0.7155 | 0.63 | 0.66 | 0.69 | 0.60697 | 0.554 | [14 empty] | 135
| 1 | 6 | 0.69 | 0.64 | 0.3 | 0.71594 | 0.66 | 0.63 | 0.7 | 0.60205 | 0.625
| Mean |  | 0.67333 | 0.66167 | 0.35 | 0.72241 | 0.62333 | 0.69333 | 0.69 | 0.59885 | 0.61667
| Median |  | 0.67 | 0.66 | 0.35 | 0.72083 | 0.62 | 0.675 | 0.69 | 0.60451 | 0.6255
| Std Dev. |  | 0.035024 | 0.030605 | 0.054772 | 0.0080125 | 0.020656 | 0.060553 | 0.014142 | 0.017852 | 0.037914
| Rel.Std.Dev. |  | 5.202% | 4.625% | 15.65% | 1.109% | 3.314% | 8.734% | 2.05% | 2.981% | 6.148%
| PDM3 |  | 2.028% | 0.2604% | -46.97% | 9.464% | -5.548% | 5.059% | 4.554% | -9.258% | -6.558%
| Z-Score (Absolute) |  | 0.11 | 0.11 | NA | 1.06 | 0.85 | 0.5 | 0.43 | 1.32 | 0.98
Results from laboratory 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 702 (Certified Value 1.13 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.08 | 1.18 | 1.25 | 1.2 | 1.16 | 1 | < 20 | 1.04 | 1.1 | 0.12 | 0.92 | < 20 | < 1 | 1.06 | < 5 | 1.5 | 1.06 | [6 empty] | 1
| 1 | 2 | 1.01 | 1.15 | 1.2 | 1.19 | 1.14 | 1.42 | < 20 | 1.13 | 1.1 | 0.19 | 0.9 | < 20 | < 1 | 0.94 | < 5 | 1.3 | 1.09 | [6 empty] | 12
| 1 | 3 | 1.2 | 1.22 | 1.14 | 1.36 | 1.13 | 1.23 | < 20 | 1.03 | 1.1 | 0.16 | 1.02 | < 20 | < 1 | 1.03 | < 5 | 1.8 | 1.12 | [6 empty] | 16
| 1 | 4 | 1.25 | 1.25 | 1.26 | 1.26 | 1.19 | 1.4 | < 20 | 1.13 | 1.1 | 0.17 | 0.96 | < 20 | < 1 | 0.95 | < 5 | 1.3 | 1.06 | [6 empty] | 1.1282
| 1 | 5 | 1.06 | 1.2 | 1.3 | 1.36 | 1.08 | 1.2 | < 20 | 1.03 | 1.1 | 0.19 | 0.8 | < 20 | < 1 | 1.26 | < 5 | 1.6 | 1.07 | [6 empty] | 136
| 1 | 6 | 1.12 | 1.21 | 1.14 | 1.26 | 1.21 | 1.02 | < 20 | 0.98 | 1.1 | 0.21 | 0.91 | < 20 | < 1 | 1.04 | < 5 | 1 | 1.13
| Mean |  | 1.12 | 1.2017 | 1.215 | 1.2717 | 1.1517 | 1.2117 |  | 1.0567 | 1.1 | 0.17333 | 0.91833 |  |  | 1.0467 |  | 1.4167 | 1.0883
| Median |  | 1.1 | 1.205 | 1.225 | 1.26 | 1.15 | 1.215 |  | 1.035 | 1.1 | 0.18 | 0.915 |  |  | 1.035 |  | 1.4 | 1.08
| Std Dev. |  | 0.090111 | 0.034303 | 0.066257 | 0.074409 | 0.046224 | 0.17938 |  | 0.060553 | 2.4324e-16 | 0.031411 | 0.072778 |  |  | 0.11553 |  | 0.27869 | 0.030605
| Rel.Std.Dev. |  | 8.046% | 2.855% | 5.453% | 5.851% | 4.014% | 14.8% |  | 5.731% | 2.211e-14% | 18.12% | 7.925% |  |  | 11.04% |  | 19.67% | 2.812%
| PDM3 |  | -0.7239% | 6.515% | 7.697% | 12.72% | 2.083% | 7.401% |  | -6.338% | -2.497% | -84.64% | -18.6% |  |  | -7.224% |  | 25.57% | -3.531%
| Z-Score (Absolute) |  | 0.2 | 0.32 | 0.41 | 0.77 | 0 | 0.39 | 56.82 | 0.61 | NA | 6.28 | 1.5 | 56.82 | 4.18 | 0.67 | 8.66 | NA | 0.41
Results from laboratories 10 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 702 (Certified Value 10 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 10.45 | 10.15 | 11.05 | 9.77 | 9.83 | 10.4 | 10.6 | 5.65 | 10 | 19.064 | 10.2 | 9.56 | 13.3 | 9.4623 | 9.75 | 9.5067 | 12.6 | 8.996 | 9.5 |  |  |  |  | 1
| 1 | 2 | 9.05 | 9.98 | 10.8 | 10.4 | 9.38 | 10.5 | 11.1 | 5.96 | 9.85 | 19.33 | 11 | 9.87 | 13.9 | 9.8507 | 9.11 | 10.0378 | 12.5 | 9.319 | 9.8 |  |  |  |  | 1
| 1 | 3 | 10.15 | 9.73 | 10.65 | 10.85 | 9.49 | 10.4 | 10.6 | 6.18 | 9.56 | 19.598 | 10.1 | 9.7 | 13 | 10.1243 | 10.08 | 9.7401 | 12.4 | 8.113 | 9.9 |  |  |  |  | 16
| 1 | 4 | 10.2 | 10.25 | 10.95 | 9.58 | 9.97 | 10.2 | 10.5 | 6.22 | 9.7 | 19.48 | 9.6 | 9.8 | 13.1 | 10.2498 | 9.1 | 9.8674 | 12.4 | 7.833 | 9.8 |  |  |  |  | 10.0105
| 1 | 5 | 9.35 | 9.9 | 10.7 | 10.35 | 9.29 | 9.94 | 10.4 | 5.82 | 9.76 | 18.58 | 10.2 | 9.19 | 12.4 | 9.2292 | 10.75 | 9.7634 | 12.8 | 8 | 9.4 |  |  |  |  | 137
| 1 | 6 | 10.25 | 10.3 | 10.7 | 10.2 | 9.83 | 10.3 | 10.4 | 5.47 | 9.91 | 18.568 | 10.1 | 9.51 | 11.9 | 9.6999 | 10.82 | 9.6393 | 12.5 | 8.591 | 9.2
| Mean |  | 9.9083 | 10.0517 | 10.8083 | 10.1917 | 9.6317 | 10.29 | 10.6 | 5.8833 | 9.7967 | 19.1033 | 10.2 | 9.605 | 12.9333 | 9.7694 | 9.935 | 9.7591 | 12.5333 | 8.4753 | 9.6
| Median |  | 10.175 | 10.065 | 10.75 | 10.275 | 9.66 | 10.35 | 10.55 | 5.89 | 9.805 | 19.197 | 10.15 | 9.63 | 13.05 | 9.7753 | 9.915 | 9.7517 | 12.5 | 8.352 | 9.65
| Std Dev. |  | 0.56605 | 0.22013 | 0.15943 | 0.4591 | 0.28046 | 0.1995 | 0.26077 | 0.29548 | 0.15731 | 0.4472 | 0.45166 | 0.24517 | 0.70048 | 0.38844 | 0.75902 | 0.18324 | 0.15055 | 0.59432 | 0.27568
| Rel.Std.Dev. |  | 5.713% | 2.19% | 1.475% | 4.505% | 2.912% | 1.939% | 2.46% | 5.022% | 1.606% | 2.341% | 4.428% | 2.553% | 5.416% | 3.976% | 7.64% | 1.878% | 1.201% | 7.012% | 2.872%
| PDM3 |  | -1.02% | 0.4114% | 7.97% | 1.81% | -3.784% | 2.792% | 5.889% | -41.23% | -2.136% | 90.83% | 1.893% | -4.051% | 29.2% | -2.409% | -0.7541% | -2.511% | 25.2% | -15.34% | -4.101%
| Z-Score (Absolute) |  | 0.05 | 0.24 | 1.78 | 0.52 | 0.62 | 0.73 | 1.36 | 8.28 | 0.28 | 18.73 | 0.54 | 0.67 | 6.13 | 0.34 | 0 | 0.36 | 5.31 | 2.98 | 0.68
| Analytical results for Ti in OREAS 702 (Certified Value 0.23 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.235 | 0.235 | 0.225 | 0.224 | 0.238 | 0.24 | 0.2337 | 0.2121 | 0.251 | 0.23277 | 0.21902 | 0.237 | 0.23 | 0.24 | 0.238 | 0.237 | 0.22863 | 0.228 | 0.2248 | 0.235 | 0.23 |  |  | 1
| 1 | 2 | 0.23 | 0.236 | 0.226 | 0.227 | 0.227 | 0.24 | 0.2352 | 0.2151 | 0.249 | 0.23245 | 0.20738 | 0.241 | 0.22 | 0.24 | 0.239 | 0.235 | 0.23413 | 0.229 | 0.217 | 0.233 | 0.23 |  |  | #N/A
| 1 | 3 | 0.243 | 0.233 | 0.224 | 0.224 | 0.225 | 0.24 | 0.2265 | 0.2144 | 0.253 | 0.23108 | 0.21605 | 0.241 | 0.22 | 0.27 | 0.241 | 0.239 | 0.23536 | 0.228 | 0.2204 | 0.234 | 0.24 |  |  | 16
| 1 | 4 | 0.234 | 0.234 | 0.226 | 0.223 | 0.23 | 0.24 | 0.231 | 0.2113 | 0.251 | 0.23213 | 0.2183 | 0.241 | 0.22 | 0.24 | 0.234 | 0.232 | 0.23071 | 0.229 | 0.2237 | 0.223 | 0.24 |  |  | 0.23015
| 1 | 5 | 0.223 | 0.233 | 0.227 | 0.229 | 0.23 | 0.24 | 0.2307 | 0.21 | 0.247 | 0.23192 | 0.21206 | 0.239 | 0.22 | 0.24 | 0.245 | 0.238 | 0.22922 | 0.229 | 0.2246 | 0.23 | 0.24 |  |  | 138
| 1 | 6 | 0.235 | 0.23 | 0.223 | 0.23 | 0.233 | 0.25 | 0.2299 | 0.2095 | 0.254 | 0.23174 | 0.20774 | 0.242 | 0.22 | 0.23 | 0.242 | 0.237 | 0.22715 | 0.229 | 0.2286 | 0.227 | 0.23
| Mean |  | 0.23333 | 0.2335 | 0.22517 | 0.22617 | 0.2305 | 0.24167 | 0.23117 | 0.21207 | 0.25083 | 0.23201 | 0.21342 | 0.24017 | 0.22167 | 0.24333 | 0.23983 | 0.23633 | 0.23087 | 0.22867 | 0.22318 | 0.23033 | 0.235
| Median |  | 0.2345 | 0.2335 | 0.2255 | 0.2255 | 0.23 | 0.24 | 0.23085 | 0.2117 | 0.251 | 0.23202 | 0.21406 | 0.241 | 0.22 | 0.24 | 0.24 | 0.237 | 0.22997 | 0.229 | 0.22415 | 0.2315 | 0.235
| Std Dev. |  | 0.0065929 | 0.0020736 | 0.001472 | 0.0029269 | 0.0045935 | 0.0040825 | 0.0030435 | 0.0022844 | 0.0025626 | 0.00058901 | 0.0051504 | 0.0018348 | 0.0040825 | 0.013663 | 0.0037639 | 0.0025033 | 0.0032345 | 0.0005164 | 0.0040052 | 0.0046332 | 0.0054772
| Rel.Std.Dev. |  | 2.826% | 0.8881% | 0.6537% | 1.294% | 1.993% | 1.689% | 1.317% | 1.077% | 1.022% | 0.2539% | 2.413% | 0.764% | 1.842% | 5.615% | 1.569% | 1.059% | 1.401% | 0.2258% | 1.795% | 2.012% | 2.331%
| PDM3 |  | 1.382% | 1.454% | -2.166% | -1.732% | 0.1509% | 5.003% | 0.4406% | -7.858% | 8.986% | 0.8087% | -7.269% | 4.351% | -3.687% | 5.727% | 4.206% | 2.685% | 0.3098% | -0.6456% | -3.028% | 0.07851% | 2.106%
| Z-Score (Absolute) |  | 0.28 | 0.3 | 0.78 | 0.65 | 0.09 | 1.37 | 0 | 2.49 | 2.57 | 0.11 | 2.32 | 1.17 | 1.24 | 1.59 | 1.13 | 0.67 | 0.04 | 0.33 | 1.04 | 0.11 | 0.5
| Analytical results for Tl in OREAS 702 (Certified Value 0.71 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 14 | 15 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.7 | 0.67 | 1.04 | 0.67 | 1.13 | 0.73 | < 20 | 0.662 | 0.68 | 0.72513 | 0.73 | 0.75 | < 1 | 0.56 | 0.61567 | 0.5 | 0.6 | 0.75 | 0.69 |  |  |  |  | 1
| 1 | 2 | 0.62 | 0.66 | 1.01 | 0.68 | 1.02 | 0.73 | < 20 | 0.718 | 0.69 | 0.73523 | 0.75 | 0.87 | < 1 | 0.56 | 0.6661 | 0.5 | 0.7 | 0.7 | 0.7 |  |  |  |  | 21
| 1 | 3 | 0.7 | 0.66 | 1.05 | 0.73 | 1.05 | 0.73 | < 20 | 0.694 | 0.66 | 0.73362 | 0.75 | 0.83 | < 1 | 0.57 | 0.67011 | 0.5 | 0.7 | 0.7 | 0.72 |  |  |  |  | 16
| 1 | 4 | 0.68 | 0.66 | 1.1 | 0.68 | 1.08 | 0.73 | < 20 | 0.68 | 0.71 | 0.72593 | 0.72 | 0.76 | < 1 | 0.58 | 0.66108 | 0.5 | 0.7 | 0.7 | 0.71 |  |  |  |  | 0.70667
| 1 | 5 | 0.66 | 0.67 | 1.06 | 0.73 | 1.03 | 0.72 | < 20 | 0.669 | 0.7 | 0.71073 | 0.77 | 0.75 | < 1 | 0.56 | 0.66272 | 0.6 | 0.8 | 0.75 | 0.71 |  |  |  |  | 139
| 1 | 6 | 0.69 | 0.67 | 1.04 | 0.67 | 1.06 | 0.8 | < 20 | 0.659 | 0.68 | 0.74582 | 0.78 | 0.77 | < 1 | 0.61 | 0.68286 | 0.5 | 0.7 | 0.75 | 0.71
| Mean |  | 0.675 | 0.665 | 1.05 | 0.69333 | 1.0617 | 0.74 |  | 0.68033 | 0.68667 | 0.72941 | 0.75 | 0.78833 |  | 0.57333 | 0.65976 | 0.51667 | 0.7 | 0.725 | 0.70667
| Median |  | 0.685 | 0.665 | 1.045 | 0.68 | 1.055 | 0.73 |  | 0.6745 | 0.685 | 0.72978 | 0.75 | 0.765 |  | 0.565 | 0.66441 | 0.5 | 0.7 | 0.725 | 0.71
| Std Dev. |  | 0.030822 | 0.0054772 | 0.029665 | 0.028752 | 0.039707 | 0.029665 |  | 0.022474 | 0.017512 | 0.011838 | 0.022804 | 0.049967 |  | 0.019664 | 0.022956 | 0.040825 | 0.063246 | 0.027386 | 0.010328
| Rel.Std.Dev. |  | 4.566% | 0.8236% | 2.825% | 4.147% | 3.74% | 4.009% |  | 3.303% | 2.55% | 1.623% | 3.04% | 6.338% |  | 3.43% | 3.479% | 7.902% | 9.035% | 3.777% | 1.462%
| PDM3 |  | -4.481% | -5.896% | 48.59% | -1.887% | 50.24% | 4.717% |  | -3.726% | -2.83% | 3.218% | 6.132% | 11.56% |  | -18.87% | -6.638% | -26.89% | -0.9432% | 2.595% | 0.0002017%
| Z-Score (Absolute) |  | 0.51 | 0.67 | 5.56 | 0.22 | 5.75 | 0.54 | 150.4 | 0.43 | 0.32 | 0.37 | 0.7 | 1.32 | 3.34 | 2.16 | 0.76 | NA | NA | 0.3 | 0
Results from laboratories 21 and 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 702 (Certified Value 0.17 ppm) | [24 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 18 | 19 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.19 | 0.2 | 0.15 | 0.16569 | 0.17 | 0.17 | 0.14381 | 0.14 | [15 empty] | 1
| 1 | 2 | 0.17 | 0.2 | 0.19 | 0.16322 | 0.2 | 0.18 | 0.15634 | 0.132 | [15 empty] | #N/A
| 1 | 3 | 0.18 | 0.2 | 0.16 | 0.16247 | 0.17 | 0.17 | 0.16071 | 0.129 | [15 empty] | 16
| 1 | 4 | 0.16 | 0.4 | 0.16 | 0.16297 | 0.17 | 0.16 | 0.15355 | 0.126 | [15 empty] | 0.16657
| 1 | 5 | 0.17 | 0.1 | 0.16 | 0.16513 | 0.2 | 0.17 | 0.15662 | 0.114 | [15 empty] | 140
| 1 | 6 | 0.15 | 0.2 | 0.15 | 0.16388 | 0.16 | 0.18 | 0.15231 | 0.142
| Mean |  | 0.17 | 0.21667 | 0.16167 | 0.16389 | 0.17833 | 0.17167 | 0.15389 | 0.1305
| Median |  | 0.17 | 0.2 | 0.16 | 0.16355 | 0.17 | 0.17 | 0.15494 | 0.1305
| Std Dev. |  | 0.014142 | 0.098319 | 0.01472 | 0.0012718 | 0.017224 | 0.0075277 | 0.0057327 | 0.010193
| Rel.Std.Dev. |  | 8.319% | 45.38% | 9.105% | 0.776% | 9.658% | 4.385% | 3.725% | 7.811%
| PDM3 |  | 2.056% | 30.07% | -2.947% | -1.609% | 7.059% | 3.057% | -7.616% | -21.66%
| Z-Score (Absolute) |  | 0.53 | NA | 0.19 | 0 | 1.25 | 0.67 | 0.87 | 2.9
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 702 (Certified Value 3.93 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 3.9 | 3.9 | 4.1 | 3.7 | 3.8 | 3.9 | 3.9 | 3.52 | 3.78 | 3.4975 | < 20 | 4 | 4.3 | 4 | 4.0859 | 3.6 | 3.621 | 4.7 | 3.646 | 4.1 |  |  |  | 1
| 1 | 2 | 3.3 | 3.3 | 4.5 | 3.8 | 3.4 | 4.3 | 4.8 | 3.74 | 3.81 | 3.4908 | < 20 | 4.5 | 4.2 | 4 | 4.213 | 4.9 | 4.1093 | 5.2 | 3.629 | 3.6 |  |  |  | #N/A
| 1 | 3 | 3.8 | 3.8 | 4.2 | 4 | 3.7 | 3.7 | 4.6 | 3.83 | 4.04 | 3.5265 | < 20 | 4 | 4.2 | 3.9 | 4.1208 | 4 | 3.9109 | 4.6 | 3.439 | 4.1 |  |  |  | 16
| 1 | 4 | 4.1 | 3.7 | 4.2 | 3.8 | 3.9 | 3.8 | 4.4 | 3.72 | 4.03 | 3.4876 | < 20 | 3.8 | 4.2 | 4.1 | 4.4129 | 4 | 3.9257 | 4.6 | 3.039 | 3.7 |  |  |  | 3.931
| 1 | 5 | 3.8 | 3.6 | 4.1 | 4.7 | 3.6 | 3.8 | 3.9 | 3.57 | 3.91 | 3.4055 | < 20 | 3.8 | 4 | 3.9 | 3.8107 | 3.7 | 3.703 | 4.4 | 3.185 | 4.1 |  |  |  | 141
| 1 | 6 | 3.4 | 3.9 | 4 | 3.3 | 4.8 | 4.1 | 4.5 | 3.67 | 3.81 | 3.4825 | < 20 | 4.2 | 4 | 3.6 | 3.8331 | 4.7 | 3.6397 | 4.4 | 3.35 | 3.9
| Mean |  | 3.7167 | 3.7 | 4.1833 | 3.8833 | 3.8667 | 3.9333 | 4.35 | 3.675 | 3.8967 | 3.4817 |  | 4.05 | 4.15 | 3.9167 | 4.0794 | 4.15 | 3.8183 | 4.65 | 3.3813 | 3.9167
| Median |  | 3.8 | 3.75 | 4.15 | 3.8 | 3.75 | 3.85 | 4.45 | 3.695 | 3.86 | 3.4892 |  | 4 | 4.2 | 3.95 | 4.1034 | 4 | 3.8069 | 4.6 | 3.3945 | 4
| Std Dev. |  | 0.30605 | 0.22804 | 0.17224 | 0.46224 | 0.48854 | 0.22509 | 0.37283 | 0.11432 | 0.11587 | 0.04046 |  | 0.26646 | 0.12247 | 0.17224 | 0.22964 | 0.53198 | 0.19438 | 0.29496 | 0.24143 | 0.22286
| Rel.Std.Dev. |  | 8.235% | 6.163% | 4.117% | 11.9% | 12.63% | 5.723% | 8.571% | 3.111% | 2.974% | 1.162% |  | 6.579% | 2.951% | 4.398% | 5.629% | 12.82% | 5.091% | 6.343% | 7.14% | 5.69%
| PDM3 |  | -5.452% | -5.876% | 6.419% | -1.213% | -1.637% | 0.05928% | 10.66% | -6.512% | -0.8735% | -11.43% |  | 3.027% | 5.571% | -0.3647% | 3.775% | 5.571% | -2.868% | 18.29% | -13.98% | -0.3647%
| Z-Score (Absolute) |  | 0.65 | 0.7 | 0.86 | 0.11 | 0.16 | 0.05 | 1.4 | 0.78 | 0.06 | 1.41 | 19.69 | 0.43 | 0.76 | 0 | 0.53 | 0.76 | 0.32 | 2.37 | 1.73 | 0
| Analytical results for V in OREAS 702 (Certified Value 38.5 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 40 | 38 | 38 | 39 | 39 | 36.9 | 40 | 38 | 36 | 37.8427 | 38.4309 | 40 | 39.32 | 32 | 37.911 | 38.8454 | 41 | 39.4556 | 37 | 36.6108 | 40 | 37 |  | 1
| 1 | 2 | 39 | 38 | 38 | 40 | 38 | 36.5 | 41 | 38 | 36 | 37.6447 | 38.0517 | 40 | 38.85 | 33 | 37.449 | 38.6867 | 41 | 39.9238 | 38 | 34.0892 | 41 | 37 |  | #N/A
| 1 | 3 | 41 | 38 | 38 | 39 | 36 | 36.5 | 39 | 38 | 37 | 37.4893 | 39.72 | 40 | 38.64 | 36 | 37.5 | 36.6193 | 41 | 39.6271 | 37 | 37.0002 | 42 | 38 |  | 16
| 1 | 4 | 39 | 38 | 39 | 38 | 37 | 37.5 | 41 | 37 | 37 | 37.9315 | 39.6525 | 40 | 38.57 | 31 | 38.567 | 35.909 | 41 | 39.6191 | 37 | 37.5888 | 40 | 37 |  | 38.4563
| 1 | 5 | 37 | 38 | 39 | 40 | 36 | 37.5 | 41 | 37 | 37 | 38.1562 | 38.03 | 40 | 39.23 | 30 | 38.527 | 37.8816 | 41 | 39.2407 | 38 | 36.9494 | 42 | 38 |  | 142
| 1 | 6 | 40 | 38 | 38 | 40 | 38 | 37.1 | 41 | 37 | 36 | 37.9452 | 38.1125 | 40 | 39.02 | 32 | 37.698 | 37.8762 | 41 | 39.5342 | 38 | 34.7475 | 42 | 37
| Mean |  | 39.3333 | 38 | 38.3333 | 39.3333 | 37.3333 | 37 | 40.5 | 37.5 | 36.5 | 37.8349 | 38.6663 | 40 | 38.9383 | 32.3333 | 37.942 | 37.6364 | 41 | 39.5668 | 37.5 | 36.1643 | 41.1667 | 37.3333
| Median |  | 39.5 | 38 | 38 | 39.5 | 37.5 | 37 | 41 | 37.5 | 36.5 | 37.8871 | 38.2717 | 40 | 38.935 | 32 | 37.8045 | 37.8789 | 41 | 39.5767 | 37.5 | 36.7801 | 41.5 | 37
| Std Dev. |  | 1.3663 | 0 | 0.5164 | 0.8165 | 1.2111 | 0.45166 | 0.83666 | 0.54772 | 0.54772 | 0.23684 | 0.80345 | 0 | 0.30656 | 2.0656 | 0.49634 | 1.1576 | 0 | 0.2252 | 0.54772 | 1.4041 | 0.98319 | 0.5164
| Rel.Std.Dev. |  | 3.474% | 0% | 1.347% | 2.076% | 3.244% | 1.221% | 2.066% | 1.461% | 1.501% | 0.626% | 2.078% | 0% | 0.7873% | 6.388% | 1.308% | 3.076% | 0% | 0.5692% | 1.461% | 3.882% | 2.388% | 1.383%
| PDM3 |  | 2.281% | -1.187% | -0.3198% | 2.281% | -2.92% | -3.787% | 5.314% | -2.487% | -5.087% | -1.616% | 0.546% | 4.014% | 1.253% | -15.92% | -1.337% | -2.132% | 6.615% | 2.888% | -2.487% | -5.96% | 7.048% | -2.92%
| Z-Score (Absolute) |  | 0.95 | 0.02 | 0.25 | 0.95 | 0.44 | 0.68 | 1.76 | 0.33 | 1.02 | 0.09 | 0.48 | 1.41 | 0.67 | 3.92 | 0.02 | 0.23 | 2.11 | 1.11 | 0.33 | 1.26 | 2.22 | 0.44
| Analytical results for W in OREAS 702 (Certified Value 0.215 wt.%) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 24 | 25 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.213 | 0.218 | 0.209 | 0.206 | 0.221 | 0.241 | 0.2272 | 0.21239 | 0.195 | 0.22304 | 0.21069 | 0.2301 | 0.1846 | 0.21813 | 0.2154 | 0.058625 | 0.19726 | 0.21826 | > 0.1 | 0.19871 | 0.23125 | 0.2256 |  | 1
| 1 | 2 | 0.21 | 0.217 | 0.207 | 0.21 | 0.21 | 0.232 | 0.2265 | 0.22768 | 0.197 | 0.22894 | 0.21061 | 0.24135 | 0.183 | 0.20433 | 0.2218 | 0.066975 | 0.20315 | 0.22837 | > 0.1 | 0.19924 | 0.22725 | 0.22474 |  | #N/A
| 1 | 3 | 0.22 | 0.215 | 0.203 | 0.205 | 0.208 | 0.236 | 0.2205 | 0.21984 | 0.197 | 0.22792 | 0.21421 | 0.2337 | 0.1835 | 0.21136 | 0.2156 | 0.10826 | 0.20744 | 0.22705 | > 0.1 | 0.1976 | 0.22855 | 0.22923 |  | 16
| 1 | 4 | 0.211 | 0.212 | 0.209 | 0.203 | 0.213 | 0.236 | 0.2245 | 0.21193 | 0.201 | 0.2268 | 0.21388 | 0.2233 | 0.1824 | 0.2225 | 0.2142 | 0.091908 | 0.19971 | 0.22495 | > 0.1 | 0.19875 | 0.2257 | 0.22485 |  | 0.21492
| 1 | 5 | 0.203 | 0.215 | 0.209 | 0.21 | 0.214 | 0.23 | 0.2244 | 0.21317 | 0.197 | 0.22322 | 0.21606 | 0.22905 | 0.1844 | 0.25541 | 0.223 | 0.079603 | 0.19726 | 0.22575 | > 0.1 | 0.19876 | 0.2212 | 0.22969 |  | 143
| 1 | 6 | 0.213 | 0.212 | 0.203 | 0.21 | 0.218 | 0.232 | 0.2256 | 0.20766 | 0.2 | 0.22327 | 0.20758 | 0.23051 | 0.184 | 0.22796 | 0.2114 | 0.068262 | 0.20248 | 0.22505 | 0.09546 | 0.20294 | 0.22395 | 0.22624
| Mean |  | 0.21167 | 0.21483 | 0.20667 | 0.20733 | 0.214 | 0.2345 | 0.22478 | 0.21544 | 0.19783 | 0.22553 | 0.21217 | 0.23133 | 0.18365 | 0.22328 | 0.2169 | 0.078939 | 0.20122 | 0.22491 | 0.09546 | 0.19933 | 0.22632 | 0.22672
| Median |  | 0.212 | 0.215 | 0.208 | 0.208 | 0.2135 | 0.234 | 0.22505 | 0.21278 | 0.197 | 0.22503 | 0.21228 | 0.2303 | 0.18375 | 0.22032 | 0.2155 | 0.073933 | 0.20109 | 0.2254 | 0.09546 | 0.19876 | 0.22647 | 0.22592
| Std Dev. |  | 0.0055015 | 0.0024833 | 0.0029439 | 0.0030768 | 0.004858 | 0.0039875 | 0.0023676 | 0.0071618 | 0.0022286 | 0.0026706 | 0.0030968 | 0.0059646 | 0.00084794 | 0.017793 | 0.0045321 | 0.01843 | 0.0039407 | 0.0035097 |  | 0.001848 | 0.0035315 | 0.0021919
| Rel.Std.Dev. |  | 2.599% | 1.156% | 1.424% | 1.484% | 2.27% | 1.7% | 1.053% | 3.324% | 1.127% | 1.184% | 1.46% | 2.578% | 0.4617% | 7.969% | 2.089% | 23.35% | 1.958% | 1.561% |  | 0.9271% | 1.56% | 0.9668%
| PDM3 |  | -1.514% | -0.04097% | -3.841% | -3.531% | -0.4287% | 9.11% | 4.589% | 0.2436% | -7.951% | 4.936% | -1.28% | 7.637% | -14.55% | 3.89% | 0.9206% | -63.27% | -6.377% | 4.646% | -55.58% | -7.253% | 5.302% | 5.492%
| Z-Score (Absolute) |  | 0.17 | 0.03 | 0.48 | 0.44 | 0.03 | 1.25 | 0.65 | 0.06 | 1.04 | 0.69 | 0.14 | 1.06 | 1.92 | 0.55 | 0.16 | 8.46 | 0.82 | 0.65 | 7.43 | 0.94 | 0.74 | 0.77
| Analytical results for Y in OREAS 702 (Certified Value 13.1 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 14.1 | 13.1 | 14.4 | 13.4 | 14.3 | 12 | 12 | 12.05 | 12.3 | 14.7121 | 10.122 | 12.9 | 13.39 | 15.7 | 10.83 | 14.8 | 13.6124 | 12.5 | 11.564 | 12.6 | 13 |  |  | 1
| 1 | 2 | 13.9 | 13.2 | 13.6 | 13.8 | 13.5 | 12.2 | 13 | 12.02 | 11.9 | 14.6186 | 9.8654 | 13.7 | 13.01 | 16 | 11.09 | 14.5 | 13.3787 | 12.2 | 12 | 13.3 | 13.4 |  |  | #N/A
| 1 | 3 | 15.2 | 13.2 | 14 | 14.4 | 13.6 | 12.3 | 12 | 12.4 | 12.2 | 14.6139 | 9.92 | 13.2 | 12.55 | 15.6 | 11.03 | 14.7 | 13.4476 | 12.4 | 11.533 | 12.4 | 13.2 |  |  | 16
| 1 | 4 | 14.2 | 13.3 | 13.8 | 13.8 | 14.1 | 12.4 | 13 | 12.18 | 12 | 14.7099 | 10.35 | 13.3 | 12.52 | 15.4 | 10.84 | 14.5 | 13.3168 | 12.1 | 11.71 | 13 | 13.4 |  |  | 13.0604
| 1 | 5 | 14 | 13.3 | 13.7 | 14.2 | 13.7 | 12.4 | 13 | 12.1 | 12.2 | 14.7301 | 9.87 | 13 | 13.44 | 15 | 10.96 | 14.7 | 13.2495 | 12.3 | 10.703 | 13.7 | 13 |  |  | 144
| 1 | 6 | 14.3 | 13.2 | 13.4 | 13.9 | 14 | 13.3 | 12 | 11.87 | 12.1 | 14.6226 | 10.01 | 13.3 | 12.56 | 14.3 | 10.89 | 14.8 | 13.2537 | 12.2 | 12.304 | 13.9 | 13.4
| Mean |  | 14.2833 | 13.2167 | 13.8167 | 13.9167 | 13.8667 | 12.4333 | 12.5 | 12.1033 | 12.1167 | 14.6679 | 10.0229 | 13.2333 | 12.9117 | 15.3333 | 10.94 | 14.6667 | 13.3764 | 12.2833 | 11.6357 | 13.15 | 13.2333
| Median |  | 14.15 | 13.2 | 13.75 | 13.85 | 13.85 | 12.35 | 12.5 | 12.075 | 12.15 | 14.6663 | 9.965 | 13.25 | 12.785 | 15.5 | 10.925 | 14.7 | 13.3478 | 12.25 | 11.637 | 13.15 | 13.3
| Std Dev. |  | 0.47081 | 0.075277 | 0.34881 | 0.34881 | 0.31411 | 0.45019 | 0.54772 | 0.17784 | 0.1472 | 0.054753 | 0.18748 | 0.28048 | 0.43023 | 0.60553 | 0.10545 | 0.13663 | 0.13818 | 0.1472 | 0.54231 | 0.59582 | 0.19664
| Rel.Std.Dev. |  | 3.296% | 0.5696% | 2.525% | 2.506% | 2.265% | 3.621% | 4.382% | 1.469% | 1.215% | 0.3733% | 1.871% | 2.119% | 3.332% | 3.949% | 0.9639% | 0.9315% | 1.033% | 1.198% | 4.661% | 4.531% | 1.486%
| PDM3 |  | 9.364% | 1.197% | 5.791% | 6.556% | 6.173% | -4.801% | -4.291% | -7.328% | -7.226% | 12.31% | -23.26% | 1.324% | -1.139% | 17.4% | -16.24% | 12.3% | 2.42% | -5.95% | -10.91% | 0.6861% | 1.324%
| Z-Score (Absolute) |  | 0.82 | 0.01 | 0.46 | 0.54 | 0.5 | 0.62 | NA | 0.87 | 0.86 | 1.12 | 2.49 | 0.01 | 0.24 | 1.64 | 1.78 | 1.12 | 0.12 | 0.73 | 1.24 | 0.06 | 0.01
Results from laboratory 7 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 702 (Certified Value 1.03 ppm) | [24 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 06 | 07 | 09 | 11 | 14 | 15 | 18 | 19 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 0.99 | 1 | 1.01 | 1.0686 | 0.93 | 1.2 | 1 | 0.95309 | 0.903 | [14 empty] | 1
| 1 | 2 | 1.07 | 1.1 | 1.14 | 1.045 | 0.94 | 1.3 | 0.99 | 0.9723 | 0.876 | [14 empty] | #N/A
| 1 | 3 | 1.13 | 1.1 | 1.07 | 1.0371 | 0.98 | 1.2 | 1.03 | 1.0196 | 0.997 | [14 empty] | 16
| 1 | 4 | 0.99 | 1 | 1.03 | 1.0686 | 0.99 | 1.3 | 1.08 | 0.98101 | 0.869 | [14 empty] | 1.028
| 1 | 5 | 1.03 | 1 | 1.01 | 1.0809 | 1.05 | 1.2 | 1.07 | 1.0415 | 0.82 | [14 empty] | 145
| 1 | 6 | 1.12 | 1 | 0.99 | 1.0487 | 1.03 | 1.1 | 1.04 | 0.94813 | 0.872
| Mean |  | 1.055 | 1.0333 | 1.0417 | 1.0582 | 0.98667 | 1.2167 | 1.035 | 0.98595 | 0.8895
| Median |  | 1.05 | 1 | 1.02 | 1.0586 | 0.985 | 1.2 | 1.035 | 0.97666 | 0.874
| Std Dev. |  | 0.061887 | 0.05164 | 0.055287 | 0.016993 | 0.04761 | 0.075277 | 0.036194 | 0.037265 | 0.059123
| Rel.Std.Dev. |  | 5.866% | 4.997% | 5.308% | 1.606% | 4.825% | 6.187% | 3.497% | 3.78% | 6.647%
| PDM3 |  | 2.63% | 0.522% | 1.333% | 2.937% | -4.018% | 18.36% | 0.6841% | -4.088% | -13.47%
| Z-Score (Absolute) |  | 0.58 | 0.05 | 0.19 | 0.67 | 1.41 | 5.29 | 0 | 1.43 | 4.24
| Analytical results for Zn in OREAS 702 (Certified Value 57 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 62 | 60 | 59 | 58 | 60 | 53 | 76 | 53 | 55 | 59.8151 | 56 | 61 | 62 | 56.64 | 53.5521 | 59 | 56.1042 | 57 | 58.9398 | 57 | 52 |  |  | 1
| 1 | 2 | 59 | 59 | 58 | 59 | 57 | 53 | 73 | 53 | 54 | 60.05 | 57 | 60 | 61 | 58.87 | 58.5274 | 59 | 56.5289 | 57 | 53.3211 | 56 | 53 |  |  | 22
| 1 | 3 | 62 | 59 | 58 | 58 | 56 | 54 | 70 | 53 | 56 | 59.2422 | 60 | 61 | 62 | 58.04 | 52.3051 | 59 | 56.3511 | 56 | 53.1329 | 58 | 55 |  |  | 16
| 1 | 4 | 60 | 59 | 59 | 57 | 57 | 54 | 71 | 53 | 55 | 59.99 | 57 | 62 | 58 | 53.23 | 55.8095 | 59 | 55.4782 | 56 | 56.8329 | 62 | 53 |  |  | 57.156
| 1 | 5 | 58 | 59 | 59 | 59 | 58 | 55 | 69 | 52 | 54 | 59.8 | 57 | 60 | 57 | 55.23 | 49.9956 | 59 | 54.8614 | 57 | 55.0401 | 59 | 54 |  |  | 146
| 1 | 6 | 61 | 59 | 58 | 60 | 59 | 59 | 69 | 52 | 54 | 60.7583 | 58 | 59 | 56 | 56.53 | 51.8797 | 60 | 54.9739 | 59 | 54.895 | 61 | 53
| Mean |  | 60.3333 | 59.1667 | 58.5 | 58.5 | 57.8333 | 54.6667 | 71.3333 | 52.6667 | 54.6667 | 59.9426 | 57.5 | 60.5 | 59.3333 | 56.4233 | 53.6782 | 59.1667 | 55.7163 | 57 | 55.3603 | 58.8333 | 53.3333
| Median |  | 60.5 | 59 | 58.5 | 58.5 | 57.5 | 54 | 70.5 | 53 | 54.5 | 59.9025 | 57 | 60.5 | 59.5 | 56.585 | 52.9286 | 59 | 55.7912 | 57 | 54.9676 | 58.5 | 53
| Std Dev. |  | 1.633 | 0.40825 | 0.54772 | 1.0488 | 1.472 | 2.2509 | 2.7325 | 0.5164 | 0.8165 | 0.49119 | 1.3784 | 1.0488 | 2.6583 | 2.0139 | 3.0579 | 0.40825 | 0.71449 | 1.0954 | 2.21 | 2.3166 | 1.0328
| Rel.Std.Dev. |  | 2.707% | 0.69% | 0.9363% | 1.793% | 2.545% | 4.118% | 3.831% | 0.9805% | 1.494% | 0.8194% | 2.397% | 1.734% | 4.48% | 3.569% | 5.697% | 0.69% | 1.282% | 1.922% | 3.992% | 3.938% | 1.936%
| PDM3 |  | 5.559% | 3.518% | 2.351% | 2.351% | 1.185% | -4.355% | 24.8% | -7.855% | -4.355% | 4.875% | 0.6018% | 5.851% | 3.809% | -1.282% | -6.085% | 3.518% | -2.519% | -0.273% | -3.142% | 2.935% | -6.688%
| Z-Score (Absolute) |  | 0.8 | 0.43 | 0.21 | 0.21 | 0 | 1.01 | 4.32 | 1.65 | 1.01 | 0.67 | 0.11 | 0.85 | 0.48 | 0.45 | 1.33 | 0.43 | 0.68 | 0.27 | 0.79 | 0.32 | 1.44
| Analytical results for Zr in OREAS 702 (Certified Value 40 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 40 | 37.6 | 43.3 | 43.4 | 42.4 | 36.8 | 41 | 42.6 | 33.3 | 30.3624 | 34 | 37.39 | 43.2 | 37.911 | 39.3632 | 49.2 | 41.3166 | 30.9 | 42.7 |  |  |  |  | 1
| 1 | 2 | 37.3 | 40.5 | 53.4 | 40.5 | 40.5 | 37.4 | 39 | 42.5 | 33.1 | 30.7699 | 34.5 | 37.13 | 41.3 | 34.987 | 39.5702 | 44.5 | 40.9307 | 33 | 42 |  |  |  |  | 4
| 1 | 3 | 41.5 | 44.6 | 42.1 | 46.1 | 41.6 | 38 | 40 | 42.5 | 33.5 | 29.9711 | 33.2 | 37.34 | 43.9 | 35.578 | 42.239 | 50.8 | 43.0465 | 31.4 | 43 |  |  |  |  | 16
| 1 | 4 | 40.3 | 40.7 | 41.3 | 43.1 | 40.6 | 38.8 | 37 | 43.3 | 33.2 | 30.2433 | 34.3 | 37.13 | 43.8 | 35.156 | 37.2331 | 44.6 | 41.0353 | 30.2 | 42.2 |  |  |  |  | 40.0151
| 1 | 5 | 38.3 | 49.3 | 40.5 | 44.3 | 39.8 | 36.1 | 36 | 41 | 34.9 | 30.5767 | 33.9 | 37.16 | 42.1 | 36.83 | 40.7608 | 48.4 | 41.6694 | 32 | 42.4 |  |  |  |  | 147
| 1 | 6 | 43.1 | 41.4 | 40.4 | 41.7 | 40.9 | 36.9 | 35 | 41.1 | 33.7 | 30.26 | 34 | 36.58 | 38.3 | 36.648 | 39.7967 | 51.3 | 40.1216 | 32.4 | 41.7
| Mean |  | 40.0833 | 42.35 | 43.5 | 43.1833 | 40.9667 | 37.3333 | 38 | 42.1667 | 33.6167 | 30.3639 | 33.9833 | 37.1217 | 42.1 | 36.185 | 39.8272 | 48.1333 | 41.3533 | 31.65 | 42.3333
| Median |  | 40.15 | 41.05 | 41.7 | 43.25 | 40.75 | 37.15 | 38 | 42.5 | 33.4 | 30.3112 | 34 | 37.145 | 42.65 | 36.113 | 39.6835 | 48.8 | 41.1759 | 31.7 | 42.3
| Std Dev. |  | 2.1018 | 4.0737 | 4.9691 | 1.96 | 0.9136 | 0.95847 | 2.3664 | 0.91579 | 0.66458 | 0.27917 | 0.4446 | 0.28799 | 2.1175 | 1.1377 | 1.6553 | 2.9676 | 0.9762 | 1.0232 | 0.47188
| Rel.Std.Dev. |  | 5.244% | 9.619% | 11.42% | 4.539% | 2.23% | 2.567% | 6.227% | 2.172% | 1.977% | 0.9194% | 1.308% | 0.7758% | 5.03% | 3.144% | 4.156% | 6.165% | 2.361% | 3.233% | 1.115%
| PDM3 |  | 0.1704% | 5.835% | 8.709% | 7.918% | 2.378% | -6.702% | -5.036% | 5.377% | -15.99% | -24.12% | -15.07% | -7.231% | 5.21% | -9.572% | -0.4697% | 20.29% | 3.344% | -20.9% | 5.793%
| Z-Score (Absolute) |  | 0 | 0.56 | 0.84 | 0.76 | 0.22 | 0.67 | 0.51 | 0.51 | 1.59 | 2.38 | 1.5 | 0.73 | 0.49 | 0.96 | 0.06 | 1.97 | 0.31 | 2.07 | 0.55
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C987: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1116: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 702 (Indicative Value 0.117 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 3
|  |  | 0.5g | 3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.1 | 11
| 1 | 3 | 0.1 | 16
| 1 | 4 | 0.2 | 0.11667
| 1 | 5 | 0.1 | 25
| 1 | 6 | 0.1
| Mean |  | 0.11667
| Median |  | 0.1
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 34.99%
| PDM3 |  | -2.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 702 (Indicative Value 3.1 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 2
|  |  | 0.5g | 2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3 | 12
| 1 | 3 | 3.2 | 16
| 1 | 4 | 3 | 3.1
| 1 | 5 | 2.8 | 26
| 1 | 6 | 3.2
| Mean |  | 3.1
| Median |  | 3.1
| Std Dev. |  | 0.20976
| Rel.Std.Dev. |  | 6.767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Au in OREAS 702 (Indicative Value 0.057 ppm) |  |  | Indicative
Au | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 3
|  |  | 0.5g | 3
| 1 | 1 | 0.0571 | 1
| 1 | 2 | 0.0525 | 21
| 1 | 3 | 0.0606 | 16
| 1 | 4 | 0.0549 | 0.056583
| 1 | 5 | 0.0582 | 27
| 1 | 6 | 0.0562
| Mean |  | 0.056583
| Median |  | 0.05665
| Std Dev. |  | 0.0027809
| Rel.Std.Dev. |  | 4.915%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 702 (Indicative Value 268 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 0
|  |  | 0.5g | 0
| 1 | 1 | 259 | 1
| 1 | 2 | 261.6 | 22
| 1 | 3 | 275.5 | 16
| 1 | 4 | 266.7 | 267.7667
| 1 | 5 | 274.3 | 28
| 1 | 6 | 269.5
| Mean |  | 267.7667
| Median |  | 268.1
| Std Dev. |  | 6.656
| Rel.Std.Dev. |  | 2.486%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 702 (Indicative Value < 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 2
|  |  | 0.5g | 2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 1
| 1 | 3 | < 0.1 | 16
| 1 | 4 | < 0.1 | < 0.1
| 1 | 5 | < 0.1 | 29
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 702 (Indicative Value 328 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 0
|  |  | 0.5g | 0
| 1 | 1 | 320.5 | 1
| 1 | 2 | 326.3 | 19
| 1 | 3 | 332.9 | 16
| 1 | 4 | 327.4 | 327.95
| 1 | 5 | 327.3 | 25
| 1 | 6 | 333.3
| Mean |  | 327.95
| Median |  | 327.35
| Std Dev. |  | 4.7344
| Rel.Std.Dev. |  | 1.444%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hg in OREAS 702 (Indicative Value < 0.01 ppm) |  |  | Indicative
Hg | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 3
|  |  | 0.5g | 3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 17
| 1 | 3 | < 0.01 | 16
| 1 | 4 | < 0.01 | < 0.01
| 1 | 5 | < 0.01 | 26
| 1 | 6 | < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 702 (Indicative Value 595 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 0
|  |  | 0.5g | 0
| 1 | 1 | 584 | 1
| 1 | 2 | 599.4 | 21
| 1 | 3 | 589.3 | 16
| 1 | 4 | 596.4 | 595.4833
| 1 | 5 | 599.1 | 27
| 1 | 6 | 604.7
| Mean |  | 595.4833
| Median |  | 597.75
| Std Dev. |  | 7.5407
| Rel.Std.Dev. |  | 1.266%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 702 (Indicative Value 10.8 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 1
|  |  | 0.5g | 1
| 1 | 1 | 10.5 | 1
| 1 | 2 | 10.7 | 12
| 1 | 3 | 10.8 | 16
| 1 | 4 | 10.9 | 10.75
| 1 | 5 | 10.7 | 28
| 1 | 6 | 10.9
| Mean |  | 10.75
| Median |  | 10.75
| Std Dev. |  | 0.15166
| Rel.Std.Dev. |  | 1.411%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 702 (Indicative Value 4.43 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 2
|  |  | 0.5g | 2
| 1 | 1 | 4.3 | 1
| 1 | 2 | 4.3 | 23
| 1 | 3 | 4.6 | 16
| 1 | 4 | 4.4 | 4.4333
| 1 | 5 | 4.5 | 29
| 1 | 6 | 4.5
| Mean |  | 4.4333
| Median |  | 4.45
| Std Dev. |  | 0.12111
| Rel.Std.Dev. |  | 2.732%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 702 (Indicative Value 0.2 ppm) |  |  | Indicative
Sb | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 2
|  |  | 0.5g | 2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 3
| 1 | 3 | 0.2 | 16
| 1 | 4 | 0.2 | 0.2
| 1 | 5 | 0.2 | 25
| 1 | 6 | 0.2
| Mean |  | 0.2
| Median |  | 0.2
| Std Dev. |  | 3.0405e-17
| Rel.Std.Dev. |  | 1.52e-14%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 702 (Indicative Value < 0.5 ppm) |  |  | Indicative
Se | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 2
|  |  | 0.5g | 2
| 1 | 1 | < 0.5 | 1
| 1 | 2 | < 0.5 | 11
| 1 | 3 | < 0.5 | 16
| 1 | 4 | < 0.5 | < 0.5
| 1 | 5 | < 0.5 | 26
| 1 | 6 | < 0.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 702 (Indicative Value 0.5 ppm) |  |  | Indicative
Tl | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 2
|  |  | 0.5g | 2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.5 | 21
| 1 | 3 | 0.5 | 16
| 1 | 4 | 0.5 | 0.5
| 1 | 5 | 0.5 | 27
| 1 | 6 | 0.5
| Mean |  | 0.5
| Median |  | 0.5
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 702 (Indicative Value 52 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 13 | ppm
|  |  | AR*MS | 0
|  |  | 0.5g | 0
| 1 | 1 | 51 | 1
| 1 | 2 | 52 | 22
| 1 | 3 | 51 | 16
| 1 | 4 | 52 | 51.5
| 1 | 5 | 52 | 28
| 1 | 6 | 51
| Mean |  | 51.5
| Median |  | 51.5
| Std Dev. |  | 0.54772
| Rel.Std.Dev. |  | 1.064%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.