## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 84
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 178.717 | 168.5349 | 188.899 | 177.4899 | 179.944
Pd, Palladium (ppb) | 581.1081 | 548.2351 | 613.9812 | 569.4708 | 592.7454
Pt, Platinum (ppb) | 517.0015 | 485.5994 | 548.4035 | 504.9622 | 529.0407
NiS Fire Assay
Pd, Palladium (ppb) | 556.22 | 494.298 | 618.142 | 539.8906 | 572.5494
Pt, Platinum (ppb) | 486.8667 | 392.2636 | 581.4698 | 418.3541 | 555.3792
Infrared Combustion
S, Sulphur (wt.%) | 24.2134 | 23.7391 | 24.6876 | 23.9451 | 24.4817
4-Acid Digestion
Ag, Silver (ppm) | 2.3765 | 2.2804 | 2.4725 | 2.2688 | 2.4841
Al, Aluminium (wt.%) | 2.2 | 2.1257 | 2.2742 | 2.1547 | 2.2453
As, Arsenic (ppm) | 167.1961 | 160.6838 | 173.7084 | 162.865 | 171.5271
Ba, Barium (ppm) | 161.1092 | 151.1492 | 171.0692 | 157.1177 | 165.1007
Be, Beryllium (ppm) | 0.44543 | 0.39215 | 0.4987 | 0.41661 | 0.47424
Bi, Bismuth (ppm) | 1.5616 | 1.4819 | 1.6412 | 1.5101 | 1.613
Ca, Calcium (wt.%) | 2.0344 | 1.9728 | 2.096 | 2.0004 | 2.0684
Cd, Cadmium (ppm) | 2.4285 | 2.2467 | 2.6104 | 2.2937 | 2.5634
Ce, Cerium (ppm) | 15.7988 | 15.0376 | 16.56 | 15.3228 | 16.2748
Co, Cobalt (ppm) | 1578.0701 | 1519.453 | 1636.6871 | 1543.1176 | 1613.0225
Cr, Chromium (wt.%) | 0.55884 | 0.48627 | 0.6314 | 0.53475 | 0.58292
Cs, Caesium (ppm) | 1.8484 | 1.7278 | 1.9689 | 1.7677 | 1.9291
Cu, Copper (wt.%) | 0.39955 | 0.38696 | 0.41214 | 0.39287 | 0.40623
Dy, Dysprosium (ppm) | 0.93845 | 0.88184 | 0.99506 | 0.90001 | 0.97689
Er, Erbium (ppm) | 0.52575 | 0.47152 | 0.57997 | 0.49901 | 0.55248
Eu, Europium (ppm) | 0.33739 | 0.29135 | 0.38344 | 0.31602 | 0.35877
Fe, Iron (wt.%) | 31.5125 | 30.6052 | 32.4197 | 31.0615 | 31.9634
Ga, Gallium (ppm) | 6.5811 | 6.0587 | 7.1035 | 6.3155 | 6.8467
Gd, Gadolinium (ppm) | 1.1618 | 1.0334 | 1.2901 | 1.1114 | 1.2121
Hf, Hafnium (ppm) | 0.83852 | 0.75183 | 0.92522 | 0.79013 | 0.88692
Ho, Holmium (ppm) | 0.18467 | 0.15965 | 0.20969 | 0.16191 | 0.20743
In, Indium (ppm) | 0.13206 | 0.11686 | 0.14726 | 0.12354 | 0.14058
K, Potassium (wt.%) | 0.36758 | 0.35482 | 0.38034 | 0.35868 | 0.37649
La, Lanthanum (ppm) | 8.1498 | 7.6405 | 8.659 | 7.8558 | 8.4437
Li, Lithium (ppm) | 10.277 | 9.5915 | 10.9625 | 9.9333 | 10.6207
Lu, Lutetium (ppm) | 0.07291 | 0.061077 | 0.084742 | IND | IND
Mg, Magnesium (wt.%) | 1.5128 | 1.4594 | 1.5661 | 1.4911 | 1.5344
Mn, Manganese (wt.%) | 0.063789 | 0.061722 | 0.065857 | 0.062586 | 0.064993
Mo, Molybdenum (ppm) | 5.0522 | 4.6532 | 5.4512 | 4.7031 | 5.4013
Na, Sodium (wt.%) | 0.44587 | 0.43254 | 0.4592 | 0.43592 | 0.45582
Nb, Niobium (ppm) | 2.3659 | 2.1728 | 2.559 | 2.2277 | 2.5041
Nd, Neodymium (ppm) | 6.7201 | 6.335 | 7.1053 | 6.494 | 6.9463
Ni, Nickel (wt.%) | 10.4663 | 10.1802 | 10.7524 | 10.2512 | 10.6814
P, Phosphorus (wt.%) | 0.021907 | 0.020866 | 0.022949 | 0.021366 | 0.022449
Pb, Lead (ppm) | 74.4077 | 71.247 | 77.5684 | 72.5007 | 76.3146
Pr, Praseodymium (ppm) | 1.8212 | 1.6702 | 1.9722 | 1.7246 | 1.9178
Rb, Rubidium (ppm) | 21.4196 | 20.0739 | 22.7653 | 20.8578 | 21.9814
Re, Rhenium (ppm) | 0.029452 | 0.023439 | 0.035465 | 0.025365 | 0.033539
S, Sulphur (wt.%) | 22.1417 | 20.978 | 23.3054 | 21.6909 | 22.5926
Sb, Antimony (ppm) | 1.5204 | 1.3825 | 1.6583 | 1.4119 | 1.629
Sc, Scandium (ppm) | 5.2099 | 4.8528 | 5.567 | 4.9921 | 5.4277
Se, Selenium (ppm) | 10.7094 | 9.5205 | 11.8983 | 9.9012 | 11.5176
Sm, Samarium (ppm) | 1.3108 | 1.1444 | 1.4772 | 1.2395 | 1.3821
Sn, Tin (ppm) | 1.4943 | 1.3323 | 1.6562 | IND | IND
Sr, Strontium (ppm) | 86.8664 | 82.7926 | 90.9402 | 85.0952 | 88.6376
Ta, Tantalum (ppm) | 0.20525 | 0.17251 | 0.23799 | 0.18544 | 0.22507
Tb, Terbium (ppm) | 0.175 | 0.15946 | 0.19053 | IND | IND
Te, Tellurium (ppm) | 1.2653 | 1.1112 | 1.4195 | 1.1199 | 1.4108
Th, Thorium (ppm) | 3.1385 | 2.9193 | 3.3576 | 2.9869 | 3.2901
Ti, Titanium (wt.%) | 0.11712 | 0.11048 | 0.12376 | 0.11442 | 0.11982
Tl, Thallium (ppm) | 1.2381 | 1.1853 | 1.2909 | 1.1997 | 1.2765
Tm, Thulium (ppm) | 0.068555 | 0.0498 | 0.08731 | IND | IND
U, Uranium (ppm) | 0.82271 | 0.75215 | 0.89327 | 0.79233 | 0.8531
V, Vanadium (ppm) | 100.2302 | 95.4131 | 105.0473 | 96.9989 | 103.4616
W, Tungsten (ppm) | 1.3059 | 1.1765 | 1.4353 | IND | IND
Y, Yttrium (ppm) | 4.7093 | 4.3599 | 5.0588 | 4.526 | 4.8926
Yb, Ytterbium (ppm) | 0.47047 | 0.41646 | 0.52449 | 0.43803 | 0.50292
Zn, Zinc (ppm) | 813.8033 | 782.3036 | 845.303 | 801.3835 | 826.2232
Zr, Zirconium (ppm) | 28.5666 | 26.2585 | 30.8748 | 27.0329 | 30.1004
Borate / Peroxide Fusion ICP
Al2O3, Aluminium(III) oxide (wt.%) | 4.1663 | 4.0055 | 4.3271 | 4.0755 | 4.2572
As, Arsenic (ppm) | 175.4004 | 164.9601 | 185.8406 | 167.5791 | 183.2216
Ba, Barium (ppm) | 175.3227 | 166.7841 | 183.8614 | 170.4294 | 180.216
Bi, Bismuth (ppm) | 1.3902 | 1.1913 | 1.5891 | IND | IND
CaO, Calcium oxide (wt.%) | 2.9309 | 2.7949 | 3.0669 | 2.8612 | 3.0006
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (ppm) | 15.829 | 14.4388 | 17.2192 | 15.3172 | 16.3407
Co, Cobalt (ppm) | 1653.8378 | 1580.8372 | 1726.8385 | 1621.9104 | 1685.7652
Cr, Chromium (wt.%) | 0.74674 | 0.7244 | 0.76909 | 0.73075 | 0.76273
Cs, Caesium (ppm) | 1.7278 | 1.4361 | 2.0194 | 1.5333 | 1.9222
Cu, Copper (wt.%) | 0.40483 | 0.39477 | 0.41489 | 0.39638 | 0.41328
Dy, Dysprosium (ppm) | 0.93436 | 0.81368 | 1.055 | 0.86323 | 1.0055
Er, Erbium (ppm) | 0.49702 | 0.41629 | 0.57776 | 0.43528 | 0.55877
Eu, Europium (ppm) | 0.29785 | 0.26776 | 0.32794 | 0.26917 | 0.32653
Fe2O3, Iron(III) oxide (wt.%) | 44.6442 | 43.398 | 45.8904 | 43.8132 | 45.4753
Ga, Gallium (ppm) | 7.0535 | 6.1792 | 7.9277 | 6.5922 | 7.5147
Gd, Gadolinium (ppm) | 1.1203 | 0.98652 | 1.2541 | 1.0045 | 1.2362
Ho, Holmium (ppm) | 0.17549 | 0.14801 | 0.20298 | IND | IND
K2O, Potassium oxide (wt.%) | 0.45348 | 0.42281 | 0.48415 | 0.43218 | 0.47477
La, Lanthanum (ppm) | 8.3036 | 7.6602 | 8.9471 | 7.8575 | 8.7497
Li, Lithium (ppm) | 12.846 | 9.9937 | 15.6983 | 10.7411 | 14.951
Lu, Lutetium (ppm) | 0.075278 | 0.047337 | 0.10322 | IND | IND
MgO, Magnesium oxide (wt.%) | 2.6427 | 2.5513 | 2.734 | 2.587 | 2.6983
MnO, Manganese oxide (wt.%) | 0.085184 | 0.081878 | 0.088491 | 0.083729 | 0.08664
Mo, Molybdenum (ppm) | 5.6239 | 4.2486 | 6.9992 | IND | IND
Nd, Neodymium (ppm) | 6.6726 | 6.1708 | 7.1745 | 6.1701 | 7.1752
Ni, Nickel (wt.%) | 10.5288 | 10.2427 | 10.815 | 10.3337 | 10.7239
P2O5, Phosphorus(V) oxide (wt.%) | 0.057341 | 0.045052 | 0.06963 | 0.055686 | 0.058996
Pb, Lead (ppm) | 73.7087 | 66.4192 | 80.9982 | 69.068 | 78.3494
Pr, Praseodymium (ppm) | 1.8068 | 1.6426 | 1.9711 | 1.6798 | 1.9339
Rb, Rubidium (ppm) | 21.9991 | 20.7299 | 23.2683 | 21.385 | 22.6132
S, Sulphur (wt.%) | 24.0762 | 23.0687 | 25.0836 | 23.6482 | 24.5042
Sb, Antimony (ppm) | 1.551 | 1.1453 | 1.9567 | IND | IND
Sc, Scandium (ppm) | 5.0133 | 4.7682 | 5.2585 | IND | IND
Se, Selenium (ppm) | < 20 | IND | IND | IND | IND
SiO2, Silicon dioxide (wt.%) | 16.2663 | 15.5532 | 16.9793 | 15.9151 | 16.6174
Sm, Samarium (ppm) | 1.2625 | 1.1253 | 1.3997 | 1.1407 | 1.3842
Sn, Tin (ppm) | 1.4046 | 0.89716 | 1.912 | IND | IND
Sr, Strontium (ppm) | 88.926 | 84.1713 | 93.6808 | 86.4142 | 91.4378
Tb, Terbium (ppm) | 0.16228 | 0.14047 | 0.18409 | IND | IND
Th, Thorium (ppm) | 3.0386 | 2.7098 | 3.3674 | 2.8174 | 3.2598
TiO2, Titanium dioxide (wt.%) | 0.20028 | 0.19021 | 0.21035 | 0.19625 | 0.20432
Tl, Thallium (ppm) | 1.1963 | 1.0737 | 1.3189 | IND | IND
Tm, Thulium (ppm) | 0.075375 | 0.062064 | 0.088686 | IND | IND
U, Uranium (ppm) | 0.78191 | 0.64537 | 0.91846 | 0.66416 | 0.89966
V, Vanadium (ppm) | 100.1436 | 94.4236 | 105.8637 | 96.0367 | 104.2506
Y, Yttrium (ppm) | 4.8956 | 4.5512 | 5.24 | 4.6704 | 5.1208
Yb, Ytterbium (ppm) | 0.45833 | 0.36111 | 0.55555 | IND | IND
Zn, Zinc (ppm) | 819.9599 | 780.152 | 859.7678 | 802.1051 | 837.8147
Zr, Zirconium (ppm) | 34.7352 | 30.7939 | 38.6765 | 31.9828 | 37.4877
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 4.1549 | 4.0586 | 4.2512 | 4.1083 | 4.2015
CaO, Calcium oxide (wt.%) | 2.9053 | 2.836 | 2.9746 | 2.8755 | 2.935
Co, Cobalt (ppm) | 1614.9264 | 1469.8918 | 1759.9611 | IND | IND
Cr2O3, Chromium(III) oxide (wt.%) | 1.0973 | 1.0595 | 1.1351 | 1.082 | 1.1127
Cu, Copper (wt.%) | 0.39752 | 0.37577 | 0.41927 | 0.3833 | 0.41174
Fe2O3, Iron(III) oxide (wt.%) | 44.9255 | 44.2091 | 45.6419 | 44.6658 | 45.1853
Hf, Hafnium (ppm) | < 80 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 0.43279 | 0.41713 | 0.44846 | 0.42258 | 0.44301
MgO, Magnesium oxide (wt.%) | 2.6988 | 2.6082 | 2.7895 | 2.666 | 2.7316
MnO, Manganese oxide (wt.%) | 0.086806 | 0.078639 | 0.094972 | 0.082408 | 0.091204
Na2O, Sodium oxide (wt.%) | 0.4613 | 0.34873 | 0.57387 | 0.44403 | 0.47858
Ni, Nickel (wt.%) | 10.371 | 10.0814 | 10.6605 | 10.2907 | 10.4512
P2O5, Phosphorus(V) oxide (wt.%) | 0.060518 | 0.048836 | 0.0722 | IND | IND
S, Sulphur (wt.%) | 22.6331 | 20.5062 | 24.7601 | 22.4503 | 22.8159
SiO2, Silicon dioxide (wt.%) | 16.0654 | 15.8045 | 16.3263 | 15.9199 | 16.2109
TiO2, Titanium dioxide (wt.%) | 0.19867 | 0.18456 | 0.21277 | 0.18697 | 0.21036
V2O5, Vanadium(V) oxide (ppm) | 249.5149 | 131.4439 | 367.5859 | IND | IND
Zn, Zinc (ppm) | 804.5695 | 771.1314 | 838.0076 | 784.7571 | 824.3818
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 12.3141 | 12.0107 | 12.6175 | 12.2239 | 12.4043
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 84
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Ir | ppb | 178.7966 | Rh | ppb | 243.2777 | Ru | ppb | 420.237
NiS Fire Assay
Au | ppb | 153.5987 | Os | ppb | 119.0417 | Ru | ppb | 305.5667
Ir | ppb | 122.5627 | Rh | ppb | 119.9833
Infrared Combustion
C | wt.% | 0.125
4-Acid Digestion
B | ppm | 216.1667 | Ge | ppm | 0.49662 | Hg | ppm | 0.043333
Borate / Peroxide Fusion ICP
Ag | ppm | 1.6342 | Hg | ppm | 0.76667 | Ta | ppm | 0.27894
B | ppm | 24.75 | In | ppm | 0.13818 | Te | ppm | 1.0475
Be | ppm | 1.0784 | Na2O | wt.% | 0.63597 | W | ppm | 1.2116
Ge | ppm | 1.347 | Nb | ppm | 1.9886
Hf | ppm | 0.93028 | Re | ppm | < 0.1
Borate Fusion XRF
As | ppm | 122.8147 | Nb | ppm | 145.8333 | Sr | ppm | 84.1532
BaO | ppm | 248.798 | Pb | ppm | 177.0833 | W | ppm | 22.5
Bi | ppm | < 100 | Rb | ppm | 50 | Y | ppm | 41.6918
Ce | ppm | 111.9342 | Sb | ppm | 49.75 | Zr | ppm | 272.2222
La | ppm | 94.1479 | Sc | ppm | < 1
Mo | ppm | < 50 | Sn | ppm | < 50
Thermogravimetry
H2O- | wt.% | 0.21167
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 84
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 178.717 | 21.4072 | 135.9026 | 221.5314 | 114.4954 | 242.9386 | 11.98% | 23.96% | 35.93% | 169.7811 | 187.6528
Pd, ppb | 581.1081 | 35.2484 | 510.6114 | 651.6049 | 475.363 | 686.8533 | 6.066% | 12.13% | 18.2% | 552.0527 | 610.1635
Pt, ppb | 517.0015 | 18.8181 | 479.3653 | 554.6377 | 460.5472 | 573.4558 | 3.64% | 7.28% | 10.92% | 491.1514 | 542.8515
NiS Fire Assay
Pd, ppb | 556.22 | 30.7888 | 494.6424 | 617.7976 | 463.8536 | 648.5864 | 5.535% | 11.07% | 16.61% | 528.409 | 584.031
Pt, ppb | 486.8667 | 47.1913 | 392.4841 | 581.2492 | 345.2929 | 628.4405 | 9.693% | 19.39% | 29.08% | 462.5233 | 511.21
Infrared Combustion
S, wt.% | 24.2134 | 0.70532 | 22.8027 | 25.624 | 22.0974 | 26.3293 | 2.913% | 5.826% | 8.739% | 23.0027 | 25.424
4-Acid Digestion
Ag, ppm | 2.3765 | 0.071116 | 2.2342 | 2.5187 | 2.1631 | 2.5898 | 2.993% | 5.985% | 8.978% | 2.2576 | 2.4953
Al, wt.% | 2.2 | 0.099426 | 2.0011 | 2.3988 | 1.9017 | 2.4982 | 4.519% | 9.039% | 13.56% | 2.09 | 2.31
As, ppm | 167.1961 | 8.9718 | 149.2524 | 185.1397 | 140.2805 | 194.1116 | 5.366% | 10.73% | 16.1% | 158.8363 | 175.5559
Ba, ppm | 161.1092 | 13.8003 | 133.5087 | 188.7098 | 119.7084 | 202.51 | 8.566% | 17.13% | 25.7% | 153.0538 | 169.1647
Be, ppm | 0.44543 | 0.045003 | 0.35542 | 0.53543 | 0.31041 | 0.58044 | 10.1% | 20.21% | 30.31% | 0.42315 | 0.4677
Bi, ppm | 1.5616 | 0.094979 | 1.3716 | 1.7515 | 1.2766 | 1.8465 | 6.082% | 12.16% | 18.25% | 1.4835 | 1.6396
Ca, wt.% | 2.0344 | 0.078786 | 1.8768 | 2.192 | 1.798 | 2.2707 | 3.873% | 7.745% | 11.62% | 1.9327 | 2.1361
Cd, ppm | 2.4285 | 0.20715 | 2.0143 | 2.8428 | 1.8071 | 3.05 | 8.53% | 17.06% | 25.59% | 2.3071 | 2.55
Ce, ppm | 15.7988 | 0.47043 | 14.858 | 16.7397 | 14.3875 | 17.2101 | 2.978% | 5.955% | 8.933% | 15.0089 | 16.5887
Co, ppm | 1578.0701 | 58.0184 | 1462.0332 | 1694.1069 | 1404.0148 | 1752.1253 | 3.677% | 7.353% | 11.03% | 1499.1665 | 1656.9736
Cr, wt.% | 0.55884 | 0.10692 | 0.34499 | 0.77268 | 0.23807 | 0.8796 | 19.13% | 38.27% | 57.4% | 0.53089 | 0.58678
Cs, ppm | 1.8484 | 0.13283 | 1.5827 | 2.114 | 1.4499 | 2.2469 | 7.186% | 14.37% | 21.56% | 1.756 | 1.9408
Cu, wt.% | 0.39955 | 0.013839 | 0.37187 | 0.42723 | 0.35803 | 0.44107 | 3.464% | 6.927% | 10.39% | 0.37957 | 0.41953
Dy, ppm | 0.93845 | 0.044804 | 0.84884 | 1.0281 | 0.80404 | 1.0729 | 4.774% | 9.549% | 14.32% | 0.89153 | 0.98537
Er, ppm | 0.52575 | 0.035644 | 0.45446 | 0.59703 | 0.41881 | 0.63268 | 6.78% | 13.56% | 20.34% | 0.49946 | 0.55203
Eu, ppm | 0.33739 | 0.043089 | 0.25122 | 0.42357 | 0.20813 | 0.46666 | 12.77% | 25.54% | 38.31% | 0.32052 | 0.35426
Fe, wt.% | 31.5125 | 1.0212 | 29.47 | 33.5549 | 28.4487 | 34.5762 | 3.241% | 6.482% | 9.722% | 29.9368 | 33.0881
Ga, ppm | 6.5811 | 0.65759 | 5.2659 | 7.8963 | 4.6083 | 8.5539 | 9.992% | 19.98% | 29.98% | 6.252 | 6.9102
Gd, ppm | 1.1618 | 0.079464 | 1.0028 | 1.3207 | 0.92336 | 1.4001 | 6.84% | 13.68% | 20.52% | 1.1037 | 1.2198
Hf, ppm | 0.83852 | 0.073286 | 0.69195 | 0.9851 | 0.61867 | 1.0584 | 8.74% | 17.48% | 26.22% | 0.7966 | 0.88045
Ho, ppm | 0.18467 | 0.021076 | 0.14252 | 0.22682 | 0.12144 | 0.2479 | 11.41% | 22.83% | 34.24% | 0.17544 | 0.1939
In, ppm | 0.13206 | 0.012818 | 0.10643 | 0.1577 | 0.093607 | 0.17051 | 9.706% | 19.41% | 29.12% | 0.12546 | 0.13866
K, wt.% | 0.36758 | 0.011886 | 0.34381 | 0.39136 | 0.33193 | 0.40324 | 3.233% | 6.467% | 9.7% | 0.34921 | 0.38596
La, ppm | 8.1498 | 0.58274 | 6.9843 | 9.3152 | 6.4015 | 9.898 | 7.15% | 14.3% | 21.45% | 7.7423 | 8.5572
Li, ppm | 10.277 | 0.66933 | 8.9384 | 11.6157 | 8.269 | 12.285 | 6.513% | 13.03% | 19.54% | 9.7632 | 10.7909
Lu, ppm | 0.07291 | 0.0081524 | 0.056605 | 0.089214 | 0.048452 | 0.097367 | 11.18% | 22.36% | 33.54% | 0.069264 | 0.076555
Mg, wt.% | 1.5128 | 0.077981 | 1.3568 | 1.6687 | 1.2788 | 1.7467 | 5.155% | 10.31% | 15.46% | 1.4371 | 1.5884
Mn, wt.% | 0.063789 | 0.0036554 | 0.056478 | 0.0711 | 0.052823 | 0.074756 | 5.73% | 11.46% | 17.19% | 0.0606 | 0.066979
Mo, ppm | 5.0522 | 0.32598 | 4.4002 | 5.7042 | 4.0743 | 6.0302 | 6.452% | 12.9% | 19.36% | 4.7996 | 5.3048
Na, wt.% | 0.44587 | 0.013336 | 0.4192 | 0.47254 | 0.40586 | 0.48588 | 2.991% | 5.982% | 8.973% | 0.42358 | 0.46816
Nb, ppm | 2.3659 | 0.15956 | 2.0468 | 2.685 | 1.8872 | 2.8446 | 6.744% | 13.49% | 20.23% | 2.2476 | 2.4842
Nd, ppm | 6.7201 | 0.30401 | 6.1121 | 7.3282 | 5.8081 | 7.6322 | 4.524% | 9.048% | 13.57% | 6.3841 | 7.0562
Ni, wt.% | 10.4663 | 0.2042 | 10.0579 | 10.8747 | 9.8537 | 11.0789 | 1.951% | 3.902% | 5.853% | 9.943 | 10.9896
P, wt.% | 0.021907 | 0.001296 | 0.019315 | 0.024499 | 0.018019 | 0.025795 | 5.916% | 11.83% | 17.75% | 0.020812 | 0.023003
Pb, ppm | 74.4077 | 4.2729 | 65.8618 | 82.9536 | 61.5888 | 87.2265 | 5.743% | 11.49% | 17.23% | 70.6873 | 78.1281
Pr, ppm | 1.8212 | 0.11584 | 1.5895 | 2.0529 | 1.4737 | 2.1687 | 6.36% | 12.72% | 19.08% | 1.7301 | 1.9123
Rb, ppm | 21.4196 | 1.6221 | 18.1754 | 24.6639 | 16.5533 | 26.286 | 7.573% | 15.15% | 22.72% | 20.3486 | 22.4906
Re, ppm | 0.029452 | 0.0052261 | 0.019 | 0.039905 | 0.013774 | 0.045131 | 17.74% | 35.49% | 53.23% | 0.02798 | 0.030925
S, wt.% | 22.1417 | 1.2179 | 19.706 | 24.5775 | 18.4881 | 25.7953 | 5.5% | 11% | 16.5% | 21.0346 | 23.2488
Sb, ppm | 1.5204 | 0.12911 | 1.2622 | 1.7786 | 1.1331 | 1.9077 | 8.491% | 16.98% | 25.47% | 1.4444 | 1.5964
Sc, ppm | 5.2099 | 0.43846 | 4.333 | 6.0868 | 3.8945 | 6.5253 | 8.416% | 16.83% | 25.25% | 4.9494 | 5.4704
Se, ppm | 10.7094 | 1.4414 | 7.8267 | 13.5921 | 6.3853 | 15.0335 | 13.46% | 26.92% | 40.38% | 10.1739 | 11.2449
Sm, ppm | 1.3108 | 0.075551 | 1.1597 | 1.4619 | 1.0842 | 1.5375 | 5.764% | 11.53% | 17.29% | 1.2453 | 1.3764
Sn, ppm | 1.4943 | 0.15374 | 1.1868 | 1.8018 | 1.033 | 1.9555 | 10.29% | 20.58% | 30.87% | 1.4196 | 1.569
Sr, ppm | 86.8664 | 5.7654 | 75.3357 | 98.3971 | 69.5703 | 104.1625 | 6.637% | 13.27% | 19.91% | 82.5231 | 91.2097
Ta, ppm | 0.20525 | 0.043477 | 0.1183 | 0.29221 | 0.074823 | 0.33568 | 21.18% | 42.36% | 63.55% | 0.19499 | 0.21552
Tb, ppm | 0.175 | 0.015335 | 0.14433 | 0.20567 | 0.12899 | 0.221 | 8.763% | 17.53% | 26.29% | 0.16625 | 0.18375
Te, ppm | 1.2653 | 0.15299 | 0.95934 | 1.5713 | 0.80635 | 1.7243 | 12.09% | 24.18% | 36.27% | 1.2021 | 1.3286
Th, ppm | 3.1385 | 0.21782 | 2.7028 | 3.5741 | 2.485 | 3.7919 | 6.94% | 13.88% | 20.82% | 2.9816 | 3.2954
Ti, wt.% | 0.11712 | 0.0058037 | 0.10551 | 0.12872 | 0.099706 | 0.13453 | 4.955% | 9.911% | 14.87% | 0.11126 | 0.12297
Tl, ppm | 1.2381 | 0.061868 | 1.1144 | 1.3618 | 1.0525 | 1.4237 | 4.997% | 9.994% | 14.99% | 1.1762 | 1.3
Tm, ppm | 0.068555 | 0.014669 | 0.039217 | 0.097893 | 0.024548 | 0.11256 | 21.4% | 42.79% | 64.19% | 0.065127 | 0.071983
U, ppm | 0.82271 | 0.046747 | 0.72922 | 0.91621 | 0.68247 | 0.96295 | 5.682% | 11.36% | 17.05% | 0.78158 | 0.86385
V, ppm | 100.2302 | 5.2697 | 89.6908 | 110.7696 | 84.4212 | 116.0393 | 5.258% | 10.52% | 15.77% | 95.2187 | 105.2417
W, ppm | 1.3059 | 0.11391 | 1.0781 | 1.5337 | 0.96416 | 1.6476 | 8.723% | 17.45% | 26.17% | 1.2406 | 1.3712
Y, ppm | 4.7093 | 0.4338 | 3.8417 | 5.5769 | 3.4079 | 6.0107 | 9.212% | 18.42% | 27.63% | 4.4739 | 4.9448
Yb, ppm | 0.47047 | 0.031451 | 0.40757 | 0.53337 | 0.37612 | 0.56483 | 6.685% | 13.37% | 20.06% | 0.44695 | 0.494
Zn, ppm | 813.8033 | 45.5484 | 722.7065 | 904.9001 | 677.1581 | 950.4485 | 5.597% | 11.19% | 16.79% | 773.1132 | 854.4935
Zr, ppm | 28.5666 | 2.5971 | 23.3724 | 33.7608 | 20.7753 | 36.3579 | 9.091% | 18.18% | 27.27% | 27.1383 | 29.995
Borate / Peroxide Fusion ICP
Al2O3, wt.% | 4.1663 | 0.13859 | 3.8891 | 4.4435 | 3.7506 | 4.5821 | 3.327% | 6.653% | 9.98% | 3.958 | 4.3747
As, ppm | 175.4004 | 12.6301 | 150.1403 | 200.6605 | 137.5102 | 213.2906 | 7.201% | 14.4% | 21.6% | 166.6304 | 184.1704
Ba, ppm | 175.3227 | 11.1468 | 153.0292 | 197.6163 | 141.8824 | 208.7631 | 6.358% | 12.72% | 19.07% | 166.5566 | 184.0889
Bi, ppm | 1.3902 | 0.1855 | 1.0192 | 1.7612 | 0.83367 | 1.9467 | 13.34% | 26.69% | 40.03% | 1.3207 | 1.4597
CaO, wt.% | 2.9309 | 0.075065 | 2.7807 | 3.081 | 2.7057 | 3.1561 | 2.561% | 5.122% | 7.684% | 2.7843 | 3.0774
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 15.829 | 1.5998 | 12.6293 | 19.0286 | 11.0294 | 20.6285 | 10.11% | 20.21% | 30.32% | 15.0375 | 16.6204
Co, ppm | 1653.8378 | 94.8964 | 1464.0451 | 1843.6306 | 1369.1487 | 1938.527 | 5.738% | 11.48% | 17.21% | 1571.1459 | 1736.5297
Cr, wt.% | 0.74674 | 0.030199 | 0.68635 | 0.80714 | 0.65615 | 0.83734 | 4.044% | 8.088% | 12.13% | 0.70941 | 0.78408
Cs, ppm | 1.7278 | 0.15894 | 1.4099 | 2.0457 | 1.251 | 2.2046 | 9.199% | 18.4% | 27.6% | 1.6414 | 1.8142
Cu, wt.% | 0.40483 | 0.012167 | 0.38049 | 0.42916 | 0.36833 | 0.44133 | 3.006% | 6.011% | 9.017% | 0.38459 | 0.42507
Dy, ppm | 0.93436 | 0.10903 | 0.71629 | 1.1524 | 0.60726 | 1.2615 | 11.67% | 23.34% | 35.01% | 0.88764 | 0.98107
Er, ppm | 0.49702 | 0.061294 | 0.37444 | 0.61961 | 0.31314 | 0.68091 | 12.33% | 24.66% | 37% | 0.47217 | 0.52188
Eu, ppm | 0.29785 | 0.03623 | 0.22539 | 0.37031 | 0.18916 | 0.40654 | 12.16% | 24.33% | 36.49% | 0.28295 | 0.31274
Fe2O3, wt.% | 44.6442 | 1.7362 | 41.1718 | 48.1167 | 39.4355 | 49.8529 | 3.889% | 7.778% | 11.67% | 42.412 | 46.8764
Ga, ppm | 7.0535 | 0.75344 | 5.5466 | 8.5603 | 4.7931 | 9.3138 | 10.68% | 21.36% | 32.05% | 6.7008 | 7.4061
Gd, ppm | 1.1203 | 0.09509 | 0.93015 | 1.3105 | 0.83506 | 1.4056 | 8.488% | 16.98% | 25.46% | 1.0643 | 1.1764
Ho, ppm | 0.17549 | 0.024229 | 0.12703 | 0.22395 | 0.10281 | 0.24818 | 13.81% | 27.61% | 41.42% | 0.16672 | 0.18427
K2O, wt.% | 0.45348 | 0.035529 | 0.38242 | 0.52453 | 0.34689 | 0.56006 | 7.835% | 15.67% | 23.5% | 0.4308 | 0.47615
La, ppm | 8.3036 | 0.5254 | 7.2528 | 9.3544 | 6.7274 | 9.8798 | 6.327% | 12.65% | 18.98% | 7.8885 | 8.7188
Li, ppm | 12.846 | 2.0238 | 8.7984 | 16.8936 | 6.7746 | 18.9174 | 15.75% | 31.51% | 47.26% | 12.2037 | 13.4883
Lu, ppm | 0.075278 | 0.013833 | 0.047612 | 0.10294 | 0.033779 | 0.11678 | 18.38% | 36.75% | 55.13% | 0.071514 | 0.079042
MgO, wt.% | 2.6427 | 0.065587 | 2.5115 | 2.7739 | 2.4459 | 2.8394 | 2.482% | 4.964% | 7.445% | 2.5105 | 2.7748
MnO, wt.% | 0.085184 | 0.0043734 | 0.076438 | 0.093931 | 0.072064 | 0.098305 | 5.134% | 10.27% | 15.4% | 0.080925 | 0.089444
Mo, ppm | 5.6239 | 0.79813 | 4.0276 | 7.2202 | 3.2295 | 8.0183 | 14.19% | 28.38% | 42.58% | 5.3427 | 5.9051
Nd, ppm | 6.6726 | 0.5085 | 5.6556 | 7.6896 | 5.1471 | 8.1981 | 7.621% | 15.24% | 22.86% | 6.339 | 7.0063
Ni, wt.% | 10.5288 | 0.2614 | 10.006 | 11.0516 | 9.7446 | 11.313 | 2.483% | 4.965% | 7.448% | 10.0024 | 11.0553
P2O5, wt.% | 0.057341 | 0.0094847 | 0.038371 | 0.07631 | 0.028887 | 0.085795 | 16.54% | 33.08% | 49.62% | 0.054474 | 0.060208
Pb, ppm | 73.7087 | 7.0389 | 59.6308 | 87.7865 | 52.5919 | 94.8255 | 9.55% | 19.1% | 28.65% | 70.0233 | 77.3941
Pr, ppm | 1.8068 | 0.12803 | 1.5508 | 2.0629 | 1.4227 | 2.1909 | 7.086% | 14.17% | 21.26% | 1.7165 | 1.8972
Rb, ppm | 21.9991 | 1.0373 | 19.9244 | 24.0738 | 18.887 | 25.1111 | 4.715% | 9.431% | 14.15% | 20.8991 | 23.099
S, wt.% | 24.0762 | 0.71113 | 22.6539 | 25.4984 | 21.9428 | 26.2096 | 2.954% | 5.907% | 8.861% | 22.8724 | 25.28
Sb, ppm | 1.551 | 0.249 | 1.053 | 2.049 | 0.80402 | 2.298 | 16.05% | 32.11% | 48.16% | 1.4735 | 1.6286
Sc, ppm | 5.0133 | 0.10003 | 4.8133 | 5.2134 | 4.7132 | 5.3134 | 1.995% | 3.991% | 5.986% | 4.7627 | 5.264
Se, ppm | < 20 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SiO2, wt.% | 16.2663 | 0.7419 | 14.7825 | 17.7501 | 14.0406 | 18.492 | 4.561% | 9.122% | 13.68% | 15.453 | 17.0796
Sm, ppm | 1.2625 | 0.088817 | 1.0848 | 1.4401 | 0.99602 | 1.5289 | 7.035% | 14.07% | 21.11% | 1.1993 | 1.3256
Sn, ppm | 1.4046 | 0.33841 | 0.72776 | 2.0814 | 0.38935 | 2.4198 | 24.09% | 48.19% | 72.28% | 1.3344 | 1.4748
Sr, ppm | 88.926 | 3.6898 | 81.5464 | 96.3056 | 77.8566 | 99.9954 | 4.149% | 8.299% | 12.45% | 84.4797 | 93.3723
Tb, ppm | 0.16228 | 0.020224 | 0.12183 | 0.20273 | 0.10161 | 0.22295 | 12.46% | 24.92% | 37.39% | 0.15417 | 0.1704
Th, ppm | 3.0386 | 0.28501 | 2.4686 | 3.6086 | 2.1836 | 3.8936 | 9.38% | 18.76% | 28.14% | 2.8867 | 3.1905
TiO2, wt.% | 0.20028 | 0.0060051 | 0.18827 | 0.21229 | 0.18227 | 0.2183 | 2.998% | 5.997% | 8.995% | 0.19027 | 0.2103
Tl, ppm | 1.1963 | 0.091309 | 1.0137 | 1.3789 | 0.92235 | 1.4702 | 7.633% | 15.27% | 22.9% | 1.1365 | 1.2561
Tm, ppm | 0.075375 | 0.016958 | 0.041458 | 0.10929 | 0.0245 | 0.12625 | 22.5% | 45% | 67.5% | 0.071606 | 0.079144
U, ppm | 0.78191 | 0.097286 | 0.58734 | 0.97648 | 0.49005 | 1.0738 | 12.44% | 24.88% | 37.33% | 0.74282 | 0.82101
V, ppm | 100.1436 | 8.0218 | 84.1001 | 116.1872 | 76.0783 | 124.2089 | 8.01% | 16.02% | 24.03% | 95.1365 | 105.1508
Y, ppm | 4.8956 | 0.31722 | 4.2612 | 5.53 | 3.9439 | 5.8472 | 6.48% | 12.96% | 19.44% | 4.6508 | 5.1404
Yb, ppm | 0.45833 | 0.050795 | 0.35674 | 0.55992 | 0.30595 | 0.61072 | 11.08% | 22.17% | 33.25% | 0.43542 | 0.48125
Zn, ppm | 819.9599 | 39.5698 | 740.8203 | 899.0995 | 701.2505 | 938.6693 | 4.826% | 9.652% | 14.48% | 778.9619 | 860.9579
Zr, ppm | 34.7352 | 3.816 | 27.1032 | 42.3672 | 23.2872 | 46.1832 | 10.99% | 21.97% | 32.96% | 32.9985 | 36.472
Borate Fusion XRF
Al2O3, wt.% | 4.1549 | 0.090666 | 3.9736 | 4.3362 | 3.8829 | 4.4269 | 2.182% | 4.364% | 6.546% | 3.9472 | 4.3626
CaO, wt.% | 2.9053 | 0.068276 | 2.7687 | 3.0418 | 2.7005 | 3.1101 | 2.35% | 4.7% | 7.05% | 2.76 | 3.0506
Co, ppm | 1614.9264 | 83.9953 | 1446.9358 | 1782.9171 | 1362.9405 | 1866.9124 | 5.201% | 10.4% | 15.6% | 1534.1801 | 1695.6728
Cr2O3, wt.% | 1.0973 | 0.041355 | 1.0146 | 1.18 | 0.97325 | 1.2214 | 3.769% | 7.537% | 11.31% | 1.0424 | 1.1522
Cu, wt.% | 0.39752 | 0.027259 | 0.343 | 0.45204 | 0.31574 | 0.4793 | 6.857% | 13.71% | 20.57% | 0.37764 | 0.41739
Fe2O3, wt.% | 44.9255 | 0.76073 | 43.4041 | 46.447 | 42.6433 | 47.2077 | 1.693% | 3.387% | 5.08% | 42.6792 | 47.1718
Hf, ppm | < 80 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 0.43279 | 0.01384 | 0.40512 | 0.46047 | 0.39128 | 0.47431 | 3.198% | 6.395% | 9.593% | 0.41116 | 0.45443
MgO, wt.% | 2.6988 | 0.10395 | 2.4909 | 2.9067 | 2.387 | 3.0107 | 3.852% | 7.704% | 11.56% | 2.5639 | 2.8338
MnO, wt.% | 0.086806 | 0.010063 | 0.066681 | 0.10693 | 0.056618 | 0.11699 | 11.59% | 23.18% | 34.78% | 0.082466 | 0.091146
Na2O, wt.% | 0.4613 | 0.095869 | 0.26957 | 0.65304 | 0.1737 | 0.74891 | 20.78% | 41.56% | 62.35% | 0.43824 | 0.48437
Ni, wt.% | 10.371 | 0.43212 | 9.5067 | 11.2352 | 9.0746 | 11.6673 | 4.167% | 8.333% | 12.5% | 9.8524 | 10.8895
P2O5, wt.% | 0.060518 | 0.012852 | 0.034814 | 0.086222 | 0.021962 | 0.099074 | 21.24% | 42.47% | 63.71% | 0.057492 | 0.063544
S, wt.% | 22.6331 | 2.0903 | 18.4525 | 26.8137 | 16.3622 | 28.904 | 9.236% | 18.47% | 27.71% | 21.5015 | 23.7648
SiO2, wt.% | 16.0654 | 0.19818 | 15.669 | 16.4618 | 15.4708 | 16.6599 | 1.234% | 2.467% | 3.701% | 15.2621 | 16.8687
TiO2, wt.% | 0.19867 | 0.01775 | 0.16317 | 0.23416 | 0.14542 | 0.25191 | 8.935% | 17.87% | 26.8% | 0.18873 | 0.2086
V2O5, ppm | 249.5149 | 85.3599 | 78.7951 | 420.2347 | 0 | 505.5946 | 34.21% | 68.42% | 102.6% | 237.0391 | 261.9906
Zn, ppm | 804.5695 | 18.3505 | 767.8684 | 841.2705 | 749.5179 | 859.621 | 2.281% | 4.562% | 6.842% | 764.341 | 844.7979
Thermogravimetry
LOI1000, wt.% | 12.3141 | 0.41377 | 11.4866 | 13.1416 | 11.0728 | 13.5554 | 3.36% | 6.72% | 10.08% | 11.6984 | 12.9298
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 84
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*NAA | fire assay with neutron activation analysis
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
FA-NS*MS | Fire Assay Nickel Sulphide Collection with Mass Spectrometry: ICP-MS finish
FA-NS*OES | Fire Assay Nickel Sulphide Collection with Optical Emmision Spectrometry [aka: A(tomic)ES, ICP-OES] finish
FA-NS*OES/MS | Fire Assay Nickel Sulphide Collection with ICP-OES or ICP-MS finish
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Text Values:
N.A. | Not Applicable (Lab 23)
NA | Not Applicable (Lab 11)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 84
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AGAT Laboratories, Thunder Bay, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Inspectorate (BV), Shanghai, Bao Shan District, China
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
MINTEK Analytical Services, Randburg, South Africa
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Ankara, Anatolia, Turkey
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 84 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppb) | Upscaled Value (ppb)
2 | 1 | ANSLu | 1 | 10 | 5 | 311897 | 0.08695 | 193.9552 | 191.5569
2 | 1 | ANSLu | 1 | 12 | 12 | 311898 | 0.084002 | 193.2298 | 191.5182
2 | 1 | ANSLu | 1 | 5 | 3 | 311899 | 0.083916 | 201.1397 | 191.9402
2 | 1 | ANSLu | 1 | 2 | 1 | 311900 | 0.087545 | 197.9711 | 191.7711
2 | 1 | ANSLu | 1 | 15 | 2 | 311901 | 0.08389 | 179.5924 | 190.7905
2 | 1 | ANSLu | 1 | 14 | 7 | 311902 | 0.086893 | 175.7864 | 190.5875
2 | 1 | ANSLu | 1 | 13 | 7 | 311903 | 0.083605 | 189.4892 | 191.3186
2 | 1 | ANSLu | 1 | 7 | 10 | 311904 | 0.086878 | 191.2817 | 191.4142
2 | 1 | ANSLu | 1 | 6 | 3 | 311905 | 0.08487 | 189.5659 | 191.3227
2 | 1 | ANSLu | 1 | 11 | 12 | 311906 | 0.087767 | 201.6053 | 191.965
2 | 1 | ANSLu | 1 | 8 | 10 | 311907 | 0.084446 | 185.5371 | 191.1077
2 | 1 | ANSLu | 1 | 19 | 4 | 311908 | 0.087146 | 195.1033 | 191.6181
2 | 1 | ANSLu | 1 | 9 | 5 | 311909 | 0.082947 | 185.8569 | 191.1248
2 | 1 | ANSLu | 1 | 18 | 9 | 311910 | 0.082699 | 194.0396 | 191.5614
2 | 1 | ANSLu | 1 | 20 | 4 | 311911 | 0.083109 | 186.7386 | 191.1718
2 | 1 | ANSLu | 1 | 1 | 1 | 311912 | 0.08279 | 182.074 | 190.9229
2 | 1 | ANSLu | 1 | 17 | 9 | 311913 | 0.087179 | 199.8521 | 191.8715
2 | 1 | ANSLu | 1 | 4 | 8 | 311914 | 0.087542 | 189.4274 | 191.3153
2 | 1 | ANSLu | 1 | 16 | 2 | 311915 | 0.086128 | 206.6245 | 192.2328
2 | 1 | ANSLu | 1 | 3 | 8 | 311916 | 0.087692 | 189.5634 | 191.3225
Mean | [6 empty] | 0.0854 | 191.4217 | 191.4217
Median | [7 empty] | 190.4238 | 191.3684
SD | [7 empty] | 7.7706 | 0.41459
RSD | [7 empty] | 4.059% | 0.2166%
RSD Upscale Factor | [7 empty] | 0.053354
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 84 (Certified Value 179 ppb) | [21 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 23 | ppb
|  |  | INAA | FA*MS | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | 0
|  |  | 0.085g | 30g | 30g | 25g | 30g | 50g | 30g | 40g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 1
| 1 | 1 | 182.074 | 171 | 183 | 196 | 184 | 207 | 70.018 | 144 | 188 | 125.5 | 198 | 176 | 198 | 193 | 180 | 201 | 150 | 162.0707 | 202 | 148 | 1
| 1 | 2 | 197.9711 | 162 | 197 | 190 | 191 | 201 | 69.97 | 144 | 208 | 145.4 | 184 | 183 | 192 | 189 | 180 | 200 | 130 | 152.7972 | 225 | 146 | 3
| 1 | 3 | 189.5634 | 160 | 181 | 198 | 189 | 202 | 70.078 | 146 | 193 | 133.536 | 191 | 184 | 175 | 195 | 180 | 189 | 140 | 151.5286 | 182 | 151 | 16
| 1 | 4 | 189.4274 | 173 | 192 | 195 | 175 | 202 | 69.86 | 147 | 187 | 136.58 | 194 | 152 | 193 | 190 | 187 | 174 | 140 | 158.0019 | 145 | 141 | 178.717
| 1 | 5 | 201.1397 | 180 | 186 | 199 | 195 | 219 | 69.889 | 145 | 199 | 137.2 | 195 | 168 | 193 | 191 | 188 | 193 | 140 | 163.4227 | 155 | 142 | 7
| 1 | 6 | 189.5659 | 172 | 194 | 198 | 180 | 216 | 69.914 | 146 | 202 | 123 | 188 | 192 | 185 | 182 | 189 | 197 | 160 | 156.3109 | 166 | 146
|  | 7 | 191.2817
|  | 8 | 185.5371
|  | 9 | 185.8569
|  | 10 | 193.9552
|  | 11 | 201.6053
|  | 12 | 193.2298
|  | 13 | 189.4892
|  | 14 | 175.7864
|  | 15 | 179.5924
|  | 16 | 206.6245
|  | 17 | 199.8521
|  | 18 | 194.0396
|  | 19 | 195.1033
|  | 20 | 186.7386
| Mean |  | 191.4217 | 169.6667 | 188.8333 | 196 | 185.6667 | 207.8333 | 69.9548 | 145.3333 | 196.1667 | 133.536 | 191.6667 | 175.8333 | 189.3333 | 190 | 184 | 192.3333 | 143.3333 | 157.3553 | 179.1667 | 145.6667
| Median |  | 190.4238 | 171.5 | 189 | 197 | 186.5 | 204.5 | 69.942 | 145.5 | 196 | 135.058 | 192.5 | 179.5 | 192.5 | 190.5 | 183.5 | 195 | 140 | 157.1564 | 174 | 146
| Std Dev. |  | 7.7706 | 7.4476 | 6.4317 | 3.2863 | 7.4207 | 7.8337 | 0.082927 | 1.2111 | 8.2805 | 8.2344 | 5.0859 | 14.2045 | 8.165 | 4.4721 | 4.4272 | 10.0333 | 10.328 | 4.8017 | 30.1822 | 3.7238
| Rel.Std.Dev. |  | 4.059% | 4.39% | 3.406% | 1.677% | 3.997% | 3.769% | 0.1185% | 0.8333% | 4.221% | 6.166% | 2.654% | 8.078% | 4.312% | 2.354% | 2.406% | 5.217% | 7.206% | 3.052% | 16.85% | 2.556%
| PDM3 |  | 7.109% | -5.064% | 5.661% | 9.671% | 3.889% | 16.29% | -60.86% | -18.68% | 9.764% | -25.28% | 7.246% | -1.614% | 5.94% | 6.313% | 2.956% | 7.619% | -19.8% | -11.95% | 0.2516% | -18.49%
| Z-Score (Absolute) |  | NA | 0.81 | 0.27 | 0.67 | 0.09 | 1.34 | 6.41 | 2.17 | 0.68 | 2.84 | 0.43 | 0.46 | 0.3 | 0.34 | 0 | 0.47 | 2.29 | 1.5 | 0.27 | 2.15
| Analytical results for Ir in OREAS 84 (Indicative Value 179 ppb) | [21 empty] | Indicative
Ir | Round | Replicate | Lab | [19 empty] | 1
| No | No | 06 | [19 empty] | ppb
|  |  | FA*OES | [19 empty] | 0
|  |  | 30g | [19 empty] | 0
| 1 | 1 | 179.979 | [19 empty] | 1
| 1 | 2 | 178.4979 | [19 empty] | 1
| 1 | 3 | 178.1179 | [19 empty] | 16
| 1 | 4 | 178.479 | [19 empty] | 178.7966
| 1 | 5 | 179.0079 | [19 empty] | 7
| 1 | 6 | 178.6979
| Mean |  | 178.7966
| Median |  | 178.5979
| Std Dev. |  | 0.64857
| Rel.Std.Dev. |  | 0.3627%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pd in OREAS 84 (Certified Value 581 ppb) | [21 empty] | Certified
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 23 |  | ppb
|  |  | FA*MS | FA*OES | FA*MS | FA*OES | FA*MS | FA*MS | FA*MS | FA*OES | FA*OES | FA*OES | FA*AAS | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS |  | 0
|  |  | 30g | 30g | 25g | 30g | 50g | 30g | 40g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g |  | 0
| 1 | 1 | 538 | 558 | 608.7 | 566 | 623 | 654.9 | 604 | 604 | 518.5 | 595 | 520 | 596 | 601 | 551 | 605 | 630 | 534 | 556 | 549 |  | 1
| 1 | 2 | 566 | 604 | 603.4 | 573 | 596 | 653.8 | 600 | 645 | 519.9 | 601 | 542 | 607 | 598 | 556 | 599 | 570 | 534.8269 | 536 | 554 |  | #N/A
| 1 | 3 | 569 | 561 | 602.7 | 562 | 595 | 654.4 | 601 | 619 | 523.8 | 606 | 551 | 560 | 606 | 568 | 585 | 600 | 535.5852 | 565 | 566 |  | 16
| 1 | 4 | 538 | 569 | 612.1 | 544 | 602 | 653.9 | 605 | 610 | 551.2 | 612 | 514 | 562 | 600 | 560 | 538 | 600 | 535.41 | 552 | 562 |  | 581.1081
| 1 | 5 | 548 | 571 | 602.9 | 592 | 603 | 653.5 | 606 | 639 | 535 | 607 | 512 | 622 | 603 | 572 | 589 | 620 | 538.85 | 573 | 566 |  | 8
| 1 | 6 | 525 | 589 | 594.3 | 559 | 586 | 652.9 | 600 | 639 | 545.8 | 593 | 560 | 606 | 578 | 556 | 600 | 630 | 539.9558 | 550 | 567
| Mean |  | 547.3333 | 575.3333 | 604.0167 | 566 | 600.8333 | 653.9 | 602.6667 | 626 | 532.3667 | 602.3333 | 533.1667 | 592.1667 | 597.6667 | 560.5 | 586 | 608.3333 | 536.438 | 555.3333 | 560.6667
| Median |  | 543 | 570 | 603.15 | 564 | 599 | 653.85 | 602.5 | 629 | 529.4 | 603.5 | 531 | 601 | 600.5 | 558 | 594 | 610 | 535.4976 | 554 | 564
| Std Dev. |  | 17.2704 | 17.7388 | 6.0836 | 15.9625 | 12.4486 | 0.6957 | 2.6583 | 17.2511 | 13.8793 | 7.3666 | 20.5175 | 25.5376 | 10.0133 | 7.9937 | 24.6577 | 23.1661 | 2.3881 | 12.801 | 7.4744
| Rel.Std.Dev. |  | 3.155% | 3.083% | 1.007% | 2.82% | 2.072% | 0.1064% | 0.4411% | 2.756% | 2.607% | 1.223% | 3.848% | 4.313% | 1.675% | 1.426% | 4.208% | 3.808% | 0.4452% | 2.305% | 1.333%
| PDM3 |  | -5.812% | -0.9938% | 3.942% | -2.6% | 3.394% | 12.53% | 3.71% | 7.725% | -8.388% | 3.653% | -8.25% | 1.903% | 2.849% | -3.546% | 0.8418% | 4.685% | -7.687% | -4.435% | -3.518%
| Z-Score (Absolute) |  | 1.17 | 0.32 | 0.54 | 0.6 | 0.45 | 2.05 | 0.5 | 1.21 | 1.62 | 0.49 | 1.6 | 0.19 | 0.35 | 0.77 | 0 | 0.67 | 1.5 | 0.93 | 0.76
| Analytical results for Pt in OREAS 84 (Certified Value 517 ppb) | [21 empty] | Certified
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 23 |  | ppb
|  |  | FA*MS | FA*OES | FA*MS | FA*OES | FA*MS | FA*MS | FA*MS | FA*OES | FA*OES | FA*OES | FA*AAS | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS |  | 0
|  |  | 30g | 30g | 25g | 30g | 50g | 30g | 40g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g |  | 0
| 1 | 1 | 489 | 494 | 526.3 | 497 | 535 | 574.358 | 517 | 524 | 507.8 | 525 | 420 | 521 | 534 | 500 | 530 | 530 | 499 | 413 | 520 |  | 1
| 1 | 2 | 527 | 538 | 514.8 | 510 | 505 | 575.5 | 514 | 572 | 511.6 | 530 | 460 | 531 | 522 | 515 | 530 | 500 | 500 | 416 | 510 |  | #N/A
| 1 | 3 | 529 | 500 | 522.7 | 517 | 502 | 574.43 | 516 | 546 | 508.3 | 520 | 440 | 488 | 531 | 517 | 520 | 540 | 501 | 429 | 520 |  | 16
| 1 | 4 | 494 | 510 | 526 | 492 | 507 | 575.03 | 518 | 526 | 510.3 | 529 | 450 | 475 | 534 | 509 | 480 | 520 | 509 | 422 | 490 |  | 517.0015
| 1 | 5 | 500 | 512 | 532.3 | 542 | 512 | 574.1 | 520 | 565 | 491.2 | 518 | 420 | 558 | 533 | 506 | 520 | 550 | 502 | 427 | 520 |  | 9
| 1 | 6 | 475 | 529 | 520 | 505 | 498 | 573.6 | 522 | 567 | 505.84 | 516 | 460 | 522 | 510 | 506 | 530 | 560 | 508 | 403 | 490
| Mean |  | 502.3333 | 513.8333 | 523.6833 | 510.5 | 509.8333 | 574.503 | 517.8333 | 550 | 505.84 | 523 | 441.6667 | 515.8333 | 527.3333 | 508.8333 | 518.3333 | 533.3333 | 503.1667 | 418.3333 | 508.3333
| Median |  | 497 | 511 | 524.35 | 507.5 | 506 | 574.394 | 517.5 | 555.5 | 508.05 | 522.5 | 445 | 521.5 | 532 | 507.5 | 525 | 535 | 501.5 | 419 | 515
| Std Dev. |  | 21.5376 | 16.8335 | 5.9985 | 17.8298 | 13.1972 | 0.67467 | 2.8577 | 21.2885 | 7.447 | 5.8652 | 18.3485 | 30.0494 | 9.6264 | 6.3061 | 19.4079 | 21.6025 | 4.2622 | 9.7091 | 14.7196
| Rel.Std.Dev. |  | 4.288% | 3.276% | 1.145% | 3.493% | 2.589% | 0.1174% | 0.5519% | 3.871% | 1.472% | 1.121% | 4.154% | 5.825% | 1.825% | 1.239% | 3.744% | 4.05% | 0.8471% | 2.321% | 2.896%
| PDM3 |  | -2.837% | -0.6128% | 1.292% | -1.258% | -1.386% | 11.12% | 0.1609% | 6.383% | -2.159% | 1.16% | -14.57% | -0.2259% | 1.998% | -1.58% | 0.2576% | 3.159% | -2.676% | -19.08% | -1.677%
| Z-Score (Absolute) |  | 0.99 | 0.08 | 0.7 | 0.34 | 0.4 | 4.73 | 0.24 | 2.79 | 0.71 | 0.65 | 5.8 | 0.08 | 0.99 | 0.48 | 0.28 | 1.47 | 0.92 | NA | 0.52
Results from Laboratory 1.18 were excluded due to a low bias observed in the Control sample.
| Analytical results for Rh in OREAS 84 (Indicative Value 243 ppb) | [21 empty] | Indicative
Rh | Round | Replicate | Lab | [19 empty] | 1
| No | No | 06 | [19 empty] | ppb
|  |  | FA*OES | [19 empty] | 0
|  |  | 30g | [19 empty] | 0
| 1 | 1 | 243.086 | [19 empty] | 1
| 1 | 2 | 244.156 | [19 empty] | 1
| 1 | 3 | 243.316 | [19 empty] | 16
| 1 | 4 | 244.096 | [19 empty] | 243.2777
| 1 | 5 | 241.946 | [19 empty] | 7
| 1 | 6 | 243.066
| Mean |  | 243.2777
| Median |  | 243.201
| Std Dev. |  | 0.81209
| Rel.Std.Dev. |  | 0.3338%
| PDM3 |  | -1.443e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 84 (Indicative Value 420 ppb) | [21 empty] | Indicative
Ru | Round | Replicate | Lab | [19 empty] | 1
| No | No | 06 | [19 empty] | ppb
|  |  | FA*OES | [19 empty] | 0
|  |  | 30g | [19 empty] | 0
| 1 | 1 | 421.822 | [19 empty] | 1
| 1 | 2 | 420.252 | [19 empty] | 1
| 1 | 3 | 418.702 | [19 empty] | 16
| 1 | 4 | 420.442 | [19 empty] | 420.237
| 1 | 5 | 420.542 | [19 empty] | 7
| 1 | 6 | 419.662
| Mean |  | 420.237
| Median |  | 420.347
| Std Dev. |  | 1.0329
| Rel.Std.Dev. |  | 0.2458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C104: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C122: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fire Assay (NiS)
| Analytical results for Au in OREAS 84 (Indicative Value 154 ppb) | [7 empty] | Indicative
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 03 | 07 | 17 | 21 | ppb
|  |  | INAA | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*OES/MS | FA-NS*OES | 0
|  |  | 0.085g | 10g | 25g | 10g | 15g | 200g | 0
| 1 | 1 | 182.074 | 181 | 186 | 114 | 82.0626 | 230 | 1
| 1 | 2 | 197.9711 | 170 | 177 | 117 | 89.9674 | 220 | 3
| 1 | 3 | 189.5634 | 168 | 176 | 137 | 83.562 | 210 | 16
| 1 | 4 | 189.4274 | 182 | 176 | 116 | 72.1653 | 210 | 153.5987
| 1 | 5 | 201.1397 | 170 | 176 | 115 | 83.5041 | 220 | 9
| 1 | 6 | 189.5659 | 170 | 177 | 112 | 86.7 | 200
|  | 7 | 191.2817
|  | 8 | 185.5371
|  | 9 | 185.8569
|  | 10 | 193.9552
|  | 11 | 201.6053
|  | 12 | 193.2298
|  | 13 | 189.4892
|  | 14 | 175.7864
|  | 15 | 179.5924
|  | 16 | 206.6245
|  | 17 | 199.8521
|  | 18 | 194.0396
|  | 19 | 195.1033
|  | 20 | 186.7386
| Mean |  | 191.4217 | 173.5 | 178 | 118.5 | 82.9936 | 215
| Median |  | 190.4238 | 170 | 176.5 | 115.5 | 83.533 | 215
| Std Dev. |  | 7.7706 | 6.253 | 3.9497 | 9.225 | 6.0179 | 10.4881
| Rel.Std.Dev. |  | 4.059% | 3.604% | 2.219% | 7.785% | 7.251% | 4.878%
| PDM3 |  | 24.62% | 12.96% | 15.89% | -22.85% | -45.97% | 39.98%
| Z-Score (Absolute) |  | NA | 0 | 0.07 | 0.89 | 1.47 | 0.67
| Analytical results for Ir in OREAS 84 (Indicative Value 123 ppb) | [7 empty] | Indicative
Ir | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 07 | 10 | 17 | 21 | ppb
|  |  | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*MS | FA-NS*OES/MS | FA-NS*OES | 0
|  |  | 10g | 25g | 10g | 30g | 15g | 200g | 0
| 1 | 1 | 139 | 157 | 74 | 131 | 84.536 | 165 | 1
| 1 | 2 | 127 | 149 | 76 | 143 | 73.7053 | 170 | 1
| 1 | 3 | 126 | 146 | 74 | 142 | 87.4978 | 160 | 16
| 1 | 4 | 140 | 145 | 78 | 130 | 78.3336 | 170 | 122.5627
| 1 | 5 | 135 | 148 | 69 | 132 | 87.762 | 210 | 10
| 1 | 6 | 126 | 148 | 73 | 116 | 74.4233 | 170
| Mean |  | 132.1667 | 148.8333 | 74 | 132.3333 | 81.043 | 174.1667
| Median |  | 131 | 148 | 74 | 131.5 | 81.4348 | 170
| Std Dev. |  | 6.6156 | 4.2622 | 3.0332 | 9.8116 | 6.3877 | 18.0046
| Rel.Std.Dev. |  | 5.006% | 2.864% | 4.099% | 7.414% | 7.882% | 10.34%
| PDM3 |  | 7.836% | 21.43% | -39.62% | 7.972% | -33.88% | 42.1%
| Z-Score (Absolute) |  | 0 | 0.38 | 1.34 | 0 | 1.18 | 0.97
| Analytical results for Os in OREAS 84 (Indicative Value 119 ppb) | [7 empty] | Indicative
Os | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 07 | 10 | 21 |  | ppb
|  |  | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*MS | FA-NS*OES |  | 0
|  |  | 10g | 25g | 10g | 30g | 200g |  | 0
| 1 | 1 | 99 | 127 | 83 | 196 | 260 |  | 1
| 1 | 2 | 77 | 137 | 82 | 201 | 260 |  | 3
| 1 | 3 | 98 | 134 | 83 | 188 | 280 |  | 16
| 1 | 4 | 63 | 119 | 82 | 197 | 270 |  | 119.0417
| 1 | 5 | 85 | 111 | 88 | 159 | 340 |  | 9
| 1 | 6 | 111 | 119 | 85 | 138 | 320
| Mean |  | 88.8333 | 124.5 | 83.8333 | 179.8333 | 288.3333
| Median |  | 91.5 | 123 | 83 | 192 | 275
| Std Dev. |  | 17.3253 | 9.9549 | 2.3166 | 25.514 | 33.7145
| Rel.Std.Dev. |  | 19.5% | 7.996% | 2.763% | 14.19% | 11.69%
| PDM3 |  | -25.38% | 4.585% | -29.58% | 51.07% | 142.2%
| Z-Score (Absolute) |  | 0.59 | 0 | 0.67 | 0.92 | 2.72
| Analytical results for Pd in OREAS 84 (Certified Value 556 ppb) | [7 empty] | Certified
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 07 | 10 | 17 | 21 | ppb
|  |  | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*MS | FA-NS*OES/MS | FA-NS*OES | 0
|  |  | 10g | 25g | 10g | 30g | 15g | 200g | 0
| 1 | 1 | 571 | 620 | 525 | 549 | 324.4411 | 515 | 1
| 1 | 2 | 538 | 583 | 526 | 586 | 359.3455 | 590 | #N/A
| 1 | 3 | 530 | 584 | 522 | 556 | 340.7664 | 560 | 16
| 1 | 4 | 575 | 584 | 524 | 573 | 334.4655 | 580 | 556.22
| 1 | 5 | 529 | 613 | 526 | 549 | 339.4515 | 350 | 11
| 1 | 6 | 504 | 576 | 527 | 451 | 345.3848 | 550
| Mean |  | 541.1667 | 593.3333 | 525 | 544 | 340.6425 | 524.1667
| Median |  | 534 | 584 | 525.5 | 552.5 | 340.109 | 555
| Std Dev. |  | 27.1987 | 18.3267 | 1.7889 | 47.8498 | 11.6135 | 89.2422
| Rel.Std.Dev. |  | 5.026% | 3.089% | 0.3407% | 8.796% | 3.409% | 17.03%
| PDM3 |  | -2.706% | 6.672% | -5.613% | -2.197% | -38.76% | -5.763%
| Z-Score (Absolute) |  | 0.55 | 4.1 | 0.55 | 0.74 | 13.09 | 0.61
| Analytical results for Pt in OREAS 84 (Certified Value 487 ppb) | [7 empty] | Certified
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 07 | 10 | 17 | 21 | ppb
|  |  | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*MS | FA-NS*OES/MS | FA-NS*OES | 0
|  |  | 10g | 25g | 10g | 30g | 15g | 200g | 0
| 1 | 1 | 521 | 580 | 472 | 512 | 265.5392 | 515 | 1
| 1 | 2 | 486 | 506 | 471 | 530 | 268.4637 | 490 | #N/A
| 1 | 3 | 469 | 525 | 415 | 521 | 279.0966 | 450 | 16
| 1 | 4 | 506 | 509 | 417 | 504 | 276.07 | 480 | 486.8667
| 1 | 5 | 399 | 527 | 414 | 585 | 290.5161 | 480 | 12
| 1 | 6 | 505 | 519 | 416 | 432 | 292.9378 | 450
| Mean |  | 481 | 527.6667 | 434.1667 | 514 | 278.7705 | 477.5
| Median |  | 495.5 | 522 | 416.5 | 516.5 | 277.5833 | 480
| Std Dev. |  | 44.0318 | 26.979 | 28.9373 | 49.3437 | 11.2015 | 24.8495
| Rel.Std.Dev. |  | 9.154% | 5.113% | 6.665% | 9.6% | 4.018% | 5.204%
| PDM3 |  | -1.205% | 8.38% | -10.82% | 5.573% | -42.74% | -1.924%
| Z-Score (Absolute) |  | 0.03 | 0.82 | 0.76 | 0.59 | 3.39 | 0.03
| Analytical results for Rh in OREAS 84 (Indicative Value 120 ppb) | [7 empty] | Indicative
Rh | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 07 | 10 | 17 | 21 | ppb
|  |  | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*MS | FA-NS*OES/MS | FA-NS*OES | 0
|  |  | 10g | 25g | 10g | 30g | 15g | 200g | 0
| 1 | 1 | 110 | 137 | 61 | 129 | 60.9368 | 120 | 1
| 1 | 2 | 109 | 123 | 61 | 116 | 60.6556 | 130 | 1
| 1 | 3 | 108 | 122 | 61 | 123 | 62.8043 | 120 | 16
| 1 | 4 | 117 | 127 | 62 | 121 | 58.0438 | 120 | 119.9833
| 1 | 5 | 91.8 | 127 | 62 | 121 | 59.1756 | 120 | 10
| 1 | 6 | 104 | 124 | 61 | 100 | 63.9434 | 120
| Mean |  | 106.6333 | 126.6667 | 61.3333 | 118.3333 | 60.9266 | 121.6667
| Median |  | 108.5 | 125.5 | 61 | 121 | 60.7962 | 120
| Std Dev. |  | 8.4052 | 5.465 | 0.5164 | 9.913 | 2.1944 | 4.0825
| Rel.Std.Dev. |  | 7.882% | 4.315% | 0.842% | 8.377% | 3.602% | 3.355%
| PDM3 |  | -11.13% | 5.57% | -48.88% | -1.375% | -49.22% | 1.403%
| Z-Score (Absolute) |  | 0.34 | 0.82 | 2.95 | 0.34 | 2.98 | 0.53
| Analytical results for Ru in OREAS 84 (Indicative Value 306 ppb) | [7 empty] | Indicative
Ru | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 07 | 10 | 17 | 21 | ppb
|  |  | FA*NAA | FA-NS*MS | FA-NS*OES/MS | FA-NS*MS | FA-NS*OES/MS | FA-NS*OES | 0
|  |  | 10g | 25g | 10g | 30g | 15g | 200g | 0
| 1 | 1 | 321 | 332 | 184 | 317 | 189.1972 | 310 | 1
| 1 | 2 | 325 | 297 | 156 | 307 | 188.4518 | 340 | 1
| 1 | 3 | 228 | 283 | 167 | 324 | 189.1831 | 330 | 16
| 1 | 4 | 294 | 313 | 160 | 328 | 183.8408 | 320 | 305.5667
| 1 | 5 | 242 | 307 | 178 | 312 | 154.6233 | 310 | 9
| 1 | 6 | 262 | 314 | 151 | 262 | 188.5198 | 300
| Mean |  | 278.6667 | 307.6667 | 166 | 308.3333 | 182.3027 | 318.3333
| Median |  | 278 | 310 | 163.5 | 314.5 | 188.4858 | 315
| Std Dev. |  | 40.8982 | 16.6333 | 12.8841 | 23.9555 | 13.7102 | 14.7196
| Rel.Std.Dev. |  | 14.68% | 5.406% | 7.762% | 7.769% | 7.521% | 4.624%
| PDM3 |  | -8.803% | 0.6872% | -45.67% | 0.9054% | -40.34% | 4.178%
| Z-Score (Absolute) |  | 0.48 | 0.48 | 4.25 | 0.51 | 3.71 | 0.84
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 84 (Indicative Value 0.125 wt.%) | [22 empty] | Indicative
C | Round | Replicate | Lab | [20 empty] | 1
| No | No | 05 | [20 empty] | wt.%
|  |  | IRC | [20 empty] | 3
3
| 1 | 1 | 0.12 | [20 empty] | 1
| 1 | 2 | 0.11 | [20 empty] | 6
| 1 | 3 | 0.1 | [20 empty] | 16
| 1 | 4 | 0.11 | [20 empty] | 0.125
| 1 | 5 | 0.13 | [20 empty] | 12
| 1 | 6 | 0.18
| Mean |  | 0.125
| Median |  | 0.115
| Std Dev. |  | 0.02881
| Rel.Std.Dev. |  | 23.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 84 (Certified Value 24.21 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 23.9 | 24 | 25.08 | 23.9 | 25.25 | 24.563 | 24.59 | 23.8 | 23.53 | 23.7 | 24.7 | 28.5 | 23.5 | 23.84 | 23.2 | 24.5984 | 23.4 | 24.9694 | 21.4 | 23.883 | 24.8 | 1
| 1 | 2 | 23.9 | 24.2 | 25.27 | 23.8 | 25.23 | 24.1693 | 24.43 | 23.6 | 24.04 | 23.7 | 24.6 | 27.67 | 23.8 | 23.866 | 23.6 | 24.3834 | 23.1 | 24.9375 | 24.6 | 23.898 | 24.8 | #N/A
| 1 | 3 | 24.5 | 24.1 | 25.22 | 23.9 | 25.32 | 24.2612 | 24.69 | 24.1 | 24.17 | 23.9 | 24.5 | 27.62 | 22.5 | 23.886 | 23 | 24.437 | 23.5 | 24.9949 | 25.7 |  | 24.9 | 16
| 1 | 4 | 24.6 | 24.1 | 25.15 | 23.5 | 25.48 | 24.2859 | 24.79 | 23.9 | 23.68 | 23.8 | 24.5 | 27.55 | 22.4 | 23.944 | 23 | 24.4342 | 22.5 | 24.9973 | 24.7 |  | 24.6 | 24.2134
| 1 | 5 | 24.2 | 24.3 | 25.15 | 24 | 25.68 | 23.4881 | 24.38 | 23.9 | 23.76 | 23.8 | 24.7 | 25.84 | 23.4 | 23.699 | 23.1 | 24.3236 | 23.1 | 25.0057 | 25.9 |  | 24.8 | 14
| 1 | 6 | 24.2 | 23.9 | 25.25 | 23.8 | 26.48 | 24.8842 | 24.38 | 24.1 | 23.8 | 23.7 | 24.8 | 25.23 | 23 | 23.719 | 23.4 | 24.4625 | 23.5 | 24.9205 | 25 |  | 24.9
| Mean |  | 24.2167 | 24.1 | 25.1867 | 23.8167 | 25.5733 | 24.2753 | 24.5433 | 23.9 | 23.83 | 23.7667 | 24.6333 | 27.0683 | 23.1 | 23.8257 | 23.2167 | 24.4399 | 23.1833 | 24.9709 | 24.55 | 23.8905 | 24.8
| Median |  | 24.2 | 24.1 | 25.185 | 23.85 | 25.4 | 24.2735 | 24.51 | 23.9 | 23.78 | 23.75 | 24.65 | 27.585 | 23.2 | 23.853 | 23.15 | 24.4356 | 23.25 | 24.9822 | 24.85 | 23.8905 | 24.8
| Std Dev. |  | 0.29269 | 0.14142 | 0.072296 | 0.17224 | 0.47513 | 0.46579 | 0.17363 | 0.18974 | 0.2358 | 0.08165 | 0.12111 | 1.2519 | 0.56569 | 0.096848 | 0.24014 | 0.092088 | 0.38166 | 0.035043 | 1.6306 | 0.010607 | 0.10954
| Rel.Std.Dev. |  | 1.209% | 0.5868% | 0.287% | 0.7232% | 1.858% | 1.919% | 0.7074% | 0.7939% | 0.9895% | 0.3435% | 0.4916% | 4.625% | 2.449% | 0.4065% | 1.034% | 0.3768% | 1.646% | 0.1403% | 6.642% | 0.0444% | 0.4417%
| PDM3 |  | 0.01359% | -0.4682% | 4.02% | -1.638% | 5.617% | 0.2556% | 1.363% | -1.294% | -1.583% | -1.845% | 1.734% | 11.79% | -4.598% | -1.601% | -4.116% | 0.9354% | -4.254% | 3.128% | 1.39% | -1.333% | 2.423%
| Z-Score (Absolute) |  | 0 | 0.2 | 1.64 | 0.67 | 2.29 | 0.1 | 0.55 | 0.53 | 0.65 | 0.76 | 0.7 | 4.81 | 1.88 | 0.66 | 1.69 | 0.38 | 1.74 | 1.27 | 0.56 | 0.55 | 0.98
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 84 (Certified Value 2.38 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.41 | 2.3 | 2.7 | 2.2 | 2.08 | 2.2 | 2.5 | 2.97 | 2.35 | 2.4 | 2.4 | 2.7 | 2.43 | 2.35 | 2.3499 | 2.39 | 2.3715 | 2.35 | 2.36 |  |  | 1
| 1 | 2 | 2.34 | 2.4 | 3 | 2.2 | 2.19 | 2.2 | 2.4 | 2.75 | 2.4 | 2.3 | 2.37 | 2.6 | 2.42 | 2.34 | 2.4108 | 2.48 | 2.2987 | 2.34 | 2.4 |  |  | 10
| 1 | 3 | 2.45 | 2.3 | 3.6 | 2.2 | 2.09 | 2.2 | 2.4 | 2.82 | 2.31 | 2.4 | 2.39 | 2.7 | 2.4 | 2.37 | 2.4382 | 2.46 | 2.4751 | 2.52 | 2.35 |  |  | 16
| 1 | 4 | 2.44 | 2.3 | 2.6 | 2.2 | 1.89 | 2.3 | 2.4 | 3.05 | 2.35 | 2.4 | 2.33 | 2.9 | 2.46 | 2.33 | 2.3455 | 2.46 | 2.3839 | 2.38 | 2.39 |  |  | 2.3765
| 1 | 5 | 2.45 | 3.1 | 2.6 | 2.2 | 1.99 | 2.3 | 2.5 | 2.64 | 2.43 | 2.4 | 2.31 | 2.7 | 2.45 | 2.3 | 2.3484 | 2.44 | 2.5179 | 2.4 | 2.28 |  |  | 16
| 1 | 6 | 2.45 | 2.8 | 3.2 | 2 | 1.96 | 2.2 | 2.8 | 2.74 | 2.36 | 2.3 | 2.33 | 2.7 | 2.46 | 2.4 | 2.3661 | 2.46 | 2.3964 | 2.42 | 2.27
| Mean |  | 2.4233 | 2.5333 | 2.95 | 2.1667 | 2.0333 | 2.2333 | 2.5 | 2.8283 | 2.3667 | 2.3667 | 2.355 | 2.7167 | 2.4367 | 2.3483 | 2.3765 | 2.4483 | 2.4072 | 2.4017 | 2.3417
| Median |  | 2.445 | 2.35 | 2.85 | 2.2 | 2.035 | 2.2 | 2.45 | 2.785 | 2.355 | 2.4 | 2.35 | 2.7 | 2.44 | 2.345 | 2.358 | 2.46 | 2.3901 | 2.39 | 2.355
| Std Dev. |  | 0.043665 | 0.33862 | 0.39875 | 0.08165 | 0.10746 | 0.05164 | 0.15492 | 0.15407 | 0.042269 | 0.05164 | 0.036742 | 0.098319 | 0.024221 | 0.034303 | 0.038842 | 0.031252 | 0.078219 | 0.065243 | 0.054924
| Rel.Std.Dev. |  | 1.802% | 13.37% | 13.52% | 3.768% | 5.285% | 2.312% | 6.197% | 5.447% | 1.786% | 2.182% | 1.56% | 3.619% | 0.994% | 1.461% | 1.634% | 1.276% | 3.249% | 2.717% | 2.346%
| PDM3 |  | 1.973% | 6.601% | 24.13% | -8.828% | -14.44% | -6.022% | 5.199% | 19.01% | -0.4119% | -0.4119% | -0.9028% | 14.32% | 2.534% | -1.183% | 0.0003131% | 3.025% | 1.296% | 1.061% | -1.464%
| Z-Score (Absolute) |  | 0.27 | 1.66 | 6.93 | 2.97 | 4.66 | 2.13 | 1.24 | 5.39 | 0.44 | 0.44 | 0.59 | 3.98 | 0.44 | 0.67 | 0.32 | 0.59 | 0.07 | 0 | 0.76
| Analytical results for Al in OREAS 84 (Certified Value 2.2 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.2 | 2.2 | 2.05 | 2.2 | 1.9845 | 2.157 | 1.97 | 2.22 | 2.24 | 2.18 | 2.23 | 2.41 | 2.15 | 2.01 | 2.35 | 2.281 | 2.366 | 2.2698 | 2.06 | 2.24 |  | 1
| 1 | 2 | 2.19 | 2.25 | 2.04 | 2.2 | 2.0005 | 2.1787 | 1.96 | 2.22 | 2.17 | 2.22 | 2.23 | 2.42 | 2.2 | 2.01 | 2.34 | 2.2706 | 2.329 | 2.2431 | 2.04 | 2.17 |  | #N/A
| 1 | 3 | 2.23 | 2.18 | 2.03 | 2.17 | 1.9347 | 2.151 | 1.93 | 2.19 | 2.14 | 2.18 | 2.22 | 2.37 | 2.22 | 2.08 | 2.3 | 2.2721 | 2.313 | 2.2561 | 2.13 | 2.25 |  | 16
| 1 | 4 | 2.21 | 2.19 | 2.02 | 2.18 | 1.928 | 2.1057 | 1.99 | 2.23 | 2.15 | 2.17 | 2.24 | 2.33 | 2.19 | 2.02 | 2.35 | 2.2123 | 2.35 | 2.2229 | 2.05 | 2.23 |  | 2.2
| 1 | 5 | 2.12 | 2.23 | 1.96 | 2.17 | 1.9595 | 2.1642 | 1.95 | 2.28 | 2.06 | 2.24 | 2.23 | 2.37 | 2.22 | 2.04 | 2.32 | 2.2532 | 2.334 | 2.2105 | 2.05 | 2.19 |  | 17
| 1 | 6 | 2.2 | 2.09 | 2.02 | 2.16 | 1.9573 | 2.1171 | 1.98 | 2.22 | 2.08 | 2.34 | 2.23 | 2.35 | 2.16 | 2.07 | 2.27 | 2.2478 | 2.371 | 2.2906 | 2.1 | 2.17
| Mean |  | 2.1917 | 2.19 | 2.02 | 2.18 | 1.9607 | 2.1456 | 1.9633 | 2.2267 | 2.14 | 2.2217 | 2.23 | 2.375 | 2.19 | 2.0383 | 2.3217 | 2.2562 | 2.3438 | 2.2488 | 2.0717 | 2.2083
| Median |  | 2.2 | 2.195 | 2.025 | 2.175 | 1.9584 | 2.154 | 1.965 | 2.22 | 2.145 | 2.2 | 2.23 | 2.37 | 2.195 | 2.03 | 2.33 | 2.2619 | 2.342 | 2.2496 | 2.055 | 2.21
| Std Dev. |  | 0.037639 | 0.055498 | 0.031623 | 0.016733 | 0.02796 | 0.028331 | 0.021602 | 0.029439 | 0.064807 | 0.064005 | 0.0063246 | 0.034496 | 0.029665 | 0.030605 | 0.031885 | 0.024837 | 0.022516 | 0.029696 | 0.035449 | 0.036009
| Rel.Std.Dev. |  | 1.717% | 2.534% | 1.565% | 0.7676% | 1.426% | 1.32% | 1.1% | 1.322% | 3.028% | 2.881% | 0.2836% | 1.452% | 1.355% | 1.501% | 1.373% | 1.101% | 0.9606% | 1.32% | 1.711% | 1.631%
| PDM3 |  | -0.377% | -0.4528% | -8.18% | -0.9074% | -10.87% | -2.47% | -10.76% | 1.214% | -2.726% | 0.9866% | 1.365% | 7.956% | -0.4528% | -7.347% | 5.532% | 2.556% | 6.54% | 2.222% | -5.832% | 0.3805%
| Z-Score (Absolute) |  | 0.01 | 0.01 | 2.12 | 0.13 | 2.85 | 0.56 | 2.82 | 0.44 | 0.63 | 0.38 | 0.49 | 2.28 | 0.01 | 1.89 | 1.62 | 0.81 | 1.89 | 0.72 | 1.48 | 0.22
| Analytical results for As in OREAS 84 (Certified Value 167 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 170 | 165 | 188 | 166 | 148.5 | 172 | 161 | 243 | 169 | 171 | 176 | 161 | 170 | 170.8 | 157.3272 | 183.2307 | 159 | 161.32 | 177.09 | 202 |  | 1
| 1 | 2 | 165 | 163 | 182 | 168 | 149.1 | 169 | 163 | 236 | 171.5 | 172 | 175 | 165 | 169 | 170.2 | 156.0025 | 185.8667 | 156 | 163.04 | 172.7 | 198 |  | 16
| 1 | 3 | 171 | 158 | 186 | 163 | 147.5 | 170 | 159 | 222 | 168 | 170 | 170 | 168 | 166 | 166.5 | 157.2655 | 184.0344 | 165 | 163.29 |  | 195 |  | 16
| 1 | 4 | 171 | 164 | 179 | 163 | 147.1 | 173 | 159 | 235 | 167.5 | 171 | 171 | 163 | 173 | 165.8 | 151.7525 | 187.7138 | 159 | 161.93 |  | 199 |  | 167.1961
| 1 | 5 | 166 | 165 | 173 | 163 | 147.5 | 174 | 166 | 228 | 169 | 172 | 169 | 167 | 172 | 168.7 | 154.6853 | 183.9851 | 158 | 161.98 |  | 186 |  | 18
| 1 | 6 | 166 | 156 | 178 | 164 | 147.7 | 169 | 160 | 260 | 174.5 | 170 | 167 | 162 | 172 | 168.5 | 154.7071 | 187.6839 | 164 | 163.59 |  | 187
| Mean |  | 168.1667 | 161.8333 | 181 | 164.5 | 147.9 | 171.1667 | 161.3333 | 237.3333 | 169.9167 | 171 | 171.3333 | 164.3333 | 170.3333 | 168.4167 | 155.29 | 185.4191 | 160.1667 | 162.525 | 174.895 | 194.5
| Median |  | 168 | 163.5 | 180.5 | 163.5 | 147.6 | 171 | 160.5 | 235.5 | 169 | 171 | 170.5 | 164 | 171 | 168.6 | 155.3548 | 184.9505 | 159 | 162.51 | 174.895 | 196.5
| Std Dev. |  | 2.7869 | 3.8687 | 5.5136 | 2.0736 | 0.74833 | 2.137 | 2.7325 | 13.2313 | 2.6347 | 0.89443 | 3.5024 | 2.8048 | 2.582 | 1.9732 | 2.0871 | 1.9676 | 3.5449 | 0.90418 | 3.1042 | 6.5955
| Rel.Std.Dev. |  | 1.657% | 2.391% | 3.046% | 1.261% | 0.506% | 1.248% | 1.694% | 5.575% | 1.551% | 0.5231% | 2.044% | 1.707% | 1.516% | 1.172% | 1.344% | 1.061% | 2.213% | 0.5563% | 1.775% | 3.391%
| PDM3 |  | 0.5805% | -3.207% | 8.256% | -1.613% | -11.54% | 2.375% | -3.506% | 41.95% | 1.627% | 2.275% | 2.475% | -1.712% | 1.876% | 0.73% | -7.121% | 10.9% | -4.204% | -2.794% | 4.605% | 16.33%
| Z-Score (Absolute) |  | 0.11 | 0.8 | 1.29 | 0.51 | 2.32 | 0.22 | 0.85 | 7.43 | 0.08 | 0.2 | 0.24 | 0.53 | 0.13 | 0.08 | 1.51 | 1.77 | 0.98 | 0.72 | 0.62 | 2.76
| Analytical results for B in OREAS 84 (Indicative Value 216 ppm) | [22 empty] | Indicative
B | Round | Replicate | Lab | [20 empty] | 1
| No | No | 07 | [20 empty] | ppm
|  |  | 4A*OES | [20 empty] | 0
0
| 1 | 1 | 214 | [20 empty] | 1
| 1 | 2 | 215 | [20 empty] | 8
| 1 | 3 | 217 | [20 empty] | 16
| 1 | 4 | 213 | [20 empty] | 216.1667
| 1 | 5 | 220 | [20 empty] | 14
| 1 | 6 | 218
| Mean |  | 216.1667
| Median |  | 216
| Std Dev. |  | 2.6394
| Rel.Std.Dev. |  | 1.221%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 84 (Certified Value 161 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 159 | 140 | 181 | 140 | 165 | 169.4198 | 127 | 170 | 146 | 80 | 175 | 170 | 160 | 23 | 170 | 153.7915 | 150 | 221.514 | 154 | 170 |  | 1
| 1 | 2 | 159 | 70 | 178 | 160 | 166 | 169.957 | 131 | 170 | 155 | 80 | 176 | 170 | 150 | 29 | 170 | 151.4907 | 148 | 206.8882 | 153 | 180 |  | #N/A
| 1 | 3 | 154 | 150 | 179 | 150 | 162 | 169.1049 | 130 | 160 | 150 | 60 | 177 | 166 | 130 | 20 | 166 | 145.3521 | 151 | 219.3821 | 160 | 188 |  | 16
| 1 | 4 | 159 | 130 | 176 | 150 | 161 | 169.2782 | 121 | 170 | 154 | 80 | 176 | 152 | 150 | 42 | 169 | 145.4722 | 147 | 216.0894 | 156 | 183 |  | 161.1092
| 1 | 5 | 158 | 160 | 173 | 160 | 163 | 169.0499 | 130 | 180 | 148 | 140 | 174 | 165 | 140 | 47 | 170 | 143.5488 | 150 | 218.2659 | 154 | 187 |  | 19
| 1 | 6 | 159 | 170 | 179 | 130 | 163 | 168.9906 | 133 | 170 | 152 | 110 | 172 | 166 | 170 | 39 | 169 | 152.8648 | 151 | 206.6454 | 158 | 185
| Mean |  | 158 | 136.6667 | 177.6667 | 148.3333 | 163.3333 | 169.3 | 128.6667 | 170 | 150.8333 | 91.6667 | 175 | 164.8333 | 150 | 33.3333 | 169 | 148.7533 | 149.5 | 214.7975 | 155.8333 | 182.1667
| Median |  | 159 | 145 | 178.5 | 150 | 163 | 169.1915 | 130 | 170 | 151 | 80 | 175.5 | 166 | 150 | 34 | 169.5 | 148.4814 | 150 | 217.1776 | 155 | 184
| Std Dev. |  | 2 | 35.5903 | 2.8048 | 11.6905 | 1.8619 | 0.35858 | 4.2269 | 6.3246 | 3.4881 | 28.5774 | 1.7889 | 6.6458 | 14.1421 | 10.9301 | 1.5492 | 4.4542 | 1.6432 | 6.4628 | 2.7142 | 6.6156
| Rel.Std.Dev. |  | 1.266% | 26.04% | 1.579% | 7.881% | 1.14% | 0.2118% | 3.285% | 3.72% | 2.313% | 31.18% | 1.022% | 4.032% | 9.428% | 32.79% | 0.9167% | 2.994% | 1.099% | 3.009% | 1.742% | 3.632%
| PDM3 |  | -1.93% | -15.17% | 10.28% | -7.93% | 1.38% | 5.084% | -20.14% | 5.518% | -6.378% | -43.1% | 8.622% | 2.312% | -6.895% | -79.31% | 4.898% | -7.669% | -7.206% | 33.32% | -3.275% | 13.07%
| Z-Score (Absolute) |  | 0.16 | NA | 0.99 | NA | 0.16 | 0.5 | 1.87 | NA | 0.57 | 4.03 | 0.84 | 0.24 | NA | 7.44 | 0.49 | 0.7 | 0.65 | 3.16 | 0.28 | 1.26
Results from laboratories 2, 4, 8, 10 and 13 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 84 (Certified Value 0.45 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 0.44 | 0.48 | 0.5 | < 0.5 | 0.32 | 0.41434 | 0.46 | < 0.5 | 0.5 | 0.44 | < 0.5 | 0.43 | 0.5 | 0.43 | 0.42 | 0.5066 | 0.5 | 0.61801 | 0.45 | < 1 |  | 1
| 1 | 2 | 0.48 | 0.5 | 0.6 | < 0.5 | 0.32 | 0.42966 | 0.42 | < 0.5 | 0.5 | 0.44 | < 0.5 | 0.41 | 0.5 | 0.43 | 0.46 | 0.46649 | 0.5 | 0.61566 | 0.57 | < 1 |  | 12
| 1 | 3 | 0.47 | 0.49 | 0.5 | < 0.5 | 0.31 | 0.42278 | 0.41 | < 0.5 | 0.5 | 0.45 | < 0.5 | 0.42 | 0.5 | 0.46 | 0.44 | 0.51657 | 0.4 | 0.52324 | 0.6 | < 1 |  | 16
| 1 | 4 | 0.43 | 0.49 | < 0.5 | < 0.5 | 0.31 | 0.42026 | 0.38 | < 0.5 | 0.5 | 0.44 | < 0.5 | 0.39 | 0.5 | 0.42 | 0.4 | 0.49462 | 0.4 | 0.56391 | 0.36 | < 1 |  | 0.44543
| 1 | 5 | 0.47 | 0.48 | < 0.5 | < 0.5 | 0.31 | 0.42134 | 0.38 | < 0.5 | 0.5 | 0.43 | < 0.5 | 0.41 | 0.5 | 0.44 | 0.41 | 0.47867 | 0.5 | 0.54336 | 0.58 | < 1 |  | 20
| 1 | 6 | 0.36 | 0.49 | < 0.5 | < 0.5 | 0.31 | 0.4385 | 0.39 | < 0.5 | 0.5 | 0.44 | < 0.5 | 0.37 | 0.5 | 0.45 | 0.41 | 0.50367 | 0.5 | 0.52812 | 0.5 | < 1
| Mean |  | 0.44167 | 0.48833 | 0.53333 |  | 0.31333 | 0.42448 | 0.40667 |  | 0.5 | 0.44 |  | 0.405 | 0.5 | 0.43833 | 0.42333 | 0.49444 | 0.46667 | 0.56538 | 0.51
| Median |  | 0.455 | 0.49 | 0.5 |  | 0.31 | 0.42206 | 0.4 |  | 0.5 | 0.44 |  | 0.41 | 0.5 | 0.435 | 0.415 | 0.49915 | 0.5 | 0.55364 | 0.535
| Std Dev. |  | 0.04446 | 0.0075277 | 0.057735 |  | 0.005164 | 0.0084471 | 0.030768 |  | 0 | 0.0063246 |  | 0.021679 | 0 | 0.01472 | 0.022509 | 0.018723 | 0.05164 | 0.042306 | 0.09252
| Rel.Std.Dev. |  | 10.07% | 1.542% | 10.83% |  | 1.648% | 1.99% | 7.566% |  | 0% | 1.437% |  | 5.353% | 0% | 3.358% | 5.317% | 3.787% | 11.07% | 7.483% | 18.14%
| PDM3 |  | -0.8438% | 9.633% | 19.74% |  | -29.66% | -4.702% | -8.701% |  | 12.25% | -1.218% |  | -9.076% | 12.25% | -1.592% | -4.96% | 11% | 4.769% | 26.93% | 14.5%
| Z-Score (Absolute) |  | 0.12 | 0.64 | NA | 2.04 | 1.33 | 0.08 | 0.28 | 2.04 | NA | 0.1 | 2.04 | 0.3 | NA | 0.08 | 0.09 | 0.71 | NA | 1.51 | 0.88 | 0.77
Results from laboratories 3, 9, 13 and 18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 84 (Certified Value 1.56 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.49 | 1.56 | 1.7 | 5 | 1.69 | 1.4619 | 1 | < 2 | 1.48 | 1.44 | 1.46 | 1.73 | < 2 | 1.56 | 1.49 | 1.5677 | 1.7 | 1.611 | 1.62 | 1.6 |  | 1
| 1 | 2 | 1.45 | 1.66 | 1.7 | < 2 | 1.74 | 1.416 | 1.08 | < 2 | 1.44 | 1.47 | 1.47 | 1.73 | < 2 | 1.56 | 1.51 | 1.5445 | 1.7 | 1.4527 | 1.58 | 1.57 |  | 18
| 1 | 3 | 1.56 | 1.62 | 1.7 | 12 | 1.54 | 1.4618 | 1.11 | < 2 | 1.47 | 1.42 | 1.49 | 1.67 | 3 | 1.57 | 1.54 | 1.5937 | 1.7 | 1.4808 | 1.63 | 1.56 |  | 16
| 1 | 4 | 1.54 | 1.67 | 1.7 | 6 | 1.53 | 1.4857 | 1.12 | < 2 | 1.44 | 1.45 | 1.48 | 1.64 | < 2 | 1.58 | 1.5 | 1.6356 | 1.7 | 1.5836 | 1.57 | 1.57 |  | 1.5616
| 1 | 5 | 1.49 | 1.65 | 1.7 | 10 | 1.48 | 1.45 | 1.11 | 4 | 1.4 | 1.46 | 1.49 | 1.7 | < 2 | 1.58 | 1.5 | 1.5403 | 1.8 | 1.5734 | 1.57 | 1.54 |  | 21
| 1 | 6 | 1.49 | 1.65 | 1.7 | 12 | 1.5 | 1.4032 | 1.04 | < 2 | 1.46 | 1.43 | 1.49 | 1.66 | < 2 | 1.59 | 1.51 | 1.5487 | 1.7 | 1.4886 | 1.6 | 1.53
| Mean |  | 1.5033 | 1.635 | 1.7 | 9 | 1.58 | 1.4464 | 1.0767 | 4 | 1.4483 | 1.445 | 1.48 | 1.6883 | 3 | 1.5733 | 1.5083 | 1.5718 | 1.7167 | 1.5317 | 1.595 | 1.5617
| Median |  | 1.49 | 1.65 | 1.7 | 10 | 1.535 | 1.4559 | 1.095 | 4 | 1.45 | 1.445 | 1.485 | 1.685 | 3 | 1.575 | 1.505 | 1.5582 | 1.7 | 1.531 | 1.59 | 1.565
| Std Dev. |  | 0.039833 | 0.040373 | 0 | 3.3166 | 0.10789 | 0.031095 | 0.04761 |  | 0.028577 | 0.018708 | 0.012649 | 0.037639 |  | 0.012111 | 0.017224 | 0.03696 | 0.040825 | 0.065424 | 0.025884 | 0.024833
| Rel.Std.Dev. |  | 2.65% | 2.469% | 0% | 36.85% | 6.828% | 2.15% | 4.422% |  | 1.973% | 1.295% | 0.8547% | 2.229% |  | 0.7697% | 1.142% | 2.352% | 2.378% | 4.271% | 1.623% | 1.59%
| PDM3 |  | -3.728% | 4.704% | 8.866% | 476.3% | 1.181% | -7.374% | -31.05% | 156.2% | -7.25% | -7.464% | -5.222% | 8.119% | 92.12% | 0.7545% | -3.408% | 0.6531% | 9.933% | -1.914% | 2.142% | 0.007374%
| Z-Score (Absolute) |  | 0.54 | 0.67 | 1.27 | NA | 0.17 | 1.06 | 4.46 | NA | 1.04 | 1.07 | 0.75 | 1.16 | NA | 0.11 | 0.49 | 0.09 | 1.43 | 0.28 | 0.31 | 0
Results from laboratories 4, 7, 8 and 13 were removed due to their high Z score reading.
| Analytical results for Ca in OREAS 84 (Certified Value 2.03 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.09 | 1.99 | 2.16 | 2.04 | 2.1048 | 2.0447 | 1.94 | 2.06 | 1.96 | 1.99 | 2.04 | 2.141 | 2.07 | 1.816 | 2.04 | 2.0807 | 1.98 | 2.0278 | 1.84 | 2.15 |  | 1
| 1 | 2 | 2.06 | 2.03 | 2.16 | 2.04 | 2.1248 | 2.0379 | 1.93 | 2.08 | 1.92 | 1.99 | 2.05 | 2.15 | 2.09 | 1.82 | 2.07 | 2.0763 | 1.958 | 2.0321 | 1.75 | 2.09 |  | #N/A
| 1 | 3 | 2.16 | 2.05 | 2.15 | 2.01 | 2.0728 | 2.0339 | 1.93 | 2.04 | 1.85 | 2.04 | 2.04 | 2.104 | 2.14 | 1.843 | 2.03 | 2.0704 | 1.958 | 2.0296 | 1.83 | 2.11 |  | 16
| 1 | 4 | 2.1 | 2.04 | 2.14 | 2.04 | 2.0565 | 2.0284 | 1.96 | 2.07 | 1.9 | 2.01 | 2.05 | 2.082 | 2.14 | 1.831 | 2.02 | 2.0062 | 1.987 | 1.992 | 1.76 | 2.11 |  | 2.0344
| 1 | 5 | 2.08 | 2 | 2.07 | 2 | 2.0977 | 2.0376 | 1.99 | 2.14 | 1.84 | 2.01 | 2.06 | 2.116 | 2.11 | 1.838 | 2.04 | 2.042 | 1.965 | 2.048 | 1.81 | 2.08 |  | 22
| 1 | 6 | 2.14 | 2.03 | 2.14 | 1.99 | 2.0738 | 2.0538 | 1.99 | 2.06 | 1.89 | 2.04 | 2.06 | 2.087 | 2.08 | 1.886 | 2.02 | 2.0582 | 1.987 | 2.1091 | 1.84 | 2.08
| Mean |  | 2.105 | 2.0233 | 2.1367 | 2.02 | 2.0884 | 2.0394 | 1.9567 | 2.075 | 1.8933 | 2.0133 | 2.05 | 2.1133 | 2.105 | 1.839 | 2.0367 | 2.0556 | 1.9725 | 2.0398 | 1.805 | 2.1033
| Median |  | 2.095 | 2.03 | 2.145 | 2.025 | 2.0857 | 2.0377 | 1.95 | 2.065 | 1.895 | 2.01 | 2.05 | 2.11 | 2.1 | 1.8345 | 2.035 | 2.0643 | 1.9725 | 2.0309 | 1.82 | 2.1
| Std Dev. |  | 0.037815 | 0.023381 | 0.033862 | 0.022804 | 0.025118 | 0.0088578 | 0.028048 | 0.034496 | 0.044572 | 0.022509 | 0.0089443 | 0.027854 | 0.030166 | 0.025219 | 0.018619 | 0.027966 | 0.013809 | 0.038635 | 0.040373 | 0.026583
| Rel.Std.Dev. |  | 1.796% | 1.156% | 1.585% | 1.129% | 1.203% | 0.4343% | 1.433% | 1.662% | 2.354% | 1.118% | 0.4363% | 1.318% | 1.433% | 1.371% | 0.9142% | 1.36% | 0.7001% | 1.894% | 2.237% | 1.264%
| PDM3 |  | 3.471% | -0.5435% | 5.027% | -0.7073% | 2.655% | 0.2455% | -3.82% | 1.996% | -6.934% | -1.035% | 0.7673% | 3.88% | 3.471% | -9.604% | 0.1119% | 1.043% | -3.042% | 0.2653% | -11.28% | 3.389%
| Z-Score (Absolute) |  | 0.78 | 0.19 | 1.16 | 0.23 | 0.58 | 0 | 0.99 | 0.42 | 1.75 | 0.31 | 0.12 | 0.88 | 0.78 | 2.4 | 0.03 | 0.19 | 0.8 | 0 | 2.81 | 0.76
| Analytical results for Cd in OREAS 84 (Certified Value 2.43 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.26 | 2.55 | 2.9 | 4.4 | 2.65 | 2.2 | 4.4 | 2.42 | 2.39 | 2.41 | 2.59 | 0.8 | 2.46 | 1.99 | 2.4378 | 2.5 | 2.6062 | 2.49 | 11.66 | 2.37 |  | 1
| 1 | 2 | 2.31 | 2.7 | 2.4 | 4.2 | 2.78 | 2.34 | 4.2 | 2.35 | 2.46 | 2.4 | 2.66 | < 0.5 | 2.42 | 2.11 | 2.3847 | 2.5 | 2.2119 | 2.42 | 11.99 | 2.38 |  | #N/A
| 1 | 3 | 2.28 | 2.75 | 2.1 | 4.1 | 2.55 | 2.45 | 3.9 | 2.34 | 2.36 | 2.41 | 2.56 | 2.9 | 2.41 | 2.11 | 2.411 | 2.5 | 2.2501 | 2.68 | 11.97 | 2.33 |  | 16
| 1 | 4 | 2.32 | 2.86 | 2.6 | 4.9 | 2.52 | 2.36 | 3.7 | 2.39 | 2.54 | 2.38 | 2.54 | 3.4 | 2.48 | 2.04 | 2.359 | 2.4 | 2.188 | 2.38 | 11.8 | 2.36 |  | 2.4285
| 1 | 5 | 2.29 | 2.66 | 2.1 | 4.2 | 2.59 | 2.29 | 4.6 | 2.29 | 2.55 | 2.45 | 2.61 | 3 | 2.46 | 1.98 | 2.3728 | 2.6 | 2.5854 | 2.41 | 11.91 | 2.28 |  | 23
| 1 | 6 | 2.35 | 2.75 | 2.5 | 3.8 | 2.5 | 2.29 | 3.8 | 2.33 | 2.51 | 2.44 | 2.55 | 1.6 | 2.48 | 2.19 | 2.3591 | 2.5 | 2.4054 | 2.61 | 11.71 | 2.17
| Mean |  | 2.3017 | 2.7117 | 2.4333 | 4.2667 | 2.5983 | 2.3217 | 4.1 | 2.3533 | 2.4683 | 2.415 | 2.585 | 2.34 | 2.4517 | 2.07 | 2.3874 | 2.5 | 2.3745 | 2.4983 | 11.84 | 2.315
| Median |  | 2.3 | 2.725 | 2.45 | 4.2 | 2.57 | 2.315 | 4.05 | 2.345 | 2.485 | 2.41 | 2.575 | 2.9 | 2.46 | 2.075 | 2.3788 | 2.5 | 2.3278 | 2.455 | 11.855 | 2.345
| Std Dev. |  | 0.031885 | 0.10381 | 0.30768 | 0.36697 | 0.10381 | 0.083766 | 0.35777 | 0.045898 | 0.079352 | 0.025884 | 0.045056 | 1.0945 | 0.029944 | 0.08124 | 0.031381 | 0.063246 | 0.18752 | 0.12123 | 0.13799 | 0.079687
| Rel.Std.Dev. |  | 1.385% | 3.828% | 12.64% | 8.601% | 3.995% | 3.608% | 8.726% | 1.95% | 3.215% | 1.072% | 1.743% | 46.77% | 1.221% | 3.925% | 1.314% | 2.53% | 7.897% | 4.852% | 1.165% | 3.442%
| PDM3 |  | -5.224% | 11.66% | 0.1973% | 75.69% | 6.991% | -4.401% | 68.83% | -3.097% | 1.638% | -0.5576% | 6.442% | -3.646% | 0.9522% | -14.76% | -1.695% | 2.942% | -2.225% | 2.874% | 387.5% | -4.675%
| Z-Score (Absolute) |  | 0.76 | 1.46 | 0.05 | 9.91 | 0.85 | 0.66 | 9 | 0.48 | 0.14 | 0.15 | 0.77 | 2.45 | 0.05 | 2.02 | 0.3 | 0.31 | 0.37 | 0.3 | 51.03 | 0.69
| Analytical results for Ce in OREAS 84 (Certified Value 15.8 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 12 | 15 | 16 | 17 | 18 | 20 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 15.8 | 16.9 | 16.1 | 1.6 | 15.4414 | 16.22 | 16.57 | 15.3 | 16.41 | 16.49 | 15.02 | 16.0196 | < 1 | 15 | 20.2 | [6 empty] | 1
| 1 | 2 | 15.1 | 17.05 | 16.6 | 2 | 15.49 | 15.68 | 16.31 | 16.1 | 16.57 | 16.3 | 15.96 | 16.1748 | < 1 | 15 | 19.7 | [6 empty] | 19
| 1 | 3 | 15.8 | 17.9 | 16.5 | 0.9 | 15.3119 | 16.03 | 15.84 | 16.05 | 15.93 | 16.13 | 16 | 15.7832 | < 1 | 15.5 | 20.7 | [6 empty] | 16
| 1 | 4 | 16 | 18.2 | 16.6 | 1.1 | 15.3524 | 15.54 | 16.07 | 16.45 | 15.77 | 16.21 | 15.99 | 15.2058 | < 1 | 15 | 20.2 | [6 empty] | 15.7988
| 1 | 5 | 15.7 | 17.25 | 15.6 | 1.4 | 15.3982 | 15.23 | 15.3 | 16.45 | 15.68 | 16.47 | 15.5 | 15.6033 | < 1 | 15 | 19.3 | [6 empty] | 24
| 1 | 6 | 15.6 | 17.7 | 16.1 | 1.6 | 15.157 | 15.16 | 15.64 | 15.75 | 15.53 | 16.36 | 15.8 | 15.5738 | < 1 | 15.4 | 19.7
| Mean |  | 15.6667 | 17.5 | 16.25 | 1.4333 | 15.3585 | 15.6433 | 15.955 | 16.0167 | 15.9817 | 16.3267 | 15.7117 | 15.7268 |  | 15.15 | 19.9667
| Median |  | 15.75 | 17.475 | 16.3 | 1.5 | 15.3753 | 15.61 | 15.955 | 16.075 | 15.85 | 16.33 | 15.88 | 15.6932 |  | 15 | 19.95
| Std Dev. |  | 0.30768 | 0.51284 | 0.3937 | 0.39328 | 0.11709 | 0.42392 | 0.46021 | 0.44008 | 0.41763 | 0.14236 | 0.38815 | 0.34631 |  | 0.23452 | 0.49666
| Rel.Std.Dev. |  | 1.964% | 2.93% | 2.423% | 27.44% | 0.7624% | 2.71% | 2.884% | 2.748% | 2.613% | 0.872% | 2.47% | 2.202% |  | 1.548% | 2.487%
| PDM3 |  | -0.8364% | 10.77% | 2.856% | -90.93% | -2.787% | -0.9841% | 0.9886% | 1.379% | 1.157% | 3.341% | -0.5516% | -0.4561% |  | -4.107% | 26.38%
| Z-Score (Absolute) |  | 0.11 | 3.25 | 0.96 | 26.17 | 0.67 | 0.15 | 0.42 | 0.53 | 0.47 | 1.1 | 0.03 | 0 | 27.88 | 1.06 | 7.76
| Analytical results for Co in OREAS 84 (Certified Value 1578 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 1580 | 1675 | 1628 | 1620 | 1633.4 | 1624.7007 | 1566 | 1620 | 1521.9 | 1600 | 1540 | 1591 | 1620 | 1463 | 1620.3 | 1608.5684 | 1730 | 1603.3359 | 1530 | 1511 | 1506 | 1
| 1 | 2 | 1550 | 1690 | 1648 | 1610 | 1646 | 1618.284 | 1562 | 1650 | 1489.7 | 1600 | 1560 | 1616 | 1640 | 1483 | 1629.4 | 1587.9937 | 1760 | 1534.7337 | 1520 | 1287.5 | 1476 | #N/A
| 1 | 3 | 1600 | 1625 | 1606 | 1600 | 1614 | 1629.7968 | 1567 | 1605 | 1440.7 | 1585 | 1540 | 1594 | 1670 | 1465 | 1597.2 | 1606.4519 | 1740 | 1569.5367 | 1590 | 1611.2 | 1480 | 16
| 1 | 4 | 1570 | 1655 | 1594 | 1615 | 1602.2 | 1627.272 | 1564 | 1630 | 1478.4 | 1595 | 1550 | 1570 | 1665 | 1487 | 1593.1 | 1526.4782 | 1760 | 1588.0305 | 1530 | 1595.58 | 1474 | 1578.0701
| 1 | 5 | 1550 | 1625 | 1545 | 1585 | 1623.4 | 1621.214 | 1594 | 1685 | 1431.8 | 1620 | 1550 | 1614 | 1665 | 1492 | 1603.1 | 1572.1646 | 1760 | 1590.5915 | 1510 |  | 1443 | 25
| 1 | 6 | 1540 | 1545 | 1594 | 1585 | 1607.5 | 1631.2618 | 1584 | 1630 | 1460.1 | 1690 | 1550 | 1556 | 1635 | 1548 | 1593.8 | 1557.8887 | 1750 | 1534.5431 | 1550 |  | 1459
| Mean |  | 1565 | 1635.8333 | 1602.5 | 1602.5 | 1621.0833 | 1625.4216 | 1572.8333 | 1636.6667 | 1470.4333 | 1615 | 1548.3333 | 1590.1667 | 1649.1667 | 1489.6667 | 1606.15 | 1576.5909 | 1750 | 1570.1286 | 1538.3333 | 1501.32 | 1473
| Median |  | 1560 | 1640 | 1600 | 1605 | 1618.7 | 1625.9864 | 1566.5 | 1630 | 1469.25 | 1600 | 1550 | 1592.5 | 1652.5 | 1485 | 1600.15 | 1580.0792 | 1755 | 1578.7836 | 1530 | 1553.29 | 1475
| Std Dev. |  | 22.5832 | 51.6156 | 35.1667 | 15.0831 | 16.5479 | 5.0168 | 13.0295 | 27.8687 | 33.3673 | 38.4708 | 7.5277 | 23.769 | 20.3511 | 30.9171 | 15.1856 | 31.3844 | 12.6491 | 29.5344 | 28.5774 | 149.1881 | 21.1849
| Rel.Std.Dev. |  | 1.443% | 3.155% | 2.194% | 0.9412% | 1.021% | 0.3086% | 0.8284% | 1.703% | 2.269% | 2.382% | 0.4862% | 1.495% | 1.234% | 2.075% | 0.9455% | 1.991% | 0.7228% | 1.881% | 1.858% | 9.937% | 1.438%
| PDM3 |  | -0.8282% | 3.66% | 1.548% | 1.548% | 2.726% | 3.001% | -0.3318% | 3.713% | -6.821% | 2.34% | -1.884% | 0.7665% | 4.505% | -5.602% | 1.779% | -0.09373% | 10.89% | -0.5032% | -2.518% | -4.864% | -6.658%
| Z-Score (Absolute) |  | 0.48 | 0.87 | 0.24 | 0.24 | 0.59 | 0.67 | 0.33 | 0.89 | 2.29 | 0.47 | 0.8 | 0 | 1.13 | 1.92 | 0.31 | 0.26 | 3.06 | 0.38 | 0.99 | 1.7 | 2.24
| Analytical results for Cr in OREAS 84 (Certified Value 0.559 wt.%) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.53 | 0.523 | 0.0961 | 0.59 | 0.12516 | 0.65393 | 0.2316 | 0.596 | 0.2106 | 0.592 | 0.622 | 0.4682 | 0.528 | 0.2238 | 0.4639 | 0.58997 | 0.571 | 0.71452 | 0.421 | 0.74418 | 0.6242 | 1
| 1 | 2 | 0.392 | 0.502 | 0.0989 | 0.581 | 0.12928 | 0.65611 | 0.2362 | 0.59 | 0.2045 | 0.612 | 0.617 | 0.4832 | 0.558 | 0.2462 | 0.464 | 0.57622 | 0.558 | 0.68063 | 0.405 | 0.7459 | 0.6321 | #N/A
| 1 | 3 | 0.532 | 0.626 | 0.1148 | 0.562 | 0.12101 | 0.65237 | 0.2352 | 0.589 | 0.2231 | 0.601 | 0.606 | 0.4478 | 0.575 | 0.3296 | 0.4605 | 0.56848 | 0.547 | 0.72095 | 0.453 | 0.77176 | 0.6516 | 16
| 1 | 4 | 0.542 | 0.519 | 0.1017 | 0.596 | 0.12458 | 0.65242 | 0.2308 | 0.586 | 0.1951 | 0.587 | 0.61 | 0.444 | 0.561 | 0.2696 | 0.4628 | 0.59598 | 0.596 | 0.761 | 0.437 | 0.67905 | 0.6276 | 0.55884
| 1 | 5 | 0.512 | 0.573 | 0.1018 | 0.553 | 0.12554 | 0.65557 | 0.2239 | 0.618 | 0.1716 | 0.557 | 0.612 | 0.4198 | 0.569 | 0.2901 | 0.4592 | 0.59518 | 0.581 | 0.74417 | 0.441 |  | 0.6399 | 26
| 1 | 6 | 0.593 | 0.523 | 0.0962 | 0.551 | 0.12668 | 0.65563 | 0.2355 | 0.593 | 0.1926 | 0.555 | 0.63 | 0.4449 | 0.557 | 0.2261 | 0.4633 | 0.58167 | 0.568 | 0.69837 | 0.41 |  | 0.6369
| Mean |  | 0.51683 | 0.54433 | 0.10158 | 0.57217 | 0.12537 | 0.65434 | 0.2322 | 0.59533 | 0.19958 | 0.584 | 0.61617 | 0.45132 | 0.558 | 0.26423 | 0.46228 | 0.58458 | 0.57017 | 0.71994 | 0.42783 | 0.73522 | 0.63538
| Median |  | 0.531 | 0.523 | 0.1003 | 0.5715 | 0.12535 | 0.65475 | 0.2334 | 0.5915 | 0.1998 | 0.5895 | 0.6145 | 0.44635 | 0.5595 | 0.2579 | 0.46305 | 0.58582 | 0.5695 | 0.71773 | 0.429 | 0.74504 | 0.6345
| Std Dev. |  | 0.067 | 0.046552 | 0.0069425 | 0.019405 | 0.0027096 | 0.0016785 | 0.0046282 | 0.011622 | 0.01761 | 0.023307 | 0.0087731 | 0.021918 | 0.016248 | 0.040951 | 0.0019773 | 0.011038 | 0.017175 | 0.029371 | 0.018851 | 0.039515 | 0.0098195
| Rel.Std.Dev. |  | 12.96% | 8.552% | 6.834% | 3.392% | 2.161% | 0.2565% | 1.993% | 1.952% | 8.823% | 3.991% | 1.424% | 4.856% | 2.912% | 15.5% | 0.4277% | 1.888% | 3.012% | 4.08% | 4.406% | 5.375% | 1.545%
| PDM3 |  | -7.516% | -2.595% | -81.82% | 2.385% | -77.56% | 17.09% | -58.45% | 6.531% | -64.29% | 4.503% | 10.26% | -19.24% | -0.1497% | -52.72% | -17.28% | 4.607% | 2.027% | 28.83% | -23.44% | 31.56% | 13.7%
| Z-Score (Absolute) |  | 0.29 | 0.1 | 3.22 | 0.1 | 3.05 | 0.68 | 2.3 | 0.26 | 2.52 | 0.18 | 0.41 | 0.75 | 0 | 2.07 | 0.67 | 0.19 | 0.09 | 1.14 | 0.92 | 1.25 | 0.55
| Analytical results for Cs in OREAS 84 (Certified Value 1.85 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.74 | 2 | 1.8 | 2 | 1.8098 | 1.86 | < 5 | 1.79 | 1.7 | 2.14 | 1.83 | 1.76 | 1.9769 | 1.9 | 1.8852 | 1.7 | 2 |  |  |  |  | 1
| 1 | 2 | 1.62 | 2.08 | 1.9 | 2 | 1.7203 | 1.83 | < 5 | 1.92 | 1.7 | 2.15 | 1.8 | 1.78 | 1.9179 | 1.9 | 1.756 | 1.75 | 2 |  |  |  |  | #N/A
| 1 | 3 | 1.67 | 2.07 | 2.1 | 1.9 | 1.7501 | 1.76 | < 5 | 1.85 | 1.7 | 2.09 | 1.78 | 1.8 | 1.9366 | 1.9 | 1.869 | 1.81 | 2 |  |  |  |  | 16
| 1 | 4 | 1.81 | 2.13 | 1.9 | 1.9 | 1.743 | 1.65 | < 5 | 1.84 | 1.7 | 2.08 | 1.81 | 1.82 | 1.9373 | 1.9 | 1.8479 | 1.74 | 2 |  |  |  |  | 1.8484
| 1 | 5 | 1.72 | 2.08 | 1.6 | 1.9 | 1.8054 | 1.75 | < 5 | 1.86 | 1.7 | 2.12 | 1.87 | 1.81 | 1.894 | 1.9 | 1.7682 | 1.77 | 2 |  |  |  |  | 27
| 1 | 6 | 1.69 | 2.09 | 1.7 | 1.8 | 1.6995 | 1.6 | < 5 | 1.8 | 1.7 | 2.08 | 1.86 | 1.78 | 1.8729 | 1.9 | 1.9255 | 1.8 | 2
| Mean |  | 1.7083 | 2.075 | 1.8333 | 1.9167 | 1.7547 | 1.7417 |  | 1.8433 | 1.7 | 2.11 | 1.825 | 1.7917 | 1.9226 | 1.9 | 1.842 | 1.7617 | 2
| Median |  | 1.705 | 2.08 | 1.85 | 1.9 | 1.7465 | 1.755 |  | 1.845 | 1.7 | 2.105 | 1.82 | 1.79 | 1.9273 | 1.9 | 1.8585 | 1.76 | 2
| Std Dev. |  | 0.064936 | 0.042308 | 0.17512 | 0.075277 | 0.044702 | 0.10068 |  | 0.046762 | 0 | 0.030984 | 0.035071 | 0.022286 | 0.03651 | 2.4324e-16 | 0.067003 | 0.040702 | 0
| Rel.Std.Dev. |  | 3.801% | 2.039% | 9.552% | 3.928% | 2.548% | 5.781% |  | 2.537% | 0% | 1.468% | 1.922% | 1.244% | 1.899% | 1.28e-14% | 3.638% | 2.31% | 0%
| PDM3 |  | -7.577% | 12.26% | -0.8149% | 3.694% | -5.07% | -5.774% |  | -0.2739% | -8.028% | 14.15% | -1.266% | -3.069% | 4.015% | 2.792% | -0.3471% | -4.692% | 8.202%
| Z-Score (Absolute) |  | 1.08 | 1.98 | 0.04 | 0.66 | 0.69 | 0.8 | 5.51 | 0.05 | 1.15 | 2.27 | 0.11 | 0.38 | 0.71 | 0.52 | 0.04 | 0.63 | NA
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 84 (Certified Value 0.4 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.407 | 0.399 | 0.4123 | 0.404 | 0.38405 | 0.3948 | 0.4019 | 0.401 | 0.39266 | 0.4 | 0.425 | 0.3913 | 0.4 | 0.3813 | 0.40034 | 0.39376 | 0.442 | 0.40547 | 0.384 | 0.402 | 0.4354 | 1
| 1 | 2 | 0.39 | 0.407 | 0.4088 | 0.404 | 0.38775 | 0.39204 | 0.4008 | 0.403 | 0.38229 | 0.407 | 0.423 | 0.3951 | 0.41 | 0.3811 | 0.4001 | 0.3955 | 0.44 | 0.40834 | 0.368 | 0.39576 | 0.4284 | #N/A
| 1 | 3 | 0.411 | 0.394 | 0.4121 | 0.401 | 0.37844 | 0.39652 | 0.4034 | 0.401 | 0.37449 | 0.401 | 0.425 | 0.3872 | 0.413 | 0.3843 | 0.39199 | 0.39274 | 0.437 | 0.40365 | 0.388 |  | 0.4226 | 16
| 1 | 4 | 0.409 | 0.396 | 0.4062 | 0.4 | 0.37589 | 0.39922 | 0.4016 | 0.406 | 0.38459 | 0.4 | 0.428 | 0.3812 | 0.412 | 0.3813 | 0.39887 | 0.38197 | 0.438 | 0.40086 | 0.377 |  | 0.4289 | 0.39955
| 1 | 5 | 0.4 | 0.404 | 0.4123 | 0.398 | 0.38054 | 0.39621 | 0.4095 | 0.413 | 0.37409 | 0.4 | 0.429 | 0.3877 | 0.413 | 0.3843 | 0.3958 | 0.38936 | 0.436 | 0.41195 | 0.384 |  | 0.4165 | 28
| 1 | 6 | 0.414 | 0.383 | 0.4098 | 0.398 | 0.37663 | 0.39529 | 0.4007 | 0.404 | 0.38009 | 0.421 | 0.428 | 0.3845 | 0.405 | 0.3946 | 0.3954 | 0.38893 | 0.441 | 0.41464 | 0.393 |  | 0.4114
| Mean |  | 0.40517 | 0.39717 | 0.41025 | 0.40083 | 0.38055 | 0.39568 | 0.40298 | 0.40467 | 0.38137 | 0.40483 | 0.42633 | 0.38783 | 0.40883 | 0.38448 | 0.39708 | 0.39038 | 0.439 | 0.40749 | 0.38233 | 0.39888 | 0.42387
| Median |  | 0.408 | 0.3975 | 0.41095 | 0.4005 | 0.37949 | 0.39575 | 0.40175 | 0.4035 | 0.38119 | 0.4005 | 0.4265 | 0.38745 | 0.411 | 0.3828 | 0.39733 | 0.39105 | 0.439 | 0.4069 | 0.384 | 0.39888 | 0.4255
| Std Dev. |  | 0.0087958 | 0.0084715 | 0.0024712 | 0.0027142 | 0.0045962 | 0.0023518 | 0.0033379 | 0.0045019 | 0.0069376 | 0.0083766 | 0.0023381 | 0.0049062 | 0.0052694 | 0.0051794 | 0.0032656 | 0.0048388 | 0.0023664 | 0.0051856 | 0.0087788 | 0.0044095 | 0.0088339
| Rel.Std.Dev. |  | 2.171% | 2.133% | 0.6024% | 0.6771% | 1.208% | 0.5944% | 0.8283% | 1.112% | 1.819% | 2.069% | 0.5484% | 1.265% | 1.289% | 1.347% | 0.8224% | 1.24% | 0.5391% | 1.273% | 2.296% | 1.105% | 2.084%
| PDM3 |  | 1.406% | -0.5966% | 2.678% | 0.3211% | -4.755% | -0.9686% | 0.8592% | 1.281% | -4.551% | 1.322% | 6.703% | -2.933% | 2.323% | -3.771% | -0.6175% | -2.296% | 9.874% | 1.986% | -4.309% | -0.1675% | 6.086%
| Z-Score (Absolute) |  | 0.37 | 0.31 | 0.79 | 0 | 1.71 | 0.43 | 0.18 | 0.32 | 1.64 | 0.34 | 2.15 | 1.1 | 0.67 | 1.38 | 0.32 | 0.88 | 3.22 | 0.56 | 1.56 | 0.16 | 1.94
| Analytical results for Dy in OREAS 84 (Certified Value 0.94 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.88 | 0.9 | 1.03 | 0.91493 | 1.01 | 0.98 | 0.93 | 0.98351 | 1.08 | [12 empty] | 1
| 1 | 2 | 0.93 | 1.1 | 1.03 | 0.90064 | 1.01 | 0.97 | 0.92 | 0.94347 | 1.06 | [12 empty] | #N/A
| 1 | 3 | 0.89 | 1 | 0.92 | 0.90891 | 0.88 | 0.99 | 0.92 | 0.99821 | 1.07 | [12 empty] | 16
| 1 | 4 | 0.87 | 0.9 | 0.9 | 0.91627 | 0.99 | 0.95 | 0.91 | 0.9686 | 1.02 | [12 empty] | 0.93845
| 1 | 5 | 0.96 | 0.9 | 0.9 | 0.89325 | 0.87 | 0.97 | 0.93 | 0.96812 | 1.1 | [12 empty] | 29
| 1 | 6 | 0.97 | 0.8 | 0.89 | 0.8964 | 0.89 | 0.96 | 0.91 | 0.96256 | 1.04
| Mean |  | 0.91667 | 0.93333 | 0.945 | 0.90507 | 0.94167 | 0.97 | 0.92 | 0.97075 | 1.0617
| Median |  | 0.91 | 0.9 | 0.91 | 0.90477 | 0.94 | 0.97 | 0.92 | 0.96836 | 1.065
| Std Dev. |  | 0.04274 | 0.10328 | 0.066558 | 0.0097172 | 0.06824 | 0.014142 | 0.0089443 | 0.018651 | 0.028577
| Rel.Std.Dev. |  | 4.662% | 11.07% | 7.043% | 1.074% | 7.247% | 1.458% | 0.9722% | 1.921% | 2.692%
| PDM3 |  | -2.321% | -0.5451% | 0.698% | -3.557% | 0.3428% | 3.362% | -1.966% | 3.441% | 13.13%
| Z-Score (Absolute) |  | 0.67 | NA | 0.04 | 0.97 | 0.04 | 0.67 | 0.59 | 0.69 | 2.99
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 84 (Certified Value 0.53 ppm) | [22 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.56 | 0.5 | 0.55 | 0.49007 | 0.38 | 0.57 | 0.52 | 0.56029 | 0.56 | [12 empty] | 1
| 1 | 2 | 0.51 | 0.6 | 0.56 | 0.51137 | 0.45 | 0.56 | 0.53 | 0.52755 | 0.56 | [12 empty] | #N/A
| 1 | 3 | 0.57 | 0.6 | 0.53 | 0.49554 | 0.43 | 0.55 | 0.52 | 0.53872 | 0.57 | [12 empty] | 16
| 1 | 4 | 0.5 | 0.5 | 0.51 | 0.50855 | 0.51 | 0.55 | 0.53 | 0.5982 | 0.57 | [12 empty] | 0.52575
| 1 | 5 | 0.5 | 0.6 | 0.5 | 0.48885 | 0.48 | 0.54 | 0.51 | 0.53301 | 0.57 | [12 empty] | 30
| 1 | 6 | 0.54 | 0.6 | 0.49 | 0.50968 | 0.45 | 0.53 | 0.51 | 0.55993 | 0.56
| Mean |  | 0.53 | 0.56667 | 0.52333 | 0.50068 | 0.45 | 0.55 | 0.52 | 0.55295 | 0.565
| Median |  | 0.525 | 0.6 | 0.52 | 0.50205 | 0.45 | 0.55 | 0.52 | 0.54932 | 0.565
| Std Dev. |  | 0.030984 | 0.05164 | 0.028048 | 0.010355 | 0.044272 | 0.014142 | 0.0089443 | 0.026058 | 0.0054772
| Rel.Std.Dev. |  | 5.846% | 9.113% | 5.359% | 2.068% | 9.838% | 2.571% | 1.72% | 4.713% | 0.9694%
| PDM3 |  | 0.8093% | 7.784% | -0.4587% | -4.768% | -14.41% | 4.613% | -1.093% | 5.174% | 7.467%
| Z-Score (Absolute) |  | 0.09 | NA | 0.09 | 0.71 | 2.1 | 0.64 | 0.18 | 0.72 | 1.05
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 84 (Certified Value 0.34 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 0.32 | 0.4 | 0.36 | 0.37712 | 0.3 | 0.37 | 0.33 | 0.33348 | 0.4 | [12 empty] | 1
| 1 | 2 | 0.31 | 0.3 | 0.36 | 0.36645 | 0.24 | 0.37 | 0.32 | 0.33231 | 0.41 | [12 empty] | #N/A
| 1 | 3 | 0.28 | 0.3 | 0.33 | 0.39327 | 0.3 | 0.37 | 0.32 | 0.33154 | 0.4 | [12 empty] | 16
| 1 | 4 | 0.34 | 0.3 | 0.33 | 0.35696 | 0.27 | 0.35 | 0.32 | 0.30332 | 0.4 | [12 empty] | 0.33739
| 1 | 5 | 0.33 | 0.4 | 0.32 | 0.38067 | 0.26 | 0.36 | 0.31 | 0.3158 | 0.4 | [12 empty] | 31
| 1 | 6 | 0.33 | 0.3 | 0.32 | 0.37125 | 0.21 | 0.35 | 0.31 | 0.33277 | 0.4
| Mean |  | 0.31833 | 0.33333 | 0.33667 | 0.37429 | 0.26333 | 0.36167 | 0.31833 | 0.32487 | 0.40167
| Median |  | 0.325 | 0.3 | 0.33 | 0.37419 | 0.265 | 0.365 | 0.32 | 0.33193 | 0.4
| Std Dev. |  | 0.02137 | 0.05164 | 0.018619 | 0.01249 | 0.035024 | 0.0098319 | 0.0075277 | 0.012515 | 0.0040825
| Rel.Std.Dev. |  | 6.713% | 15.49% | 5.53% | 3.337% | 13.3% | 2.719% | 2.365% | 3.852% | 1.016%
| PDM3 |  | -5.65% | -1.204% | -0.2157% | 10.93% | -21.95% | 7.194% | -5.65% | -3.712% | 19.05%
| Z-Score (Absolute) |  | 0.39 | NA | 0.18 | 1.35 | 2.1 | 0.96 | 0.39 | 0.18 | 2.21
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 84 (Certified Value 31.51 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 32.1 | 31.6 | 32.26 | 31.5 | > 25 | 33.3157 | 30.91 | 31.2 | 36.6801 | 31.6 | 30.5 | 31.92 | 31 | > 15 | 29.91 | 30.8149 | 31.825 | 32.6765 | 19.3 | > 15 |  | 1
| 1 | 2 | 32 | 32.1 | 32.07 | 31.4 | > 25 | 33.0655 | 30.46 | 31.5 | 35.5272 | 31.8 | 30.7 | 31.69 | 31.5 | > 15 | 29.19 | 30.8929 | 31.475 | 33.2412 | 18.4 | > 15 |  | #N/A
| 1 | 3 | 33.3 | 30.9 | 31.6 | 31.2 | > 25 | 33.1286 | 30.2 | 31 | 34.6374 | 31.5 | 30.9 | 32.2 | 31.6 | > 15 | 29.27 | 30.8534 | 31.405 | 32.941 | 19.4 | > 15 |  | 16
| 1 | 4 | 33 | 31.2 | 31.62 | 31.3 | > 25 | 33.1627 | 30.7 | 31.8 | 34.738 | 31.5 | 30.5 | 32.71 | 31.7 | > 15 | 29.29 | 30.911 | 32.035 | 32.6872 | 18.9 | > 15 |  | 31.5125
| 1 | 5 | 32.9 | 31.2 | 30.42 | 31.2 | > 25 | 33.0763 | 30.39 | 32.3 | 36.04 | 30.8 | 30.6 | 31.62 | 32 | > 15 | 29.27 | 30.7756 | 31.615 | 33.313 | 18.9 | > 15 |  | 32
| 1 | 6 | 32.8 | 29.5 | 31.44 | 31.4 | > 25 | 33.1264 | 30.48 | 31.7 | 35.5403 | 32.2 | 30.4 | 31.55 | 31.1 | > 15 | 29.05 | 30.4719 | 32.105 | 33.087 | 19.2 | > 15
| Mean |  | 32.6833 | 31.0833 | 31.5683 | 31.3333 |  | 33.1459 | 30.5233 | 31.5833 | 35.5272 | 31.5667 | 30.6 | 31.9483 | 31.4833 |  | 29.33 | 30.7866 | 31.7433 | 32.991 | 19.0167
| Median |  | 32.85 | 31.2 | 31.61 | 31.35 |  | 33.1275 | 30.47 | 31.6 | 35.5337 | 31.55 | 30.55 | 31.805 | 31.55 |  | 29.27 | 30.8342 | 31.72 | 33.014 | 19.05
| Std Dev. |  | 0.51929 | 0.87958 | 0.64325 | 0.12111 |  | 0.09067 | 0.24857 | 0.46224 | 0.77548 | 0.45898 | 0.17889 | 0.44215 | 0.37639 |  | 0.29772 | 0.16199 | 0.29171 | 0.27176 | 0.3656
| Rel.Std.Dev. |  | 1.589% | 2.83% | 2.038% | 0.3865% |  | 0.2735% | 0.8144% | 1.464% | 2.183% | 1.454% | 0.5846% | 1.384% | 1.196% |  | 1.015% | 0.5262% | 0.919% | 0.8237% | 1.923%
| PDM3 |  | 3.716% | -1.362% | 0.1773% | -0.5684% |  | 5.183% | -3.139% | 0.2249% | 12.74% | 0.172% | -2.896% | 1.383% | -0.0924% |  | -6.926% | -2.303% | 0.7327% | 4.692% | -39.65%
| Z-Score (Absolute) |  | 0.97 | 0.42 | 0 | 0.2 | NA | 1.37 | 0.9 | 0.01 | 3.42 | 0 | 0.84 | 0.33 | 0.07 | NA | 1.93 | 0.67 | 0.15 | 1.23 | 10.85 | NA
| Analytical results for Ga in OREAS 84 (Certified Value 6.58 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.58 | 7.36 | 5.1 | < 10 | < 0.05 | 6.593 | 5.8 | < 10 | 6.6 | 7 | 6.4 | 7.46 | < 10 | 5.56 | 3.95 | 6.9341 | 7.2 | 6.7684 | 6.5 | 7 |  | 1
| 1 | 2 | 6.12 | 7.62 | 5.3 | < 10 | < 0.05 | 6.495 | 5.8 | < 10 | 6.5 | 7.26 | 6.6 | 7.61 | < 10 | 5.63 | 4.05 | 6.6821 | 7.3 | 6.3969 | 6.38 | 6.9 |  | #N/A
| 1 | 3 | 6.63 | 7.65 | 5.7 | < 10 | < 0.05 | 6.6006 | 5.5 | < 10 | 6.2 | 7.16 | 6.5 | 7.37 | < 10 | 5.9 | 3.92 | 6.7413 | 7.3 | 6.4674 | 6.78 | 7.1 |  | 16
| 1 | 4 | 6.78 | 7.9 | 5.2 | < 10 | < 0.05 | 6.5055 | 5.7 | < 10 | 6.3 | 7.06 | 6.3 | 7.22 | < 10 | 5.79 | 4.01 | 6.262 | 7.5 | 6.97 | 6.3 | 7 |  | 6.5811
| 1 | 5 | 6.98 | 7.59 | 5.8 | < 10 | < 0.05 | 6.6678 | 5.5 | < 10 | 6.1 | 6.99 | 6.3 | 7.56 | < 10 | 5.82 | 3.91 | 6.3757 | 7.5 | 6.9004 | 6.52 | 6.8 |  | 33
| 1 | 6 | 6.5 | 7.59 | 5.3 | < 10 | < 0.05 | 6.5802 | 5.3 | < 10 | 6.2 | 6.77 | 6.4 | 7.44 | < 10 | 5.58 | 4.02 | 6.3974 | 7.3 | 6.6909 | 6.41 | 6.6
| Mean |  | 6.5983 | 7.6183 | 5.4 |  |  | 6.5737 | 5.6 |  | 6.3167 | 7.04 | 6.4167 | 7.4433 |  | 5.7133 | 3.9767 | 6.5654 | 7.35 | 6.699 | 6.4817 | 6.9
| Median |  | 6.605 | 7.605 | 5.3 |  |  | 6.5866 | 5.6 |  | 6.25 | 7.03 | 6.4 | 7.45 |  | 5.71 | 3.98 | 6.5398 | 7.3 | 6.7296 | 6.455 | 6.95
| Std Dev. |  | 0.2893 | 0.17244 | 0.28284 |  |  | 0.064535 | 0.2 |  | 0.19408 | 0.16745 | 0.1169 | 0.13924 |  | 0.14166 | 0.05785 | 0.25954 | 0.12247 | 0.22971 | 0.1669 | 0.17889
| Rel.Std.Dev. |  | 4.384% | 2.264% | 5.238% |  |  | 0.9817% | 3.571% |  | 3.072% | 2.379% | 1.822% | 1.871% |  | 2.479% | 1.455% | 3.953% | 1.666% | 3.429% | 2.575% | 2.593%
| PDM3 |  | 0.2619% | 15.76% | -17.95% |  |  | -0.1127% | -14.91% |  | -4.018% | 6.973% | -2.499% | 13.1% |  | -13.19% | -39.57% | -0.2379% | 11.68% | 1.791% | -1.511% | 4.846%
| Z-Score (Absolute) |  | 0.13 | 0.99 | 0.89 | 1.23 | 5.47 | 0.11 | 0.72 | 1.23 | 0.11 | 0.5 | 0.03 | 0.85 | 1.23 | 0.63 | 2.1 | 0.1 | 0.77 | 0.21 | 0.03 | 0.38
| Analytical results for Gd in OREAS 84 (Certified Value 1.16 ppm) | [22 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 1.14 | 1 | 1.27 | 1.4817 | 1.05 | 1.3 | 1.13 | 1.1377 | 1.2 | [12 empty] | 1
| 1 | 2 | 1.15 | 1.2 | 1.3 | 1.4747 | 0.97 | 1.3 | 1.12 | 1.103 | 1.3 | [12 empty] | #N/A
| 1 | 3 | 1.19 | 1.2 | 1.17 | 1.4467 | 1.08 | 1.2 | 1.13 | 1.0847 | 1.2 | [12 empty] | 16
| 1 | 4 | 1.16 | 1.1 | 1.18 | 1.495 | 1.13 | 1.2 | 1.15 | 1.1127 | 1.3 | [12 empty] | 1.1618
| 1 | 5 | 1.08 | 1.2 | 1.15 | 1.4576 | 1.13 | 1.3 | 1.15 | 1.0833 | 1.3 | [12 empty] | 34
| 1 | 6 | 1.13 | 1.1 | 1.15 | 1.4611 | 1.05 | 1.2 | 1.16 | 1.1227 | 1.2
| Mean |  | 1.1417 | 1.1333 | 1.2033 | 1.4695 | 1.0683 | 1.25 | 1.14 | 1.1073 | 1.25
| Median |  | 1.145 | 1.15 | 1.175 | 1.4679 | 1.065 | 1.25 | 1.14 | 1.1079 | 1.25
| Std Dev. |  | 0.03656 | 0.08165 | 0.065013 | 0.017674 | 0.060139 | 0.054772 | 0.015492 | 0.021438 | 0.054772
| Rel.Std.Dev. |  | 3.202% | 7.204% | 5.403% | 1.203% | 5.629% | 4.382% | 1.359% | 1.936% | 4.382%
| PDM3 |  | -1.729% | -2.446% | 3.579% | 26.49% | -8.041% | 7.596% | -1.872% | -4.683% | 7.596%
| Z-Score (Absolute) |  | 0 | 0.09 | 0.67 | 3.58 | 0.8 | 1.18 | 0.02 | 0.38 | 1.18
| Analytical results for Ge in OREAS 84 (Indicative Value 0.5 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 16 | 17 | 19 | 20 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 0.37 | 0.67 | 0.5 | 0.09 | 0.5 | 1.6 | 1.2 | 0.9 | 2.5 | 0.23 | < 1 | 0.33266 | 0.27 | [8 empty] | 1
| 1 | 2 | 0.36 | 0.68 | < 0.5 | 0.1 | 0.48 | 1.6 | 1.3 | 0.8 | 2.5 | 0.25 | < 1 | 0.34459 | 0.26 | [8 empty] | 13
| 1 | 3 | 0.39 | 0.66 | < 0.5 | 0.09 | 0.48 | 1.5 | 1.13 | 0.8 | 2.5 | 0.24 | < 1 | 0.33889 | 0.32 | [8 empty] | 16
| 1 | 4 | 0.36 | 0.67 | 0.6 | 0.08 | 0.48 | 1.6 | 1.21 | 0.9 | 2.4 | 0.24 | < 1 | 0.32076 | 0.23 | [8 empty] | 0.49662
| 1 | 5 | 0.38 | 0.66 | 0.6 | 0.08 | 0.47 | 1.6 | 0.96 | 0.8 | 2.5 | 0.24 | < 1 | 0.32472 | 0.28 | [8 empty] | 14
| 1 | 6 | 0.35 | 0.68 | 0.5 | 0.09 | 0.46 | 1.6 | 0.97 | 0.8 | 2.5 | 0.25 | < 1 | 0.33582 | 0.29
| Mean |  | 0.36833 | 0.67 | 0.55 | 0.088333 | 0.47833 | 1.5833 | 1.1283 | 0.83333 | 2.4833 | 0.24167 |  | 0.33291 | 0.275
| Median |  | 0.365 | 0.67 | 0.55 | 0.09 | 0.48 | 1.6 | 1.165 | 0.8 | 2.5 | 0.24 |  | 0.33424 | 0.275
| Std Dev. |  | 0.01472 | 0.0089443 | 0.057735 | 0.0075277 | 0.013292 | 0.040825 | 0.13761 | 0.05164 | 0.040825 | 0.0075277 |  | 0.0088946 | 0.030166
| Rel.Std.Dev. |  | 3.996% | 1.335% | 10.5% | 8.522% | 2.779% | 2.578% | 12.2% | 6.197% | 1.644% | 3.115% |  | 2.672% | 10.97%
| PDM3 |  | -25.83% | 34.91% | 10.75% | -82.21% | -3.683% | 218.8% | 127.2% | 67.8% | 400% | -51.34% |  | -32.97% | -44.63%
| Z-Score (Absolute) |  | 0.36 | 0.64 | 0.09 | 1.29 | 0 | 3.66 | 2.16 | 1.18 | 6.65 | 0.78 | 0.07 | 0.48 | 0.67
| Analytical results for Hf in OREAS 84 (Certified Value 0.84 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.8 | 0.8 | 0.7 | 0.8 | 0.9 | 0.84 | 0.8 | 0.8 | 0.94 | 0.73 | 0.9 | 0.81589 | 0.9 | 0.91521 | 0.8 | 0.89 | [5 empty] | 1
| 1 | 2 | 0.9 | 0.9 | 0.9 | 0.77 | 1 | 0.81 | 0.9 | 0.8 | 0.9 | 0.76 | 1.1 | 0.80758 | 0.9 | 0.80404 | 0.8 | 0.84 | [5 empty] | 14
| 1 | 3 | 1 | 0.9 | 0.8 | 0.72 | 1 | 0.8 | 0.8 | 0.8 | 0.93 | 0.77 | 0.8 | 0.80549 | 0.9 | 0.88222 | 0.8 | 0.84 | [5 empty] | 16
| 1 | 4 | 0.8 | 0.9 | 0.8 | 0.73 | 0.9 | 0.85 | 0.9 | 0.8 | 0.97 | 0.76 | 0.8 | 0.79156 | 0.9 | 0.92119 | 0.8 | 0.87 | [5 empty] | 0.83852
| 1 | 5 | 0.8 | 0.8 | 0.7 | 0.69 | 1 | 0.78 | 0.8 | 0.8 | 0.91 | 0.76 | 0.8 | 0.80375 | 0.9 | 0.87292 | 0.9 | 0.82 | [5 empty] | 35
| 1 | 6 | 1 | 0.9 | 0.8 | 0.7 | 1 | 0.77 | 0.8 | 0.8 | 0.91 | 0.8 | 0.8 | 0.80985 | 0.9 | 0.79865 | 0.8 | 0.79
| Mean |  | 0.88333 | 0.86667 | 0.78333 | 0.735 | 0.96667 | 0.80833 | 0.83333 | 0.8 | 0.92667 | 0.76333 | 0.86667 | 0.80569 | 0.9 | 0.86571 | 0.81667 | 0.84167
| Median |  | 0.85 | 0.9 | 0.8 | 0.725 | 1 | 0.805 | 0.8 | 0.8 | 0.92 | 0.76 | 0.8 | 0.80653 | 0.9 | 0.87757 | 0.8 | 0.84
| Std Dev. |  | 0.098319 | 0.05164 | 0.075277 | 0.042308 | 0.05164 | 0.031885 | 0.05164 | 1.2162e-16 | 0.02582 | 0.022509 | 0.12111 | 0.0081031 | 0 | 0.053204 | 0.040825 | 0.035449
| Rel.Std.Dev. |  | 11.13% | 5.958% | 9.61% | 5.756% | 5.342% | 3.945% | 6.197% | 1.52e-14% | 2.786% | 2.949% | 13.97% | 1.006% | 0% | 6.146% | 4.999% | 4.212%
| PDM3 |  | 5.344% | 3.356% | -6.582% | -12.35% | 15.28% | -3.601% | -0.6191% | -4.594% | 10.51% | -8.967% | 3.356% | -3.916% | 7.331% | 3.242% | -2.607% | 0.3747%
| Z-Score (Absolute) |  | 0.89 | 0.57 | 1.05 | 1.99 | 2.51 | 0.57 | 0.08 | 0.73 | 1.73 | 1.44 | 0.57 | 0.62 | 1.22 | 0.55 | 0.41 | 0.08
| Analytical results for Hg in OREAS 84 (Indicative Value 0.043 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 05 | 17 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | [19 empty] | 3
3
| 1 | 1 | 0.06 | < 1 | [19 empty] | 1
| 1 | 2 | 0.05 | < 1 | [19 empty] | 10
| 1 | 3 | 0.03 | < 1 | [19 empty] | 16
| 1 | 4 | 0.04 | < 1 | [19 empty] | 0.043333
| 1 | 5 | 0.04 | < 1 | [19 empty] | 14
| 1 | 6 | 0.04 | < 1
| Mean |  | 0.043333
| Median |  | 0.04
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 23.83%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 84 (Certified Value 0.18 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 0.21 | 0.2 | 0.19 | 0.16638 | 0.15 | 0.22 | 0.2 | 0.20884 | 0.2 | [12 empty] | 1
| 1 | 2 | 0.2 | 0.2 | 0.18 | 0.16694 | 0.15 | 0.21 | 0.19 | 0.20193 | 0.2 | [12 empty] | #N/A
| 1 | 3 | 0.17 | 0.2 | 0.17 | 0.17131 | 0.15 | 0.21 | 0.19 | 0.20498 | 0.2 | [12 empty] | 16
| 1 | 4 | 0.19 | 0.2 | 0.17 | 0.16827 | 0.14 | 0.21 | 0.18 | 0.18729 | 0.19 | [12 empty] | 0.18467
| 1 | 5 | 0.19 | 0.2 | 0.16 | 0.16146 | 0.15 | 0.21 | 0.18 | 0.20561 | 0.21 | [12 empty] | 36
| 1 | 6 | 0.19 | 0.2 | 0.16 | 0.16819 | 0.15 | 0.2 | 0.18 | 0.20291 | 0.2
| Mean |  | 0.19167 | 0.2 | 0.17167 | 0.16709 | 0.14833 | 0.21 | 0.18667 | 0.20193 | 0.2
| Median |  | 0.19 | 0.2 | 0.17 | 0.16757 | 0.15 | 0.21 | 0.185 | 0.20395 | 0.2
| Std Dev. |  | 0.013292 | 3.0405e-17 | 0.01169 | 0.0032428 | 0.0040825 | 0.0063246 | 0.008165 | 0.0075614 | 0.0063246
| Rel.Std.Dev. |  | 6.935% | 1.52e-14% | 6.81% | 1.941% | 2.752% | 3.012% | 4.374% | 3.745% | 3.162%
| PDM3 |  | 3.789% | 8.302% | -7.041% | -9.518% | -19.68% | 13.72% | 1.082% | 9.346% | 8.302%
| Z-Score (Absolute) |  | 0.11 | NA | 0.78 | 0.98 | 1.82 | 0.93 | 0.11 | 0.57 | 0.48
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 84 (Certified Value 0.13 ppm) | [22 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 20 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 0.12 | 0.128 | 0.14 | 0.159 | 0.12 | 0.12 | 0.132 | 0.14 | 0.151 | 0.129 | 0.114 | 0.13473 | 0.14655 | 0.127 | 0.15 | [6 empty] | 1
| 1 | 2 | 0.11 | 0.142 | 0.14 | 0.158 | 0.1 | 0.12 | 0.126 | 0.13 | 0.148 | 0.129 | 0.113 | 0.13161 | 0.13859 | 0.134 | 0.14 | [6 empty] | 11
| 1 | 3 | 0.15 | 0.145 | 0.2 | 0.151 | 0.12 | 0.11 | 0.134 | 0.13 | 0.149 | 0.128 | 0.12 | 0.12827 | 0.1351 | 0.139 | 0.14 | [6 empty] | 16
| 1 | 4 | 0.15 | 0.149 | 0.15 | 0.142 | 0.11 | 0.12 | 0.128 | 0.13 | 0.142 | 0.133 | 0.116 | 0.12354 | 0.14649 | 0.135 | 0.14 | [6 empty] | 0.13206
| 1 | 5 | 0.13 | 0.14 | 0.12 | 0.141 | 0.1 | 0.11 | 0.139 | 0.13 | 0.147 | 0.134 | 0.111 | 0.12026 | 0.13447 | 0.13 | 0.13 | [6 empty] | 37
| 1 | 6 | 0.13 | 0.136 | 0.12 | 0.139 | 0.13 | 0.11 | 0.145 | 0.13 | 0.148 | 0.13 | 0.116 | 0.13834 | 0.13746 | 0.129 | 0.13
| Mean |  | 0.13167 | 0.14 | 0.145 | 0.14833 | 0.11333 | 0.115 | 0.134 | 0.13167 | 0.1475 | 0.1305 | 0.115 | 0.12946 | 0.13978 | 0.13233 | 0.13833
| Median |  | 0.13 | 0.141 | 0.14 | 0.1465 | 0.115 | 0.115 | 0.133 | 0.13 | 0.148 | 0.1295 | 0.115 | 0.12994 | 0.13803 | 0.132 | 0.14
| Std Dev. |  | 0.016021 | 0.0073485 | 0.029496 | 0.0088919 | 0.012111 | 0.0054772 | 0.0070711 | 0.0040825 | 0.0030166 | 0.002429 | 0.0030984 | 0.006817 | 0.0054364 | 0.0044572 | 0.0075277
| Rel.Std.Dev. |  | 12.17% | 5.249% | 20.34% | 5.995% | 10.69% | 4.763% | 5.277% | 3.101% | 2.045% | 1.861% | 2.694% | 5.266% | 3.889% | 3.368% | 5.442%
| PDM3 |  | -0.2981% | 6.012% | 9.798% | 12.32% | -14.18% | -12.92% | 1.469% | -0.2981% | 11.69% | -1.182% | -12.92% | -1.969% | 5.845% | 0.2067% | 4.75%
| Z-Score (Absolute) |  | 0.06 | 0.69 | 1.15 | 1.45 | 1.72 | 1.57 | 0.15 | 0.06 | 1.37 | 0.17 | 1.57 | 0.26 | 0.67 | 0 | 0.54
| Analytical results for K in OREAS 84 (Certified Value 0.368 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.37 | 0.36 | 0.31 | 0.36 | 0.3804 | 0.36455 | 0.36 | 0.37 | 0.4 | 0.36 | 0.361 | 0.4 | 0.35 | 0.34 | 0.4 | 0.37482 | 0.382 | 0.39258 | 0.36 | 0.374 | 0.39 | 1
| 1 | 2 | 0.36 | 0.36 | 0.31 | 0.36 | 0.3827 | 0.36286 | 0.36 | 0.37 | 0.39 | 0.37 | 0.365 | 0.4 | 0.36 | 0.34 | 0.41 | 0.37403 | 0.374 | 0.38855 | 0.35 | 0.374 | 0.38 | #N/A
| 1 | 3 | 0.37 | 0.36 | 0.3 | 0.36 | 0.3774 | 0.36522 | 0.36 | 0.36 | 0.38 | 0.37 | 0.359 | 0.4 | 0.37 | 0.35 | 0.39 | 0.37324 | 0.374 | 0.39182 | 0.37 |  | 0.38 | 16
| 1 | 4 | 0.37 | 0.36 | 0.29 | 0.36 | 0.3706 | 0.36556 | 0.36 | 0.36 | 0.39 | 0.36 | 0.362 | 0.39 | 0.36 | 0.34 | 0.41 | 0.36216 | 0.374 | 0.37559 | 0.35 |  | 0.38 | 0.36758
| 1 | 5 | 0.37 | 0.36 | 0.3 | 0.36 | 0.3781 | 0.36822 | 0.37 | 0.37 | 0.37 | 0.36 | 0.36 | 0.4 | 0.36 | 0.34 | 0.4 | 0.3703 | 0.374 | 0.39383 | 0.35 |  | 0.37 | 38
| 1 | 6 | 0.37 | 0.36 | 0.3 | 0.36 | 0.3783 | 0.36827 | 0.37 | 0.36 | 0.38 | 0.38 | 0.358 | 0.39 | 0.36 | 0.35 | 0.4 | 0.36825 | 0.382 | 0.37732 | 0.36 |  | 0.36
| Mean |  | 0.36833 | 0.36 | 0.30167 | 0.36 | 0.37792 | 0.36578 | 0.36333 | 0.365 | 0.385 | 0.36667 | 0.36083 | 0.39667 | 0.36 | 0.34333 | 0.40167 | 0.37047 | 0.37667 | 0.38662 | 0.35667 | 0.374 | 0.37667
| Median |  | 0.37 | 0.36 | 0.3 | 0.36 | 0.3782 | 0.36539 | 0.36 | 0.365 | 0.385 | 0.365 | 0.3605 | 0.4 | 0.36 | 0.34 | 0.4 | 0.37177 | 0.374 | 0.39018 | 0.355 | 0.374 | 0.38
| Std Dev. |  | 0.0040825 | 6.0809e-17 | 0.0075277 | 6.0809e-17 | 0.004075 | 0.002124 | 0.005164 | 0.0054772 | 0.010488 | 0.008165 | 0.0024833 | 0.005164 | 0.0063246 | 0.005164 | 0.0075277 | 0.0047592 | 0.0041312 | 0.0080806 | 0.008165 | 0 | 0.010328
| Rel.Std.Dev. |  | 1.108% | 1.689e-14% | 2.495% | 1.689e-14% | 1.078% | 0.5807% | 1.421% | 1.501% | 2.724% | 2.227% | 0.6882% | 1.302% | 1.757% | 1.504% | 1.874% | 1.285% | 1.097% | 2.09% | 2.289% | 0% | 2.742%
| PDM3 |  | 0.2038% | -2.063% | -17.93% | -2.063% | 2.811% | -0.4905% | -1.156% | -0.703% | 4.738% | -0.2496% | -1.837% | 7.912% | -2.063% | -6.597% | 9.272% | 0.7839% | 2.471% | 5.178% | -2.97% | 1.745% | 2.471%
| Z-Score (Absolute) |  | 0.16 | 0.65 | 6.35 | 0.65 | 1.1 | 0.09 | 0.33 | 0.16 | 1.79 | 0 | 0.57 | 2.93 | 0.65 | 2.28 | 3.42 | 0.37 | 0.94 | 1.95 | 0.98 | 0.67 | 0.98
| Analytical results for La in OREAS 84 (Certified Value 8.15 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 23 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 8 | 8.7 | 8.6 | 10 | 7.2 | 8.6008 | 7.9 | 10 | 7.5 | 8 | 8.31 | 10 | 8.53 | 8.6 | 8.3981 | 3 | 7.8 | 10.5 |  |  |  | 1
| 1 | 2 | 7.6 | 9 | 8.7 | 10 | 7.3 | 8.8377 | 7.7 | 10 | 7.5 | 8.4 | 8.19 | 10 | 8.51 | 8.9 | 8.561 | 3 | 7.8 | 10.2 |  |  |  | 20
| 1 | 3 | 8 | 9.1 | 8.5 | 10 | 7 | 8.6331 | 7.7 | 10 | 7.2 | 8.4 | 8.02 | 10 | 8.35 | 8.6 | 8.2597 | 3 | 7.9 | 10.8 |  |  |  | 16
| 1 | 4 | 7.9 | 9.3 | 8.6 | 10 | 7 | 8.7737 | 7.6 | < 10 | 7.3 | 8.5 | 7.92 | 10 | 8.74 | 8.9 | 7.9157 | 3 | 7.8 | 10.4 |  |  |  | 8.1498
| 1 | 5 | 8 | 8.8 | 8 | 10 | 7.2 | 8.6297 | 7.7 | < 10 | 7 | 8.5 | 7.97 | 10 | 8.61 | 8.6 | 8.1069 | 3 | 7.9 | 10 |  |  |  | 39
| 1 | 6 | 7.9 | 9.5 | 8.4 | 10 | 7.1 | 8.6116 | 7.4 | < 10 | 7.1 | 8.2 | 7.78 | 10 | 8.69 | 8.8 | 8.1326 | 3 | 8 | 10.2
| Mean |  | 7.9 | 9.0667 | 8.4667 | 10 | 7.1333 | 8.6811 | 7.6667 | 10 | 7.2667 | 8.3333 | 8.0317 | 10 | 8.5717 | 8.7333 | 8.229 | 3 | 7.8667 | 10.35
| Median |  | 7.95 | 9.05 | 8.55 | 10 | 7.15 | 8.6314 | 7.7 | 10 | 7.25 | 8.4 | 7.995 | 10 | 8.57 | 8.7 | 8.1961 | 3 | 7.85 | 10.3
| Std Dev. |  | 0.15492 | 0.30111 | 0.25033 | 0 | 0.12111 | 0.099298 | 0.1633 | 0 | 0.20656 | 0.19664 | 0.19094 | 0 | 0.14034 | 0.15055 | 0.22895 | 0 | 0.08165 | 0.28107
| Rel.Std.Dev. |  | 1.961% | 3.321% | 2.957% | 0% | 1.698% | 1.144% | 2.13% | 0% | 2.843% | 2.36% | 2.377% | 0% | 1.637% | 1.724% | 2.782% | 0% | 1.038% | 2.716%
| PDM3 |  | -3.065% | 11.25% | 3.889% | 22.7% | -12.47% | 6.52% | -5.928% | 22.7% | -10.84% | 2.253% | -1.449% | 22.7% | 5.177% | 7.161% | 0.9723% | -63.19% | -3.474% | 27%
| Z-Score (Absolute) |  | 0.63 | 1.3 | 0.31 | NA | 1.9 | 0.66 | 1.02 | NA | 1.68 | 0.09 | 0.41 | NA | 0.48 | 0.75 | 0.09 | NA | 0.69 | 3.43
Results from laboratory 18 were removed due to their 1 ppm reading resolution. Results from laboratories 4, 8 and 13 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 84 (Certified Value 10.3 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 10 | 10 | 11 | 10 | 10 | 10.1309 | 10 | 10 | < 1 | 10 | 10.8 | 9.4 | 10 | 11 | 12.4 | 10.9291 | 13 | 11.6091 | 9.5 | 12 |  | 1
| 1 | 2 | 10.4 | 10 | 11 | 10 | 10 | 10.1914 | 10 | 10 | < 1 | 10 | 10.9 | 9.3 | 10 | 11 | 12.3 | 10.4781 | 12 | 11.2778 | 9.7 | 12 |  | #N/A
| 1 | 3 | 9.5 | 10 | 11 | 10 | 9 | 10.3264 | 10 | 10 | < 1 | 10 | 10.7 | 9.1 | 10 | 11 | 12.1 | 10.9327 | 12 | 11.3432 | 11 | 11 |  | 16
| 1 | 4 | 9.3 | 10 | 10 | 10 | 9 | 10.1924 | 10 | 10 | < 1 | 10 | 10.6 | 8.8 | 10 | 11 | 12.5 | 10.3082 | 12 | 11.4543 | 10.4 | 11 |  | 10.277
| 1 | 5 | 9.8 | 10 | 10 | 10 | 10 | 10.2342 | 10 | 10 | < 1 | 10 | 10.6 | 9.1 | 10 | 11 | 12.4 | 10.4327 | 12 | 11.7112 | 10.3 | 11 |  | 40
| 1 | 6 | 10.1 | 10 | 10 | 10 | 9 | 10.299 | 10 | 10 | < 1 | 10 | 10.5 | 8.9 | 10 | 11 | 12.3 | 10.9513 | 12 | 11.6539 | 11.1 | 11
| Mean |  | 9.85 | 10 | 10.5 | 10 | 9.5 | 10.2291 | 10 | 10 |  | 10 | 10.6833 | 9.1 | 10 | 11 | 12.3333 | 10.672 | 12.1667 | 11.5083 | 10.3333 | 11.3333
| Median |  | 9.9 | 10 | 10.5 | 10 | 9.5 | 10.2133 | 10 | 10 |  | 10 | 10.65 | 9.1 | 10 | 11 | 12.35 | 10.7036 | 12 | 11.5317 | 10.35 | 11
| Std Dev. |  | 0.40373 | 0 | 0.54772 | 0 | 0.54772 | 0.073204 | 0 | 0 |  | 0 | 0.1472 | 0.22804 | 0 | 0 | 0.13663 | 0.29641 | 0.40825 | 0.17651 | 0.6532 | 0.5164
| Rel.Std.Dev. |  | 4.099% | 0% | 5.216% | 0% | 5.766% | 0.7157% | 0% | 0% |  | 0% | 1.378% | 2.506% | 0% | 0% | 1.108% | 2.777% | 3.355% | 1.534% | 6.321% | 4.556%
| PDM3 |  | -4.155% | -2.696% | 2.17% | -2.696% | -7.561% | -0.4665% | -2.696% | -2.696% |  | -2.696% | 3.954% | -11.45% | -2.696% | 7.035% | 20.01% | 3.844% | 18.39% | 11.98% | 0.5479% | 10.28%
| Z-Score (Absolute) |  | 0.38 | 0.16 | 0.55 | 0.16 | 0.88 | 0.16 | 0.16 | 0.16 | 13.75 | 0.16 | 0.81 | 1.45 | 0.16 | 1.27 | 3.17 | 0.8 | 2.93 | 1.99 | 0.31 | 1.74
| Analytical results for Lu in OREAS 84 (Certified Value 0.073 ppm) | [22 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 23 | [11 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [11 empty] | 3
3
| 1 | 1 | 0.08 | 0.09 | 0.12 | 0.06 | 0.07 | 0.08 | 0.08 | 0.08 | 0.069989 | 0.08 | [11 empty] | 1
| 1 | 2 | 0.06 | 0.08 | 0.13 | 0.07 | 0.07 | 0.08 | 0.08 | 0.07 | 0.069409 | 0.08 | [11 empty] | #N/A
| 1 | 3 | 0.06 | 0.1 | 0.11 | 0.06 | 0.07 | 0.08 | 0.08 | 0.07 | 0.072049 | 0.08 | [11 empty] | 16
| 1 | 4 | 0.09 | 0.08 | 0.12 | 0.07 | 0.07 | 0.08 | 0.08 | 0.07 | 0.062249 | 0.08 | [11 empty] | 0.07291
| 1 | 5 | 0.07 | 0.06 | 0.11 | 0.07 | 0.06 | 0.08 | 0.08 | 0.07 | 0.067226 | 0.07 | [11 empty] | 41
| 1 | 6 | 0.06 | 0.08 | 0.1 | 0.06 | 0.06 | 0.08 | 0.08 | 0.07 | 0.068194 | 0.07
| Mean |  | 0.07 | 0.081667 | 0.115 | 0.065 | 0.066667 | 0.08 | 0.08 | 0.071667 | 0.068186 | 0.076667
| Median |  | 0.065 | 0.08 | 0.115 | 0.065 | 0.07 | 0.08 | 0.08 | 0.07 | 0.068802 | 0.08
| Std Dev. |  | 0.012649 | 0.013292 | 0.010488 | 0.0054772 | 0.005164 | 0 | 0 | 0.0040825 | 0.0033418 | 0.005164
| Rel.Std.Dev. |  | 18.07% | 16.28% | 9.12% | 8.427% | 7.746% | 0% | 0% | 5.696% | 4.901% | 6.736%
| PDM3 |  | -3.991% | 12.01% | 57.73% | -10.85% | -8.563% | 9.725% | 9.725% | -1.705% | -6.479% | 5.153%
| Z-Score (Absolute) |  | 0.48 | 0.86 | 4.66 | 1.05 | 0.86 | 0.67 | 0.67 | 0.29 | 0.68 | 0.29
| Analytical results for Mg in OREAS 84 (Certified Value 1.51 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.57 | 1.43 | 1.52 | 1.46 | 1.3957 | 1.5934 | 1.45 | 1.47 | 1.55 | 1.49 | 1.59 | 1.638 | 1.5 | 1.377 | 1.65 | 1.4466 | 1.683 | 1.5295 | 1.5 | 1.54 |  | 1
| 1 | 2 | 1.57 | 1.46 | 1.52 | 1.45 | 1.4061 | 1.5705 | 1.45 | 1.48 | 1.51 | 1.51 | 1.6 | 1.654 | 1.5 | 1.386 | 1.6 | 1.4575 | 1.652 | 1.5372 | 1.47 | 1.5 |  | #N/A
| 1 | 3 | 1.62 | 1.46 | 1.51 | 1.44 | 1.3715 | 1.5832 | 1.43 | 1.45 | 1.46 | 1.49 | 1.59 | 1.611 | 1.55 | 1.432 | 1.6 | 1.4498 | 1.646 | 1.532 | 1.52 | 1.53 |  | 16
| 1 | 4 | 1.6 | 1.46 | 1.51 | 1.46 | 1.3655 | 1.5684 | 1.45 | 1.46 | 1.5 | 1.49 | 1.6 | 1.599 | 1.53 | 1.39 | 1.6 | 1.4126 | 1.677 | 1.528 | 1.48 | 1.53 |  | 1.5128
| 1 | 5 | 1.59 | 1.42 | 1.46 | 1.43 | 1.3834 | 1.5876 | 1.45 | 1.52 | 1.46 | 1.44 | 1.59 | 1.621 | 1.53 | 1.405 | 1.61 | 1.437 | 1.658 | 1.5343 | 1.48 | 1.49 |  | 42
| 1 | 6 | 1.6 | 1.47 | 1.5 | 1.42 | 1.3738 | 1.588 | 1.47 | 1.47 | 1.48 | 1.5 | 1.59 | 1.601 | 1.5 | 1.429 | 1.66 | 1.445 | 1.67 | 1.5259 | 1.52 | 1.47
| Mean |  | 1.5917 | 1.45 | 1.5033 | 1.4433 | 1.3827 | 1.5819 | 1.45 | 1.475 | 1.4933 | 1.4867 | 1.5933 | 1.6207 | 1.5183 | 1.4032 | 1.62 | 1.4414 | 1.6643 | 1.5311 | 1.495 | 1.51
| Median |  | 1.595 | 1.46 | 1.51 | 1.445 | 1.3786 | 1.5854 | 1.45 | 1.47 | 1.49 | 1.49 | 1.59 | 1.616 | 1.515 | 1.3975 | 1.605 | 1.4458 | 1.664 | 1.5307 | 1.49 | 1.515
| Std Dev. |  | 0.019408 | 0.02 | 0.022509 | 0.01633 | 0.015605 | 0.010152 | 0.012649 | 0.02429 | 0.034448 | 0.024221 | 0.005164 | 0.021732 | 0.02137 | 0.023043 | 0.027568 | 0.015617 | 0.014624 | 0.004193 | 0.021679 | 0.027568
| Rel.Std.Dev. |  | 1.219% | 1.379% | 1.497% | 1.131% | 1.129% | 0.6418% | 0.8724% | 1.647% | 2.307% | 1.629% | 0.3241% | 1.341% | 1.407% | 1.642% | 1.702% | 1.083% | 0.8787% | 0.2739% | 1.45% | 1.826%
| PDM3 |  | 5.216% | -4.149% | -0.6237% | -4.59% | -8.6% | 4.568% | -4.149% | -2.497% | -1.285% | -1.725% | 5.326% | 7.133% | 0.3679% | -7.245% | 7.088% | -4.718% | 10.02% | 1.214% | -1.175% | -0.183%
| Z-Score (Absolute) |  | 1.19 | 0.63 | 0.05 | 0.72 | 1.5 | 1.06 | 0.63 | 0.31 | 0.07 | 0.16 | 1.21 | 1.56 | 0.25 | 1.23 | 1.55 | 0.74 | 2.12 | 0.41 | 0.05 | 0.14
| Analytical results for Mn in OREAS 84 (Certified Value 0.064 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.0668 | 0.0637 | 0.0634 | 0.063 | 0.0572 | 0.062141 | 0.0583 | 0.0629 | 0.07 | 0.0637 | 0.0648 | 0.0681 | 0.0631 | 0.0584 | 0.0681 | 0.066926 | 0.062 | 0.06802 | 0.061 | 0.069749 | 0.0693 | 1
| 1 | 2 | 0.0668 | 0.0647 | 0.0635 | 0.0629 | 0.0579 | 0.062 | 0.0571 | 0.0628 | 0.06 | 0.0648 | 0.0652 | 0.0682 | 0.0641 | 0.0596 | 0.0685 | 0.066431 | 0.062 | 0.070858 | 0.0578 | 0.069226 | 0.068 | #N/A
| 1 | 3 | 0.0685 | 0.0636 | 0.0629 | 0.0626 | 0.0565 | 0.062003 | 0.0572 | 0.062 | 0.06 | 0.0638 | 0.0657 | 0.067 | 0.0648 | 0.0602 | 0.0675 | 0.06659 | 0.062 | 0.070465 | 0.0617 | 0.069701 | 0.0684 | 16
| 1 | 4 | 0.0684 | 0.0641 | 0.0628 | 0.0622 | 0.056 | 0.061903 | 0.0564 | 0.0627 | 0.06 | 0.0632 | 0.0661 | 0.0661 | 0.0644 | 0.0591 | 0.0676 | 0.062774 | 0.062 | 0.067835 | 0.0604 | 0.068152 | 0.069 | 0.063789
| 1 | 5 | 0.0672 | 0.0629 | 0.0647 | 0.0616 | 0.0571 | 0.061936 | 0.0583 | 0.0644 | 0.06 | 0.061 | 0.0644 | 0.067 | 0.065 | 0.0595 | 0.0675 | 0.064215 | 0.062 | 0.071492 | 0.0602 |  | 0.0649 | 43
| 1 | 6 | 0.0674 | 0.0642 | 0.0627 | 0.0614 | 0.0566 | 0.062037 | 0.0566 | 0.0627 | 0.06 | 0.0642 | 0.0646 | 0.0662 | 0.0635 | 0.0606 | 0.0682 | 0.063942 | 0.062 | 0.06998 | 0.0613 |  | 0.0653
| Mean |  | 0.067517 | 0.063867 | 0.063333 | 0.062283 | 0.056883 | 0.062003 | 0.057317 | 0.062917 | 0.061667 | 0.06345 | 0.065133 | 0.0671 | 0.06415 | 0.059567 | 0.0679 | 0.065146 | 0.062 | 0.069775 | 0.0604 | 0.069207 | 0.067483
| Median |  | 0.0673 | 0.0639 | 0.06315 | 0.0624 | 0.05685 | 0.062001 | 0.05715 | 0.06275 | 0.06 | 0.06375 | 0.065 | 0.067 | 0.06425 | 0.05955 | 0.06785 | 0.065323 | 0.062 | 0.070223 | 0.0607 | 0.069464 | 0.0682
| Std Dev. |  | 0.00076004 | 0.00061536 | 0.00074476 | 0.00067057 | 0.00066156 | 8.3299e-05 | 0.00081833 | 0.00079352 | 0.0040825 | 0.0013142 | 0.00066232 | 0.00089889 | 0.00073959 | 0.00078145 | 0.00042426 | 0.0017234 | 0 | 0.0015155 | 0.00139 | 0.00074185 | 0.0019052
| Rel.Std.Dev. |  | 1.126% | 0.9635% | 1.176% | 1.077% | 1.163% | 0.1343% | 1.428% | 1.261% | 6.62% | 2.071% | 1.017% | 1.34% | 1.153% | 1.312% | 0.6248% | 2.645% | 0% | 2.172% | 2.301% | 1.072% | 2.823%
| PDM3 |  | 5.843% | 0.1213% | -0.7148% | -2.361% | -10.83% | -2.8% | -10.15% | -1.368% | -3.328% | -0.5319% | 2.107% | 5.19% | 0.5655% | -6.62% | 6.444% | 2.127% | -2.805% | 9.384% | -5.313% | 8.493% | 5.791%
| Z-Score (Absolute) |  | 1.54 | 0.16 | 0.04 | 0.44 | 2.48 | 0.55 | 2.32 | 0.2 | 0.67 | 0 | 0.64 | 1.38 | 0.26 | 1.47 | 1.68 | 0.64 | 0.56 | 2.39 | 1.15 | 2.18 | 1.53
| Analytical results for Mo in OREAS 84 (Certified Value 5.05 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.38 | 9.47 | 6 | 6 | 3.6 | 4.791 | 5.5 | 5 | 4.89 | 9.05 | 5 | 5.3 | 4 | 5.02 | 6.08 | < 10 | 5 | 5.7619 | 4.74 | 6.17 |  | 1
| 1 | 2 | 5.43 | 9.89 | 5 | 6 | 3.6 | 4.4789 | 4.3 | 5 | 4.89 | 9.42 | 5.1 | 5.1 | 4 | 5.04 | 5.7 | < 10 | 4.94 | 5.5198 | 5.02 | 5.19 |  | 21
| 1 | 3 | 5.21 | 10.1 | 6 | 7 | 3.4 | 4.8838 | 4.6 | 5 | 4.96 | 9.69 | 5 | 4.7 | 6 | 5.05 | 5.82 | < 10 | 4.83 | 5.4741 | 5.1 | 5.34 |  | 16
| 1 | 4 | 4.99 | 10.2 | 5 | 7 | 3.5 | 4.8661 | 4.9 | 5 | 5.04 | 9.61 | 5.1 | 5.1 | 5 | 5.29 | 6.01 | < 10 | 5.23 | 5.7779 | 4.81 | 5.13 |  | 5.0522
| 1 | 5 | 5.52 | 10.4 | 6 | 6 | 3.5 | 4.4856 | 4.8 | 5 | 4.77 | 7.44 | 5 | 4.8 | 5 | 5.32 | 5.78 | < 10 | 5.02 | 5.9309 | 4.76 | 4.96 |  | 44
| 1 | 6 | 5.08 | 9.71 | 5 | 7 | 4.1 | 4.8523 | 4.8 | 5 | 4.66 | 7.49 | 5 | 5.1 | 5 | 4.88 | 5.73 | < 10 | 4.96 | 6.0194 | 4.95 | 4.9
| Mean |  | 5.2683 | 9.9617 | 5.5 | 6.5 | 3.6167 | 4.7263 | 4.8167 | 5 | 4.8683 | 8.7833 | 5.0333 | 5.0167 | 4.8333 | 5.1 | 5.8533 |  | 4.9967 | 5.7473 | 4.8967 | 5.2817
| Median |  | 5.295 | 9.995 | 5.5 | 6.5 | 3.55 | 4.8216 | 4.8 | 5 | 4.89 | 9.235 | 5 | 5.1 | 5 | 5.045 | 5.8 |  | 4.98 | 5.7699 | 4.88 | 5.16
| Std Dev. |  | 0.20894 | 0.34008 | 0.54772 | 0.54772 | 0.24833 | 0.1916 | 0.39707 | 0 | 0.13556 | 1.0449 | 0.05164 | 0.22286 | 0.75277 | 0.17053 | 0.15565 |  | 0.13216 | 0.2169 | 0.14841 | 0.46327
| Rel.Std.Dev. |  | 3.966% | 3.414% | 9.959% | 8.427% | 6.866% | 4.054% | 8.244% | 0% | 2.785% | 11.9% | 1.026% | 4.442% | 15.57% | 3.344% | 2.659% |  | 2.645% | 3.774% | 3.031% | 8.771%
| PDM3 |  | 4.278% | 97.17% | 8.863% | 28.66% | -28.41% | -6.451% | -4.662% | -1.033% | -3.639% | 73.85% | -0.3736% | -0.7035% | -4.332% | 0.946% | 15.86% |  | -1.099% | 13.76% | -3.079% | 4.542%
| Z-Score (Absolute) |  | 0.73 | 14.74 | NA | NA | 4.21 | 0.89 | 0.62 | NA | 0.47 | 11.22 | 0.02 | 0.02 | NA | 0.22 | 2.47 | 0.07 | 0.08 | 2.16 | 0.38 | 0.77
Results from laboratories 3, 4, 8 and 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 84 (Certified Value 0.446 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.46 | 0.44 | 0.44 | 0.44 | 0.4572 | 0.44069 | 0.43 | 0.44 | 0.47 | 0.43 | 0.44 | 0.461 | 0.43 | 0.432 | 0.49 | 0.4611 | 0.453 | 0.48124 | 0.44 | 0.34125 | 0.46 | 1
| 1 | 2 | 0.46 | 0.45 | 0.44 | 0.43 | 0.4619 | 0.43989 | 0.43 | 0.44 | 0.46 | 0.43 | 0.442 | 0.462 | 0.44 | 0.43 | 0.5 | 0.45812 | 0.453 | 0.45907 | 0.43 | 0.3635 | 0.45 | #N/A
| 1 | 3 | 0.47 | 0.44 | 0.44 | 0.43 | 0.4508 | 0.44041 | 0.43 | 0.43 | 0.44 | 0.43 | 0.435 | 0.454 | 0.44 | 0.436 | 0.47 | 0.45811 | 0.453 | 0.47337 | 0.45 |  | 0.45 | 16
| 1 | 4 | 0.47 | 0.44 | 0.44 | 0.43 | 0.4429 | 0.43962 | 0.43 | 0.44 | 0.46 | 0.43 | 0.443 | 0.449 | 0.43 | 0.431 | 0.49 | 0.44572 | 0.453 | 0.47493 | 0.43 |  | 0.45 | 0.44587
| 1 | 5 | 0.47 | 0.45 | 0.42 | 0.43 | 0.4562 | 0.44023 | 0.44 | 0.45 | 0.43 | 0.47 | 0.443 | 0.458 | 0.44 | 0.432 | 0.48 | 0.45354 | 0.445 | 0.45454 | 0.43 |  | 0.46 | 45
| 1 | 6 | 0.47 | 0.43 | 0.45 | 0.43 | 0.4526 | 0.43957 | 0.44 | 0.44 | 0.45 | 0.48 | 0.441 | 0.45 | 0.43 | 0.445 | 0.5 | 0.45345 | 0.46 | 0.46489 | 0.44 |  | 0.44
| Mean |  | 0.46667 | 0.44167 | 0.43833 | 0.43167 | 0.4536 | 0.44007 | 0.43333 | 0.44 | 0.45167 | 0.445 | 0.44067 | 0.45567 | 0.435 | 0.43433 | 0.48833 | 0.45501 | 0.45283 | 0.46801 | 0.43667 | 0.35237 | 0.45167
| Median |  | 0.47 | 0.44 | 0.44 | 0.43 | 0.4544 | 0.44006 | 0.43 | 0.44 | 0.455 | 0.43 | 0.4415 | 0.456 | 0.435 | 0.432 | 0.49 | 0.45582 | 0.453 | 0.46913 | 0.435 | 0.35237 | 0.45
| Std Dev. |  | 0.005164 | 0.0075277 | 0.0098319 | 0.0040825 | 0.0065106 | 0.00045061 | 0.005164 | 0.0063246 | 0.01472 | 0.023452 | 0.0030111 | 0.0055377 | 0.0054772 | 0.0056095 | 0.01169 | 0.0054262 | 0.0047504 | 0.010225 | 0.008165 | 0.015737 | 0.0075277
| Rel.Std.Dev. |  | 1.107% | 1.704% | 2.243% | 0.9457% | 1.435% | 0.1024% | 1.192% | 1.437% | 3.259% | 5.27% | 0.6833% | 1.215% | 1.259% | 1.292% | 2.394% | 1.193% | 1.049% | 2.185% | 1.87% | 4.466% | 1.667%
| PDM3 |  | 4.664% | -0.9428% | -1.69% | -3.186% | 1.734% | -1.301% | -2.812% | -1.317% | 1.3% | -0.1952% | -1.167% | 2.197% | -2.438% | -2.588% | 9.524% | 2.049% | 1.562% | 4.965% | -2.064% | -20.97% | 1.3%
| Z-Score (Absolute) |  | 1.69 | 0 | 0.22 | 0.67 | 0.8 | 0.11 | 0.56 | 0.11 | 0.67 | 0.22 | 0.07 | 0.94 | 0.45 | 0.49 | 3.15 | 0.9 | 0.73 | 1.78 | 0.34 | 6.02 | 0.67
| Analytical results for Nb in OREAS 84 (Certified Value 2.37 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.3 | 2.5 | 2.3 | 2.71 | 1.9196 | 2.2 | 2.7 | 2.3 | 2.2 | 2.4 | 2.3 | 2.5 | 2.2926 | 2 | 2.4798 | 1.9 | 2.5 |  |  |  |  | 1
| 1 | 2 | 2.1 | 2.6 | 2.4 | 2.65 | 1.7771 | 2.1 | 2.8 | 2.4 | 2.2 | 2.4 | 2.3 | 2.3 | 2.2579 | 2 | 2.3722 | 1.9 | 2.5 |  |  |  |  | 13
| 1 | 3 | 2.2 | 2.6 | 2.2 | 2.47 | 1.8005 | 2.3 | 2.5 | 2.4 | 2.2 | 2.4 | 2.3 | 2.4 | 2.33 | 2 | 2.2471 | 2 | 2.6 |  |  |  |  | 16
| 1 | 4 | 2.3 | 2.7 | 2.2 | 2.49 | 1.9285 | 2.2 | 2.5 | 2.3 | 2.2 | 2.3 | 2.4 | 2.3 | 2.177 | 2 | 2.4549 | 1.8 | 2.6 |  |  |  |  | 2.3659
| 1 | 5 | 2.3 | 2.6 | 2.3 | 2.43 | 1.923 | 2.2 | 2.4 | 2.4 | 2.2 | 2.3 | 2.3 | 2.2 | 2.2303 | 2 | 2.4668 | 1.8 | 2.5 |  |  |  |  | 46
| 1 | 6 | 2.2 | 2.7 | 2.2 | 2.48 | 1.7675 | 2 | 2.4 | 2.4 | 2.3 | 2.3 | 2.4 | 2.2 | 2.2099 | 2 | 2.2887 | 1.9 | 2.5
| Mean |  | 2.2333 | 2.6167 | 2.2667 | 2.5383 | 1.8527 | 2.1667 | 2.55 | 2.3667 | 2.2167 | 2.35 | 2.3333 | 2.3167 | 2.2496 | 2 | 2.3849 | 1.8833 | 2.5333
| Median |  | 2.25 | 2.6 | 2.25 | 2.485 | 1.8601 | 2.2 | 2.5 | 2.4 | 2.2 | 2.35 | 2.3 | 2.3 | 2.2441 | 2 | 2.4135 | 1.9 | 2.5
| Std Dev. |  | 0.08165 | 0.075277 | 0.08165 | 0.11321 | 0.078566 | 0.10328 | 0.16432 | 0.05164 | 0.040825 | 0.054772 | 0.05164 | 0.1169 | 0.055857 | 0 | 0.099025 | 0.075277 | 0.05164
| Rel.Std.Dev. |  | 3.656% | 2.877% | 3.602% | 4.46% | 4.241% | 4.767% | 6.444% | 2.182% | 1.842% | 2.331% | 2.213% | 5.046% | 2.483% | 0% | 4.152% | 3.997% | 2.038%
| PDM3 |  | -5.604% | 10.6% | -4.195% | 7.288% | -21.69% | -8.422% | 7.781% | 0.03175% | -6.308% | -0.6727% | -1.377% | -2.082% | -4.917% | -15.47% | 0.802% | -20.4% | 7.076%
| Z-Score (Absolute) |  | 0.62 | 1.97 | 0.39 | 1.44 | 3.18 | 1.07 | 1.52 | 0.28 | 0.73 | 0.17 | 0.06 | 0.06 | 0.51 | NA | 0.4 | 2.98 | 1.4
Results from laboratory 18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 84 (Certified Value 6.72 ppm) | [22 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 6.5 | 6.4 | 7.95 | 6.592 | 6.5 | 7.1 | 6.5 | 7.232 | 8.2 | [12 empty] | 1
| 1 | 2 | 6.3 | 6.4 | 7.96 | 6.6324 | 6.6 | 7 | 6.8 | 7.2483 | 8 | [12 empty] | #N/A
| 1 | 3 | 6.7 | 6.8 | 7.27 | 6.5385 | 6.3 | 7 | 6.9 | 7.0123 | 8.1 | [12 empty] | 16
| 1 | 4 | 6.5 | 6.6 | 7.2 | 6.5594 | 6.3 | 6.8 | 6.7 | 6.8607 | 8.1 | [12 empty] | 6.7201
| 1 | 5 | 6.4 | 6.3 | 6.99 | 6.5617 | 6.3 | 6.8 | 6.6 | 7.0777 | 8 | [12 empty] | 47
| 1 | 6 | 6.5 | 6.5 | 7.13 | 6.6351 | 6 | 6.8 | 6.8 | 7.0318 | 8.1
| Mean |  | 6.4833 | 6.5 | 7.4167 | 6.5865 | 6.3333 | 6.9167 | 6.7167 | 7.0771 | 8.0833
| Median |  | 6.5 | 6.45 | 7.235 | 6.5768 | 6.3 | 6.9 | 6.75 | 7.0547 | 8.1
| Std Dev. |  | 0.13292 | 0.17889 | 0.42716 | 0.040372 | 0.20656 | 0.13292 | 0.1472 | 0.14587 | 0.075277
| Rel.Std.Dev. |  | 2.05% | 2.752% | 5.759% | 0.6129% | 3.261% | 1.922% | 2.192% | 2.061% | 0.9313%
| PDM3 |  | -3.524% | -3.276% | 10.36% | -1.988% | -5.756% | 2.924% | -0.05176% | 5.312% | 20.29%
| Z-Score (Absolute) |  | 0.67 | 0.63 | 2.02 | 0.38 | 1.11 | 0.58 | 0 | 1.04 | 3.95
| Analytical results for Ni in OREAS 84 (Certified Value 10.47 wt.%) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 10.8 | 10.1 | 10.6551 | 10.35 | 10.8866 | 10.9938 | 10.69 | 10.4 | 11.5314 | 10.5 | 10.221 | 10.4606 | 10.15 | > 1 | > 1 | 9.7679 | 10.6 | 10.3805 | 12.2 | 10.46 |  | 1
| 1 | 2 | 10.8 | 10.3 | 10.7544 | 10.4 | 11.2175 | 10.6815 | 10.68 | 10.35 | 11.0402 | 10.35 | 10.181 | 10.3362 | 10.6 | > 1 | > 1 | 9.6245 | 10.5 | 10.4974 | 11.9 | 10.414 |  | #N/A
| 1 | 3 | 10.6 | 10.25 | 10.7389 | 10.55 | 11.4022 | 10.7311 | 10.66 | 10.35 | 10.926 | 10.5 | 10.52 | 10.5232 | 10.5 | > 1 | > 1 | 9.6488 | 10.4 | 10.2541 | 11.9 | 10.474 |  | 16
| 1 | 4 | 10.6 | 10.3 | 10.6182 | 10.4 | 10.7889 | 10.5571 | 10.22 | 10.4 | 10.7821 | 10.6 | 10.111 | 10.6923 | 10.15 | > 1 | > 1 | 9.7124 | 10.4 | 10.5213 | 12.2 | 10.509 |  | 10.4663
| 1 | 5 | 10 | 10.25 | 10.4151 | 10.4 | 11.2222 | 10.834 | 10.57 | 10.15 | 11.0727 | 10.85 | 10.411 | 10.3496 | 10.4 | > 1 | > 1 | 9.5026 | 10.4 | 10.3431 | 12.2 | 10.496 |  | 48
| 1 | 6 | 10.2 | 10.45 | 10.6673 | 10.35 | 11.423 | 10.8537 | 10.73 | 10.3 | 11.2343 | 10.85 | 10.177 | 10.2593 | 10.25 | > 1 | > 1 | 9.3595 | 10.5 | 10.4475 | 12.1 | 10.532
| Mean |  | 10.5 | 10.275 | 10.6415 | 10.4083 | 11.1567 | 10.7752 | 10.5917 | 10.325 | 11.0978 | 10.6083 | 10.2702 | 10.4369 | 10.3417 |  |  | 9.6026 | 10.4667 | 10.4073 | 12.0833 | 10.4808
| Median |  | 10.6 | 10.275 | 10.6612 | 10.4 | 11.2199 | 10.7825 | 10.67 | 10.35 | 11.0564 | 10.55 | 10.201 | 10.4051 | 10.325 |  |  | 9.6366 | 10.45 | 10.414 | 12.15 | 10.485
| Std Dev. |  | 0.32863 | 0.11292 | 0.12238 | 0.073598 | 0.26359 | 0.15211 | 0.18968 | 0.093541 | 0.26069 | 0.20351 | 0.15914 | 0.15655 | 0.18819 |  |  | 0.14905 | 0.08165 | 0.10104 | 0.1472 | 0.041446
| Rel.Std.Dev. |  | 3.13% | 1.099% | 1.15% | 0.7071% | 2.363% | 1.412% | 1.791% | 0.906% | 2.349% | 1.918% | 1.55% | 1.5% | 1.82% |  |  | 1.552% | 0.7801% | 0.9708% | 1.218% | 0.3954%
| PDM3 |  | 0.3218% | -1.828% | 1.674% | -0.554% | 6.597% | 2.951% | 1.198% | -1.35% | 6.033% | 1.357% | -1.874% | -0.2814% | -1.191% |  |  | -8.252% | 0.003294% | -0.564% | 15.45% | 0.1386%
| Z-Score (Absolute) |  | 0.12 | 0.95 | 0.8 | 0.31 | 3.25 | 1.43 | 0.56 | 0.71 | 2.97 | 0.64 | 0.97 | 0.18 | 0.63 | NA | NA | 4.15 | 0.03 | 0.32 | 7.66 | 0.03
| Analytical results for P in OREAS 84 (Certified Value 0.022 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.023 | 0.022 | 0.02 | 0.023 | 0.0215 | 0.086192 | 0.0217 | 0.022 | 0.0224 | 0.023 | 0.023 | 0.021 | 0.023 | 0.021 | 0.0242 | 0.022595 | < 0.0009 | 0.023614 | 0.0298 | 0.02 |  | 1
| 1 | 2 | 0.023 | 0.023 | 0.02 | 0.022 | 0.0211 | 0.086076 | 0.0201 | 0.022 | 0.0223 | 0.023 | 0.023 | 0.021 | 0.023 | 0.021 | 0.0238 | 0.022141 | < 0.0009 | 0.022521 | 0.0283 | 0.02 |  | #N/A
| 1 | 3 | 0.023 | 0.023 | 0.02 | 0.022 | 0.0205 | 0.086353 | 0.0189 | 0.023 | 0.0215 | 0.023 | 0.023 | 0.021 | 0.024 | 0.021 | 0.024 | 0.022215 | < 0.0009 | 0.022947 | 0.0301 | 0.02 |  | 16
| 1 | 4 | 0.023 | 0.023 | 0.02 | 0.022 | 0.0205 | 0.086205 | 0.0193 | 0.022 | 0.0215 | 0.023 | 0.023 | 0.021 | 0.023 | 0.021 | 0.0239 | 0.020979 | < 0.0009 | 0.022366 | 0.0293 | 0.02 |  | 0.021907
| 1 | 5 | 0.023 | 0.022 | 0.02 | 0.022 | 0.0206 | 0.086067 | 0.0186 | 0.023 | 0.0206 | 0.023 | 0.023 | 0.022 | 0.023 | 0.021 | 0.0238 | 0.021221 | < 0.0009 | 0.022871 | 0.0278 | 0.02 |  | 49
| 1 | 6 | 0.022 | 0.023 | 0.02 | 0.022 | 0.0205 | 0.086189 | 0.0173 | 0.023 | 0.0213 | 0.023 | 0.023 | 0.021 | 0.023 | 0.022 | 0.0235 | 0.02148 | < 0.0009 | 0.023779 | 0.0277 | 0.02
| Mean |  | 0.022833 | 0.022667 | 0.02 | 0.022167 | 0.020783 | 0.08618 | 0.019317 | 0.0225 | 0.0216 | 0.023 | 0.023 | 0.021167 | 0.023167 | 0.021167 | 0.023867 | 0.021772 |  | 0.023016 | 0.028833 | 0.02
| Median |  | 0.023 | 0.023 | 0.02 | 0.022 | 0.02055 | 0.086191 | 0.0191 | 0.0225 | 0.0215 | 0.023 | 0.023 | 0.021 | 0.023 | 0.021 | 0.02385 | 0.021811 |  | 0.022909 | 0.0288 | 0.02
| Std Dev. |  | 0.00040825 | 0.0005164 | 0 | 0.00040825 | 0.00042151 | 0.00010444 | 0.0014865 | 0.00054772 | 0.00066933 | 3.8006e-18 | 3.8006e-18 | 0.00040825 | 0.00040825 | 0.00040825 | 0.00023381 | 0.00063668 |  | 0.00057161 | 0.0010386 | 0
| Rel.Std.Dev. |  | 1.788% | 2.278% | 0% | 1.842% | 2.028% | 0.1212% | 7.695% | 2.434% | 3.099% | 1.652e-14% | 1.652e-14% | 1.929% | 1.762% | 1.929% | 0.9796% | 2.924% |  | 2.483% | 3.602% | 0%
| PDM3 |  | 4.227% | 3.466% | -8.706% | 1.184% | -5.131% | 293.4% | -11.83% | 2.705% | -1.403% | 4.988% | 4.988% | -3.381% | 5.748% | -3.381% | 8.944% | -0.6176% |  | 5.062% | 31.61% | -8.706%
| Z-Score (Absolute) |  | 0.34 | 0.22 | 1.57 | 0.11 | 1.05 | 43.05 | 2.03 | 0.11 | 0.49 | 0.45 | 0.45 | 0.79 | 0.56 | 0.79 | 1.03 | 0.38 | 14.77 | 0.46 | 4.38 | 1.57
| Analytical results for Pb in OREAS 84 (Certified Value 74 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 72.5 | 73 | 66 | 90 | 71 | 68 | 80 | 71.9 | 74 | 83.4 | 75 | 74.4 | 89.3 | 69.2392 | 77.5 | 75.1872 | 74.4 | 104.25 | 83 |  |  | 1
| 1 | 2 | 71.7 | 72 | 70 | 91 | 73 | 67 | 79.6 | 73.5 | 74 | 83.2 | 74 | 74.1 | 92.3 | 68.4299 | 78 | 74.2372 | 74.6 | 101.22 | 85 |  |  | 17
| 1 | 3 | 71.9 | 74 | 67 | 88 | 73 | 66 | 76.7 | 72 | 74 | 81.6 | 85 | 74.7 | 88.5 | 69.2687 | 77.9 | 73.6901 | 78.2 | 107.23 | 82 |  |  | 16
| 1 | 4 | 73.9 | 72 | 68 | 89 | 73 | 68 | 78.9 | 72.3 | 74 | 80 | 76 | 76.4 | 88.2 | 67.546 | 78.8 | 75.5762 | 74.6 | 98.31 | 82 |  |  | 74.4077
| 1 | 5 | 73.6 | 75 | 71 | 91 | 74 | 70 | 75.9 | 73.9 | 75 | 82.2 | 74 | 74.9 | 88.5 | 68.5405 | 78.7 | 75.9624 | 75 | 102.03 | 80 |  |  | 50
| 1 | 6 | 73.5 | 72 | 59 | 89 | 72 | 66 | 77 | 73.2 | 74 | 81.2 | 74 | 76.5 | 89 | 70.1731 | 79.6 | 73.5856 | 76.8 | 102.86 | 79
| Mean |  | 72.85 | 73 | 66.8333 | 89.6667 | 72.6667 | 67.5 | 78.0167 | 72.8 | 74.1667 | 81.9333 | 76.3333 | 75.1667 | 89.3 | 68.8662 | 78.4167 | 74.7064 | 75.6 | 102.65 | 81.8333
| Median |  | 73 | 72.5 | 67.5 | 89.5 | 73 | 67.5 | 77.95 | 72.75 | 74 | 81.9 | 74.5 | 74.8 | 88.75 | 68.8899 | 78.35 | 74.7122 | 74.8 | 102.445 | 82
| Std Dev. |  | 0.94181 | 1.2649 | 4.2622 | 1.2111 | 1.0328 | 1.5166 | 1.7011 | 0.8438 | 0.40825 | 1.2817 | 4.3205 | 1.0309 | 1.5218 | 0.89982 | 0.76267 | 1.0074 | 1.5492 | 2.9948 | 2.137
| Rel.Std.Dev. |  | 1.293% | 1.733% | 6.377% | 1.351% | 1.421% | 2.247% | 2.18% | 1.159% | 0.5504% | 1.564% | 5.66% | 1.371% | 1.704% | 1.307% | 0.9726% | 1.349% | 2.049% | 2.918% | 2.611%
| PDM3 |  | -2.093% | -1.892% | -10.18% | 20.51% | -2.34% | -9.284% | 4.85% | -2.161% | -0.3239% | 10.11% | 2.588% | 1.02% | 20.01% | -7.447% | 5.388% | 0.4015% | 1.602% | 37.96% | 9.98%
| Z-Score (Absolute) |  | 0.55 | 0.51 | 1.97 | 3.43 | 0.59 | 1.81 | 0.67 | 0.56 | 0.24 | 1.6 | 0.28 | 0 | 3.34 | 1.49 | 0.77 | 0.11 | 0.1 | 6.5 | 1.58
| Analytical results for Pr in OREAS 84 (Certified Value 1.82 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 1.68 | 1.83 | 2.08 | 1.7863 | 1.78 | 1.94 | 1.88 | 1.9311 | 2 | [12 empty] | 1
| 1 | 2 | 1.56 | 1.81 | 2.05 | 1.7812 | 1.67 | 1.91 | 1.88 | 1.9366 | 1.9 | [12 empty] | #N/A
| 1 | 3 | 1.75 | 1.82 | 1.92 | 1.7587 | 1.66 | 1.86 | 1.87 | 1.8906 | 1.9 | [12 empty] | 16
| 1 | 4 | 1.74 | 1.63 | 1.92 | 1.7491 | 1.68 | 1.86 | 1.88 | 1.8442 | 1.9 | [12 empty] | 1.8212
| 1 | 5 | 1.71 | 1.77 | 1.86 | 1.7251 | 1.61 | 1.87 | 1.8 | 1.8495 | 2 | [12 empty] | 51
| 1 | 6 | 1.64 | 1.73 | 1.88 | 1.7929 | 1.56 | 1.84 | 1.88 | 1.8891 | 1.9
| Mean |  | 1.68 | 1.765 | 1.9517 | 1.7655 | 1.66 | 1.88 | 1.865 | 1.8902 | 1.9333
| Median |  | 1.695 | 1.79 | 1.92 | 1.7699 | 1.665 | 1.865 | 1.88 | 1.8898 | 1.9
| Std Dev. |  | 0.071274 | 0.075829 | 0.091305 | 0.02598 | 0.074027 | 0.037417 | 0.032094 | 0.038983 | 0.05164
| Rel.Std.Dev. |  | 4.243% | 4.296% | 4.678% | 1.472% | 4.459% | 1.99% | 1.721% | 2.062% | 2.671%
| PDM3 |  | -7.753% | -3.085% | 7.164% | -3.056% | -8.851% | 3.229% | 2.405% | 3.788% | 6.158%
| Z-Score (Absolute) |  | 1.44 | 0.78 | 0.67 | 0.77 | 1.59 | 0.12 | 0 | 0.2 | 0.53
| Analytical results for Rb in OREAS 84 (Certified Value 21.4 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 21.3 | 23.2 | 22.5 | 23.3 | 20.4109 | 19.9 | 19.8 | 22.2 | 20.3 | 23.2 | 20.8 | 20.4 | 23.8997 | 22.9 | 22.6339 | 18.8 | 22.7 |  |  |  |  | 1
| 1 | 2 | 19.7 | 23.8 | 21.8 | 23.1 | 20.7281 | 19.8 | 19.6 | 22.6 | 20.5 | 23.7 | 20.1 | 20.7 | 24.8206 | 22.9 | 21.1172 | 19 | 21.9 |  |  |  |  | 18
| 1 | 3 | 21.4 | 24.7 | 21.1 | 21.6 | 20.5093 | 19.7 | 18.7 | 22.3 | 20.2 | 23.2 | 20.9 | 20.5 | 23.367 | 23.2 | 22.2071 | 19.6 | 22.3 |  |  |  |  | 16
| 1 | 4 | 20.9 | 26 | 21.7 | 21.4 | 20.7349 | 18.9 | 18.7 | 22.1 | 20.2 | 22.8 | 21.1 | 20.5 | 22.5006 | 23 | 22.5563 | 18.8 | 22.8 |  |  |  |  | 21.4196
| 1 | 5 | 21.5 | 24.4 | 21.4 | 21.4 | 20.8533 | 19.2 | 18.3 | 22.1 | 20 | 23.5 | 20.7 | 20 | 22.5847 | 23.1 | 21.6049 | 18.5 | 21.3 |  |  |  |  | 52
| 1 | 6 | 20.5 | 23.7 | 21.2 | 21.5 | 20.7873 | 18.8 | 18.6 | 22 | 20.2 | 23.4 | 20.9 | 20.1 | 22.5951 | 22.9 | 21.6896 | 19 | 22.2
| Mean |  | 20.8833 | 24.3 | 21.6167 | 22.05 | 20.6706 | 19.3833 | 18.95 | 22.2167 | 20.2333 | 23.3 | 20.75 | 20.3667 | 23.2946 | 23 | 21.9682 | 18.95 | 22.2
| Median |  | 21.1 | 24.1 | 21.55 | 21.55 | 20.7315 | 19.45 | 18.7 | 22.15 | 20.2 | 23.3 | 20.85 | 20.45 | 22.9811 | 22.95 | 21.9483 | 18.9 | 22.25
| Std Dev. |  | 0.68823 | 0.98793 | 0.51153 | 0.8961 | 0.172 | 0.47924 | 0.60249 | 0.2137 | 0.1633 | 0.30984 | 0.34496 | 0.26583 | 0.92995 | 0.12649 | 0.59664 | 0.36742 | 0.55136
| Rel.Std.Dev. |  | 3.296% | 4.066% | 2.366% | 4.064% | 0.8321% | 2.472% | 3.179% | 0.9619% | 0.8071% | 1.33% | 1.662% | 1.305% | 3.992% | 0.55% | 2.716% | 1.939% | 2.484%
| PDM3 |  | -2.504% | 13.45% | 0.92% | 2.943% | -3.497% | -9.507% | -11.53% | 3.721% | -5.538% | 8.779% | -3.126% | -4.916% | 8.754% | 7.378% | 2.561% | -11.53% | 3.643%
| Z-Score (Absolute) |  | 0.4 | 1.45 | 0 | 0.23 | 0.51 | 1.2 | 1.44 | 0.32 | 0.75 | 0.91 | 0.47 | 0.67 | 0.91 | 0.75 | 0.19 | 1.44 | 0.31
| Analytical results for Re in OREAS 84 (Certified Value 0.029 ppm) | [22 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 10 | 11 | 12 | 16 | 17 | 19 | 20 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | [9 empty] | 3
3
| 1 | 1 | 0.028 | 0.029 | 0.04 | 0.028 | 0.02 | 0.032 | < 0.05 | < 0.1 | 0.018 | 0.039694 | < 0.05 | 0.028 | [9 empty] | 1
| 1 | 2 | 0.031 | 0.033 | 0.03 | 0.029 | 0.02 | 0.03 | < 0.05 | < 0.1 | 0.016 | 0.043373 | < 0.05 | 0.029 | [9 empty] | 14
| 1 | 3 | 0.028 | 0.035 | 0.03 | 0.024 | 0.02 | 0.033 | < 0.05 | < 0.1 | 0.018 | 0.042364 | < 0.05 | 0.031 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.032 | 0.04 | 0.026 | 0.02 | 0.036 | < 0.05 | < 0.1 | 0.019 | 0.040359 | < 0.05 | 0.031 | [9 empty] | 0.029452
| 1 | 5 | 0.03 | 0.032 | 0.04 | 0.026 | 0.02 | 0.033 | < 0.05 | < 0.1 | 0.013 | 0.038973 | < 0.05 | 0.029 | [9 empty] | 53
| 1 | 6 | 0.03 | 0.032 | 0.05 | 0.026 | 0.02 | 0.029 | < 0.05 | < 0.1 | 0.018 | 0.041968 | < 0.05 | 0.031
| Mean |  | 0.0295 | 0.032167 | 0.038333 | 0.0265 | 0.02 | 0.032167 |  |  | 0.017 | 0.041122 |  | 0.029833
| Median |  | 0.03 | 0.032 | 0.04 | 0.026 | 0.02 | 0.0325 |  |  | 0.018 | 0.041163 |  | 0.03
| Std Dev. |  | 0.0012247 | 0.0019408 | 0.0075277 | 0.0017607 | 0 | 0.0024833 |  |  | 0.0021909 | 0.0017068 |  | 0.0013292
| Rel.Std.Dev. |  | 4.152% | 6.034% | 19.64% | 6.644% | 0% | 7.72% |  |  | 12.89% | 4.151% |  | 4.455%
| PDM3 |  | 0.1617% | 9.216% | 30.15% | -10.02% | -32.09% | 9.216% |  |  | -42.28% | 39.62% |  | 1.293%
| Z-Score (Absolute) |  | 0.02 | 0.36 | 1.25 | 0.46 | 1.4 | 0.36 | 0.67 | 2.94 | 1.83 | 1.66 | 0.67 | 0.02
| Analytical results for S in OREAS 84 (Certified Value 22.14 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | > 20 | 20 | 23.17 | > 10 | 16.5986 | 23.1916 | > 10 | > 10 | 22.5927 | > 10 | 22.6 | > 10 | > 10 | > 10 | > 10 | 20.248 | 16 | 21.4437 | 23.8 | 23.3 |  | 1
| 1 | 2 | > 20 | 20.4 | 23.18 | > 10 | 16.3783 | 23.4068 | > 10 | > 10 | 22.2052 | > 10 | 23.2 | > 10 | > 10 | > 10 | > 10 | 20.6886 | 15.5 | 21.248 | 23.4 | 22.9 |  | #N/A
| 1 | 3 | > 20 | 19.9 | 22.93 | > 10 | 15.8477 | 23.1229 | > 10 | > 10 | 22.2759 | > 10 | 22.6 | > 10 | > 10 | > 10 | > 10 | 20.5087 | 15.9 | 21.2942 | 22.6 | 23.9 |  | 16
| 1 | 4 | > 20 | 20.3 | 22.84 | > 10 | 16.1311 | 23.1316 | > 10 | > 10 | 20.6758 | > 10 | 22.7 | > 10 | > 10 | > 10 | > 10 | 20.5044 | 15.9 | 21.229 | 23.2 | 23.8 |  | 22.1417
| 1 | 5 | > 20 | 20.1 | 22.04 | > 10 | 16.2557 | 23.2251 | > 10 | > 10 | 22.8835 | > 10 | 22.4 | > 10 | > 10 | > 10 | > 10 | 20.3479 | 16 | 21.4922 | 22.1 | 23.5 |  | 54
| 1 | 6 | > 20 | 20.2 | 22.73 | > 10 | 16.5075 | 23.1022 | > 10 | > 10 | 22.5981 | > 10 | 22.4 | > 10 | > 10 | > 10 | > 10 | 20.2819 | 15.9 | 21.3645 | 23 | 23.4
| Mean |  |  | 20.15 | 22.815 |  | 16.2865 | 23.1967 |  |  | 22.2052 |  | 22.65 |  |  |  |  | 20.4299 | 15.8667 | 21.3453 | 23.0167 | 23.4667
| Median |  |  | 20.15 | 22.885 |  | 16.317 | 23.1616 |  |  | 22.4343 |  | 22.6 |  |  |  |  | 20.4262 | 15.9 | 21.3293 | 23.1 | 23.45
| Std Dev. |  |  | 0.18708 | 0.41975 |  | 0.27288 | 0.11278 |  |  | 0.78852 |  | 0.29496 |  |  |  |  | 0.16757 | 0.18619 | 0.10699 | 0.60139 | 0.36148
| Rel.Std.Dev. |  |  | 0.9285% | 1.84% |  | 1.676% | 0.4862% |  |  | 3.551% |  | 1.302% |  |  |  |  | 0.8202% | 1.173% | 0.5012% | 2.613% | 1.54%
| PDM3 |  |  | -8.995% | 3.041% |  | -26.44% | 4.765% |  |  | 0.2868% |  | 2.296% |  |  |  |  | -7.731% | -28.34% | -3.597% | 3.952% | 5.984%
| Z-Score (Absolute) |  | NA | 1.4 | 0.41 | NA | 4.03 | 0.67 | NA | NA | 0 | NA | 0.3 | NA | NA | NA | NA | 1.21 | 4.31 | 0.58 | 0.55 | 0.86
| Analytical results for Sb in OREAS 84 (Certified Value 1.52 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 19 | 20 | 23 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.5 | 1.58 | 1.7 | < 5 | 1.53 | 2.3508 | 2.79 | 10 | 1.47 | 1.6 | 1.5 | 1.4 | < 5 | 1.6 | 1.46 | 1.4115 | 2.2802 | 1.74 | 1.61 |  |  | 1
| 1 | 2 | 1.5 | 1.72 | 1.7 | < 5 | 1.51 | 2.3745 | 2.89 | 13 | 1.7 | 1.62 | 1.5 | 1.32 | < 5 | 1.57 | 1.51 | 1.441 | 2.4227 | 1.76 | 1.58 |  |  | 19
| 1 | 3 | 1.5 | 1.69 | 1.7 | < 5 | 1.44 | 2.3279 | 2.83 | 9 | 1.4 | 1.6 | 1.5 | 1.3 | < 5 | 1.6 | 1.38 | 1.4007 | 2.4569 | 1.76 | 1.54 |  |  | 16
| 1 | 4 | 1.5 | 1.76 | 1.4 | < 5 | 1.4 | 2.3737 | 2.54 | 10 | 1.44 | 1.63 | 1.5 | 1.42 | < 5 | 1.6 | 1.37 | 1.3726 | 2.4296 | 1.68 | 1.59 |  |  | 1.5204
| 1 | 5 | 1.5 | 1.67 | 1.5 | < 5 | 1.38 | 2.4858 | 2.36 | 13 | 1.39 | 1.5 | 1.5 | 1.26 | < 5 | 1.62 | 1.36 | 1.3491 | 2.1185 | 1.64 | 1.49 |  |  | 55
| 1 | 6 | 1.5 | 1.67 | 1.4 | < 5 | 1.37 | 2.2182 | 2.38 | 11 | 1.42 | 1.74 | 1.5 | 1.26 | < 5 | 1.62 | 1.38 | 1.3468 | 2.0588 | 1.7 | 1.52
| Mean |  | 1.5 | 1.6817 | 1.5667 |  | 1.4383 | 2.3552 | 2.6317 | 11 | 1.47 | 1.615 | 1.5 | 1.3267 |  | 1.6017 | 1.41 | 1.387 | 2.2945 | 1.7133 | 1.555
| Median |  | 1.5 | 1.68 | 1.6 |  | 1.42 | 2.3623 | 2.665 | 10.5 | 1.43 | 1.61 | 1.5 | 1.31 |  | 1.6 | 1.38 | 1.3866 | 2.3515 | 1.72 | 1.56
| Std Dev. |  | 0 | 0.06047 | 0.15055 |  | 0.067946 | 0.086402 | 0.23524 | 1.6733 | 0.11628 | 0.076877 | 0 | 0.068896 |  | 0.018348 | 0.060663 | 0.037315 | 0.17192 | 0.048442 | 0.045935
| Rel.Std.Dev. |  | 0% | 3.596% | 9.61% |  | 4.724% | 3.669% | 8.939% | 15.21% | 7.91% | 4.76% | 0% | 5.193% |  | 1.146% | 4.302% | 2.69% | 7.493% | 2.827% | 2.954%
| PDM3 |  | -1.342% | 10.61% | 3.043% |  | -5.398% | 54.9% | 73.09% | 623.5% | -3.315% | 6.222% | -1.342% | -12.74% |  | 5.345% | -7.262% | -8.777% | 50.91% | 12.69% | 2.275%
| Z-Score (Absolute) |  | 0.42 | 0.33 | 0.14 | 3.71 | 0.67 | 3.11 | 4.25 | 38.8 | 0.54 | 0.06 | 0.42 | 1.14 | 3.71 | 0 | 0.79 | 0.89 | 2.86 | 0.46 | 0.19
| Analytical results for Sc in OREAS 84 (Certified Value 5.21 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.6 | 5.8 | < 20 | 3 | 4.67 | 5.0126 | 3.8 | 6 | 4.9 | 5.8 | 5 | 4.9 | 5 | 4.8 | 5.6 | 5.2213 | 6.4 | 5.5144 | 5 | 5.8 |  | 1
| 1 | 2 | 5.4 | 6 | < 20 | 3 | 4.7 | 4.9542 | 4 | 6 | 4.8 | 5.9 | 5 | 4.9 | 5 | 4.8 | 5.6 | 5.2053 | 6.4 | 4.8105 | 5.2 | 5.6 |  | 20
| 1 | 3 | 5.2 | 6 | < 20 | 3 | 4.55 | 5.0827 | 3.7 | 6 | 4.7 | 5.6 | 5 | 4.8 | 6 | 4.8 | 5.5 | 5.2026 | 6.6 | 5.324 | 5.2 | 5.6 |  | 16
| 1 | 4 | 5.3 | 6.3 | < 20 | 3 | 4.53 | 4.8699 | 3.6 | 6 | 4.8 | 5.8 | 5 | 4.9 | 6 | 4.9 | 5.6 | 5.0012 | 6.6 | 5.0903 | 5.2 | 5.8 |  | 5.2099
| 1 | 5 | 5.3 | 6 | < 20 | 3 | 4.61 | 4.889 | 3.4 | 6 | 4.6 | 5.9 | 5 | 4.8 | 6 | 4.9 | 5.6 | 5.0471 | 6.7 | 5.438 | 4.9 | 5.4 |  | 56
| 1 | 6 | 5.3 | 6 | < 20 | 3 | 4.59 | 4.7783 | 3.4 | 6 | 4.7 | 5.9 | 5 | 4.6 | 5 | 4.9 | 5.4 | 4.9901 | 6.7 | 4.9911 | 5.3 | 5.4
| Mean |  | 5.35 | 6.0167 |  | 3 | 4.6083 | 4.9311 | 3.65 | 6 | 4.75 | 5.8167 | 5 | 4.8167 | 5.5 | 4.85 | 5.55 | 5.1113 | 6.5667 | 5.1947 | 5.1333 | 5.6
| Median |  | 5.3 | 6 |  | 3 | 4.6 | 4.9216 | 3.65 | 6 | 4.75 | 5.85 | 5 | 4.85 | 5.5 | 4.85 | 5.6 | 5.1248 | 6.6 | 5.2071 | 5.2 | 5.6
| Std Dev. |  | 0.13784 | 0.16021 |  | 0 | 0.066458 | 0.10863 | 0.23452 | 0 | 0.10488 | 0.1169 | 0 | 0.1169 | 0.54772 | 0.054772 | 0.083666 | 0.10972 | 0.13663 | 0.275 | 0.15055 | 0.17889
| Rel.Std.Dev. |  | 2.576% | 2.663% |  | 0% | 1.442% | 2.203% | 6.425% | 0% | 2.208% | 2.01% | 0% | 2.427% | 9.959% | 1.129% | 1.507% | 2.147% | 2.081% | 5.294% | 2.933% | 3.194%
| PDM3 |  | 2.689% | 15.49% |  | -42.42% | -11.55% | -5.351% | -29.94% | 15.17% | -8.828% | 11.65% | -4.029% | -7.548% | 5.568% | -6.908% | 6.528% | -1.893% | 26.04% | -0.2917% | -1.47% | 7.488%
| Z-Score (Absolute) |  | 0.31 | 1.44 | 8.15 | NA | 0.94 | 0.39 | 2.55 | NA | 0.7 | 1.1 | NA | 0.59 | NA | 0.53 | 0.65 | 0.09 | 2.36 | 0.05 | 0.05 | 0.73
Results from laboratories 4, 8, 11 and 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 84 (Certified Value 10.7 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 20 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 1
1
| 1 | 1 | 11.5 | 13 | 12 | 11.38 | 8 | 13 | 10 | 11 | 10 | 11 | 11 | 9.1622 | 11.4719 | 10.1 | 15 | [6 empty] | 1
| 1 | 2 | 11.6 | 14 | 9 | 11.51 | 8 | 13 | 10 | 11 | 10 | 11 | 10 | 9.397 | 10.8219 | 10.1 | 11 | [6 empty] | #N/A
| 1 | 3 | 11.5 | 13 | 9 | 10.9 | 7 | 11 | 9 | 12 | 10 | 11 | 10 | 8.8905 | 12.0385 | 10 | 14 | [6 empty] | 16
| 1 | 4 | 12.7 | 13 | 10 | 10.68 | 9 | 12 | 9 | 11 | 10 | 11 | 10 | 8.4905 | 12.0785 | 9.9 | 15 | [6 empty] | 10.7094
| 1 | 5 | 12.4 | 13 | 11 | 10.65 | 8 | 12 | 13 | 11 | 10 | 11 | 10 | 8.4555 | 11.399 | 9.8 | 14 | [6 empty] | 57
| 1 | 6 | 11.4 | 13 | 12 | 10.64 | 8 | 12 | 12 | 12 | 10 | 11 | 10 | 8.4947 | 10.8296 | 10.3 | 14
| Mean |  | 11.85 | 13.1667 | 10.5 | 10.96 | 8 | 12.1667 | 10.5 | 11.3333 | 10 | 11 | 10.1667 | 8.8151 | 11.4399 | 10.0333 | 13.8333
| Median |  | 11.55 | 13 | 10.5 | 10.79 | 8 | 12 | 10 | 11 | 10 | 11 | 10 | 8.6926 | 11.4354 | 10.05 | 14
| Std Dev. |  | 0.55408 | 0.40825 | 1.3784 | 0.38972 | 0.63246 | 0.75277 | 1.6432 | 0.5164 | 0 | 0 | 0.40825 | 0.40051 | 0.55195 | 0.17512 | 1.472
| Rel.Std.Dev. |  | 4.676% | 3.101% | 13.13% | 3.556% | 7.906% | 6.187% | 15.65% | 4.556% | 0% | 0% | 4.016% | 4.543% | 4.825% | 1.745% | 10.64%
| PDM3 |  | 10.65% | 22.94% | -1.955% | 2.34% | -25.3% | 13.61% | -1.955% | 5.826% | -6.624% | 2.713% | -5.068% | -17.69% | 6.821% | -6.313% | 29.17%
| Z-Score (Absolute) |  | 0.67 | 1.67 | 0.35 | 0 | 2.24 | 0.91 | 0.35 | 0.28 | 0.73 | 0.03 | 0.6 | 1.63 | 0.36 | 0.7 | 2.18
| Analytical results for Sm in OREAS 84 (Certified Value 1.31 ppm) | [22 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 12 | 16 | 17 | 18 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 1.32 | 1.2 | 1.42 | 1.2334 | 1.15 | 1.4 | 1.27 | 1.3685 | 1.5 | [12 empty] | 1
| 1 | 2 | 1.28 | 1.5 | 1.41 | 1.2638 | 1.12 | 1.3 | 1.32 | 1.3497 | 1.5 | [12 empty] | #N/A
| 1 | 3 | 1.43 | 1.5 | 1.3 | 1.204 | 1.18 | 1.3 | 1.3 | 1.3659 | 1.6 | [12 empty] | 16
| 1 | 4 | 1.39 | 1.3 | 1.28 | 1.2319 | 1.02 | 1.3 | 1.28 | 1.3294 | 1.5 | [12 empty] | 1.3108
| 1 | 5 | 1.39 | 1.3 | 1.26 | 1.2386 | 1.01 | 1.3 | 1.22 | 1.3454 | 1.5 | [12 empty] | 58
| 1 | 6 | 1.19 | 1.2 | 1.24 | 1.2583 | 1.04 | 1.3 | 1.31 | 1.3556 | 1.6
| Mean |  | 1.3333 | 1.3333 | 1.3183 | 1.2383 | 1.0867 | 1.3167 | 1.2833 | 1.3524 | 1.5333
| Median |  | 1.355 | 1.3 | 1.29 | 1.236 | 1.08 | 1.3 | 1.29 | 1.3526 | 1.5
| Std Dev. |  | 0.088694 | 0.13663 | 0.077567 | 0.021401 | 0.072572 | 0.040825 | 0.036148 | 0.014406 | 0.05164
| Rel.Std.Dev. |  | 6.652% | 10.25% | 5.884% | 1.728% | 6.678% | 3.101% | 2.817% | 1.065% | 3.368%
| PDM3 |  | 1.717% | 1.717% | 0.5728% | -5.53% | -17.1% | 0.4456% | -2.097% | 3.175% | 16.97%
| Z-Score (Absolute) |  | 0.3 | 0.3 | 0 | 1.58 | 4.58 | 0.03 | 0.69 | 0.67 | 4.25
| Analytical results for Sn in OREAS 84 (Certified Value 1.49 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.5 | 1.9 | 2 | 1.6 | 2.2 | 2 | 1.3 | 1.4 | 1.6 | 1.5 | 1.4 | < 2.5 | 1.43 | 1.6984 | 1.9 | 1.4 | [5 empty] | 1
| 1 | 2 | 1.5 | 1.8 | 2 | 1.59 | 1.9 | 2 | 1.6 | 1.4 | 1.6 | 1.5 | 1.4 | < 2.5 | 1.44 | 1.6102 | 1.5 | 1.4 | [5 empty] | 21
| 1 | 3 | 1.3 | 1.9 | 2 | 1.47 | 2 | 1.9 | 1.3 | 1.4 | 1.6 | 1.5 | 1.3 | < 2.5 | 1.39 | 1.5591 | 1.6 | 1.4 | [5 empty] | 16
| 1 | 4 | 1.5 | 2 | 2 | 1.5 | 2 | 1.9 | 1.3 | 1.3 | 1.5 | 1.6 | 1.2 | < 2.5 | 1.5 | 1.5272 | 1.5 | 1.4 | [5 empty] | 1.4943
| 1 | 5 | 1.6 | 1.9 | 1 | 1.46 | 2 | 1.9 | 1.5 | 1.3 | 1.5 | 1.6 | 1.2 | < 2.5 | 1.44 | 1.5902 | 1.5 | 1.4 | [5 empty] | 59
| 1 | 6 | 1.5 | 1.8 | 1 | 1.45 | 1.8 | 2 | 1.4 | 1.3 | 1.5 | 1.5 | 1.2 | < 2.5 | 1.44 | 1.7126 | 1.5 | 1.3
| Mean |  | 1.4833 | 1.8833 | 1.6667 | 1.5117 | 1.9833 | 1.95 | 1.4 | 1.35 | 1.55 | 1.5333 | 1.2833 |  | 1.44 | 1.6163 | 1.5833 | 1.3833
| Median |  | 1.5 | 1.9 | 2 | 1.485 | 2 | 1.95 | 1.35 | 1.35 | 1.55 | 1.5 | 1.25 |  | 1.44 | 1.6002 | 1.5 | 1.4
| Std Dev. |  | 0.098319 | 0.075277 | 0.5164 | 0.066758 | 0.13292 | 0.054772 | 0.12649 | 0.054772 | 0.054772 | 0.05164 | 0.098319 |  | 0.035214 | 0.074752 | 0.16021 | 0.040825
| Rel.Std.Dev. |  | 6.628% | 3.997% | 30.98% | 4.416% | 6.702% | 2.809% | 9.035% | 4.057% | 3.534% | 3.368% | 7.661% |  | 2.445% | 4.625% | 10.12% | 2.951%
| PDM3 |  | -0.7321% | 26.04% | 11.54% | 1.164% | 32.73% | 30.5% | -6.309% | -9.655% | 3.729% | 2.614% | -14.12% |  | -3.632% | 8.164% | 5.96% | -7.424%
| Z-Score (Absolute) |  | 0.17 | 2.24 | NA | 0 | 2.85 | 2.65 | 0.67 | 0.98 | 0.23 | 0.13 | 1.38 | 1.58 | 0.43 | 0.63 | 0.43 | 0.77
Results from laboratory 3 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 84 (Certified Value 87 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 84.3 | 92 | 91.8 | 92 | 74 | 87.7428 | 84 | 92 | 88.1 | 87.4 | 83.1 | 90 | 92 | 80.4 | 91.7 | 84.1579 | 91 | 86.447 | 80.3 | 97.5 |  | 1
| 1 | 2 | 82.7 | 92 | 91 | 91 | 72 | 84.158 | 83 | 92 | 84.3 | 89.4 | 83.2 | 90 | 93 | 80.2 | 91.9 | 81.8278 | 90 | 85.6717 | 79.6 | 95.6 |  | 16
| 1 | 3 | 82.3 | 91 | 92.7 | 90 | 77 | 86.0914 | 83 | 90 | 81.7 | 87.9 | 82.8 | 89 | 95 | 80.5 | 89.6 | 80.6921 | 90 | 87.514 | 82.6 | 98.8 |  | 16
| 1 | 4 | 84.9 | 92 | 92.1 | 92 | 71 | 83.8861 | 83 | 92 | 85.1 | 88.3 | 83.2 | 87 | 94 | 80.8 | 91.4 | 77.0092 | 90 | 89.3726 | 79.5 | 98.6 |  | 86.8664
| 1 | 5 | 84.2 | 91 | 88.5 | 90 | 73 | 86.1316 | 83 | 94 | 81.4 | 92.9 | 82.6 | 88 | 94 | 81 | 91.3 | 78.2505 | 89 | 90.7998 | 78.9 | 97.3 |  | 60
| 1 | 6 | 81.5 | 86 | 91 | 90 | 75 | 85.5003 | 83 | 92 | 83.4 | 92.4 | 83.2 | 88 | 92 | 83.4 | 91.1 | 77.7557 | 91 | 89.459 | 81.4 | 95.7
| Mean |  | 83.3167 | 90.6667 | 91.1833 | 90.8333 | 73.6667 | 85.585 | 83.1667 | 92 | 84 | 89.7167 | 83.0167 | 88.6667 | 93.3333 | 81.05 | 91.1667 | 79.9489 | 90.1667 | 88.2107 | 80.3833 | 97.25
| Median |  | 83.45 | 91.5 | 91.4 | 90.5 | 73.5 | 85.7958 | 83 | 92 | 83.85 | 88.85 | 83.15 | 88.5 | 93.5 | 80.65 | 91.35 | 79.4713 | 90 | 88.4433 | 79.95 | 97.4
| Std Dev. |  | 1.3393 | 2.3381 | 1.4689 | 0.98319 | 2.1602 | 1.4242 | 0.40825 | 1.2649 | 2.4706 | 2.371 | 0.25626 | 1.2111 | 1.2111 | 1.1862 | 0.81894 | 2.7617 | 0.75277 | 1.9826 | 1.3819 | 1.3722
| Rel.Std.Dev. |  | 1.607% | 2.579% | 1.611% | 1.082% | 2.932% | 1.664% | 0.4909% | 1.375% | 2.941% | 2.643% | 0.3087% | 1.366% | 1.298% | 1.464% | 0.8983% | 3.454% | 0.8349% | 2.248% | 1.719% | 1.411%
| PDM3 |  | -4.086% | 4.375% | 4.97% | 4.567% | -15.2% | -1.475% | -4.259% | 5.91% | -3.3% | 3.281% | -4.432% | 2.072% | 7.445% | -6.696% | 4.95% | -7.963% | 3.799% | 1.548% | -7.463% | 11.95%
| Z-Score (Absolute) |  | 0.86 | 0.38 | 0.46 | 0.4 | 2.49 | 0.48 | 0.89 | 0.6 | 0.75 | 0.22 | 0.91 | 0.04 | 0.83 | 1.25 | 0.46 | 1.43 | 0.29 | 0.04 | 1.36 | 1.49
| Analytical results for Ta in OREAS 84 (Certified Value 0.21 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.17 | 0.2 | 0.2 | 0.25 | 0.3 | 0.28 | 0.18 | 0.18 | 0.22 | 0.19 | 0.36 | 0.32796 | 0.12 | 0.25316 | < 0.05 | 0.2 | [5 empty] | 1
| 1 | 2 | 0.17 | 0.19 | 0.2 | 0.21 | 0.32 | 0.28 | 0.18 | 0.18 | 0.22 | 0.18 | 0.37 | 0.27523 | 0.12 | 0.24002 | < 0.05 | 0.19 | [5 empty] | 10
| 1 | 3 | 0.16 | 0.19 | 0.2 | 0.23 | 0.34 | 0.27 | 0.19 | 0.17 | 0.22 | 0.19 | 0.38 | 0.26931 | 0.12 | 0.24496 | < 0.05 | 0.2 | [5 empty] | 16
| 1 | 4 | 0.18 | 0.2 | 0.2 | 0.23 | 0.34 | 0.27 | 0.18 | 0.18 | 0.22 | 0.18 | 0.43 | 0.28567 | 0.12 | 0.26711 | < 0.05 | 0.2 | [5 empty] | 0.20525
| 1 | 5 | 0.17 | 0.19 | 0.2 | 0.21 | 0.34 | 0.25 | 0.18 | 0.18 | 0.23 | 0.18 | 0.38 | 0.27364 | 0.12 | 0.2536 | < 0.05 | 0.22 | [5 empty] | 61
| 1 | 6 | 0.18 | 0.19 | 0.2 | 0.2 | 0.33 | 0.25 | 0.17 | 0.18 | 0.21 | 0.19 | 0.34 | 0.26537 | 0.11 | 0.26628 | < 0.05 | 0.19
| Mean |  | 0.17167 | 0.19333 | 0.2 | 0.22167 | 0.32833 | 0.26667 | 0.18 | 0.17833 | 0.22 | 0.185 | 0.37667 | 0.28286 | 0.11833 | 0.25419 |  | 0.2
| Median |  | 0.17 | 0.19 | 0.2 | 0.22 | 0.335 | 0.27 | 0.18 | 0.18 | 0.22 | 0.185 | 0.375 | 0.27444 | 0.12 | 0.25338 |  | 0.2
| Std Dev. |  | 0.0075277 | 0.005164 | 3.0405e-17 | 0.018348 | 0.016021 | 0.013663 | 0.0063246 | 0.0040825 | 0.0063246 | 0.0054772 | 0.030111 | 0.023128 | 0.0040825 | 0.01096 |  | 0.010954
| Rel.Std.Dev. |  | 4.385% | 2.671% | 1.52e-14% | 8.278% | 4.879% | 5.123% | 3.514% | 2.289% | 2.875% | 2.961% | 7.994% | 8.176% | 3.45% | 4.312% |  | 5.477%
| PDM3 |  | -16.36% | -5.807% | -2.559% | 7.997% | 59.97% | 29.92% | -12.3% | -13.12% | 7.185% | -9.867% | 83.51% | 37.81% | -42.35% | 23.84% |  | -2.559%
| Z-Score (Absolute) |  | 0.67 | 0.16 | NA | 0.52 | 3.05 | 1.59 | 0.48 | 0.52 | 0.48 | 0.36 | 4.2 | 1.97 | 1.94 | 1.29 | 4.16 | 0
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 84 (Certified Value 0.17 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 23 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.16 | 0.19 | < 0.2 | 0.18076 | 0.14 | 0.16 | 0.18 | 0.17 | 0.21 | 0.16603 | 0.18 | 0.2 | [9 empty] | 1
| 1 | 2 | 0.15 | 0.2 | < 0.2 | 0.16929 | 0.13 | 0.15 | 0.19 | 0.17 | 0.18 | 0.1655 | 0.18 | 0.2 | [9 empty] | #N/A
| 1 | 3 | 0.18 | 0.18 | < 0.2 | 0.17428 | 0.12 | 0.15 | 0.18 | 0.17 | 0.18 | 0.16495 | 0.18 | 0.21 | [9 empty] | 16
| 1 | 4 | 0.16 | 0.19 | < 0.2 | 0.17525 | 0.11 | 0.15 | 0.18 | 0.17 | 0.18 | 0.15247 | 0.19 | 0.21 | [9 empty] | 0.175
| 1 | 5 | 0.16 | 0.19 | < 0.2 | 0.17174 | 0.14 | 0.14 | 0.18 | 0.17 | 0.17 | 0.16374 | 0.19 | 0.19 | [9 empty] | 62
| 1 | 6 | 0.17 | 0.16 | < 0.2 | 0.16815 | 0.1 | 0.15 | 0.18 | 0.18 | 0.18 | 0.15969 | 0.19 | 0.19
| Mean |  | 0.16333 | 0.185 |  | 0.17325 | 0.12333 | 0.15 | 0.18167 | 0.17167 | 0.18333 | 0.16206 | 0.185 | 0.2
| Median |  | 0.16 | 0.19 |  | 0.17301 | 0.125 | 0.15 | 0.18 | 0.17 | 0.18 | 0.16435 | 0.185 | 0.2
| Std Dev. |  | 0.010328 | 0.013784 |  | 0.0045927 | 0.01633 | 0.0063246 | 0.0040825 | 0.0040825 | 0.013663 | 0.0052242 | 0.0054772 | 0.0089443
| Rel.Std.Dev. |  | 6.323% | 7.451% |  | 2.651% | 13.24% | 4.216% | 2.247% | 2.378% | 7.452% | 3.224% | 2.961% | 4.472%
| PDM3 |  | -6.665% | 5.716% |  | -1.001% | -29.52% | -14.28% | 3.811% | -1.903% | 4.763% | -7.391% | 5.716% | 14.29%
| Z-Score (Absolute) |  | 0.53 | 0.72 | 4.17 | 0.05 | 2.83 | 1.29 | 0.53 | 0.05 | 0.63 | 0.6 | 0.72 | 1.59
| Analytical results for Te in OREAS 84 (Certified Value 1.27 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 20 | 23 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 0.86 | 1.23 | 1.5 | 1.33 | 0.81 | 1.18 | 1.12 | 1.2 | 1.34 | 1.17 | 1.28 | 1.3811 | 2.0646 | 1.19 | 1.52 | [6 empty] | 1
| 1 | 2 | 0.77 | 1.34 | 1.2 | 1.26 | 0.93 | 1.09 | 1.25 | 1.1 | 1.39 | 1.22 | 1.13 | 1.5056 | 1.8141 | 0.9 | 1.54 | [6 empty] | 11
| 1 | 3 | 0.5 | 1.23 | 1.6 | 1.18 | 0.75 | 1.09 | 1.22 | 1.2 | 1.29 | 1.16 | 1.26 | 1.5694 | 1.6672 | 1 | 1.49 | [6 empty] | 16
| 1 | 4 | 0.74 | 1.34 | 1.3 | 1.24 | 0.79 | 1.14 | 1.32 | 1.1 | 1.28 | 1.19 | 1.15 | 1.4233 | 1.9738 | 1.12 | 1.48 | [6 empty] | 1.2653
| 1 | 5 | 0.88 | 1.28 | 1.4 | 1.21 | 0.83 | 0.98 | 1.33 | 1.2 | 1.34 | 1.27 | 1.17 | 1.3328 | 1.1945 | 1.07 | 1.35 | [6 empty] | 63
| 1 | 6 | 0.58 | 1.28 | 1.6 | 1.16 | 0.92 | 1.03 | 1.14 | 1.1 | 1.28 | 1.25 | 1.17 | 1.4406 | 1.3487 | 1.57 | 1.41
| Mean |  | 0.72167 | 1.2833 | 1.4333 | 1.23 | 0.83833 | 1.085 | 1.23 | 1.15 | 1.32 | 1.21 | 1.1933 | 1.4421 | 1.6771 | 1.1417 | 1.465
| Median |  | 0.755 | 1.28 | 1.45 | 1.225 | 0.82 | 1.09 | 1.235 | 1.15 | 1.315 | 1.205 | 1.17 | 1.4319 | 1.7406 | 1.095 | 1.485
| Std Dev. |  | 0.15237 | 0.049261 | 0.1633 | 0.061319 | 0.072226 | 0.072319 | 0.088091 | 0.054772 | 0.044272 | 0.044272 | 0.061536 | 0.085172 | 0.34576 | 0.23233 | 0.071764
| Rel.Std.Dev. |  | 21.11% | 3.839% | 11.39% | 4.985% | 8.615% | 6.665% | 7.162% | 4.763% | 3.354% | 3.659% | 5.157% | 5.906% | 20.62% | 20.35% | 4.899%
| PDM3 |  | -42.97% | 1.424% | 13.28% | -2.791% | -33.75% | -14.25% | -2.791% | -9.114% | 4.322% | -4.372% | -5.689% | 13.97% | 32.55% | -9.772% | 15.78%
| Z-Score (Absolute) |  | 3.81 | 0.4 | 1.52 | 0 | 2.93 | 1.09 | 0 | 0.6 | 0.67 | 0.15 | 0.27 | 1.59 | 3.35 | 0.66 | 1.76
| Analytical results for Th in OREAS 84 (Certified Value 3.14 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.91 | 3.13 | 3.3 | < 20 | 3.28 | 2.8931 | 2.75 | < 20 | 3.6 | 3.18 | 3.02 | 3.44 | < 20 | 2.97 | 3.34 | 3.17 | 3.44 | 3.2087 | 3.2 | 3.4 |  | 1
| 1 | 2 | 2.73 | 3.25 | 3.4 | < 20 | 3.23 | 2.9732 | 2.78 | < 20 | 3.4 | 3.45 | 3.01 | 3.34 | < 20 | 3.24 | 3.37 | 3.2582 | 3.23 | 2.8762 | 3.1 | 3.3 |  | #N/A
| 1 | 3 | 2.74 | 3.25 | 3.2 | < 20 | 2.92 | 2.8252 | 2.71 | < 20 | 3 | 3.19 | 3.08 | 3.43 | < 20 | 3.08 | 3.41 | 3.3227 | 3.21 | 3.0957 | 3.3 | 3.3 |  | 16
| 1 | 4 | 2.82 | 3.39 | 3.2 | < 20 | 2.91 | 2.8872 | 2.82 | < 20 | 3.3 | 3.15 | 3.04 | 3.36 | < 20 | 3.15 | 3.33 | 3.0774 | 3.35 | 3.3182 | 3.1 | 3.5 |  | 3.1385
| 1 | 5 | 2.73 | 3.28 | 3 | < 20 | 2.81 | 2.6546 | 2.82 | < 20 | 3.2 | 3.06 | 3.02 | 3.26 | < 20 | 3.12 | 3.26 | 3.1414 | 3.32 | 3.1497 | 3.4 | 3.3 |  | 64
| 1 | 6 | 2.79 | 3.25 | 3.2 | < 20 | 2.78 | 2.8323 | 2.8 | < 20 | 3.1 | 2.96 | 3.01 | 3.32 | < 20 | 3.14 | 3.41 | 3.2114 | 3.94 | 2.9493 | 3.3 | 3.3
| Mean |  | 2.7867 | 3.2583 | 3.2167 |  | 2.9883 | 2.8443 | 2.78 |  | 3.2667 | 3.165 | 3.03 | 3.3583 |  | 3.1167 | 3.3533 | 3.1968 | 3.415 | 3.0996 | 3.2333 | 3.35
| Median |  | 2.765 | 3.25 | 3.2 |  | 2.915 | 2.8598 | 2.79 |  | 3.25 | 3.165 | 3.02 | 3.35 |  | 3.13 | 3.355 | 3.1907 | 3.335 | 3.1227 | 3.25 | 3.3
| Std Dev. |  | 0.070616 | 0.083046 | 0.13292 |  | 0.21424 | 0.10708 | 0.043359 |  | 0.21602 | 0.16453 | 0.026833 | 0.06824 |  | 0.089144 | 0.056804 | 0.087029 | 0.27046 | 0.16415 | 0.12111 | 0.083666
| Rel.Std.Dev. |  | 2.534% | 2.549% | 4.132% |  | 7.169% | 3.765% | 1.56% |  | 6.613% | 5.198% | 0.8856% | 2.032% |  | 2.86% | 1.694% | 2.722% | 7.92% | 5.296% | 3.746% | 2.497%
| PDM3 |  | -11.21% | 3.819% | 2.491% |  | -4.784% | -9.374% | -11.42% |  | 4.085% | 0.8452% | -3.456% | 7.005% |  | -0.6949% | 6.846% | 1.86% | 8.811% | -1.238% | 3.022% | 6.74%
| Z-Score (Absolute) |  | 2.26 | 0.17 | 0.04 | 34.92 | 1.22 | 1.96 | 2.29 | 34.92 | 0.21 | 0.31 | 1.01 | 0.69 | 34.92 | 0.56 | 0.66 | 0.15 | 0.98 | 0.65 | 0.04 | 0.64
| Analytical results for Ti in OREAS 84 (Certified Value 0.117 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.11 | 0.118 | 0.1088 | 0.11 | 0.0973 | 0.1094 | 0.1 | 0.12 | 0.12 | 0.119 | 0.119 | 0.123 | 0.11 | 0.102 | 0.124 | 0.11786 | 0.126 | 0.11892 | 0.12 | 0.11666 | 0.13 | 1
| 1 | 2 | 0.11 | 0.118 | 0.1091 | 0.11 | 0.0972 | 0.10581 | 0.1 | 0.12 | 0.12 | 0.122 | 0.12 | 0.122 | 0.11 | 0.103 | 0.122 | 0.11857 | 0.126 | 0.1188 | 0.11 | 0.1178 | 0.12 | #N/A
| 1 | 3 | 0.11 | 0.118 | 0.1085 | 0.11 | 0.0953 | 0.10949 | 0.1 | 0.12 | 0.11 | 0.119 | 0.119 | 0.121 | 0.12 | 0.106 | 0.121 | 0.11752 | 0.126 | 0.11845 | 0.12 | 0.12289 | 0.13 | 16
| 1 | 4 | 0.12 | 0.118 | 0.1076 | 0.11 | 0.0949 | 0.10719 | 0.1 | 0.12 | 0.11 | 0.118 | 0.119 | 0.118 | 0.12 | 0.103 | 0.123 | 0.11701 | 0.126 | 0.11735 | 0.12 | 0.10791 | 0.13 | 0.11712
| 1 | 5 | 0.11 | 0.116 | 0.1045 | 0.11 | 0.0953 | 0.10972 | 0.1 | 0.12 | 0.11 | 0.116 | 0.12 | 0.121 | 0.12 | 0.104 | 0.121 | 0.11608 | 0.126 | 0.12033 | 0.12 |  | 0.12 | 65
| 1 | 6 | 0.12 | 0.119 | 0.1074 | 0.11 | 0.0958 | 0.10939 | 0.1 | 0.12 | 0.11 | 0.121 | 0.12 | 0.12 | 0.11 | 0.104 | 0.116 | 0.11609 | 0.126 | 0.12052 | 0.12 |  | 0.12
| Mean |  | 0.11333 | 0.11783 | 0.10765 | 0.11 | 0.095967 | 0.1085 | 0.1 | 0.12 | 0.11333 | 0.11917 | 0.1195 | 0.12083 | 0.115 | 0.10367 | 0.12117 | 0.11719 | 0.126 | 0.11906 | 0.11833 | 0.11632 | 0.125
| Median |  | 0.11 | 0.118 | 0.10805 | 0.11 | 0.09555 | 0.1094 | 0.1 | 0.12 | 0.11 | 0.119 | 0.1195 | 0.121 | 0.115 | 0.1035 | 0.1215 | 0.11726 | 0.126 | 0.11886 | 0.12 | 0.11723 | 0.125
| Std Dev. |  | 0.005164 | 0.00098319 | 0.0016814 | 0 | 0.0010347 | 0.0016164 | 1.5202e-17 | 0 | 0.005164 | 0.002137 | 0.00054772 | 0.0017224 | 0.0054772 | 0.0013663 | 0.0027869 | 0.00099357 | 0 | 0.0011953 | 0.0040825 | 0.0062236 | 0.0054772
| Rel.Std.Dev. |  | 4.556% | 0.8344% | 1.562% | 0% | 1.078% | 1.49% | 1.52e-14% | 0% | 4.556% | 1.793% | 0.4583% | 1.425% | 4.763% | 1.318% | 2.3% | 0.8479% | 0% | 1.004% | 3.45% | 5.351% | 4.382%
| PDM3 |  | -3.231% | 0.6117% | -8.083% | -6.077% | -18.06% | -7.358% | -14.62% | 2.462% | -3.231% | 1.75% | 2.035% | 3.173% | -1.808% | -11.48% | 3.458% | 0.05938% | 7.585% | 1.66% | 1.039% | -0.6847% | 6.731%
| Z-Score (Absolute) |  | 0.81 | 0.06 | 1.9 | 1.45 | 4.16 | 1.74 | NA | 0.48 | 0.81 | 0.32 | 0.38 | 0.64 | 0.48 | 2.67 | 0.71 | 0.06 | 1.62 | 0.3 | 0.16 | 0.23 | 1.45
Results from laboratory 7 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Tl in OREAS 84 (Certified Value 1.24 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 19 | 20 | 23 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.18 | 1.26 | 1.3 | < 10 | 1.303 | 1.3508 | 1.05 | < 10 | 1.19 | 1.16 | 1.22 | 1.28 | < 10 | 1.18 | 1.26 | 1.2194 | 1.2408 | 1.19 | 1.35 |  |  | 1
| 1 | 2 | 1.14 | 1.29 | 1.3 | < 10 | 1.323 | 1.3048 | 1.05 | < 10 | 1.16 | 1.17 | 1.22 | 1.3 | < 10 | 1.19 | 1.29 | 1.2692 | 1.1658 | 1.16 | 1.35 |  |  | #N/A
| 1 | 3 | 1.16 | 1.27 | 1.3 | < 10 | 1.196 | 1.351 | 1.05 | < 10 | 1.14 | 1.16 | 1.22 | 1.27 | < 10 | 1.21 | 1.29 | 1.2746 | 1.2522 | 1.22 | 1.33 |  |  | 16
| 1 | 4 | 1.21 | 1.34 | 1.3 | < 10 | 1.216 | 1.323 | 1.06 | < 10 | 1.16 | 1.16 | 1.22 | 1.28 | < 10 | 1.21 | 1.28 | 1.19 | 1.323 | 1.19 | 1.33 |  |  | 1.2381
| 1 | 5 | 1.18 | 1.3 | 1.2 | < 10 | 1.16 | 1.3095 | 1.11 | < 10 | 1.13 | 1.2 | 1.21 | 1.28 | < 10 | 1.2 | 1.26 | 1.2122 | 1.321 | 1.17 | 1.3 |  |  | 66
| 1 | 6 | 1.18 | 1.27 | 1.2 | < 10 | 1.149 | 1.3025 | 1.03 | < 10 | 1.14 | 1.16 | 1.2 | 1.27 | < 10 | 1.25 | 1.27 | 1.2379 | 1.1836 | 1.2 | 1.29
| Mean |  | 1.175 | 1.2883 | 1.2667 |  | 1.2245 | 1.3236 | 1.0583 |  | 1.1533 | 1.1683 | 1.215 | 1.28 |  | 1.2067 | 1.275 | 1.2339 | 1.2478 | 1.1883 | 1.325
| Median |  | 1.18 | 1.28 | 1.3 |  | 1.206 | 1.3163 | 1.05 |  | 1.15 | 1.16 | 1.22 | 1.28 |  | 1.205 | 1.275 | 1.2287 | 1.2465 | 1.19 | 1.33
| Std Dev. |  | 0.023452 | 0.029269 | 0.05164 |  | 0.072951 | 0.02232 | 0.027142 |  | 0.021602 | 0.016021 | 0.0083666 | 0.010954 |  | 0.024221 | 0.013784 | 0.033244 | 0.066207 | 0.02137 | 0.0251
| Rel.Std.Dev. |  | 1.996% | 2.272% | 4.077% |  | 5.958% | 1.686% | 2.565% |  | 1.873% | 1.371% | 0.6886% | 0.8558% |  | 2.007% | 1.081% | 2.694% | 5.306% | 1.798% | 1.894%
| PDM3 |  | -5.096% | 4.058% | 2.308% |  | -1.098% | 6.905% | -14.52% |  | -6.846% | -5.634% | -1.865% | 3.385% |  | -2.538% | 2.981% | -0.3401% | 0.7801% | -4.019% | 7.019%
| Z-Score (Absolute) |  | 0.83 | 0.46 | 0.21 | 42.58 | 0.26 | 0.86 | 2.15 | 42.58 | 1.07 | 0.9 | 0.37 | 0.37 | 42.58 | 0.47 | 0.31 | 0.16 | 0 | 0.67 | 0.88
| Analytical results for Tm in OREAS 84 (Certified Value 0.069 ppm) | [22 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 12 | 16 | 17 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 3
3
| 1 | 1 | 0.07 | < 0.1 | 0.07 | 0.05 | 0.09 | 0.09 | 0.078422 | [14 empty] | 1
| 1 | 2 | 0.05 | < 0.1 | 0.07 | 0.04 | 0.08 | 0.08 | 0.082379 | [14 empty] | #N/A
| 1 | 3 | 0.06 | < 0.1 | 0.06 | 0.05 | 0.08 | 0.08 | 0.076375 | [14 empty] | 16
| 1 | 4 | 0.08 | < 0.1 | 0.06 | 0.04 | 0.08 | 0.08 | 0.073788 | [14 empty] | 0.068555
| 1 | 5 | 0.06 | < 0.1 | 0.06 | 0.04 | 0.08 | 0.08 | 0.074677 | [14 empty] | 67
| 1 | 6 | 0.07 | < 0.1 | 0.06 | 0.04 | 0.08 | 0.08 | 0.07234
| Mean |  | 0.065 |  | 0.063333 | 0.043333 | 0.081667 | 0.081667 | 0.07633
| Median |  | 0.065 |  | 0.06 | 0.04 | 0.08 | 0.08 | 0.075526
| Std Dev. |  | 0.010488 |  | 0.005164 | 0.005164 | 0.0040825 | 0.0040825 | 0.0036355
| Rel.Std.Dev. |  | 16.14% |  | 8.154% | 11.92% | 4.999% | 4.999% | 4.763%
| PDM3 |  | -5.186% |  | -7.617% | -36.79% | 19.13% | 19.13% | 11.34%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.07 | 0.97 | 0.75 | 0.75 | 0.51
| Analytical results for U in OREAS 84 (Certified Value 0.82 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.8 | 0.7 | 0.9 | < 10 | 0.85 | 0.80312 | 0.7 | < 10 | 0.8 | 0.8 | 0.79 | 0.88 | < 10 | 0.81 | 0.9 | 1.0609 | 0.87 | 0.77128 | 0.82 | 0.88 |  | 1
| 1 | 2 | 0.8 | 0.8 | 0.9 | < 10 | 0.83 | 0.79763 | 0.7 | < 10 | 0.8 | 0.8 | 0.78 | 0.92 | < 10 | 0.82 | 0.8 | 0.99705 | 0.86 | 0.78454 | 0.91 | 0.88 |  | #N/A
| 1 | 3 | 0.8 | 0.8 | 0.8 | < 10 | 0.77 | 0.8166 | 0.7 | < 10 | 0.8 | 0.7 | 0.77 | 0.89 | < 10 | 0.81 | 0.9 | 1.0687 | 0.86 | 0.80469 | 1.07 | 0.95 |  | 16
| 1 | 4 | 0.8 | 0.8 | 0.8 | < 10 | 0.75 | 0.82119 | 0.7 | < 10 | 0.8 | 0.8 | 0.78 | 0.87 | < 10 | 0.82 | 0.9 | 0.94929 | 0.82 | 0.82091 | 0.81 | 0.88 |  | 0.82271
| 1 | 5 | 0.8 | 0.8 | 0.8 | < 10 | 0.75 | 0.80772 | 0.7 | < 10 | 0.8 | 0.8 | 0.79 | 0.86 | < 10 | 0.8 | 0.8 | 1.0169 | 1.61 | 0.85255 | 0.85 | 0.82 |  | 68
| 1 | 6 | 0.8 | 0.8 | 0.9 | < 10 | 0.75 | 0.82057 | 0.7 | < 10 | 0.8 | 0.8 | 0.79 | 0.85 | < 10 | 0.82 | 0.8 | 1.0532 | 0.92 | 0.85136 | 0.84 | 0.86
| Mean |  | 0.8 | 0.78333 | 0.85 |  | 0.78333 | 0.81114 | 0.7 |  | 0.8 | 0.78333 | 0.78333 | 0.87833 |  | 0.81333 | 0.85 | 1.0243 | 0.99 | 0.81422 | 0.88333 | 0.87833
| Median |  | 0.8 | 0.8 | 0.85 |  | 0.76 | 0.81216 | 0.7 |  | 0.8 | 0.8 | 0.785 | 0.875 |  | 0.815 | 0.85 | 1.0351 | 0.865 | 0.8128 | 0.845 | 0.88
| Std Dev. |  | 1.2162e-16 | 0.040825 | 0.054772 |  | 0.045019 | 0.0097804 | 1.2162e-16 |  | 1.2162e-16 | 0.040825 | 0.008165 | 0.024833 |  | 0.008165 | 0.054772 | 0.04598 | 0.30542 | 0.033792 | 0.097912 | 0.042151
| Rel.Std.Dev. |  | 1.52e-14% | 5.212% | 6.444% |  | 5.747% | 1.206% | 1.737e-14% |  | 1.52e-14% | 5.212% | 1.042% | 2.827% |  | 1.004% | 6.444% | 4.489% | 30.85% | 4.15% | 11.08% | 4.799%
| PDM3 |  | -2.761% | -4.787% | 3.317% |  | -4.787% | -1.407% | -14.92% |  | -2.761% | -4.787% | -4.787% | 6.761% |  | -1.14% | 3.317% | 24.51% | 20.33% | -1.032% | 7.368% | 6.761%
| Z-Score (Absolute) |  | 0.44 | 0.67 | 0.25 | 57.62 | 0.67 | 0.29 | 1.83 | 57.62 | 0.44 | 0.67 | 0.67 | 0.64 | 57.62 | 0.26 | 0.25 | 2.66 | 2.18 | 0.25 | 0.71 | 0.64
| Analytical results for V in OREAS 84 (Certified Value 100 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 102 | 99 | 58 | 101 | 62 | 80.4481 | 69 | 100 | 105 | 100 | 119 | 113 | 101 | 60 | 102 | 101.4336 | 96.5 | 112.0428 | 96.5 | 75.57 | 92 | 1
| 1 | 2 | 101 | 100 | 58 | 101 | 59 | 81.5027 | 69 | 101 | 102 | 99 | 120 | 113 | 103 | 63 | 101 | 100.8426 | 95.4 | 106.4564 | 95.9 | 100.18 | 91 | #N/A
| 1 | 3 | 98 | 98 | 58 | 100 | 60 | 81.3614 | 69 | 99 | 96 | 101 | 120 | 111 | 103 | 76 | 103 | 102.4219 | 92.9 | 109.5972 | 99.1 | 100.36 | 97 | 16
| 1 | 4 | 102 | 99 | 58 | 101 | 60 | 81.4965 | 69 | 100 | 100 | 101 | 119 | 108 | 103 | 65 | 103 | 96.7277 | 98.3 | 111.2456 | 95.4 | 86.02 | 93 | 100.2302
| 1 | 5 | 101 | 96 | 57 | 99 | 61 | 82.0643 | 70 | 103 | 97 | 100 | 120 | 110 | 103 | 67 | 99 | 98.0583 | 95.6 | 111.247 | 95.2 | 95.86 | 97 | 69
| 1 | 6 | 99.1 | 93 | 56 | 98 | 61 | 81.1591 | 69 | 101 | 99 | 103 | 118 | 107 | 102 | 62 | 105 | 98.6731 | 94.4 | 106.1427 | 97.9 | 69.38 | 84
| Mean |  | 100.5167 | 97.5 | 57.5 | 100 | 60.5 | 81.3387 | 69.1667 | 100.6667 | 99.8333 | 100.6667 | 119.3333 | 110.3333 | 102.5 | 65.5 | 102.1667 | 99.6929 | 95.5167 | 109.4553 | 96.6667 | 87.895 | 92.3333
| Median |  | 101 | 98.5 | 58 | 100.5 | 60.5 | 81.4289 | 69 | 100.5 | 99.5 | 100.5 | 119.5 | 110.5 | 103 | 64 | 102.5 | 99.7579 | 95.5 | 110.4214 | 96.2 | 90.94 | 92.5
| Std Dev. |  | 1.6253 | 2.5884 | 0.83666 | 1.2649 | 1.0488 | 0.52995 | 0.40825 | 1.3663 | 3.3116 | 1.3663 | 0.8165 | 2.5033 | 0.83666 | 5.6833 | 2.0412 | 2.2047 | 1.8346 | 2.5728 | 1.5371 | 13.1755 | 4.8028
| Rel.Std.Dev. |  | 1.617% | 2.655% | 1.455% | 1.265% | 1.734% | 0.6515% | 0.5902% | 1.357% | 3.317% | 1.357% | 0.6842% | 2.269% | 0.8163% | 8.677% | 1.998% | 2.212% | 1.921% | 2.351% | 1.59% | 14.99% | 5.202%
| PDM3 |  | 0.2858% | -2.724% | -42.63% | -0.2297% | -39.64% | -18.85% | -30.99% | 0.4355% | -0.396% | 0.4355% | 19.06% | 10.08% | 2.265% | -34.65% | 1.932% | -0.5361% | -4.703% | 9.204% | -3.555% | -12.31% | -7.879%
| Z-Score (Absolute) |  | 0.13 | 0.35 | 6.81 | 0.05 | 6.33 | 2.96 | 4.93 | 0.16 | 0.02 | 0.16 | 3.17 | 1.72 | 0.45 | 5.52 | 0.4 | 0 | 0.67 | 1.58 | 0.49 | 1.9 | 1.19
| Analytical results for W in OREAS 84 (Certified Value 1.31 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.3 | 1.3 | 1 | < 10 | 1.35 | 1.7 | < 10 | 1.4 | 1.2 | 1.1 | 1.4 | < 10 | 1.2 | 1.4 | 1.1655 | 1.3 | 1.3737 | 1.4 | 1 |  |  | 1
| 1 | 2 | 1.2 | 1.4 | 1 | < 10 | 1.35 | 2.3 | < 10 | 1.3 | 1.2 | 1.1 | 1.4 | < 10 | 1.3 | 1.6 | 1.0827 | 1.4 | 1.3327 | 1.4 | 1 |  |  | 12
| 1 | 3 | 1.2 | 1.4 | 2 | < 10 | 1.25 | 2.2 | < 10 | 1.3 | 1.2 | 1.2 | 1.4 | < 10 | 1.2 | 1.4 | 1.1064 | 1.4 | 1.3464 | 1.5 | 1 |  |  | 16
| 1 | 4 | 1.5 | 1.4 | 1 | < 10 | 1.23 | 1.9 | < 10 | 1.3 | 1.2 | 1.2 | 1.5 | < 10 | 1.2 | 1.3 | 1.0007 | 1.4 | 1.4498 | 1.4 | 1 |  |  | 1.3059
| 1 | 5 | 1.4 | 1.4 | 1 | < 10 | 1.2 | 2 | < 10 | 1.3 | 1.4 | 1.2 | 1.4 | < 10 | 1.3 | 1.3 | 1.1493 | 1.4 | 1.3973 | 1.4 | 1 |  |  | 70
| 1 | 6 | 1.3 | 1.4 | 1 | < 10 | 1.21 | 1.8 | < 10 | 1.3 | 1.2 | 1.1 | 1.3 | < 10 | 1.3 | 1.4 | 1.1747 | 1.3 | 1.2915 | 1.4 | 1
| Mean |  | 1.3167 | 1.3833 | 1.1667 |  | 1.265 | 1.9833 |  | 1.3167 | 1.2333 | 1.15 | 1.4 |  | 1.25 | 1.4 | 1.1132 | 1.3667 | 1.3652 | 1.4167 | 1
| Median |  | 1.3 | 1.4 | 1 |  | 1.24 | 1.95 |  | 1.3 | 1.2 | 1.15 | 1.4 |  | 1.25 | 1.4 | 1.1278 | 1.4 | 1.36 | 1.4 | 1
| Std Dev. |  | 0.1169 | 0.040825 | 0.40825 |  | 0.068044 | 0.23166 |  | 0.040825 | 0.08165 | 0.054772 | 0.063246 |  | 0.054772 | 0.10954 | 0.065501 | 0.05164 | 0.054959 | 0.040825 | 0
| Rel.Std.Dev. |  | 8.879% | 2.951% | 34.99% |  | 5.379% | 11.68% |  | 3.101% | 6.62% | 4.763% | 4.518% |  | 4.382% | 7.825% | 5.884% | 3.779% | 4.026% | 2.882% | 0%
| PDM3 |  | 0.824% | 5.929% | -10.66% |  | -3.132% | 51.87% |  | 0.824% | -5.557% | -11.94% | 7.205% |  | -4.281% | 7.205% | -14.76% | 4.653% | 4.542% | 8.482% | -23.42%
| Z-Score (Absolute) |  | 0.33 | 0.11 | NA | 24.1 | 0.67 | 4.09 | 24.1 | 0.33 | 0.88 | 1.44 | 0.22 | 24.1 | 0.77 | 0.22 | 1.68 | 0 | 0.01 | 0.33 | NA
Results from laboratories 3 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 84 (Certified Value 4.71 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 4.9 | 5.3 | 4.9 | 4.81 | 4.8 | 4.7 | 5.1 | 4.5 | 4.2 | 4.7 | 5.1 | 3.8899 | 4.8 | 4.8427 | 4.6 | 5.5 | [5 empty] | 1
| 1 | 2 | 4.48 | 5.4 | 4.8 | 4.76 | 4.6 | 4.6 | 5.3 | 4.4 | 4 | 4.7 | 5.1 | 3.9265 | 4.8 | 4.4998 | 4.6 | 5.2 | [5 empty] | 18
| 1 | 3 | 4.87 | 5.5 | 5.1 | 4.46 | 4.4 | 4.4 | 5.3 | 4.4 | 3.9 | 4.6 | 4.9 | 3.8992 | 4.9 | 4.7259 | 4.9 | 5.4 | [5 empty] | 16
| 1 | 4 | 4.88 | 5.7 | 5 | 4.51 | 4.6 | 4.5 | 5.2 | 4.3 | 3.7 | 4.8 | 5.1 | 3.7685 | 4.8 | 4.5511 | 4.7 | 5.4 | [5 empty] | 4.7093
| 1 | 5 | 4.89 | 5.4 | 4.7 | 4.62 | 4.5 | 4.3 | 5.1 | 4.3 | 4.1 | 4.7 | 5.1 | 3.8554 | 4.9 | 4.8064 | 4.7 | 5.1 | [5 empty] | 71
| 1 | 6 | 4.49 | 5.4 | 5.1 | 4.57 | 4.4 | 4.3 | 5.1 | 4.4 | 3.6 | 4.8 | 5 | 3.8748 | 4.8 | 4.5141 | 4.7 | 5.3
| Mean |  | 4.7517 | 5.45 | 4.9333 | 4.6217 | 4.55 | 4.4667 | 5.1833 | 4.3833 | 3.9167 | 4.7167 | 5.05 | 3.8691 | 4.8333 | 4.6567 | 4.7 | 5.3167
| Median |  | 4.875 | 5.4 | 4.95 | 4.595 | 4.55 | 4.45 | 5.15 | 4.4 | 3.95 | 4.7 | 5.1 | 3.8823 | 4.8 | 4.6385 | 4.7 | 5.35
| Std Dev. |  | 0.20683 | 0.13784 | 0.1633 | 0.13848 | 0.15166 | 0.1633 | 0.098319 | 0.075277 | 0.23166 | 0.075277 | 0.083666 | 0.054726 | 0.05164 | 0.15357 | 0.10954 | 0.1472
| Rel.Std.Dev. |  | 4.353% | 2.529% | 3.31% | 2.996% | 3.333% | 3.656% | 1.897% | 1.717% | 5.915% | 1.596% | 1.657% | 1.414% | 1.068% | 3.298% | 2.331% | 2.769%
| PDM3 |  | 0.8993% | 15.73% | 4.757% | -1.861% | -3.383% | -5.153% | 10.07% | -6.922% | -16.83% | 0.1561% | 7.234% | -17.84% | 2.633% | -1.118% | -0.1978% | 12.9%
| Z-Score (Absolute) |  | 0.13 | 2.14 | 0.65 | 0.25 | 0.46 | 0.7 | 1.37 | 0.94 | 2.29 | 0.02 | 0.99 | 2.43 | 0.36 | 0.15 | 0.02 | 1.76
| Analytical results for Yb in OREAS 84 (Certified Value 0.47 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 23 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.48 | 0.5 | 0.5 | 0.46079 | 0.47 | 0.5 | 0.5 | 0.5 | 0.45 | 0.51202 | 0.46 | 0.5 | [9 empty] | 1
| 1 | 2 | 0.46 | 0.4 | 0.51 | 0.45268 | 0.44 | 0.5 | 0.5 | 0.5 | 0.47 | 0.47417 | 0.46 | 0.5 | [9 empty] | #N/A
| 1 | 3 | 0.47 | 0.5 | 0.46 | 0.4544 | 0.46 | 0.4 | 0.5 | 0.5 | 0.49 | 0.48726 | 0.48 | 0.5 | [9 empty] | 16
| 1 | 4 | 0.43 | 0.5 | 0.47 | 0.44757 | 0.44 | 0.4 | 0.5 | 0.5 | 0.47 | 0.4581 | 0.46 | 0.5 | [9 empty] | 0.47047
| 1 | 5 | 0.42 | 0.4 | 0.45 | 0.43782 | 0.48 | 0.4 | 0.5 | 0.5 | 0.46 | 0.46345 | 0.46 | 0.5 | [9 empty] | 72
| 1 | 6 | 0.49 | 0.5 | 0.44 | 0.44908 | 0.43 | 0.4 | 0.5 | 0.5 | 0.48 | 0.48665 | 0.45 | 0.5
| Mean |  | 0.45833 | 0.46667 | 0.47167 | 0.45039 | 0.45333 | 0.43333 | 0.5 | 0.5 | 0.47 | 0.48028 | 0.46167 | 0.5
| Median |  | 0.465 | 0.5 | 0.465 | 0.45088 | 0.45 | 0.4 | 0.5 | 0.5 | 0.47 | 0.48041 | 0.46 | 0.5
| Std Dev. |  | 0.027869 | 0.05164 | 0.027869 | 0.0077098 | 0.019664 | 0.05164 | 0 | 0 | 0.014142 | 0.019541 | 0.0098319 | 0
| Rel.Std.Dev. |  | 6.08% | 11.07% | 5.909% | 1.712% | 4.338% | 11.92% | 0% | 0% | 3.009% | 4.069% | 2.13% | 0%
| PDM3 |  | -2.58% | -0.8089% | 0.2539% | -4.269% | -3.643% | -7.894% | 6.276% | 6.276% | -0.1003% | 2.084% | -1.872% | 6.276%
| Z-Score (Absolute) |  | 0.5 | 0.08 | 0.17 | 0.9 | 0.75 | 1.75 | 1.59 | 1.59 | 0.08 | 0.6 | 0.33 | 1.59
| Analytical results for Zn in OREAS 84 (Certified Value 814 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 770 | 820 | 868 | 825 | 702 | 874.3706 | 787 | 807 | 776 | 808 | 849 | 821 | 834 | 687 | 809 | 802.8252 | 790 | 872.576 | 627 | 898 |  | 1
| 1 | 2 | 776 | 821 | 865 | 818 | 703 | 875.2185 | 788 | 823 | 758 | 808 | 855 | 828 | 842 | 694 | 812 | 798.7895 | 788 | 896.3662 | 597 | 879 |  | #N/A
| 1 | 3 | 807 | 805 | 855 | 810 | 690 | 876.3349 | 782 | 803 | 736 | 804 | 838 | 806 | 862 | 707 | 807 | 798.0628 | 779 | 854.5606 | 636 | 876 |  | 16
| 1 | 4 | 795 | 810 | 859 | 826 | 687 | 873.4386 | 799 | 812 | 762 | 807 | 841 | 802 | 851 | 699 | 798 | 770.4005 | 789 | 885.1562 | 620 | 889 |  | 813.8033
| 1 | 5 | 784 | 810 | 837 | 806 | 694 | 875.3704 | 793 | 840 | 736 | 806 | 844 | 815 | 853 | 704 | 806 | 774.6245 | 778 | 885.4704 | 622 | 876 |  | 73
| 1 | 6 | 786 | 782 | 841 | 801 | 690 | 874.0597 | 795 | 815 | 753 | 830 | 841 | 809 | 836 | 716 | 803 | 769.8983 | 789 | 845.2369 | 636 | 883
| Mean |  | 786.3333 | 808 | 854.1667 | 814.3333 | 694.3333 | 874.7988 | 790.6667 | 816.6667 | 753.5 | 810.5 | 844.6667 | 813.5 | 846.3333 | 701.1667 | 805.8333 | 785.7668 | 785.5 | 873.2277 | 623 | 883.5
| Median |  | 785 | 810 | 857 | 814 | 692 | 874.7946 | 790.5 | 813.5 | 755.5 | 807.5 | 842.5 | 812 | 846.5 | 701.5 | 806.5 | 786.3437 | 788.5 | 878.8661 | 624.5 | 881
| Std Dev. |  | 13.2765 | 14.1845 | 12.6557 | 10.2892 | 6.7132 | 1.0433 | 6.1536 | 13.3367 | 15.566 | 9.6695 | 6.2823 | 9.7724 | 10.8566 | 10.1866 | 4.8751 | 15.6452 | 5.4681 | 19.7987 | 14.4222 | 8.6429
| Rel.Std.Dev. |  | 1.688% | 1.756% | 1.482% | 1.264% | 0.9669% | 0.1193% | 0.7783% | 1.633% | 2.066% | 1.193% | 0.7438% | 1.201% | 1.283% | 1.453% | 0.605% | 1.991% | 0.6961% | 2.267% | 2.315% | 0.9783%
| PDM3 |  | -3.376% | -0.7131% | 4.96% | 0.06513% | -14.68% | 7.495% | -2.843% | 0.3518% | -7.41% | -0.4059% | 3.792% | -0.03727% | 3.997% | -13.84% | -0.9794% | -3.445% | -3.478% | 7.302% | -23.45% | 8.564%
| Z-Score (Absolute) |  | 0.52 | 0.03 | 1.02 | 0.12 | 2.62 | 1.49 | 0.42 | 0.17 | 1.27 | 0.03 | 0.81 | 0.1 | 0.85 | 2.46 | 0.08 | 0.54 | 0.54 | 1.46 | 4.25 | 1.69
| Analytical results for Zr in OREAS 84 (Certified Value 28.6 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 28.7 | 30.7 | 25 | 30.2 | 23.2645 | 31 | 28.6 | 28.7 | 26.7 | 29 | 27.2 | 31.7 | 27.9497 | 39 | 29.4345 | 26.2 | 30.1 |  |  |  |  | 1
| 1 | 2 | 28.6 | 32.6 | 26 | 27.9 | 23.3859 | 31 | 27.8 | 28.3 | 27 | 29 | 26.8 | 33 | 28.4814 | 39 | 31.5273 | 25.2 | 29.4 |  |  |  |  | 19
| 1 | 3 | 42.1 | 31.7 | 26 | 26.8 | 23.5839 | 30 | 26.9 | 27.9 | 27 | 30 | 27.6 | 34 | 27.7369 | 39 | 27.0705 | 27.1 | 29.7 |  |  |  |  | 16
| 1 | 4 | 35.1 | 34.6 | 24 | 29.8 | 23.2496 | 31 | 28.9 | 30.1 | 26.7 | 31 | 27.5 | 31.3 | 27.1226 | 38 | 30.0795 | 26.5 | 30.5 |  |  |  |  | 28.5666
| 1 | 5 | 32.8 | 31.7 | 24 | 29.9 | 23.5066 | 30 | 26.7 | 30.6 | 26.7 | 28 | 27.5 | 31.6 | 27.2506 | 40 | 28.8377 | 27.7 | 29 |  |  |  |  | 74
| 1 | 6 | 40.7 | 32.4 | 25 | 27.6 | 23.5835 | 31 | 27 | 29.6 | 26.7 | 29 | 28.8 | 31.9 | 28.4256 | 37 | 27.0069 | 26.5 | 28.5
| Mean |  | 34.6667 | 32.2833 | 25 | 28.7 | 23.429 | 30.6667 | 27.65 | 29.2 | 26.8 | 29.3333 | 27.5667 | 32.25 | 27.8278 | 38.6667 | 28.9927 | 26.5333 | 29.5333
| Median |  | 33.95 | 32.05 | 25 | 28.85 | 23.4463 | 31 | 27.4 | 29.15 | 26.7 | 29 | 27.5 | 31.8 | 27.8433 | 39 | 29.1361 | 26.5 | 29.55
| Std Dev. |  | 5.7919 | 1.3167 | 0.89443 | 1.4394 | 0.15165 | 0.5164 | 0.93541 | 1.0658 | 0.15492 | 1.0328 | 0.67132 | 1.0368 | 0.57239 | 1.0328 | 1.7585 | 0.84538 | 0.72847
| Rel.Std.Dev. |  | 16.71% | 4.079% | 3.578% | 5.015% | 0.6473% | 1.684% | 3.383% | 3.65% | 0.5781% | 3.521% | 2.435% | 3.215% | 2.057% | 2.671% | 6.065% | 3.186% | 2.467%
| PDM3 |  | 21.35% | 13.01% | -12.49% | 0.4668% | -17.98% | 7.351% | -3.209% | 2.217% | -6.184% | 2.684% | -3.5% | 12.89% | -2.586% | 35.36% | 1.492% | -7.118% | 3.384%
| Z-Score (Absolute) |  | 2.29 | 1.33 | 1.61 | 0.12 | 2.24 | 0.67 | 0.54 | 0.08 | 0.88 | 0.14 | 0.57 | 1.31 | 0.47 | 3.9 | 0 | 0.99 | 0.22
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1077: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1095: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 84 (Indicative Value 1.63 ppm) | [22 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 09 | 12 | 15 | 16 | 18 | 20 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*OES/MS | [14 empty] | 2
2
| 1 | 1 | < 10 | < 1 | < 10 | 1 | 2.37 | 0.7 | 1.8 | [14 empty] | 1
| 1 | 2 | < 10 | < 1 | < 10 | 2 | 2.37 | 0.5 | 1.7 | [14 empty] | 10
| 1 | 3 | < 10 | < 1 | < 10 | 1 | 2.3 | 0.5 | 2 | [14 empty] | 16
| 1 | 4 | < 10 | < 1 | < 10 | 2 | 2.33 | 0.8 | 1.9 | [14 empty] | 1.6342
| 1 | 5 | < 10 | < 1 | < 10 | 2 | 2.4 | 0.8 | 2 | [14 empty] | 16
| 1 | 6 | < 10 | < 1 | < 10 | 2 | 2.35 | 0.6 | 1.8
| Mean |  |  |  |  | 1.6667 | 2.3533 | 0.65 | 1.8667
| Median |  |  |  |  | 2 | 2.36 | 0.65 | 1.85
| Std Dev. |  |  |  |  | 0.5164 | 0.035024 | 0.13784 | 0.12111
| Rel.Std.Dev. |  |  |  |  | 30.98% | 1.488% | 21.21% | 6.488%
| PDM3 |  |  |  |  | 1.989% | 44.01% | -60.22% | 14.23%
| Z-Score (Absolute) |  | 1.74 | 0.76 | 1.74 | 0.11 | 0.27 | 0.67 | 0
| Analytical results for Al2O3 in OREAS 84 (Certified Value 4.17 wt.%) | [22 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 3.949 | 3.93 | 4.232 | 4.35 | 4.1235 | 4.0316 | 4.062 | 4.27 | 3.7756 | 4.2 | 3.949 | 4.21 | 4.157 | 4.09 | 4.2219 | 3.36 | 4.4003 | 4.195 | 4.176 | 4.1078 |  | 1
| 1 | 2 | 4.214 | 4.04 | 4.138 | 4.33 | 4.1121 | 4.1197 | 4.025 | 4.25 | 3.6739 | 4.15 | 3.968 | 4.18 | 4.251 | 4.1 | 4.1139 | 3.47 | 4.45 | 4.176 | 4.157 | 4.3251 |  | #N/A
| 1 | 3 | 4.214 | 3.82 | 4.27 | 4.36 | 4.1692 | 4.0249 | 3.949 | 4.21 | 3.645 | 4.41 | 3.779 | 4.35 | 4.195 | 4.03 | 4.0776 | 3.51 | 4.3486 | 4.232 | 4.157 | 4.3213 |  | 16
| 1 | 4 | 3.968 | 3.85 | 4.251 | 4.4 | 4.1739 | 4.0362 | 3.836 | 4.25 | 3.6726 | 4.12 | 3.628 | 4.36 | 4.27 | 4.09 | 4.0325 | 3.57 | 4.3461 | 4.176 | 4.138 | 4.0227 |  | 4.1663
| 1 | 5 | 4.157 | 3.93 | 4.195 | 4.31 | 4.1499 | 3.9919 | 4.119 | 4.27 | 3.5711 | 4.14 | 3.514 | 4.42 | 4.138 | 4.08 | 4.0814 | 3.52 | 4.4107 | 4.27 | 4.138 | 4.2438 |  | 76
| 1 | 6 | 4.195 | 3.97 | 4.27 | 4.35 | 4.1669 | 4.1073 | 3.873 | 4.23 | 3.79 | 4.07 | 3.496 | 4.27 | 4.214 | 3.98 | 4.1526 | 3.62 | 4.4246 | 4.138 | 4.195 | 4.019
| Mean |  | 4.1162 | 3.9233 | 4.226 | 4.35 | 4.1492 | 4.0519 | 3.9773 | 4.2467 | 3.6881 | 4.1817 | 3.7223 | 4.2983 | 4.2042 | 4.0617 | 4.1133 | 3.5083 | 4.3967 | 4.1978 | 4.1602 | 4.1733
| Median |  | 4.176 | 3.93 | 4.2415 | 4.35 | 4.1584 | 4.0339 | 3.987 | 4.25 | 3.6733 | 4.145 | 3.7035 | 4.31 | 4.2045 | 4.085 | 4.0977 | 3.515 | 4.4055 | 4.1855 | 4.157 | 4.1758
| Std Dev. |  | 0.12403 | 0.079917 | 0.051455 | 0.030332 | 0.025928 | 0.0503 | 0.11061 | 0.023381 | 0.082467 | 0.11957 | 0.20903 | 0.093684 | 0.051596 | 0.047081 | 0.066574 | 0.089312 | 0.041727 | 0.046727 | 0.022212 | 0.14191
| Rel.Std.Dev. |  | 3.013% | 2.037% | 1.218% | 0.6973% | 0.6249% | 1.241% | 2.781% | 0.5506% | 2.236% | 2.859% | 5.616% | 2.18% | 1.227% | 1.159% | 1.618% | 2.546% | 0.949% | 1.113% | 0.5339% | 3.401%
| PDM3 |  | -1.204% | -5.833% | 1.432% | 4.408% | -0.4102% | -2.746% | -4.536% | 1.928% | -11.48% | 0.3679% | -10.66% | 3.168% | 0.908% | -2.512% | -1.272% | -15.79% | 5.53% | 0.756% | -0.1481% | 0.1667%
| Z-Score (Absolute) |  | 0.28 | 1.69 | 0.52 | 1.42 | 0.04 | 0.75 | 1.29 | 0.67 | 3.4 | 0.2 | 3.15 | 1.05 | 0.36 | 0.68 | 0.3 | 4.71 | 1.76 | 0.31 | 0.04 | 0.14
| Analytical results for As in OREAS 84 (Certified Value 175 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 18 | 19 | 20 | 23 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [5 empty] | 0
0
| 1 | 1 | 164 | 200 | 188 | 200 | 181 | 200 | 167 | 200 | 178 | 300 | 181 | 160 | 13.1 | 183.9714 | 171 | 148 | [5 empty] | 1
| 1 | 2 | 175 | 300 | 189 | 200 | 166 | 200 | 170 | 100 | 182 | 300 | 182 | 161.9 | 10.8 | 179.9227 | 162 | 181 | [5 empty] | 16
| 1 | 3 | 178 | 100 | 192 | 200 | 195 | 300 | 163 | 100 | 183 | 300 | 180 | 158 | 13.3 | 182.497 | 168 | 213 | [5 empty] | 16
| 1 | 4 | 163 | 100 | 195 | 200 | 168 | 200 | 169 | 100 | 179 | 300 | 187 | 159.5 | 13.4 | 179.2715 | 158 | 172 | [5 empty] | 175.4004
| 1 | 5 | 173 | 100 | 187 | 200 | 166 | 200 | 164 | 100 | 180 | 300 | 179 | 157.7 | 14.7 | 185.0643 | 171 | 209 | [5 empty] | 77
| 1 | 6 | 176 | 200 | 191 | 200 | 161 | 200 | 165 | 200 | 185 | 300 | 177 | 159.2 | 13.9 | 179.9953 | 165 | 180
| Mean |  | 171.5 | 166.6667 | 190.3333 | 200 | 172.8333 | 216.6667 | 166.3333 | 133.3333 | 181.1667 | 300 | 181 | 159.3833 | 13.2 | 181.787 | 165.8333 | 183.8333
| Median |  | 174 | 150 | 190 | 200 | 167 | 200 | 166 | 100 | 181 | 300 | 180.5 | 159.35 | 13.35 | 181.2462 | 166.5 | 180.5
| Std Dev. |  | 6.4109 | 81.6497 | 2.9439 | 0 | 12.7658 | 40.8248 | 2.8048 | 51.6398 | 2.6394 | 0 | 3.4059 | 1.5171 | 1.3084 | 2.4096 | 5.1929 | 24.2109
| Rel.Std.Dev. |  | 3.738% | 48.99% | 1.547% | 0% | 7.386% | 18.84% | 1.686% | 38.73% | 1.457% | 0% | 1.882% | 0.9519% | 9.912% | 1.325% | 3.131% | 13.17%
| PDM3 |  | -2.224% | -4.979% | 8.514% | 14.02% | -1.464% | 23.53% | -5.169% | -23.98% | 3.288% | 71.04% | 3.192% | -9.132% | -92.47% | 3.641% | -5.454% | 4.808%
| Z-Score (Absolute) |  | 0.11 | NA | 1.42 | NA | 0 | NA | 0.53 | NA | 0.67 | NA | 0.66 | 1.09 | 12.92 | 0.72 | 0.57 | 0.89
Results from laboratories 2, 4, 8, 10 and 13 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 84 (Indicative Value 24.8 ppm) | [22 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 03 | 05 | 09 | 12 | 19 | 20 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | [14 empty] | 1
1
| 1 | 1 | 20 | < 50 | 123 | < 1000 | 25 | < 50 | < 10 | [14 empty] | 1
| 1 | 2 | 20 | < 50 | 112 | < 1000 | 23 | < 50 | < 10 | [14 empty] | 8
| 1 | 3 | 20 | 58 | 129 | < 1000 | 26 | < 50 | < 10 | [14 empty] | 16
| 1 | 4 | 20 | 59 | 154 | < 1000 | 27 | < 50 | 10 | [14 empty] | 24.75
| 1 | 5 | 20 | 65 | 251 | < 1000 | 27 | < 50 | < 10 | [14 empty] | 17
| 1 | 6 | 20 | 59 | 77 | < 1000 | 25 | < 50 | < 10
| Mean |  | 20 | 60.25 | 141 |  | 25.5 |  | 10
| Median |  | 20 | 59 | 126 |  | 25.5 |  | 10
| Std Dev. |  | 0 | 3.2016 | 59.4542 |  | 1.5166
| Rel.Std.Dev. |  | 0% | 5.314% | 42.17% |  | 5.947%
| PDM3 |  | -19.19% | 143.4% | 469.7% |  | 3.03% |  | -59.6%
| Z-Score (Absolute) |  | 0.19 | 0.79 | 3.96 | 16.27 | 0 | 0.02 | 0.67
| Analytical results for Ba in OREAS 84 (Certified Value 175 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [8 empty] | 0
0
| 1 | 1 | 159 | 174 | 180 | 190 | 156 | 186 | 172 | 168 | 177.7846 | 187 | 173.1545 | 167 | 188 | [8 empty] | 1
| 1 | 2 | 173 | 177 | 180 | 181 | 158 | 187 | 175 | 164 | 173.5265 | 193 | 171.0289 | 167 | 199 | [8 empty] | #N/A
| 1 | 3 | 175 | 177 | 179 | 166 | 153 | 187 | 174 | 162 | 175.5774 | 191 | 161.5998 | 169 | 184 | [8 empty] | 16
| 1 | 4 | 158 | 173 | 182 | 178 | 155 | 187 | 176 | 167 | 173.9787 | 190 | 161.6567 | 167 | 177 | [8 empty] | 175.3227
| 1 | 5 | 167 | 174 | 180 | 202 | 155 | 189 | 175 | 167 | 175.2887 | 197 | 175.6892 | 171 | 193 | [8 empty] | 78
| 1 | 6 | 173 | 169 | 179 | 190 | 155 | 181 | 174 | 164 | 175.7045 | 193 | 173.1833 | 166 | 185
| Mean |  | 167.5 | 174 | 180 | 184.5 | 155.3333 | 186.1667 | 174.3333 | 165.3333 | 175.3101 | 191.8333 | 169.3854 | 167.8333 | 187.6667
| Median |  | 170 | 174 | 180 | 185.5 | 155 | 187 | 174.5 | 165.5 | 175.4331 | 192 | 172.0917 | 167 | 186.5
| Std Dev. |  | 7.4766 | 2.9665 | 1.0954 | 12.3572 | 1.633 | 2.7142 | 1.3663 | 2.3381 | 1.5036 | 3.3714 | 6.1874 | 1.8348 | 7.6333
| Rel.Std.Dev. |  | 4.464% | 1.705% | 0.6086% | 6.698% | 1.051% | 1.458% | 0.7837% | 1.414% | 0.8577% | 1.757% | 3.653% | 1.093% | 4.067%
| PDM3 |  | -4.462% | -0.7545% | 2.668% | 5.235% | -11.4% | 6.185% | -0.5643% | -5.698% | -0.007222% | 9.417% | -3.387% | -4.272% | 7.041%
| Z-Score (Absolute) |  | 0.67 | 0.03 | 0.56 | 1 | 1.87 | 1.17 | 0 | 0.89 | 0.1 | 1.73 | 0.49 | 0.64 | 1.32
| Analytical results for Be in OREAS 84 (Indicative Value 1.08 ppm) | [22 empty] | Indicative
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 3 | < 1 | 9.5 | 0.56535 | < 5 | < 5 | < 5 | < 5 | 0.43 | 2.9 | < 1 | < 20 | [9 empty] | 1
| 1 | 2 | < 3 | < 1 | 12.3 | 0.58425 | < 5 | < 5 | < 5 | < 5 | 0.45 | 1.4 | < 1 | < 20 | [9 empty] | 12
| 1 | 3 | < 3 | < 1 | 9 | 0.56552 | < 5 | < 5 | < 5 | < 5 | 0.48 | 2.6 | < 1 | < 20 | [9 empty] | 16
| 1 | 4 | < 3 | < 1 | 13 | 0.57088 | < 5 | < 5 | < 5 | < 5 | 0.43 | 2 | < 1 | < 20 | [9 empty] | 1.0784
| 1 | 5 | < 3 | < 1 | 13 | 0.56185 | < 5 | < 5 | < 5 | < 5 | 0.4 | 3.1 | < 1 | < 20 | [9 empty] | 18
| 1 | 6 | < 3 | < 1 | 12.5 | 0.58278 | < 5 | < 5 | < 5 | < 5 | 0.39 | 1.4 | < 1 | < 20
| Mean |  |  |  | 11.55 | 0.57177 |  |  |  |  | 0.43 | 2.2333
| Median |  |  |  | 12.4 | 0.5682 |  |  |  |  | 0.43 | 2.3
| Std Dev. |  |  |  | 1.8097 | 0.0095536 |  |  |  |  | 0.032863 | 0.74476
| Rel.Std.Dev. |  |  |  | 15.67% | 1.671% |  |  |  |  | 7.643% | 33.35%
| PDM3 |  |  |  | 971.1% | -46.98% |  |  |  |  | -60.12% | 107.1%
| Z-Score (Absolute) |  | 0.44 | 0.95 | 4.65 | 0.91 | 0.07 | 0.07 | 0.07 | 0.07 | 0.98 | 0.07 | 0.95 | 3.87
| Analytical results for Bi in OREAS 84 (Certified Value 1.39 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 09 | 12 | 15 | 16 | 18 | 19 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 2 | 1.7 | 1.4094 | < 5 | < 0.5 | 1.5 | 1.5 | 1.29 | 1.1 | 1.3664 | 1.5 | [10 empty] | 1
| 1 | 2 | < 2 | 1.6 | 1.3874 | < 5 | 1.2 | 1.5 | 1.5 | 1.34 | < 0.1 | 1.2193 | 1.6 | [10 empty] | 18
| 1 | 3 | < 2 | 1.7 | 1.4303 | < 5 | 1.4 | 1.5 | 1.5 | 1.3 | 1 | 1.1682 | 1.5 | [10 empty] | 16
| 1 | 4 | < 2 | 1.7 | 1.422 | < 5 | < 0.5 | 1.6 | 1.5 | 1.31 | 1.1 | 1.3047 | 1.4 | [10 empty] | 1.3902
| 1 | 5 | < 2 | 1.6 | 1.44 | < 5 | < 0.5 | 1.5 | 1.5 | 1.29 | 1.7 | 1.3729 | 1.6 | [10 empty] | 79
| 1 | 6 | < 2 | 1.5 | 1.3073 | < 5 | 1 | 1.5 | 1.5 | 1.31 | 0.9 | 1.2422 | 1.5
| Mean |  |  | 1.6333 | 1.3994 |  | 1.2 | 1.5167 | 1.5 | 1.3067 | 1.16 | 1.2789 | 1.5167
| Median |  |  | 1.65 | 1.4157 |  | 1.2 | 1.5 | 1.5 | 1.305 | 1.1 | 1.2735 | 1.5
| Std Dev. |  |  | 0.08165 | 0.048677 |  | 0.2 | 0.040825 | 0 | 0.018619 | 0.31305 | 0.082833 | 0.075277
| Rel.Std.Dev. |  |  | 4.999% | 3.478% |  | 16.67% | 2.692% | 0% | 1.425% | 26.99% | 6.477% | 4.963%
| PDM3 |  |  | 17.49% | 0.6623% |  | -13.68% | 9.098% | 7.899% | -6.008% | -16.56% | -8.003% | 9.098%
| Z-Score (Absolute) |  | 2.24 | 1.31 | 0 | 6.16 | 3.78 | 0.66 | 0.56 | 0.52 | 2.38 | 0.67 | 0.66
| Analytical results for CaO in OREAS 84 (Certified Value 2.93 wt.%) | [22 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 2.798 | 2.71 | 2.94 | 2.9 | 2.863 | 2.4841 | 3.05 | 2.95 | 2.7725 | 2.92 | 2.94 | 2.86 | 2.94 | 2.91 | 2.9976 | 2.88 | 3.0295 | 2.868 | 2.868 | 2.8152 |  | 1
| 1 | 2 | 3.008 | 2.69 | 2.94 | 2.83 | 2.8465 | 2.4668 | 3.008 | 2.97 | 2.7759 | 2.82 | 2.94 | 2.93 | 2.94 | 2.924 | 2.9794 | 3.02 | 2.8894 | 2.868 | 2.952 | 3.067 |  | #N/A
| 1 | 3 | 2.952 | 2.57 | 3.08 | 2.9 | 2.8556 | 2.4731 | 2.952 | 2.85 | 2.6299 | 3.12 | 2.94 | 2.98 | 2.94 | 2.882 | 2.9367 | 2.97 | 2.8688 | 2.896 | 2.966 | 2.8977 |  | 16
| 1 | 4 | 2.812 | 2.57 | 2.94 | 2.91 | 2.8794 | 2.5057 | 2.896 | 2.97 | 2.6911 | 2.8 | 2.94 | 3.02 | 2.94 | 2.91 | 2.9496 | 3.01 | 2.9933 | 2.854 | 2.924 | 2.748 |  | 2.9309
| 1 | 5 | 2.924 | 2.71 | 2.94 | 2.88 | 2.8576 | 2.4712 | 3.134 | 2.95 | 2.7087 | 2.82 | 2.94 | 3.13 | 2.8 | 2.91 | 2.9033 | 2.98 | 3.0328 | 2.896 | 2.952 | 2.8865 |  | 80
| 1 | 6 | 2.952 | 2.76 | 2.94 | 2.95 | 2.8467 | 2.4811 | 2.924 | 2.97 | 2.7097 | 2.84 | 2.94 | 3.02 | 2.94 | 2.868 | 2.9127 | 3.03 | 3.023 | 2.84 | 2.98 | 3.1118
| Mean |  | 2.9077 | 2.6683 | 2.9633 | 2.895 | 2.8581 | 2.4804 | 2.994 | 2.9433 | 2.7146 | 2.8867 | 2.94 | 2.99 | 2.9167 | 2.9007 | 2.9466 | 2.9817 | 2.9728 | 2.8703 | 2.9403 | 2.921
| Median |  | 2.938 | 2.7 | 2.94 | 2.9 | 2.8566 | 2.4771 | 2.98 | 2.96 | 2.7092 | 2.83 | 2.94 | 3 | 2.94 | 2.91 | 2.9432 | 2.995 | 3.0082 | 2.868 | 2.952 | 2.8921
| Std Dev. |  | 0.084194 | 0.079603 | 0.057155 | 0.03937 | 0.012242 | 0.013955 | 0.088544 | 0.046762 | 0.054565 | 0.12176 | 0 | 0.091652 | 0.057155 | 0.021078 | 0.0369 | 0.054924 | 0.074201 | 0.022429 | 0.040008 | 0.14183
| Rel.Std.Dev. |  | 2.896% | 2.983% | 1.929% | 1.36% | 0.4283% | 0.5626% | 2.957% | 1.589% | 2.01% | 4.218% | 0% | 3.065% | 1.96% | 0.7266% | 1.252% | 1.842% | 2.496% | 0.7814% | 1.361% | 4.855%
| PDM3 |  | -0.7918% | -8.958% | 1.107% | -1.224% | -2.481% | -15.37% | 2.154% | 0.4251% | -7.378% | -1.508% | 0.3114% | 2.017% | -0.4847% | -1.031% | 0.535% | 1.733% | 1.431% | -2.066% | 0.3227% | -0.3358%
| Z-Score (Absolute) |  | 0.18 | 4.49 | 0.78 | 0.41 | 1.08 | 7.87 | 1.37 | 0.46 | 3.66 | 0.56 | 0.36 | 1.29 | 0.05 | 0.31 | 0.51 | 1.14 | 0.99 | 0.85 | 0.41 | 0.05
| Analytical results for Cd in OREAS 84 (Certified Value < 10 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 07 | 09 | 12 | 15 | 16 | 18 | 19 | 20 | [11 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [11 empty] | 1
1
| 1 | 1 | < 2 | < 10 | 1.9 | < 10 | 2.7 | 2.4 | 2.19 | < 0.1 | < 10 | < 5 | [11 empty] | 1
| 1 | 2 | 2 | < 10 | 1.4 | < 10 | 2.6 | 2.4 | 2.18 | < 0.1 | < 10 | < 5 | [11 empty] | #N/A
| 1 | 3 | < 2 | < 10 | 1.5 | < 10 | 2.8 | 2.4 | 2.18 | < 0.1 | < 10 | < 5 | [11 empty] | 16
| 1 | 4 | < 2 | < 10 | 1.3 | < 10 | 2.6 | 2.6 | 2.16 | < 0.1 | < 10 | < 5 | [11 empty] | < 10
| 1 | 5 | 2 | < 10 | 2 | < 10 | 2.6 | 2.4 | 2.09 | < 0.1 | < 10 | < 5 | [11 empty] | 81
| 1 | 6 | 2 | < 10 | 2.1 | < 10 | 2.6 | 2.5 | 2.1 | 0.1 | < 10 | < 5
| Mean |  | 2 |  | 1.7 |  | 2.65 | 2.45 | 2.15 | 0.1
| Median |  | 2 |  | 1.7 |  | 2.6 | 2.4 | 2.17 | 0.1
| Std Dev. |  | 0 |  | 0.34059 |  | 0.083666 | 0.083666 | 0.043818
| Rel.Std.Dev. |  | 0% |  | 20.03% |  | 3.157% | 3.415% | 2.038%
PDM3
| Z-Score (Absolute) |  | 0.75 | 1.95 | 0.6 | 1.95 | 0.13 | 0.02 | 0.25 | 1.86 | 1.95 | 0.02
| Analytical results for Ce in OREAS 84 (Certified Value 15.8 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 1
1
| 1 | 1 | 14.1 | 10 | 16.0037 | 17.9 | 14.1 | 16.9 | 16.2 | 16.3 | 16.21 | 17.8368 | 18 | 15.8 | [9 empty] | 1
| 1 | 2 | 15.2 | 16 | 15.8669 | 17.2 | 14 | 16.6 | 21.8 | 16.2 | 15.92 | 17.0767 | 20 | 16.3 | [9 empty] | 19
| 1 | 3 | 15.3 | 17 | 15.9642 | 17.3 | 14.1 | 17 | 15.9 | 15.5 | 15.77 | 18.373 | 23 | 16 | [9 empty] | 16
| 1 | 4 | 14.5 | 14 | 16.1094 | 16.3 | 14.3 | 16.6 | 16.2 | 16.5 | 15.74 | 18.4407 | 17 | 15.2 | [9 empty] | 15.829
| 1 | 5 | 14.7 | 27 | 16.0025 | 18.9 | 13.8 | 16.7 | 17.1 | 15.9 | 15.2 | 18.2638 | 23 | 16.8 | [9 empty] | 82
| 1 | 6 | 15 | 10 | 15.9219 | 18 | 13.8 | 16.9 | 15 | 15.6 | 15.46 | 17.172 | 17 | 15.2
| Mean |  | 14.8 | 15.6667 | 15.9781 | 17.6 | 14.0167 | 16.7833 | 17.0333 | 16 | 15.7167 | 17.8605 | 19.6667 | 15.8833
| Median |  | 14.85 | 15 | 15.9834 | 17.6 | 14.05 | 16.8 | 16.2 | 16.05 | 15.755 | 18.0503 | 19 | 15.9
| Std Dev. |  | 0.45607 | 6.2823 | 0.082709 | 0.88091 | 0.19408 | 0.17224 | 2.4304 | 0.4 | 0.3524 | 0.60842 | 2.8048 | 0.62743
| Rel.Std.Dev. |  | 3.082% | 40.1% | 0.5176% | 5.005% | 1.385% | 1.026% | 14.27% | 2.5% | 2.242% | 3.406% | 14.26% | 3.95%
| PDM3 |  | -6.501% | -1.025% | 0.9422% | 11.19% | -11.45% | 6.029% | 7.609% | 1.081% | -0.7094% | 12.83% | 24.24% | 0.3435%
| Z-Score (Absolute) |  | 0.87 | 0.24 | 0.01 | 1.18 | 1.45 | 0.58 | 0.77 | 0.01 | 0.2 | 1.37 | 2.7 | 0.08
| Analytical results for Co in OREAS 84 (Certified Value 1654 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES |  | 0
0
| 1 | 1 | 1800 | 1510 | 1669 | 1720 | 1718.9842 | 1604 | 1680 | 1271 | 1690 | 1590 | 1759 | 1630 | 1713 | 1754.2 | 539.8679 | 1610 | 1613.4286 | 1580 | 1600 | 1427 |  | 1
| 1 | 2 | 1980 | 1520 | 1665 | 1710 | 1716.7265 | 1612 | 1720 | 1172 | 1680 | 1590 | 1734 | 1640 | 1685 | 1737.9 | 451.6565 | 1600 | 1500.6303 | 1560 | 1600 | 1444 |  | #N/A
| 1 | 3 | 1970 | 1470 | 1729 | 1700 | 1717.4388 | 1631 | 1680 | 1357 | 1860 | 1570 | 1747 | 1680 | 1704 | 1727.4 | 532.2183 | 1600 | 1501.6442 | 1570 | 1600 | 1469 |  | 16
| 1 | 4 | 1820 | 1520 | 1690 | 1710 | 1713.1065 | 1598 | 1680 | 1205 | 1650 | 1580 | 1758 | 1670 | 1760 | 1728.7 | 575.4207 | 1650 | 1561.0184 | 1510 | 1600 | 1423 |  | 1653.8378
| 1 | 5 | 1900 | 1520 | 1699 | 1710 | 1711.1305 | 1617 | 1700 | 1153 | 1670 | 1620 | 1751 | 1730 | 1678 | 1738.4 | 630.3209 | 1650 | 1599.8129 | 1600 | 1600 | 1440 |  | 83
| 1 | 6 | 1880 | 1530 | 1649 | 1690 | 1711.854 | 1598 | 1680 | 1260 | 1690 | 1620 | 1762 | 1730 | 1677 | 1728.1 | 728.0976 | 1700 | 1575.9842 | 1520 | 1600 | 1487
| Mean |  | 1891.6667 | 1511.6667 | 1683.5 | 1706.6667 | 1714.8734 | 1610 | 1690 | 1236.3333 | 1706.6667 | 1595 | 1751.8333 | 1680 | 1702.8333 | 1735.7833 | 576.2636 | 1635 | 1558.7531 | 1556.6667 | 1600 | 1448.3333
| Median |  | 1890 | 1520 | 1679.5 | 1710 | 1714.9165 | 1608 | 1680 | 1232.5 | 1685 | 1590 | 1754.5 | 1675 | 1694.5 | 1733.3 | 557.6443 | 1630 | 1568.5013 | 1565 | 1600 | 1442
| Std Dev. |  | 74.4088 | 21.3698 | 28.6199 | 10.328 | 3.2607 | 12.7906 | 16.7332 | 75.349 | 76.5942 | 20.7364 | 10.3425 | 42.8952 | 31.5431 | 10.2955 | 94.6235 | 39.37 | 48.2018 | 35.0238 | 0 | 24.9292
| Rel.Std.Dev. |  | 3.934% | 1.414% | 1.7% | 0.6052% | 0.1901% | 0.7944% | 0.9901% | 6.095% | 4.488% | 1.3% | 0.5904% | 2.553% | 1.852% | 0.5931% | 16.42% | 2.408% | 3.092% | 2.25% | 0% | 1.721%
| PDM3 |  | 14.38% | -8.596% | 1.794% | 3.194% | 3.691% | -2.651% | 2.187% | -25.24% | 3.194% | -3.558% | 5.925% | 1.582% | 2.963% | 4.955% | -65.16% | -1.139% | -5.749% | -5.875% | -3.255% | -12.43%
| Z-Score (Absolute) |  | 2.04 | 1.62 | 0.03 | 0.26 | 0.34 | 0.67 | 0.1 | 4.27 | 0.26 | 0.82 | 0.69 | 0 | 0.22 | 0.54 | 10.63 | 0.43 | 1.17 | 1.19 | NA | 2.23
Results from laboratory 21 were removed due to their 100 ppm reading resolution.
| Analytical results for Cr in OREAS 84 (Certified Value 0.747 wt.%) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.686 | 0.691 | 0.7845 | 0.766 | 0.754 | 0.57421 | 0.7323 | 0.78 | 0.7075 | 0.739 | 0.707 | 0.7184 | 0.746 | 0.7778 | 0.746 | 0.74968 | 0.763 | 0.8705 | 0.753 | 0.76 | 0.6212 | 1
| 1 | 2 | 0.763 | 0.691 | 0.7692 | 0.78 | 0.7623 | 0.57546 | 0.7469 | 0.78 | 0.7013 | 0.725 | 0.727 | 0.7325 | 0.753 | 0.7586 | 0.746 | 0.75742 | 0.787 | 0.82532 | 0.753 | 0.76 | 0.7284 | #N/A
| 1 | 3 | 0.753 | 0.691 | 0.7935 | 0.773 | 0.7717 | 0.57451 | 0.7483 | 0.78 | 0.5841 | 0.821 | 0.733 | 0.7183 | 0.78 | 0.7665 | 0.746 | 0.75729 | 0.779 | 0.81068 | 0.76 | 0.76 | 0.7014 | 16
| 1 | 4 | 0.7 | 0.698 | 0.7861 | 0.773 | 0.7628 | 0.57811 | 0.7323 | 0.773 | 0.6944 | 0.725 | 0.71 | 0.7122 | 0.78 | 0.7928 | 0.753 | 0.747 | 0.771 | 0.82596 | 0.753 | 0.76 | 0.6701 | 0.74674
| 1 | 5 | 0.727 | 0.698 | 0.7793 | 0.759 | 0.7719 | 0.57924 | 0.7656 | 0.787 | 0.6496 | 0.753 | 0.722 | 0.7073 | 0.801 | 0.766 | 0.753 | 0.75687 | 0.776 | 0.86211 | 0.765 | 0.76 | 0.6941 | 84
| 1 | 6 | 0.743 | 0.712 | 0.7923 | 0.773 | 0.7559 | 0.57915 | 0.7319 | 0.78 | 0.6892 | 0.759 | 0.726 | 0.7137 | 0.794 | 0.7741 | 0.759 | 0.76433 | 0.794 | 0.86729 | 0.742 | 0.76 | 0.6998
| Mean |  | 0.72867 | 0.69683 | 0.78415 | 0.77067 | 0.7631 | 0.57678 | 0.74288 | 0.78 | 0.67102 | 0.75367 | 0.72083 | 0.71707 | 0.77567 | 0.77263 | 0.7505 | 0.75543 | 0.77833 | 0.84364 | 0.75433 | 0.76 | 0.68583
| Median |  | 0.735 | 0.6945 | 0.7853 | 0.773 | 0.76255 | 0.57678 | 0.7396 | 0.78 | 0.6918 | 0.746 | 0.724 | 0.716 | 0.78 | 0.7703 | 0.7495 | 0.75708 | 0.7775 | 0.84403 | 0.753 | 0.76 | 0.69695
| Std Dev. |  | 0.030402 | 0.0081833 | 0.0089977 | 0.0072296 | 0.0075734 | 0.0023218 | 0.013462 | 0.0044272 | 0.047199 | 0.035837 | 0.010226 | 0.0086243 | 0.021951 | 0.011951 | 0.0053944 | 0.0062116 | 0.011094 | 0.025911 | 0.0077889 | 1.2162e-16 | 0.036722
| Rel.Std.Dev. |  | 4.172% | 1.174% | 1.147% | 0.9381% | 0.9924% | 0.4025% | 1.812% | 0.5676% | 7.034% | 4.755% | 1.419% | 1.203% | 2.83% | 1.547% | 0.7188% | 0.8223% | 1.425% | 3.071% | 1.033% | 1.6e-14% | 5.354%
| PDM3 |  | -2.421% | -6.684% | 5.009% | 3.203% | 2.19% | -22.76% | -0.5171% | 4.453% | -10.14% | 0.927% | -3.47% | -3.974% | 3.873% | 3.467% | 0.5029% | 1.164% | 4.23% | 12.98% | 1.016% | 1.775% | -8.157%
| Z-Score (Absolute) |  | 0.72 | 1.62 | 0.84 | 0.47 | 0.25 | 4.99 | 0.32 | 0.72 | 2.34 | 0.02 | 0.94 | 1.05 | 0.59 | 0.51 | 0.11 | 0.03 | 0.67 | 2.51 | 0 | 0.16 | 1.92
| Analytical results for Cs in OREAS 84 (Certified Value 1.73 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 06 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 1.6 | 1.8 | 1.5906 | 1.8 | 1.9 | 1.7 | 1.68 | 1.8268 | 1.6 | 3.2298 | 1.9 | [10 empty] | 1
| 1 | 2 | 1.9 | 1.8 | 1.5763 | 1.6 | 1.8 | 1.7 | 1.75 | 1.7665 | 1.6 | 2.2579 | 1.6 | [10 empty] | #N/A
| 1 | 3 | 1.4 | 1.9 | 1.5224 | 1.8 | 1.9 | 2.1 | 1.74 | 2.0824 | 1.5 | 2.141 | 1.5 | [10 empty] | 16
| 1 | 4 | 1.7 | 1.9 | 1.5711 | 1.6 | 1.8 | 1.8 | 1.73 | 1.7427 | 1.7 | 2.4713 | 2 | [10 empty] | 1.7278
| 1 | 5 | 1.9 | 1.9 | 1.6248 | 1.4 | 2.1 | 1.8 | 1.67 | 1.7108 | 1.7 | 3.2671 | 1.7 | [10 empty] | 85
| 1 | 6 | 1.5 | 1.6 | 1.5854 | 1.6 | 1.8 | 1.7 | 1.66 | 1.9368 | 1.7 | 1.7862 | 1.6
| Mean |  | 1.6667 | 1.8167 | 1.5785 | 1.6333 | 1.8833 | 1.8 | 1.705 | 1.8443 | 1.6333 | 2.5255 | 1.7167
| Median |  | 1.65 | 1.85 | 1.5809 | 1.6 | 1.85 | 1.75 | 1.705 | 1.7967 | 1.65 | 2.3646 | 1.65
| Std Dev. |  | 0.20656 | 0.1169 | 0.033301 | 0.15055 | 0.1169 | 0.15492 | 0.03937 | 0.14122 | 0.08165 | 0.6025 | 0.19408
| Rel.Std.Dev. |  | 12.39% | 6.435% | 2.11% | 9.218% | 6.207% | 8.607% | 2.309% | 7.657% | 4.999% | 23.86% | 11.31%
| PDM3 |  | -3.537% | 5.145% | -8.642% | -5.466% | 9.003% | 4.18% | -1.318% | 6.746% | -5.466% | 46.17% | -0.6432%
| Z-Score (Absolute) |  | 0.4 | 0.81 | 1.12 | 0.67 | 1.35 | 0.67 | 0.09 | 1.03 | 0.67 | 6.55 | 0
| Analytical results for Cu in OREAS 84 (Certified Value 0.405 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.389 | 0.368 | 0.4034 | 0.41 | 0.3956 | 0.40352 | 0.393 | 0.405 | 0.3389 | 0.398 | 0.42 | 0.3752 | 0.39 | 0.4115 | 0.4063 | 0.23499 | 0.226 | 0.41119 | 0.382 | 0.4 | 0.3563 | 1
| 1 | 2 | 0.422 | 0.38 | 0.3947 | 0.412 | 0.3962 | 0.40179 | 0.3998 | 0.407 | 0.3379 | 0.401 | 0.42 | 0.3813 | 0.401 | 0.4184 | 0.41084 | 0.20814 | 0.266 | 0.42066 | 0.38 | 0.39 | 0.4168 | #N/A
| 1 | 3 | 0.424 | 0.365 | 0.4083 | 0.412 | 0.3991 | 0.40289 | 0.4044 | 0.403 | 0.3405 | 0.425 | 0.425 | 0.3554 | 0.406 | 0.4171 | 0.40114 | 0.29211 | 0.285 | 0.40611 | 0.383 | 0.4 | 0.3908 | 16
| 1 | 4 | 0.4 | 0.366 | 0.4059 | 0.414 | 0.4026 | 0.40341 | 0.4002 | 0.409 | 0.3295 | 0.401 | 0.414 | 0.3504 | 0.41 | 0.4222 | 0.40497 | 0.24485 | 0.288 | 0.42653 | 0.381 | 0.39 | 0.3711 | 0.40483
| 1 | 5 | 0.408 | 0.373 | 0.4021 | 0.409 | 0.4007 | 0.40258 | 0.4028 | 0.408 | 0.334 | 0.4 | 0.42 | 0.3403 | 0.416 | 0.4123 | 0.40532 | 0.25742 | 0.281 | 0.43753 | 0.388 | 0.39 | 0.3879 | 86
| 1 | 6 | 0.414 | 0.376 | 0.4081 | 0.409 | 0.3949 | 0.40256 | 0.4045 | 0.409 | 0.3391 | 0.394 | 0.419 | 0.3477 | 0.406 | 0.4178 | 0.401 | 0.24738 | 0.27 | 0.43679 | 0.378 | 0.4 | 0.4091
| Mean |  | 0.4095 | 0.37133 | 0.40375 | 0.411 | 0.39818 | 0.40279 | 0.40078 | 0.40683 | 0.33665 | 0.40317 | 0.41967 | 0.35838 | 0.40483 | 0.41655 | 0.40493 | 0.24748 | 0.26933 | 0.42314 | 0.382 | 0.395 | 0.38867
| Median |  | 0.411 | 0.3705 | 0.40465 | 0.411 | 0.39765 | 0.40274 | 0.4015 | 0.4075 | 0.3384 | 0.4005 | 0.42 | 0.3529 | 0.406 | 0.41745 | 0.40514 | 0.24612 | 0.2755 | 0.42359 | 0.3815 | 0.395 | 0.38935
| Std Dev. |  | 0.013413 | 0.0059889 | 0.0050793 | 0.002 | 0.0031006 | 0.00063551 | 0.0043093 | 0.0024014 | 0.0041375 | 0.011017 | 0.0035024 | 0.016259 | 0.0088185 | 0.0040193 | 0.0036542 | 0.027573 | 0.022888 | 0.012992 | 0.0034059 | 0.0054772 | 0.022673
| Rel.Std.Dev. |  | 3.275% | 1.613% | 1.258% | 0.4866% | 0.7787% | 0.1578% | 1.075% | 0.5903% | 1.229% | 2.733% | 0.8346% | 4.537% | 2.178% | 0.9649% | 0.9024% | 11.14% | 8.498% | 3.071% | 0.8916% | 1.387% | 5.834%
| PDM3 |  | 1.154% | -8.274% | -0.2665% | 1.524% | -1.642% | -0.5031% | -0.9993% | 0.4951% | -16.84% | -0.4106% | 3.665% | -11.47% | 0.001101% | 2.895% | 0.02457% | -38.87% | -33.47% | 4.522% | -5.639% | -2.428% | -3.992%
| Z-Score (Absolute) |  | 0.55 | 2.58 | 0.08 | 0.67 | 0.38 | 0 | 0.17 | 0.33 | 5.43 | 0.03 | 1.39 | 3.65 | 0.17 | 1.13 | 0.18 | 12.76 | 10.96 | 1.67 | 1.71 | 0.64 | 1.16
| Analytical results for Dy in OREAS 84 (Certified Value 0.93 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.8 | < 2 | 0.79806 | 0.7 | 0.81 | 1 | 1.05 | 0.93 | 0.94 | 1.1418 | 0.79 | 0.96 | [9 empty] | 1
| 1 | 2 | 0.9 | < 2 | 0.78621 | 0.81 | 0.83 | 1 | 1 | 1.01 | 0.93 | 1.0418 | 1.04 | 0.94 | [9 empty] | #N/A
| 1 | 3 | 0.9 | < 2 | 0.81347 | 0.7 | 0.88 | 1 | 0.93 | 0.94 | 0.95 | 1.0415 | 1.23 | 0.87 | [9 empty] | 16
| 1 | 4 | 0.8 | < 2 | 0.79529 | 0.65 | 0.77 | 1 | 0.99 | 0.93 | 0.91 | 1.1253 | 1.03 | 0.99 | [9 empty] | 0.93436
| 1 | 5 | 0.9 | < 2 | 0.81935 | 0.69 | 0.82 | 1 | 0.99 | 0.93 | 0.94 | 1.0964 | 0.87 | 0.83 | [9 empty] | 87
| 1 | 6 | 0.9 | < 2 | 0.80289 | 0.7 | 0.83 | 1 | 0.94 | 0.91 | 0.92 | 1.157 | 0.97 | 0.83
| Mean |  | 0.86667 |  | 0.80255 | 0.70833 | 0.82333 | 1 | 0.98333 | 0.94167 | 0.93167 | 1.1006 | 0.98833 | 0.90333
| Median |  | 0.9 |  | 0.80048 | 0.7 | 0.825 | 1 | 0.99 | 0.93 | 0.935 | 1.1108 | 1 | 0.905
| Std Dev. |  | 0.05164 |  | 0.012178 | 0.053448 | 0.03559 | 0 | 0.043665 | 0.034881 | 0.01472 | 0.049912 | 0.15263 | 0.069186
| Rel.Std.Dev. |  | 5.958% |  | 1.517% | 7.546% | 4.323% | 0% | 4.441% | 3.704% | 1.58% | 4.535% | 15.44% | 7.659%
| PDM3 |  | -7.244% |  | -14.11% | -24.19% | -11.88% | 7.026% | 5.242% | 0.7825% | -0.2877% | 17.8% | 5.777% | -3.32%
| Z-Score (Absolute) |  | NA | 0.74 | 1.57 | 2.68 | 1.33 | NA | 0.55 | 0.06 | 0.06 | 1.92 | 0.61 | 0.39
Results from laboratories 1 and 11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 84 (Certified Value 0.5 ppm) | [22 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.5 | < 2 | 0.45951 | 0.66 | 0.46 | 0.6 | 0.59 | 0.49 | 0.51 | 0.56935 | 0.24 | 0.5 | [9 empty] | 1
| 1 | 2 | 0.5 | < 2 | 0.47085 | 0.69 | 0.45 | 0.6 | 0.55 | 0.51 | 0.51 | 0.5755 | 0.62 | 0.47 | [9 empty] | #N/A
| 1 | 3 | 0.5 | < 2 | 0.49406 | 0.6 | 0.53 | 0.5 | 0.53 | 0.5 | 0.51 | 0.5681 | 0.44 | 0.51 | [9 empty] | 16
| 1 | 4 | 0.5 | < 2 | 0.37855 | 0.8 | 0.4 | 0.6 | 0.56 | 0.47 | 0.52 | 0.59811 | 0.62 | 0.46 | [9 empty] | 0.49702
| 1 | 5 | 0.5 | < 2 | 0.40843 | 0.79 | 0.43 | 0.5 | 0.55 | 0.46 | 0.51 | 0.49253 | 0.33 | 0.54 | [9 empty] | 88
| 1 | 6 | 0.5 | < 2 | 0.49251 | 0.87 | 0.42 | 0.6 | 0.5 | 0.47 | 0.52 | 0.53365 | 0.42 | 0.47
| Mean |  | 0.5 |  | 0.45065 | 0.735 | 0.44833 | 0.56667 | 0.54667 | 0.48333 | 0.51333 | 0.55621 | 0.445 | 0.49167
| Median |  | 0.5 |  | 0.46518 | 0.74 | 0.44 | 0.6 | 0.55 | 0.48 | 0.51 | 0.56873 | 0.43 | 0.485
| Std Dev. |  | 0 |  | 0.047126 | 0.10134 | 0.04535 | 0.05164 | 0.030111 | 0.019664 | 0.005164 | 0.037431 | 0.15307 | 0.030605
| Rel.Std.Dev. |  | 0% |  | 10.46% | 13.79% | 10.12% | 9.113% | 5.508% | 4.068% | 1.006% | 6.73% | 34.4% | 6.225%
| PDM3 |  | 0.5988% |  | -9.33% | 47.88% | -9.796% | 14.01% | 9.988% | -2.755% | 3.281% | 11.91% | -10.47% | -1.078%
| Z-Score (Absolute) |  | NA | 6.36 | 0.66 | 2.97 | 0.69 | NA | 0.56 | 0.24 | 0.14 | 0.69 | 0.73 | 0.14
Results from laboratories 1 and 11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 84 (Certified Value 0.3 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.3 | < 2 | 0.26184 | 0.31 | 0.28 | 0.4 | 0.3 | 0.33 | 0.33 | 0.37968 | 0.37 | 0.32 | [9 empty] | 1
| 1 | 2 | 0.3 | < 2 | 0.25945 | 0.21 | 0.29 | 0.3 | 0.34 | 0.3 | 0.33 | 0.39884 | 0.38 | 0.32 | [9 empty] | #N/A
| 1 | 3 | 0.3 | < 2 | 0.27774 | 0.24 | 0.3 | 0.4 | 0.29 | 0.32 | 0.33 | 0.38778 | 0.34 | 0.36 | [9 empty] | 16
| 1 | 4 | 0.3 | < 2 | 0.25271 | 0.22 | 0.28 | 0.3 | 0.31 | 0.32 | 0.32 | 0.40845 | 0.29 | 0.29 | [9 empty] | 0.29785
| 1 | 5 | 0.3 | < 2 | 0.27752 | 0.46 | 0.28 | 0.3 | 0.3 | 0.32 | 0.32 | 0.3843 | 0.49 | 0.32 | [9 empty] | 89
| 1 | 6 | 0.3 | < 2 | 0.26433 | 0.4 | 0.28 | 0.4 | 0.27 | 0.33 | 0.31 | 0.40296 | 0.33 | 0.27
| Mean |  | 0.3 |  | 0.2656 | 0.30667 | 0.285 | 0.35 | 0.30167 | 0.32 | 0.32333 | 0.39367 | 0.36667 | 0.31333
| Median |  | 0.3 |  | 0.26308 | 0.275 | 0.28 | 0.35 | 0.3 | 0.32 | 0.325 | 0.39331 | 0.355 | 0.32
| Std Dev. |  | 0 |  | 0.010093 | 0.10347 | 0.0083666 | 0.054772 | 0.023166 | 0.010954 | 0.008165 | 0.0114 | 0.068313 | 0.030768
| Rel.Std.Dev. |  | 0% |  | 3.8% | 33.74% | 2.936% | 15.65% | 7.679% | 3.423% | 2.525% | 2.896% | 18.63% | 9.82%
| PDM3 |  | 0.7228% |  | -10.83% | 2.961% | -4.313% | 17.51% | 1.282% | 7.438% | 8.557% | 32.17% | 23.11% | 5.199%
| Z-Score (Absolute) |  | NA | 19.75 | 1.48 | 0.29 | 0.92 | NA | 0.43 | 0.1 | 0.19 | 2.23 | 1.44 | 0.1
Results from laboratories 1 and 11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe2O3 in OREAS 84 (Certified Value 44.64 wt.%) | [22 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | > 43 | 42.2 | 46.094 | 47.2 | 44.6821 | 41.81 | 44.02 | 45.9 | 32.6355 | 45.9 | 42.591 | 45.1 | > 35.7 | 44.664 | 39.7897 | 43.3 | 51.9942 | 42.891 | 44.607 | 45.229 | > 43 | 1
| 1 | 2 | > 43 | 42.3 | 45.221 | 47.2 | 45.0001 | 43.5867 | 44.006 | 46.6 | 30.5834 | 45.6 | 41.204 | 45.4 | > 35.7 | 44.507 | 38.585 | 45.3 | 49.153 | 42.605 | 45.179 | 45.498 | > 43 | #N/A
| 1 | 3 | > 43 | 41.5 | 46.665 | 47.5 | 45.237 | 41.91 | 43.92 | 45.8 | > 35.7 | 49.7 | 40.861 | 46.7 | > 35.7 | 44.178 | 39.8287 | 45.4 | 47.6863 | 43.034 | 45.321 |  | > 43 | 16
| 1 | 4 | > 43 | 42.6 | 46.48 | 47.6 | 45.5601 | 43.2307 | 43.963 | 45.5 | 30.6596 | 44.6 | 42.948 | 46.5 | > 35.7 | 43.978 | 39.0249 | 45.1 | 48.4301 | 42.891 | 45.179 |  | > 43 | 44.6442
| 1 | 5 | > 43 | 42.5 | 45.665 | 46.9 | 45.3262 | 42.068 | 44.306 | 46.1 | 30.1028 | 45.2 | 42.09 | 47.9 | > 35.7 | 43.906 | 41.691 | 45.1 | 51.3387 | 43.606 | 45.75 |  | > 43 | 90
| 1 | 6 | > 43 | 42.9 | 46.422 | 47 | 45.1561 | 42.6726 | 43.892 | 45.6 | 34.446 | 45.1 | 40.032 | 47.5 | > 35.7 | 44.821 | 42.9975 | 46.4 | 51.1618 | 42.462 | 45.321 |  | > 43
| Mean |  |  | 42.3333 | 46.0912 | 47.2333 | 45.1603 | 42.5464 | 44.0178 | 45.9167 | 31.6855 | 46.0167 | 41.621 | 46.5167 |  | 44.3423 | 40.3195 | 45.1 | 49.9607 | 42.9148 | 45.2262 | 45.3635
| Median |  |  | 42.4 | 46.258 | 47.2 | 45.1965 | 42.3703 | 43.9845 | 45.85 | 30.6596 | 45.4 | 41.647 | 46.6 |  | 44.3425 | 39.8092 | 45.2 | 50.1574 | 42.891 | 45.25 | 45.3635
| Std Dev. |  |  | 0.4761 | 0.55295 | 0.27325 | 0.29919 | 0.74061 | 0.1494 | 0.39707 | 1.8231 | 1.8584 | 1.1134 | 1.1107 |  | 0.37679 | 1.6887 | 1.006 | 1.7689 | 0.39852 | 0.36884 | 0.19021
| Rel.Std.Dev. |  |  | 1.125% | 1.2% | 0.5785% | 0.6625% | 1.741% | 0.3394% | 0.8648% | 5.754% | 4.039% | 2.675% | 2.388% |  | 0.8497% | 4.188% | 2.231% | 3.541% | 0.9286% | 0.8155% | 0.4193%
| PDM3 |  |  | -5.176% | 3.241% | 5.799% | 1.156% | -4.699% | -1.403% | 2.85% | -29.03% | 3.074% | -6.772% | 4.194% |  | -0.6762% | -9.687% | 1.021% | 11.91% | -3.874% | 1.303% | 1.611%
| Z-Score (Absolute) |  | NA | 1.51 | 0.52 | 1.14 | 0.02 | 1.39 | 0.6 | 0.42 | 7.26 | 0.48 | 1.89 | 0.75 | NA | 0.43 | 2.6 | 0.02 | 2.61 | 1.2 | 0.05 | 0.13 | NA
| Analytical results for Ga in OREAS 84 (Certified Value 7.05 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.5 | 7 | 7.5109 | 9 | 6 | 7 | 7 | 4.68 | 6.949 | 6.7 | 6.9392 | 6.9 | [9 empty] | 1
| 1 | 2 | 1.6 | 7 | 7.4232 | 9 | 6 | 7 | 7 | 4.93 | 7.2677 | 7.7 | 6.3411 | 6.5 | [9 empty] | #N/A
| 1 | 3 | 1.7 | 7 | 7.2366 | 8 | 5 | 7 | 7 | 4.91 | 7.1095 | 7.8 | 6.7279 | 6.9 | [9 empty] | 16
| 1 | 4 | 1.5 | 7 | 7.3718 | 9 | 5 | 7 | 7 | 4.83 | 7.4213 | 8.2 | 7.4321 | 6.2 | [9 empty] | 7.0535
| 1 | 5 | 1.6 | 7 | 7.337 | 8 | 6 | 7 | 7 | 4.72 | 6.6966 | 7.3 | 6.4703 | 6.6 | [9 empty] | 91
| 1 | 6 | 1.7 | 7 | 7.4098 | 8 | 6 | 7 | 7 | 4.64 | 7.0768 | 7.3 | 7.0861 | 6.3
| Mean |  | 1.6 | 7 | 7.3815 | 8.5 | 5.6667 | 7 | 7 | 4.785 | 7.0868 | 7.5 | 6.8328 | 6.5667
| Median |  | 1.6 | 7 | 7.3908 | 8.5 | 6 | 7 | 7 | 4.775 | 7.0932 | 7.5 | 6.8335 | 6.55
| Std Dev. |  | 0.089443 | 0 | 0.092044 | 0.54772 | 0.5164 | 0 | 0 | 0.12243 | 0.25159 | 0.51769 | 0.40467 | 0.29439
| Rel.Std.Dev. |  | 5.59% | 0% | 1.247% | 6.444% | 9.113% | 0% | 0% | 2.559% | 3.55% | 6.902% | 5.923% | 4.483%
| PDM3 |  | -77.32% | -0.7578% | 4.652% | 20.51% | -19.66% | -0.7578% | -0.7578% | -32.16% | 0.4734% | 6.331% | -3.129% | -6.901%
| Z-Score (Absolute) |  | 8.94 | 0 | 0.63 | 2.48 | 2.21 | 0 | 0 | 3.67 | 0.14 | 0.83 | 0.28 | 0.72
| Analytical results for Gd in OREAS 84 (Certified Value 1.12 ppm) | [22 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 1 | < 5 | 1.4661 | 1.44 | 1.04 | 1.1 | 1.19 | 1.02 | 1.13 | 1.3749 | 1.5 | 1.1 | [9 empty] | 1
| 1 | 2 | 1.1 | < 5 | 1.4123 | 1.27 | 1.06 | 1.1 | 1.31 | 1.02 | 1.1 | 1.3219 | 1.77 | 1.1 | [9 empty] | #N/A
| 1 | 3 | 1.1 | < 5 | 1.4195 | 1.1 | 1.1 | 1.1 | 1.26 | 1 | 1.09 | 1.4467 | 2.08 | 1.12 | [9 empty] | 16
| 1 | 4 | 1 | < 5 | 1.4439 | 1.24 | 1.05 | 1.2 | 1.12 | 1.11 | 1.12 | 1.494 | 1.11 | 1.04 | [9 empty] | 1.1203
| 1 | 5 | 1.2 | < 5 | 1.3572 | 1.3 | 0.96 | 1.1 | 1.28 | 1.05 | 1.09 | 1.4926 | 1.93 | 1.12 | [9 empty] | 92
| 1 | 6 | 1.1 | < 5 | 1.4329 | 1.42 | 1.06 | 1.1 | 1.21 | 1.06 | 1.09 | 1.376 | 1.04 | 0.98
| Mean |  | 1.0833 |  | 1.422 | 1.295 | 1.045 | 1.1167 | 1.2283 | 1.0433 | 1.1033 | 1.4177 | 1.5717 | 1.0767
| Median |  | 1.1 |  | 1.4262 | 1.285 | 1.055 | 1.1 | 1.235 | 1.035 | 1.095 | 1.4114 | 1.635 | 1.1
| Std Dev. |  | 0.075277 |  | 0.037017 | 0.12518 | 0.046368 | 0.040825 | 0.069113 | 0.039328 | 0.017512 | 0.070735 | 0.4306 | 0.055737
| Rel.Std.Dev. |  | 6.949% |  | 2.603% | 9.666% | 4.437% | 3.656% | 5.627% | 3.769% | 1.587% | 4.989% | 27.4% | 5.177%
| PDM3 |  | -3.303% |  | 26.93% | 15.59% | -6.724% | -0.3273% | 9.64% | -6.873% | -1.517% | 26.54% | 40.29% | -3.898%
| Z-Score (Absolute) |  | 0.48 | 7.16 | 1.35 | 0.66 | 0.69 | 0.3 | 0.3 | 0.7 | 0.37 | 1.32 | 2.15 | 0.52
| Analytical results for Ge in OREAS 84 (Indicative Value 1.35 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 09 | 12 | 15 | 16 | 18 | 19 | 20 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 0.7 | < 1 | 2 | 1 | 1 | 0.39 | 2.5 | < 1 | < 1 | [12 empty] | 1
| 1 | 2 | < 0.7 | < 1 | 2 | 1 | 1 | 0.41 | 1.9 | < 1 | < 1 | [12 empty] | 13
| 1 | 3 | < 0.7 | < 1 | 3 | 1 | 1 | 0.42 | 3.3 | < 1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.7 | < 1 | 2 | 1 | 1 | 0.4 | 2.5 | < 1 | < 1 | [12 empty] | 1.347
| 1 | 5 | < 0.7 | < 1 | 2 | < 1 | 1 | 0.4 | 1.7 | < 1 | < 1 | [12 empty] | 19
| 1 | 6 | < 0.7 | < 1 | 2 | < 1 | 1 | 0.39 | 2.1 | < 1 | < 1
| Mean |  |  |  | 2.1667 | 1 | 1 | 0.40167 | 2.3333
| Median |  |  |  | 2 | 1 | 1 | 0.4 | 2.3
| Std Dev. |  |  |  | 0.40825 | 0 | 0 | 0.01169 | 0.57155
| Rel.Std.Dev. |  |  |  | 18.84% | 0% | 0% | 2.91% | 24.49%
| PDM3 |  |  |  | 60.85% | -25.76% | -25.76% | -70.18% | 73.22%
| Z-Score (Absolute) |  | 0.67 | 0 | 7.49 | 1.5 | 2.25 | 0.44 | 8.24 | 0 | 0
| Analytical results for Hf in OREAS 84 (Indicative Value 0.93 ppm) | [22 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 09 | 15 | 16 | 17 | 18 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | [15 empty] | 2
2
| 1 | 1 | 20 | < 2 | < 1 | 0.8 | 0.87231 | 1.1 | [15 empty] | 1
| 1 | 2 | 30 | < 2 | < 1 | 0.9 | 0.883 | 0.7 | [15 empty] | 14
| 1 | 3 | 40 | < 2 | < 1 | 0.8 | 1.0034 | 1.5 | [15 empty] | 16
| 1 | 4 | 40 | < 2 | < 1 | 0.8 | 0.90499 | 1.1 | [15 empty] | 0.93028
| 1 | 5 | 20 | < 2 | < 1 | 1 | 1.0233 | 1.1 | [15 empty] | 20
| 1 | 6 | 20 | < 2 | < 1 | 0.9 | 0.858 | 1
| Mean |  | 28.3333 |  |  | 0.86667 | 0.92416 | 1.0833
| Median |  | 25 |  |  | 0.85 | 0.89399 | 1.1
| Std Dev. |  | 9.8319 |  |  | 0.08165 | 0.071044 | 0.25626
| Rel.Std.Dev. |  | 34.7% |  |  | 9.421% | 7.687% | 23.65%
| PDM3 |  | 2946% |  |  | -6.838% | -0.6571% | 16.45%
| Z-Score (Absolute) |  | 170.37 | 0.24 | 2.88 | 0.59 | 0.24 | 0.75
| Analytical results for Hg in OREAS 84 (Indicative Value 0.77 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | Lab | [20 empty] | 1
| No | No | 18 | [20 empty] | ppm
|  |  | BF*MS | [20 empty] | 2
2
| 1 | 1 | 0.9 | [20 empty] | 1
| 1 | 2 | 1.1 | [20 empty] | 10
| 1 | 3 | 0.5 | [20 empty] | 16
| 1 | 4 | 0.5 | [20 empty] | 0.76667
| 1 | 5 | 1.1 | [20 empty] | 16
| 1 | 6 | 0.5
| Mean |  | 0.76667
| Median |  | 0.7
| Std Dev. |  | 0.30111
| Rel.Std.Dev. |  | 39.28%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 84 (Certified Value 0.18 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.2 | < 2 | 0.12886 | 0.31 | 0.16 | 0.2 | 0.19 | 0.16 | 0.18 | 0.19784 | 0.25 | 0.18 | [9 empty] | 1
| 1 | 2 | < 0.2 | < 2 | 0.13877 | 0.14 | 0.16 | 0.2 | 0.2 | 0.18 | 0.17 | 0.20305 | 0.16 | 0.18 | [9 empty] | #N/A
| 1 | 3 | < 0.2 | < 2 | 0.12157 | 0.26 | 0.19 | 0.2 | 0.19 | 0.17 | 0.17 | 0.20445 | 0.26 | 0.19 | [9 empty] | 16
| 1 | 4 | < 0.2 | < 2 | 0.14643 | 0.15 | 0.16 | 0.2 | 0.18 | 0.17 | 0.18 | 0.19537 | 0.21 | 0.18 | [9 empty] | 0.17549
| 1 | 5 | < 0.2 | < 2 | 0.16302 | 0.13 | 0.17 | 0.2 | 0.19 | 0.17 | 0.17 | 0.20425 | 0.23 | 0.19 | [9 empty] | 93
| 1 | 6 | < 0.2 | < 2 | 0.17685 | 0.2 | 0.17 | 0.2 | 0.18 | 0.17 | 0.17 | 0.20713 | 0.23 | 0.16
| Mean |  |  |  | 0.14591 | 0.19833 | 0.16833 | 0.2 | 0.18833 | 0.17 | 0.17333 | 0.20201 | 0.22333 | 0.18
| Median |  |  |  | 0.1426 | 0.175 | 0.165 | 0.2 | 0.19 | 0.17 | 0.17 | 0.20365 | 0.23 | 0.18
| Std Dev. |  |  |  | 0.020903 | 0.073052 | 0.01169 | 3.0405e-17 | 0.0075277 | 0.0063246 | 0.005164 | 0.0044666 | 0.03559 | 0.010954
| Rel.Std.Dev. |  |  |  | 14.33% | 36.83% | 6.945% | 1.52e-14% | 3.997% | 3.72% | 2.979% | 2.211% | 15.94% | 6.086%
| PDM3 |  |  |  | -16.85% | 13.02% | -4.079% | 13.97% | 7.318% | -3.129% | -1.23% | 15.11% | 27.26% | 2.569%
| Z-Score (Absolute) |  | 2.94 | 30.16 | 1.25 | 0.67 | 0.43 | NA | 0.31 | 0.37 | 0.25 | 0.81 | 1.59 | 0
Results from laboratory 11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 84 (Indicative Value 0.14 ppm) | [22 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 09 | 12 | 15 | 16 | 19 | 20 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [13 empty] | 2
2
| 1 | 1 | < 0.2 | 0.1 | < 0.2 | < 0.2 | < 0.2 | 0.103 | 0.22206 | < 0.2 | [13 empty] | 1
| 1 | 2 | 0.2 | 0.1 | < 0.2 | < 0.2 | < 0.2 | 0.092 | 0.20388 | < 0.2 | [13 empty] | 11
| 1 | 3 | 0.2 | 0.1 | < 0.2 | < 0.2 | < 0.2 | 0.095 | 0.16272 | < 0.2 | [13 empty] | 16
| 1 | 4 | 0.2 | 0.2 | < 0.2 | < 0.2 | < 0.2 | 0.1 | 0.21049 | < 0.2 | [13 empty] | 0.13818
| 1 | 5 | < 0.2 | 0.1 | < 0.2 | < 0.2 | < 0.2 | 0.098 | 0.17299 | < 0.2 | [13 empty] | 17
| 1 | 6 | < 0.2 | 0.1 | < 0.2 | < 0.2 | < 0.2 | 0.103 | 0.15314 | < 0.2
| Mean |  | 0.2 | 0.11667 |  |  |  | 0.0985 | 0.18755
| Median |  | 0.2 | 0.1 |  |  |  | 0.099 | 0.18844
| Std Dev. |  | 3.3993e-17 | 0.040825 |  |  |  | 0.0044159 | 0.02827
| Rel.Std.Dev. |  | 1.7e-14% | 34.99% |  |  |  | 4.483% | 15.07%
| PDM3 |  | 44.74% | -15.57% |  |  |  | -28.72% | 35.73%
| Z-Score (Absolute) |  | 44.95 | 14.98 | 0 | 0 | 0 | 1.35 | 78.71 | 0
| Analytical results for K2O in OREAS 84 (Certified Value 0.453 wt.%) | [22 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES |  |  |  | 3
3
| 1 | 1 | 0.36 | 0.42 | 0.47 | 0.47 | 0.55147 | 0.494 | 0.42 | 0.44257 | 0.43 | 0.48 | 0.45 | 0.48 | 0.494 | 0.50874 | 0.5 | 0.47866 | 0.422 | 0.4794 |  |  |  | 1
| 1 | 2 | 0.48 | 0.42 | 0.458 | 0.48 | 0.42703 | 0.482 | 0.41 | 0.45185 | 0.42 | 0.48 | 0.45 | 0.48 | 0.482 | 0.48489 | 0.5 | 0.47924 | 0.422 | 0.4867 |  |  |  | #N/A
| 1 | 3 | 0.36 | 0.4 | 0.458 | 0.49 | 0.43546 | 0.458 | 0.46 | 0.40535 | 0.48 | 0.48 | 0.47 | 0.48 | 0.47 | 0.48121 | 0.51 | 0.45698 | 0.41 | 0.4662 |  |  |  | 16
| 1 | 4 | 0.36 | 0.41 | 0.47 | 0.43 | 0.43161 | 0.458 | 0.42 | 0.42041 | 0.46 | 0.36 | 0.48 | 0.48 | 0.482 | 0.48339 | 0.49 | 0.46093 | 0.41 | 0.4638 |  |  |  | 0.45348
| 1 | 5 | 0.48 | 0.41 | 0.458 | 0.47 | 0.42113 | 0.494 | 0.43 | 0.45341 | 0.45 | 0.36 | 0.46 | 0.48 | 0.482 | 0.47393 | 0.48 | 0.48507 | 0.41 | 0.4614 |  |  |  | 94
| 1 | 6 | 0.48 | 0.43 | 0.47 | 0.46 | 0.41956 | 0.446 | 0.42 | 0.45462 | 0.43 | 0.36 | 0.47 | 0.48 | 0.494 | 0.47561 | 0.47 | 0.47079 | 0.41 | 0.4903
| Mean |  | 0.42 | 0.415 | 0.464 | 0.46667 | 0.44771 | 0.472 | 0.42667 | 0.43804 | 0.445 | 0.42 | 0.46333 | 0.48 | 0.484 | 0.48463 | 0.49167 | 0.47195 | 0.414 | 0.47463
| Median |  | 0.42 | 0.415 | 0.464 | 0.47 | 0.42932 | 0.47 | 0.42 | 0.44721 | 0.44 | 0.42 | 0.465 | 0.48 | 0.482 | 0.4823 | 0.495 | 0.47473 | 0.41 | 0.4728
| Std Dev. |  | 0.065727 | 0.010488 | 0.0065727 | 0.020656 | 0.05119 | 0.020669 | 0.017512 | 0.020503 | 0.022583 | 0.065727 | 0.012111 | 0 | 0.0090333 | 0.012573 | 0.01472 | 0.011111 | 0.0061968 | 0.012469
| Rel.Std.Dev. |  | 15.65% | 2.527% | 1.417% | 4.426% | 11.43% | 4.379% | 4.104% | 4.681% | 5.075% | 15.65% | 2.614% | 0% | 1.866% | 2.594% | 2.994% | 2.354% | 1.497% | 2.627%
| PDM3 |  | -7.382% | -8.485% | 2.321% | 2.909% | -1.272% | 4.085% | -5.912% | -3.405% | -1.869% | -7.382% | 2.174% | 5.849% | 6.731% | 6.87% | 8.422% | 4.073% | -8.705% | 4.666%
| Z-Score (Absolute) |  | 1.46 | 1.69 | 0.01 | 0.11 | 0.55 | 0.29 | 1.28 | 0.89 | 0.64 | 1.46 | 0.01 | 0.63 | 0.7 | 0.73 | 0.98 | 0.29 | 1.73 | 0.38
| Analytical results for La in OREAS 84 (Certified Value 8.3 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | 23 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*OES | [8 empty] | 2
3
| 1 | 1 | 7.5 | 10 | 9.1421 | 7 | 7.5 | 8.5 | 8.4 | 8.4 | 9.8 | 8.0797 | 10 | 8 | 11 | [8 empty] | 1
| 1 | 2 | 8 | 11 | 9.0792 | 7 | 7.5 | 8.7 | 9.2 | 8.2 | 10 | 8.0875 | 11 | 8.3 | 12 | [8 empty] | 20
| 1 | 3 | 8.2 | 11 | 9.215 | 7 | 7.6 | 9.1 | 8.4 | 7.9 | 9.9 | 8.7067 | 13 | 8.3 | 11 | [8 empty] | 16
| 1 | 4 | 7.7 | 11 | 9.253 | 7 | 7.7 | 8.6 | 8.5 | 8.3 | 10.6 | 8.5757 | 10 | 8 | 11 | [8 empty] | 8.3036
| 1 | 5 | 7.9 | 12 | 9.0972 | 8 | 7.4 | 8.8 | 8.6 | 8 | 10.2 | 8.366 | 14 | 8.6 | < 10 | [8 empty] | 95
| 1 | 6 | 8.3 | 10 | 8.8834 | 7 | 7.2 | 8.6 | 7.9 | 8 | 10.2 | 8.3888 | 10 | 7.9 | 13
| Mean |  | 7.9333 | 10.8333 | 9.1117 | 7.1667 | 7.4833 | 8.7167 | 8.5 | 8.1333 | 10.1167 | 8.3674 | 11.3333 | 8.1833 | 11.6
| Median |  | 7.95 | 11 | 9.1196 | 7 | 7.5 | 8.65 | 8.45 | 8.1 | 10.1 | 8.3774 | 10.5 | 8.15 | 11
| Std Dev. |  | 0.30111 | 0.75277 | 0.13035 | 0.40825 | 0.17224 | 0.2137 | 0.41952 | 0.19664 | 0.28577 | 0.253 | 1.7512 | 0.26394 | 0.89443
| Rel.Std.Dev. |  | 3.795% | 6.949% | 1.431% | 5.696% | 2.302% | 2.452% | 4.936% | 2.418% | 2.825% | 3.024% | 15.45% | 3.225% | 7.711%
| PDM3 |  | -4.459% | 30.46% | 9.731% | -13.69% | -9.879% | 4.974% | 2.365% | -2.051% | 21.83% | 0.7681% | 36.49% | -1.449% | 39.7%
| Z-Score (Absolute) |  | 0.84 | NA | 1.44 | NA | 1.71 | 0.67 | 0.26 | 0.45 | 3.38 | 0 | NA | 0.36 | NA
Results from laboratories 5, 7, 18 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 84 (Certified Value 12.8 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 19 | 20 | 23 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | [9 empty] | 1
1
| 1 | 1 | 5 | 13 | 32 | 11.1285 | < 1 | < 10 | 15 | 12 | 10.7 | 15.1749 | 11 | 15 | [9 empty] | 1
| 1 | 2 | 4.3 | 26 | 28 | 11.0968 | < 1 | < 10 | 13 | 13 | 10.7 | 14.8966 | 11 | 12 | [9 empty] | #N/A
| 1 | 3 | 5.7 | 37 | 55 | 10.907 | < 1 | 11 | 13 | 12 | 10.7 | 13.7989 | 11 | 18 | [9 empty] | 16
| 1 | 4 | 5 | 44 | 22 | 10.8618 | < 1 | 14 | 14 | 14 | 10.8 | 12.7509 | 11 | 17 | [9 empty] | 12.846
| 1 | 5 | 5.6 | 41 | 213 | 10.8014 | < 1 | 13 | 14 | 13 | 11.5 | 13.1841 | 11 | 18 | [9 empty] | 96
| 1 | 6 | 4.7 | 36 | 27 | 10.6561 | < 1 | 14 | 15 | 13 | 11.4 | 15.5517 | 11 | 15
| Mean |  | 5.05 | 32.8333 | 62.8333 | 10.9086 |  | 13 | 14 | 12.8333 | 10.9667 | 14.2262 | 11 | 15.8333
| Median |  | 5 | 36.5 | 30 | 10.8844 |  | 13.5 | 14 | 13 | 10.75 | 14.3478 | 11 | 16
| Std Dev. |  | 0.53198 | 11.479 | 74.4672 | 0.17957 |  | 1.4142 | 0.89443 | 0.75277 | 0.37771 | 1.1447 | 0 | 2.3166
| Rel.Std.Dev. |  | 10.53% | 34.96% | 118.5% | 1.646% |  | 10.88% | 6.389% | 5.866% | 3.444% | 8.046% | 0% | 14.63%
| PDM3 |  | -60.69% | 155.6% | 389.1% | -15.08% |  | 1.199% | 8.983% | -0.09871% | -14.63% | 10.74% | -14.37% | 23.25%
| Z-Score (Absolute) |  | 2.11 | 6.42 | 15.63 | 0.31 | 3.5 | 0.49 | 0.64 | 0.28 | 0.29 | 0.71 | 0.28 | 1.2
| Analytical results for Lu in OREAS 84 (Certified Value 0.075 ppm) | [22 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 05 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [11 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [11 empty] | 3
3
| 1 | 1 | < 2 | 0.09 | 0.08 | 0.08 | < 0.1 | 0.07 | 0.07 | < 0.2 | < 0.01 | 0.06 | [11 empty] | 1
| 1 | 2 | < 2 | 0.11 | 0.07 | 0.1 | < 0.1 | 0.07 | 0.06 | < 0.2 | < 0.01 | 0.07 | [11 empty] | #N/A
| 1 | 3 | < 2 | 0.07 | 0.07 | 0.1 | < 0.1 | 0.07 | 0.07 | < 0.2 | < 0.01 | 0.08 | [11 empty] | 16
| 1 | 4 | < 2 | 0.09 | 0.08 | 0.1 | < 0.1 | 0.07 | 0.06 | < 0.2 | < 0.01 | 0.07 | [11 empty] | 0.075278
| 1 | 5 | < 2 | 0.06 | 0.08 | 0.06 | 0.1 | 0.07 | 0.06 | < 0.2 | < 0.01 | 0.06 | [11 empty] | 97
| 1 | 6 | < 2 | 0.09 | 0.07 | 0.1 | 0.1 | 0.07 | 0.06 | < 0.2 | < 0.01 | 0.07
| Mean |  |  | 0.085 | 0.075 | 0.09 | 0.1 | 0.07 | 0.063333 |  |  | 0.068333
| Median |  |  | 0.09 | 0.075 | 0.1 | 0.1 | 0.07 | 0.06 |  |  | 0.07
| Std Dev. |  |  | 0.017607 | 0.0054772 | 0.016733 | 0 | 0 | 0.005164 |  |  | 0.0075277
| Rel.Std.Dev. |  |  | 20.71% | 7.303% | 18.59% | 0% | 0% | 8.154% |  |  | 11.02%
| PDM3 |  |  | 12.92% | -0.369% | 19.56% | 32.84% | -7.011% | -15.87% |  |  | -9.225%
| Z-Score (Absolute) |  | 53.46 | 0.58 | 0 | 0.87 | NA | 0.29 | 0.67 | 1.44 | 4.05 | 0.39
Results from laboratory 12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for MgO in OREAS 84 (Certified Value 2.64 wt.%) | [22 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 2.57 | 2.44 | 2.686 | 2.66 | 2.6934 | 2.6374 | 2.686 | 2.67 | 2.4987 | 2.61 | 2.554 | 2.62 | 2.62 | 2.56 | 2.6389 | 2.58 | 2.6431 | 2.587 | 53.145 | 2.4657 |  | 1
| 1 | 2 | 2.736 | 2.42 | 2.637 | 2.71 | 2.6833 | 2.6463 | 2.653 | 2.72 | 2.4815 | 2.56 | 2.587 | 2.66 | 2.637 | 2.55 | 2.6526 | 2.67 | 2.4121 | 2.587 | 3.432 | 2.7261 |  | #N/A
| 1 | 3 | 2.753 | 2.39 | 2.719 | 2.69 | 2.6974 | 2.6608 | 2.603 | 2.65 | 2.2601 | 2.86 | 2.404 | 2.72 | 2.637 | 2.55 | 2.6274 | 2.69 | 2.4483 | 2.62 | 3.432 | 2.6117 |  | 16
| 1 | 4 | 2.57 | 2.45 | 2.703 | 2.71 | 2.7072 | 2.6324 | 2.537 | 2.65 | 2.4332 | 2.61 | 2.371 | 2.72 | 2.67 | 2.58 | 2.6021 | 2.68 | 2.5695 | 2.603 | 3.383 | 2.4657 |  | 2.6427
| 1 | 5 | 2.703 | 2.44 | 2.67 | 2.64 | 2.6995 | 2.6396 | 2.719 | 2.69 | 2.383 | 2.66 | 2.305 | 2.79 | 2.603 | 2.58 | 2.6375 | 2.64 | 2.6021 | 2.653 | 3.466 | 2.5652 |  | 98
| 1 | 6 | 2.736 | 2.49 | 2.703 | 2.66 | 2.6783 | 2.6288 | 2.587 | 2.67 | 2.4037 | 2.69 | 2.321 | 2.74 | 2.637 | 2.51 | 2.6547 | 2.74 | 2.5987 | 2.57 | 3.416 | 2.7609
| Mean |  | 2.678 | 2.4383 | 2.6863 | 2.6783 | 2.6932 | 2.6409 | 2.6308 | 2.675 | 2.4101 | 2.665 | 2.4237 | 2.7083 | 2.634 | 2.555 | 2.6355 | 2.6667 | 2.5456 | 2.6033 | 11.7123 | 2.5992
| Median |  | 2.7195 | 2.44 | 2.6945 | 2.675 | 2.6954 | 2.6385 | 2.628 | 2.67 | 2.4185 | 2.635 | 2.3875 | 2.72 | 2.637 | 2.555 | 2.6382 | 2.675 | 2.5841 | 2.595 | 3.432 | 2.5884
| Std Dev. |  | 0.085213 | 0.033116 | 0.029405 | 0.029269 | 0.010703 | 0.011478 | 0.067541 | 0.026646 | 0.08574 | 0.10559 | 0.11954 | 0.060139 | 0.022271 | 0.025884 | 0.01928 | 0.053541 | 0.093122 | 0.029629 | 20.2978 | 0.12586
| Rel.Std.Dev. |  | 3.182% | 1.358% | 1.095% | 1.093% | 0.3974% | 0.4346% | 2.567% | 0.9961% | 3.558% | 3.962% | 4.932% | 2.221% | 0.8455% | 1.013% | 0.7315% | 2.008% | 3.658% | 1.138% | 173.3% | 4.842%
| PDM3 |  | 1.336% | -7.733% | 1.652% | 1.349% | 1.911% | -0.06794% | -0.4484% | 1.223% | -8.803% | 0.8444% | -8.288% | 2.484% | -0.3286% | -3.318% | -0.2709% | 0.9075% | -3.672% | -1.489% | 343.2% | -1.645%
| Z-Score (Absolute) |  | 0.67 | 3.37 | 0.81 | 0.68 | 0.93 | 0.05 | 0.12 | 0.62 | 3.85 | 0.45 | 3.62 | 1.18 | 0.07 | 1.4 | 0.05 | 0.48 | 1.56 | 0.59 | 153.14 | 0.66
| Analytical results for MnO in OREAS 84 (Certified Value 0.085 wt.%) | [22 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.07489 | 0.08 | 0.08522 | 0.09 | 0.08303 | 0.079088 | 0.08199 | 0.09 | 0.08277 | 0.09 | 0.08729 | 0.08 | 0.08845 | 0.08651 | 0.090676 | 0.08 | 0.091034 | 0.07812 | 0.129 | 0.08406 |  | 1
| 1 | 2 | 0.08212 | 0.08 | 0.0847 | 0.09 | 0.08277 | 0.079051 | 0.08587 | 0.09 | 0.08393 | 0.09 | 0.08767 | 0.08 | 0.08948 | 0.08677 | 0.088891 | 0.09 | 0.087479 | 0.07825 | 0.129 | 0.08922 |  | #N/A
| 1 | 3 | 0.08199 | 0.08 | 0.08677 | 0.09 | 0.08264 | 0.079177 | 0.08793 | 0.09 | 0.07941 | 0.09 | 0.08871 | 0.09 | 0.08884 | 0.08561 | 0.089335 | 0.09 | 0.089027 | 0.07864 | 0.129 | 0.08496 |  | 16
| 1 | 4 | 0.07683 | 0.08 | 0.086 | 0.09 | 0.08354 | 0.079266 | 0.08561 | 0.09 | 0.08122 | 0.09 | 0.08574 | 0.09 | 0.09026 | 0.086 | 0.087505 | 0.08 | 0.088539 | 0.07812 | 0.155 | 0.08109 |  | 0.085184
| 1 | 5 | 0.07941 | 0.08 | 0.08496 | 0.09 | 0.08303 | 0.079344 | 0.08729 | 0.09 | 0.0798 | 0.09 | 0.0869 | 0.09 | 0.08832 | 0.08574 | 0.087392 | 0.08 | 0.089691 | 0.07967 | 0.129 | 0.08445 |  | 99
| 1 | 6 | 0.08135 | 0.08 | 0.08638 | 0.09 | 0.08212 | 0.079194 | 0.08522 | 0.09 | 0.08186 | 0.09 | 0.08587 | 0.08 | 0.08974 | 0.08612 | 0.087518 | 0.09 | 0.091545 | 0.07734 | 0.129 | 0.08871
| Mean |  | 0.079432 | 0.08 | 0.085672 | 0.09 | 0.082855 | 0.079187 | 0.085652 | 0.09 | 0.081498 | 0.09 | 0.08703 | 0.085 | 0.089182 | 0.086125 | 0.088553 | 0.085 | 0.089552 | 0.078357 | 0.13333 | 0.085415
| Median |  | 0.08038 | 0.08 | 0.08561 | 0.09 | 0.0829 | 0.079186 | 0.08574 | 0.09 | 0.08154 | 0.09 | 0.087095 | 0.085 | 0.08916 | 0.08606 | 0.088204 | 0.085 | 0.089359 | 0.078185 | 0.129 | 0.084705
| Std Dev. |  | 0.0029949 | 0 | 0.00083312 | 1.5202e-17 | 0.00047416 | 0.00010908 | 0.0020745 | 1.5202e-17 | 0.0017311 | 1.5202e-17 | 0.0011248 | 0.0054772 | 0.00076917 | 0.00044572 | 0.0013231 | 0.0054772 | 0.0015353 | 0.00076969 | 0.010614 | 0.0030664
| Rel.Std.Dev. |  | 3.77% | 0% | 0.9725% | 1.689e-14% | 0.5723% | 0.1377% | 2.422% | 1.689e-14% | 2.124% | 1.689e-14% | 1.292% | 6.444% | 0.8625% | 0.5175% | 1.494% | 6.444% | 1.714% | 0.9823% | 7.961% | 3.59%
| PDM3 |  | -6.753% | -6.086% | 0.5719% | 5.653% | -2.735% | -7.041% | 0.5484% | 5.653% | -4.327% | 5.653% | 2.167% | -0.2166% | 4.692% | 1.104% | 3.954% | -0.2166% | 5.128% | -8.015% | 56.52% | 0.2706%
| Z-Score (Absolute) |  | 1.13 | 1.03 | 0 | 0.79 | 0.51 | 1.18 | 0 | 0.79 | 0.76 | 0.79 | 0.25 | 0.12 | 0.64 | 0.08 | 0.53 | 0.12 | 0.71 | 1.33 | 8.69 | 0.05
| Analytical results for Mo in OREAS 84 (Certified Value 5.62 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 18 | 19 | 20 | 23 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [9 empty] | 2
2
| 1 | 1 | 5 | 5 | 8 | 5 | 5 | < 10 | 10 | 6.11 | 7.5 | 5.2395 | < 5 | N.A. | [9 empty] | 1
| 1 | 2 | 5 | 5 | 8 | 6 | 5 | < 10 | 9 | 6.01 | 8.3 | 4.9281 | < 5 | N.A. | [9 empty] | 21
| 1 | 3 | 6 | 6 | 8 | 7 | 5 | < 10 | 9 | 6.02 | 5.7 | 4.7748 | < 5 | N.A. | [9 empty] | 16
| 1 | 4 | 5 | 6 | 8 | 6 | 5 | < 10 | 8 | 5.94 | 6.4 | 4.9292 | < 5 | N.A. | [9 empty] | 5.6239
| 1 | 5 | 5 | 5 | 8 | 5 | 5 | < 10 | 9 | 6.13 | 9.6 | 5.1701 | < 5 | N.A. | [9 empty] | 100
| 1 | 6 | 5 | 6 | 8 | 5 | 5 | < 10 | 9 | 6.18 | 6.8 | 5.1325 | < 5 | N.A.
| Mean |  | 5.1667 | 5.5 | 8 | 5.6667 | 5 |  | 9 | 6.065 | 7.3833 | 5.029
| Median |  | 5 | 5.5 | 8 | 5.5 | 5 |  | 9 | 6.065 | 7.15 | 5.0309
| Std Dev. |  | 0.40825 | 0.54772 | 0 | 0.8165 | 0 |  | 0.63246 | 0.08961 | 1.4077 | 0.17871
| Rel.Std.Dev. |  | 7.902% | 9.959% | 0% | 14.41% | 0% |  | 7.027% | 1.477% | 19.07% | 3.553%
| PDM3 |  | -8.13% | -2.203% | 42.25% | 0.7603% | -11.09% |  | 60.03% | 7.843% | 31.28% | -10.58%
| Z-Score (Absolute) |  | 0.45 | 0 | 3.37 | 0.22 | 0.67 | 0.67 | 4.72 | 0.76 | 2.54 | 0.64 | 4.05 | NA
| Analytical results for Na2O in OREAS 84 (Indicative Value 0.636 wt.%) | [22 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 07 | 16 | 17 | 18 | [17 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | [17 empty] | 3
3
| 1 | 1 | 0.661 | 0.634 | 0.69726 | 0.58 | [17 empty] | 1
| 1 | 2 | 0.647 | 0.674 | 0.68212 | 0.58 | [17 empty] | 12
| 1 | 3 | 0.62 | 0.62 | 0.68797 | 0.58 | [17 empty] | 16
| 1 | 4 | 0.607 | 0.62 | 0.68857 | 0.59 | [17 empty] | 0.63597
| 1 | 5 | 0.674 | 0.634 | 0.68107 | 0.58 | [17 empty] | 18
| 1 | 6 | 0.58 | 0.661 | 0.68655 | 0.6
| Mean |  | 0.6315 | 0.6405 | 0.68726 | 0.585
| Median |  | 0.6335 | 0.634 | 0.68726 | 0.58
| Std Dev. |  | 0.035501 | 0.022215 | 0.0057864 | 0.0083666
| Rel.Std.Dev. |  | 5.622% | 3.468% | 0.842% | 1.43%
| PDM3 |  | -0.7024% | 0.7127% | 8.065% | -8.014%
| Z-Score (Absolute) |  | 0.11 | 0.11 | 1.25 | 1.24
| Analytical results for Nb in OREAS 84 (Indicative Value 1.99 ppm) | [22 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 2.4 | < 5 | 24 | < 10 | < 10 | 2 | 2.2 | 2.4975 | < 1 | 2.2208 | 8 | [10 empty] | 1
| 1 | 2 | < 2.4 | < 5 | 23 | < 10 | < 10 | 2 | 2.3 | 2.0191 | < 1 | 2.1022 | 7 | [10 empty] | 13
| 1 | 3 | < 2.4 | < 5 | 24 | < 10 | < 10 | 2 | 2.2 | 3.2512 | < 1 | 2.1408 | 6 | [10 empty] | 16
| 1 | 4 | < 2.4 | < 5 | 23 | < 10 | < 10 | 2 | 2.1 | 3.39 | < 1 | 2.1498 | < 5 | [10 empty] | 1.9886
| 1 | 5 | < 2.4 | < 5 | 23 | < 10 | < 10 | 2 | 2 | 3.0793 | 2 | 2.2298 | 6 | [10 empty] | 19
| 1 | 6 | < 2.4 | < 5 | 22 | < 10 | < 10 | 2 | 2.2 | 2.4083 | 1 | 2.1681 | 5
| Mean |  |  |  | 23.1667 |  |  | 2 | 2.1667 | 2.7743 | 1.5 | 2.1686 | 6.4
| Median |  |  |  | 23 |  |  | 2 | 2.2 | 2.7884 | 1.5 | 2.159 | 6
| Std Dev. |  |  |  | 0.75277 |  |  | 0 | 0.10328 | 0.54415 | 0.70711 | 0.048995 | 1.1402
| Rel.Std.Dev. |  |  |  | 3.249% |  |  | 0% | 4.767% | 19.61% | 47.14% | 2.259% | 17.82%
| PDM3 |  |  |  | 1065% |  |  | 0.5747% | 8.956% | 39.51% | -24.57% | 9.053% | 221.8%
| Z-Score (Absolute) |  | 0.67 | 0 | 10.72 | 1.3 | 1.3 | 0.26 | 0.17 | 0.14 | 0.86 | 0.17 | 1.69
| Analytical results for Nd in OREAS 84 (Certified Value 6.67 ppm) | [22 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 05 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 6.2 | 6 | 5.7 | 6.2 | 6.8 | 7 | 6.8 | 6.7 | 8.1066 | 7.8 | 6.59 | [10 empty] | 1
| 1 | 2 | 6.8 | 8 | 5.6 | 6.1 | 6.8 | 8.4 | 6.8 | 6.7 | 6.7381 | 9.8 | 6.65 | [10 empty] | #N/A
| 1 | 3 | 6.9 | 8 | 4.9 | 6.2 | 7.2 | 6.9 | 6.8 | 6.6 | 7.6642 | 12.2 | 6.44 | [10 empty] | 16
| 1 | 4 | 6.4 | 8 | 7.6 | 6 | 7.2 | 7.1 | 6.9 | 6.6 | 6.6385 | 7.8 | 6.55 | [10 empty] | 6.6726
| 1 | 5 | 6.5 | 9 | 7.3 | 6.2 | 6.6 | 7.5 | 6.7 | 6.4 | 7.2122 | 11.4 | 7.05 | [10 empty] | 101
| 1 | 6 | 6.6 | 7 | 5.3 | 6 | 7 | 6.7 | 6.5 | 6.5 | 6.3326 | 7.4 | 6.61
| Mean |  | 6.5667 | 7.6667 | 6.0667 | 6.1167 | 6.9333 | 7.2667 | 6.75 | 6.5833 | 7.1153 | 9.4 | 6.6483
| Median |  | 6.55 | 8 | 5.65 | 6.15 | 6.9 | 7.05 | 6.8 | 6.6 | 6.9751 | 8.8 | 6.6
| Std Dev. |  | 0.2582 | 1.0328 | 1.1112 | 0.098319 | 0.24221 | 0.61536 | 0.13784 | 0.1169 | 0.67461 | 2.0552 | 0.20942
| Rel.Std.Dev. |  | 3.932% | 13.47% | 18.32% | 1.607% | 3.493% | 8.468% | 2.042% | 1.776% | 9.481% | 21.86% | 3.15%
| PDM3 |  | -1.588% | 14.9% | -9.081% | -8.332% | 3.907% | 8.903% | 1.159% | -1.338% | 6.635% | 40.87% | -0.3641%
| Z-Score (Absolute) |  | 0.27 | NA | 1.31 | 1.21 | 0.49 | 1.18 | 0.11 | 0.24 | 0.86 | 5.6 | 0.11
Results from laboratory 5 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 84 (Certified Value 10.53 wt.%) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 18 | 19 | 20 | 21 | 23 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 9.96 | 9.95 | 10.4896 | 10.7 | 10.2057 | 12.0378 | 10.49 | 10.65 | 9.6136 | 10.65 | 10.3 | 10.949 | 10.25 | 10.64 | > 1 | 10.8 | 10.2557 | 10.4 | 10.1 | 10.718 |  | 1
| 1 | 2 | 10.6 | 10.05 | 10.2584 | 10.7 | 9.9427 | 12.1307 | 10.34 | 10.75 | 9.6613 | 10.6 | 10.4 | 11.1681 | 10.4 | 10.66 | > 1 | 10.7 | 10.5471 | 9.98 | 10.3 | 10.605 |  | #N/A
| 1 | 3 | 10.8 | 9.82 | 10.7218 | 10.75 | 10.1266 | 12.1122 | 10.22 | 10.6 | 8.6437 | 11.5 | 10.5 | 10.9511 | 10.65 | 10.52 | > 1 | 10.6 | 10.4981 | 10.9 | 10.3 | 10.764 |  | 16
| 1 | 4 | 10.1 | 10 | 10.5889 | 10.8 | 10.1629 | 12.1242 | 10.5 | 10.55 | 9.5494 | 10.3 | 10.3 | 10.7548 | 10.7 | 10.61 | > 1 | 10.6 | 10.5714 | 10.7 | 10.1 | 10.732 |  | 10.5288
| 1 | 5 | 10.5 | 10 | 10.6272 | 10.75 | 10.0243 | 12.0528 | 10.72 | 10.8 | 9.1405 | 10.7 | 10.4 | 10.9579 | 10.95 | 10.4 | > 1 | 10.6 | 10.4867 | 10.6 | 10.4 | 10.787 |  | 102
| 1 | 6 | 10.7 | 10.05 | 10.7338 | 10.75 | 9.9913 | 12.0677 | 10.45 | 10.7 | 9.3043 | 10.7 | 10.4 | 10.1985 | 10.9 | 10.4 | > 1 | 10.7 | 10.5385 | 10 | 10.2 | 10.631
| Mean |  | 10.4433 | 9.9783 | 10.57 | 10.7417 | 10.0756 | 12.0876 | 10.4533 | 10.675 | 9.3188 | 10.7417 | 10.3833 | 10.8299 | 10.6417 | 10.5383 |  | 10.6667 | 10.4829 | 10.43 | 10.2333 | 10.7062
| Median |  | 10.55 | 10 | 10.608 | 10.75 | 10.0755 | 12.0899 | 10.47 | 10.675 | 9.4268 | 10.675 | 10.4 | 10.9501 | 10.675 | 10.565 |  | 10.65 | 10.5183 | 10.5 | 10.25 | 10.725
| Std Dev. |  | 0.33833 | 0.08612 | 0.17724 | 0.037639 | 0.10445 | 0.039709 | 0.16848 | 0.093541 | 0.38604 | 0.40052 | 0.075277 | 0.33584 | 0.27462 | 0.11737 |  | 0.08165 | 0.11568 | 0.37709 | 0.12111 | 0.072898
| Rel.Std.Dev. |  | 3.24% | 0.8631% | 1.677% | 0.3504% | 1.037% | 0.3285% | 1.612% | 0.8763% | 4.143% | 3.729% | 0.725% | 3.101% | 2.581% | 1.114% |  | 0.7655% | 1.104% | 3.615% | 1.183% | 0.6809%
| PDM3 |  | -0.812% | -5.228% | 0.3906% | 2.022% | -4.305% | 14.8% | -0.717% | 1.388% | -11.49% | 2.022% | -1.382% | 2.86% | 1.072% | 0.09032% |  | 1.309% | -0.436% | -0.9386% | -2.806% | 1.684%
| Z-Score (Absolute) |  | 0.41 | 2.44 | 0.14 | 0.88 | 2.01 | 6.74 | 0.37 | 0.59 | 5.31 | 0.88 | 0.67 | 1.27 | 0.45 | 0 | NA | 0.56 | 0.24 | 0.47 | 1.33 | 0.73
| Analytical results for P2O5 in OREAS 84 (Certified Value 0.057 wt.%) | [22 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 20 | 23 | [9 empty] | wt.%
|  |  | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | PF*OES | [9 empty] | 3
3
| 1 | 1 | 0.069 | 0.05981 | 0.19934 | 0.06646 | 0.04698 | 0.046 | 0.046 | 0.05454 | 0.061976 | 0.34 | 0.092 | 0.039 | [9 empty] | 1
| 1 | 2 | 0.069 | 0.0621 | 0.19781 | 0.06944 | 0.05248 | 0.069 | 0.046 | 0.05339 | 0.056521 | 0.35 | 0.092 | 0.0458 | [9 empty] | #N/A
| 1 | 3 | 0.069 | 0.05958 | 0.19846 | 0.06646 | 0.0495 | 0.046 | 0.046 | 0.05339 | 0.061951 | 0.35 | 0.092 | 0.0756 | [9 empty] | 16
| 1 | 4 | 0.069 | 0.06004 | 0.19824 | 0.0621 | 0.0566 | 0.046 | 0.069 | 0.05385 | 0.059341 | 0.35 | 0.092 | 0.0756 | [9 empty] | 0.057341
| 1 | 5 | 0.069 | 0.05912 | 0.19737 | 0.06829 | 0.04446 | 0.046 | 0.046 | 0.05385 | 0.059511 | 0.36 | 0.092 | 0.0596 | [9 empty] | 103
| 1 | 6 | 0.069 | 0.06371 | 0.19706 | 0.06691 | 0.05454 | 0.046 | 0.046 | 0.05477 | 0.054977 | 0.35 | 0.092 | 0.0458
| Mean |  | 0.069 | 0.060727 | 0.19805 | 0.06661 | 0.05076 | 0.049833 | 0.049833 | 0.053965 | 0.059046 | 0.35 | 0.092 | 0.0569
| Median |  | 0.069 | 0.059925 | 0.19803 | 0.066685 | 0.05099 | 0.046 | 0.046 | 0.05385 | 0.059426 | 0.35 | 0.092 | 0.0527
| Std Dev. |  | 0 | 0.0017885 | 0.0008173 | 0.0025023 | 0.0046242 | 0.0093897 | 0.0093897 | 0.0005773 | 0.0028375 | 0.0063246 | 1.5202e-17 | 0.015959
| Rel.Std.Dev. |  | 0% | 2.945% | 0.4127% | 3.757% | 9.11% | 18.84% | 18.84% | 1.07% | 4.806% | 1.807% | 1.652e-14% | 28.05%
| PDM3 |  | 20.33% | 5.905% | 245.4% | 16.17% | -11.48% | -13.09% | -13.09% | -5.887% | 2.974% | 510.4% | 60.44% | -0.7688%
| Z-Score (Absolute) |  | 0.66 | 0.06 | 10.36 | 0.5 | 0.68 | 0.77 | 0.77 | 0.44 | 0.06 | 21.75 | 2.38 | 0.22
| Analytical results for Pb in OREAS 84 (Certified Value 74 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 18 | 19 | 20 | 23 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [5 empty] | 0
1
| 1 | 1 | 67.4 | < 100 | 71 | < 100 | 86 | 100 | 67 | < 100 | 70 | < 100 | 74 | 90.7 | 82.9 | 75.4232 | 72 | 174 | [5 empty] | 1
| 1 | 2 | 75.7 | 100 | 71 | 100 | 93 | 100 | 67 | < 100 | 51 | < 100 | 73 | 91.1 | 54.9 | 75.4987 | 73 | 134 | [5 empty] | 17
| 1 | 3 | 70.1 | 100 | 73 | 100 | 80 | 100 | 69 | < 100 | 60 | 100 | 72 | 89.7 | 68.7 | 72.9356 | 72 | 141 | [5 empty] | 16
| 1 | 4 | 66.7 | 100 | 71 | 100 | 86 | 100 | 67 | < 100 | 59 | < 100 | 75 | 90.6 | 74.5 | 74.8088 | 71 | 79 | [5 empty] | 73.7087
| 1 | 5 | 67.7 | 100 | 73 | 100 | 87 | 100 | 66 | < 100 | 56 | < 100 | 72 | 91.7 | 79.7 | 73.3315 | 77 | 110 | [5 empty] | 104
| 1 | 6 | 69 | < 100 | 70 | 100 | 89 | 100 | 65 | < 100 | 51 | < 100 | 74 | 91.2 | 85.8 | 76.9193 | 70 | 114
| Mean |  | 69.4333 | 100 | 71.5 | 100 | 86.8333 | 100 | 66.8333 |  | 57.8333 | 100 | 73.3333 | 90.8333 | 74.4167 | 74.8195 | 72.5 | 125.3333
| Median |  | 68.35 | 100 | 71 | 100 | 86.5 | 100 | 67 |  | 57.5 | 100 | 73.5 | 90.9 | 77.1 | 75.116 | 72 | 124
| Std Dev. |  | 3.3019 | 0 | 1.2247 | 0 | 4.2622 | 0 | 1.3292 |  | 7.0828 |  | 1.2111 | 0.6802 | 11.3406 | 1.4831 | 2.429 | 32.247
| Rel.Std.Dev. |  | 4.756% | 0% | 1.713% | 0% | 4.909% | 0% | 1.989% |  | 12.25% |  | 1.651% | 0.7488% | 15.24% | 1.982% | 3.35% | 25.73%
| PDM3 |  | -5.8% | 35.67% | -2.997% | 35.67% | 17.81% | 35.67% | -9.328% |  | -21.54% | 35.67% | -0.5092% | 23.23% | 0.9605% | 1.507% | -1.64% | 70.04%
| Z-Score (Absolute) |  | 0.49 | NA | 0.2 | NA | 1.96 | NA | 0.86 | 3.23 | 2.13 | NA | 0.06 | 2.53 | 0.21 | 0.27 | 0.06 | 7.39
Results from laboratories 2, 4, 8 and 13 were removed due to their 100 ppm reading resolution.
| Analytical results for Pr in OREAS 84 (Certified Value 1.81 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 1.6 | < 5 | 1.9236 | 1.95 | 1.61 | 1.85 | 1.82 | 1.81 | 1.82 | 1.9336 | 2.09 | 1.66 | [9 empty] | 1
| 1 | 2 | 1.8 | < 5 | 1.8156 | 1.6 | 1.66 | 1.92 | 2.1 | 1.82 | 1.79 | 1.9136 | 2.4 | 1.74 | [9 empty] | #N/A
| 1 | 3 | 1.7 | < 5 | 1.9066 | 1.78 | 1.68 | 1.93 | 1.83 | 1.8 | 1.76 | 1.9656 | 3.19 | 1.73 | [9 empty] | 16
| 1 | 4 | 1.6 | < 5 | 1.876 | 1.31 | 1.65 | 1.9 | 1.83 | 1.84 | 1.78 | 2.0924 | 1.92 | 1.68 | [9 empty] | 1.8068
| 1 | 5 | 1.7 | < 5 | 1.9546 | 2.09 | 1.63 | 1.87 | 1.87 | 1.78 | 1.81 | 2.0537 | 3.15 | 1.88 | [9 empty] | 105
| 1 | 6 | 1.7 | < 5 | 1.8511 | 1.67 | 1.62 | 1.94 | 1.71 | 1.78 | 1.74 | 1.8353 | 1.98 | 1.64
| Mean |  | 1.6833 |  | 1.8879 | 1.7333 | 1.6417 | 1.9017 | 1.86 | 1.805 | 1.7833 | 1.9657 | 2.455 | 1.7217
| Median |  | 1.7 |  | 1.8913 | 1.725 | 1.64 | 1.91 | 1.83 | 1.805 | 1.785 | 1.9496 | 2.245 | 1.705
| Std Dev. |  | 0.075277 |  | 0.050585 | 0.27471 | 0.026394 | 0.035449 | 0.12931 | 0.023452 | 0.030111 | 0.094368 | 0.57816 | 0.086814
| Rel.Std.Dev. |  | 4.472% |  | 2.679% | 15.85% | 1.608% | 1.864% | 6.952% | 1.299% | 1.688% | 4.801% | 23.55% | 5.042%
| PDM3 |  | -6.835% |  | 4.487% | -4.068% | -9.141% | 5.249% | 2.943% | -0.1011% | -1.3% | 8.793% | 35.87% | -4.713%
| Z-Score (Absolute) |  | 0.96 | 4.29 | 0.36 | 0.64 | 1.23 | 0.44 | 0.18 | 0.18 | 0.32 | 0.86 | 4 | 0.71
| Analytical results for Rb in OREAS 84 (Certified Value 22 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 06 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [10 empty] | 1
2
| 1 | 1 | 17.1 | 22 | 21.1959 | 18 | 23.3 | 21.5 | 21.1 | 21.3101 | 24 | 23.2848 | 20 | [10 empty] | 1
| 1 | 2 | 18 | 21.9 | 21.0944 | 18 | 22 | 21.5 | 21.7 | 22.0242 | 24.1 | 23.0702 | 19 | [10 empty] | 18
| 1 | 3 | 18.6 | 22 | 21.1945 | 19 | 21.9 | 21.3 | 21 | 23.8185 | 23.8 | 21.9174 | 20 | [10 empty] | 16
| 1 | 4 | 17.3 | 22.1 | 21.0543 | 18 | 21.8 | 21.8 | 20.9 | 23.2649 | 24.1 | 22.2801 | 19 | [10 empty] | 21.9991
| 1 | 5 | 17.6 | 21.7 | 20.8782 | 18 | 22.3 | 21.6 | 20.1 | 22.4604 | 23.3 | 23.0512 | 21 | [10 empty] | 106
| 1 | 6 | 18.2 | 21.9 | 21.1827 | 18 | 21.8 | 20.8 | 20 | 20.5972 | 23 | 21.9767 | 19
| Mean |  | 17.8 | 21.9333 | 21.1 | 18.1667 | 22.1833 | 21.4167 | 20.8 | 22.2459 | 23.7167 | 22.5967 | 19.6667
| Median |  | 17.8 | 21.95 | 21.1385 | 18 | 21.95 | 21.5 | 20.95 | 22.2423 | 23.9 | 22.6656 | 19.5
| Std Dev. |  | 0.56921 | 0.13663 | 0.12353 | 0.40825 | 0.57764 | 0.34303 | 0.64498 | 1.2 | 0.46224 | 0.60832 | 0.8165
| Rel.Std.Dev. |  | 3.198% | 0.6229% | 0.5854% | 2.247% | 2.604% | 1.602% | 3.101% | 5.394% | 1.949% | 2.692% | 4.152%
| PDM3 |  | -19.09% | -0.2988% | -4.087% | -17.42% | 0.8376% | -2.647% | -5.451% | 1.122% | 7.808% | 2.717% | -10.6%
| Z-Score (Absolute) |  | 4.2 | 0 | 0.85 | NA | 0.25 | 0.53 | 1.15 | 0.32 | 1.81 | 0.67 | NA
Results from laboratories 9 and 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Re in OREAS 84 (Indicative Value < 0.1 ppm) | [22 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 03 | 12 | 16 | 17 | 19 | [16 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | [16 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.05 | 0.031 | < 0.01 | < 0.1 | [16 empty] | 1
| 1 | 2 | < 0.1 | < 0.05 | 0.033 | < 0.01 | < 0.1 | [16 empty] | 14
| 1 | 3 | < 0.1 | < 0.05 | 0.032 | < 0.01 | < 0.1 | [16 empty] | 16
| 1 | 4 | < 0.1 | < 0.05 | 0.027 | < 0.01 | < 0.1 | [16 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.05 | 0.03 | < 0.01 | < 0.1 | [16 empty] | 20
| 1 | 6 | < 0.1 | < 0.05 | 0.029 | < 0.01 | < 0.1
| Mean |  |  |  | 0.030333
| Median |  |  |  | 0.0305
| Std Dev. |  |  |  | 0.0021602
| Rel.Std.Dev. |  |  |  | 7.122%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.18 | 0 | 0.87 | 0.67
| Analytical results for S in OREAS 84 (Certified Value 24.08 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 16 | 19 | 20 | [6 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 22.3 | 23.8 | 23.76 | 24.9 | > 20 | 28.7818 | 24.4 | 24.5 | 24.3 | 24.2 | 24.3 | 24.5 | > 10 | 23.6566 | 23.8 | [6 empty] | 1
| 1 | 2 | 24 | 24.4 | 23.19 | 24.6 | > 20 | 28.8036 | 24.32 | 24.2 | 24.4 | 24.7 | 24.5 | 24.7 | > 10 | 24.284 | 23.5 | [6 empty] | #N/A
| 1 | 3 | 24 | 23.2 | 23.97 | 24.9 | > 20 | 28.6922 | 24.41 | 24.3 | 26.2 | 23.3 | 23.9 | 25.2 | > 10 | 22.1168 | 23.9 | [6 empty] | 16
| 1 | 4 | 22.2 | 23.4 | 24.05 | 25.1 | > 20 | 28.5589 | 24.46 | 24.6 | 24.5 | 23.5 | 24.6 | 25.5 | > 10 | 22.5159 | 23.7 | [6 empty] | 24.0762
| 1 | 5 | 23.5 | 23.9 | 23.67 | 24.7 | > 20 | 28.4394 | 24.4 | 24.8 | 24.8 | 23 | 24.3 | 25.6 | > 10 | 23.3887 | 24.2 | [6 empty] | 107
| 1 | 6 | 23.2 | 24.1 | 23.89 | 24.7 | > 20 | 28.6084 | 24.43 | 24.4 | 24.4 | 23.8 | 23.4 | 25 | > 10 | 21.3355 | 23.7
| Mean |  | 23.2 | 23.8 | 23.755 | 24.8167 |  | 28.6474 | 24.4033 | 24.4667 | 24.7667 | 23.75 | 24.1667 | 25.0833 |  | 22.8829 | 23.8
| Median |  | 23.35 | 23.85 | 23.825 | 24.8 |  | 28.6503 | 24.405 | 24.45 | 24.45 | 23.65 | 24.3 | 25.1 |  | 22.9523 | 23.75
| Std Dev. |  | 0.7975 | 0.44272 | 0.30918 | 0.18348 |  | 0.13936 | 0.046762 | 0.21602 | 0.72296 | 0.62209 | 0.44572 | 0.43551 |  | 1.0894 | 0.23664
| Rel.Std.Dev. |  | 3.437% | 1.86% | 1.302% | 0.7394% |  | 0.4865% | 0.1916% | 0.8829% | 2.919% | 2.619% | 1.844% | 1.736% |  | 4.761% | 0.9943%
| PDM3 |  | -3.639% | -1.147% | -1.334% | 3.076% |  | 18.99% | 1.359% | 1.622% | 2.868% | -1.355% | 0.3759% | 4.183% |  | -4.956% | -1.147%
| Z-Score (Absolute) |  | 1.56 | 0.59 | 0.67 | 1.05 | NA | 7.25 | 0.38 | 0.49 | 0.97 | 0.67 | 0 | 1.48 | NA | 2.08 | 0.59
| Analytical results for Sb in OREAS 84 (Certified Value 1.55 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 07 | 09 | 12 | 15 | 16 | 18 | 19 | 20 | 23 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [10 empty] | 2
2
| 1 | 1 | < 2 | 1.4 | 12 | 1.2 | 1.8 | 1.5 | 1.48 | 1 | 2.1198 | 3 | 84 | [10 empty] | 1
| 1 | 2 | 2 | 1.2 | 12 | 1.4 | 1.9 | 1.5 | 1.4 | 1 | 1.9937 | 3 | 88 | [10 empty] | 19
| 1 | 3 | < 2 | 1.3 | 12 | 1.4 | 2.1 | 1.5 | 1.46 | < 1 | 1.4417 | 2 | 63 | [10 empty] | 16
| 1 | 4 | < 2 | 1.5 | 13 | 1.4 | 2 | 1.5 | 1.44 | 1 | 1.5982 | 2 | 61 | [10 empty] | 1.551
| 1 | 5 | < 2 | 1.6 | 12 | 1.2 | 1.9 | 1.5 | 1.38 | < 1 | 1.7176 | 3 | 84 | [10 empty] | 108
| 1 | 6 | < 2 | 1.4 | 13 | 1.3 | 1.8 | 1.5 | 1.39 | 1 | 1.6157 | 2 | 72
| Mean |  | 2 | 1.4 | 12.3333 | 1.3167 | 1.9167 | 1.5 | 1.425 | 1 | 1.7478 | 2.5 | 75.3333
| Median |  | 2 | 1.4 | 12 | 1.35 | 1.9 | 1.5 | 1.42 | 1 | 1.6666 | 2.5 | 78
| Std Dev. |  |  | 0.14142 | 0.5164 | 0.098319 | 0.1169 | 0 | 0.040866 | 0 | 0.25817 | 0.54772 | 11.6562
| Rel.Std.Dev. |  |  | 10.1% | 4.187% | 7.467% | 6.099% | 0% | 2.868% | 0% | 14.77% | 21.91% | 15.47%
| PDM3 |  | 28.95% | -9.737% | 695.2% | -15.11% | 23.57% | -3.289% | -8.125% | -35.53% | 12.69% | 61.18% | 4757%
| Z-Score (Absolute) |  | NA | 0.37 | 39.85 | 0.67 | 1.53 | 0 | 0.28 | NA | 0.91 | NA | NA
Results from laboratories 1, 18, 20 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 84 (Certified Value 5.01 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [9 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | [9 empty] | 2
3
| 1 | 1 | < 10 | 5 | 6 | 5 | < 5 | 5 | 5.7 | 4.9398 | 5 | < 20 | < 10 | < 5 | [9 empty] | 1
| 1 | 2 | < 10 | 5 | 5 | 5 | 5 | 5 | 5.8 | 5.0566 | 5 | < 20 | < 10 | < 5 | [9 empty] | 20
| 1 | 3 | < 10 | 5 | 6 | < 5 | < 5 | 5 | 5.7 | 4.8399 | 5 | < 20 | < 10 | < 5 | [9 empty] | 16
| 1 | 4 | < 10 | 5 | 6 | 5 | < 5 | 5 | 5.7 | 5.4192 | 5 | < 20 | < 10 | < 5 | [9 empty] | 5.0133
| 1 | 5 | < 10 | 5 | 6 | 5 | < 5 | 5 | 5.7 | 5.2775 | 5 | < 20 | < 10 | < 5 | [9 empty] | 109
| 1 | 6 | < 10 | 5 | 5 | 6 | < 5 | 5 | 5.6 | 4.9466 | 5 | < 20 | < 10 | < 5
| Mean |  |  | 5 | 5.6667 | 5.2 | 5 | 5 | 5.7 | 5.08 | 5
| Median |  |  | 5 | 6 | 5 | 5 | 5 | 5.7 | 5.0016 | 5
| Std Dev. |  |  | 0 | 0.5164 | 0.44721 |  | 0 | 0.063246 | 0.22349 | 0
| Rel.Std.Dev. |  |  | 0% | 9.113% | 8.6% |  | 0% | 1.11% | 4.399% | 0%
| PDM3 |  |  | -0.2658% | 13.03% | 3.724% | -0.2658% | -0.2658% | 13.7% | 1.329% | -0.2658%
| Z-Score (Absolute) |  | 0 | 0 | 2.72 | 1.02 | 8.52 | 0 | 2.86 | 0.33 | 0 | 20.44 | 0 | 10.22
| Analytical results for Se in OREAS 84 (Certified Value < 20 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 03 | 12 | 16 | 18 | 19 | 20 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [14 empty] | 1
1
| 1 | 1 | 10 | < 20 | < 50 | 11 | 57 | < 20 | 10 | [14 empty] | 1
| 1 | 2 | 11 | < 20 | < 50 | 11 | 35 | < 20 | 11 | [14 empty] | #N/A
| 1 | 3 | 10 | < 20 | < 50 | 11 | 27 | < 20 | 10 | [14 empty] | 16
| 1 | 4 | 12 | < 20 | < 50 | 10 | 49 | < 20 | 11 | [14 empty] | < 20
| 1 | 5 | 10 | < 20 | < 50 | 10 | 14 | < 20 | 9 | [14 empty] | 110
| 1 | 6 | 9 | < 20 | < 50 | 10 | 44 | < 20 | 10
| Mean |  | 10.3333 |  |  | 10.5 | 37.6667 |  | 10.1667
| Median |  | 10 |  |  | 10.5 | 39.5 |  | 10
| Std Dev. |  | 1.0328 |  |  | 0.54772 | 15.6418 |  | 0.75277
| Rel.Std.Dev. |  | 9.995% |  |  | 5.216% | 41.53% |  | 7.404%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 29.67 | 0.34 | 55.29 | 0.67 | 0.34
| Analytical results for SiO2 in OREAS 84 (Certified Value 16.27 wt.%) | [22 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 08 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES |  |  |  | 2
3
| 1 | 1 | 14.29 | 16.2 | 16.26 | 17.1 | 18.7221 | 16.7 | 17.8 | 15.62 | 16 | 16.47 | 15.78 | 17.4377 | 16 | 14.2046 | 15.19 | 15.552 | 16.591 | 15.6333 |  |  |  | 1
| 1 | 2 | 15.146 | 16.2 | 15.83 | 17.4 | 18.7473 | 16.9 | 17.4 | 15.62 | 16 | 16.69 | 15.81 | 16.5189 | 16.6 | 14.7528 | 14.97 | 15.338 | 16.408 | 17.1286 |  |  |  | #N/A
| 1 | 3 | 15.466 | 16 | 16.69 | 17.4 | 18.6939 | 16.4 | 18.8 | 14.97 | 16.7 | 16.47 | 15.62 | 16.4597 | 16.5 | 13.4953 | 15.19 | 15.402 |  | 16.2643 |  |  |  | 16
| 1 | 4 | 14.418 | 16.4 | 16.47 | 17.1 | 18.7445 | 16.4 | 16.9 | 14.55 | 16.7 | 16.69 | 15.8 | 16.1772 | 16.6 | 13.7252 | 14.97 | 15.424 |  | 15.3103 |  |  |  | 16.2663
| 1 | 5 | 15.231 | 16.4 | 16.26 | 16.9 | 18.8021 | 16.7 | 17.6 | 14.12 | 17.1 | 16.26 | 15.82 | 16.0706 | 16.6 | 14.4069 | 15.4 | 15.424 |  | 16.1125 |  |  |  | 111
| 1 | 6 | 15.188 | 16.7 | 16.47 | 17.1 | 18.7494 | 16.7 | 16.9 | 14.12 | 16.9 | 16.47 | 15.98 | 16.1964 | 17.1 | 14.6427 | 14.97 | 15.466 |  | 17.3617
| Mean |  | 14.9565 | 16.3167 | 16.33 | 17.1667 | 18.7432 | 16.6333 | 17.5667 | 14.8333 | 16.5667 | 16.5083 | 15.8017 | 16.4768 | 16.5667 | 14.2046 | 15.115 | 15.4343 | 16.4995 | 16.3018
| Median |  | 15.167 | 16.3 | 16.365 | 17.1 | 18.7459 | 16.7 | 17.5 | 14.76 | 16.7 | 16.47 | 15.805 | 16.3281 | 16.6 | 14.3058 | 15.08 | 15.424 | 16.4995 | 16.1884
| Std Dev. |  | 0.48138 | 0.24014 | 0.29278 | 0.19664 | 0.035775 | 0.19664 | 0.70616 | 0.6862 | 0.46332 | 0.16253 | 0.11462 | 0.50174 | 0.35024 | 0.50326 | 0.17638 | 0.071256 | 0.1294 | 0.80922
| Rel.Std.Dev. |  | 3.219% | 1.472% | 1.793% | 1.145% | 0.1909% | 1.182% | 4.02% | 4.626% | 2.797% | 0.9845% | 0.7253% | 3.045% | 2.114% | 3.543% | 1.167% | 0.4617% | 0.7843% | 4.964%
| PDM3 |  | -8.052% | 0.3097% | 0.3917% | 5.535% | 15.23% | 2.256% | 7.994% | -8.809% | 1.847% | 1.488% | -2.856% | 1.294% | 1.847% | -12.67% | -7.078% | -5.115% | 1.434% | 0.2182%
| Z-Score (Absolute) |  | 2.35 | 0.14 | 0.12 | 1.24 | 3.79 | 0.37 | 1.89 | 2.55 | 0.27 | 0.17 | 0.98 | 0.12 | 0.27 | 3.56 | 2.09 | 1.57 | 0.16 | 0.16
| Analytical results for Sm in OREAS 84 (Certified Value 1.26 ppm) | [22 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 1.1 | < 2 | 1.2554 | 1.47 | 1.1 | 1.3 | 1.4 | 1.2 | 1.27 | 1.3011 | 1.09 | 1.3 | [9 empty] | 1
| 1 | 2 | 1.2 | < 2 | 1.2635 | 1.14 | 1.2 | 1.3 | 1.5 | 1.2 | 1.31 | 1.2791 | 1.94 | 1.4 | [9 empty] | #N/A
| 1 | 3 | 1.2 | < 2 | 1.2255 | 1.31 | 1 | 1.3 | 1.3 | 1.2 | 1.25 | 1.3337 | 1.49 | 1.3 | [9 empty] | 16
| 1 | 4 | 1.2 | < 2 | 1.2339 | 1.87 | 1.1 | 1.4 | 1.3 | 1.3 | 1.3 | 1.4346 | 0.95 | 1.2 | [9 empty] | 1.2625
| 1 | 5 | 1.3 | < 2 | 1.2607 | 1.62 | 1.1 | 1.3 | 1.4 | 1.3 | 1.24 | 1.2754 | 2.51 | 1.3 | [9 empty] | 112
| 1 | 6 | 1.2 | < 2 | 1.2808 | 1.6 | 1.2 | 1.3 | 1.3 | 1.2 | 1.26 | 1.2998 | 1.51 | 1.2
| Mean |  | 1.2 |  | 1.2533 | 1.5017 | 1.1167 | 1.3167 | 1.3667 | 1.2333 | 1.2717 | 1.3206 | 1.5817 | 1.2833
| Median |  | 1.2 |  | 1.2581 | 1.535 | 1.1 | 1.3 | 1.35 | 1.2 | 1.265 | 1.3005 | 1.5 | 1.3
| Std Dev. |  | 0.063246 |  | 0.020333 | 0.25608 | 0.075277 | 0.040825 | 0.08165 | 0.05164 | 0.027869 | 0.059569 | 0.5735 | 0.075277
| Rel.Std.Dev. |  | 5.27% |  | 1.622% | 17.05% | 6.741% | 3.101% | 5.974% | 4.187% | 2.192% | 4.511% | 36.26% | 5.866%
| PDM3 |  | -4.948% |  | -0.7265% | 18.95% | -11.55% | 4.293% | 8.253% | -2.308% | 0.7284% | 4.605% | 25.28% | 1.653%
| Z-Score (Absolute) |  | 0.86 | 3.08 | 0.27 | 2.48 | 1.78 | 0.43 | 0.99 | 0.49 | 0.06 | 0.48 | 3.37 | 0.06
| Analytical results for Sn in OREAS 84 (Certified Value 1.4 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.2 | < 100 | < 20 | 16 | < 5 | < 10 | 2 | 1.4 | 1.4555 | 1 | < 100 | 3 | [9 empty] | 1
| 1 | 2 | 1.2 | < 100 | < 20 | 13 | < 5 | < 10 | 2 | 1.4 | 1.3033 | 1 | < 100 | 3 | [9 empty] | 21
| 1 | 3 | 1.2 | < 100 | < 20 | 15 | 36 | < 10 | 2 | 1.4 | 1.4062 | 0.8 | < 100 | 3 | [9 empty] | 16
| 1 | 4 | 1.2 | < 100 | < 20 | 15 | < 5 | < 10 | 2 | 1.4 | 1.5386 | 1.9 | < 100 | 2 | [9 empty] | 1.4046
| 1 | 5 | 1.3 | < 100 | < 20 | 14 | < 5 | < 10 | 2 | 1.3 | 1.3304 | 1.3 | < 100 | 2 | [9 empty] | 113
| 1 | 6 | 1.2 | < 100 | < 20 | 14 | < 5 | < 10 | 2 | 1.4 | 1.2635 | 1.1 | < 100 | 3
| Mean |  | 1.2167 |  |  | 14.5 | 36 |  | 2 | 1.3833 | 1.3829 | 1.1833 |  | 2.6667
| Median |  | 1.2 |  |  | 14.5 | 36 |  | 2 | 1.4 | 1.3683 | 1.05 |  | 3
| Std Dev. |  | 0.040825 |  |  | 1.0488 |  |  | 0 | 0.040825 | 0.10335 | 0.38687 |  | 0.5164
| Rel.Std.Dev. |  | 3.355% |  |  | 7.233% |  |  | 0% | 2.951% | 7.473% | 32.69% |  | 19.36%
| PDM3 |  | -13.38% |  |  | 932.3% | 2463% |  | 42.39% | -1.513% | -1.542% | -15.75% |  | 89.85%
| Z-Score (Absolute) |  | 0.67 | 11.82 | 1.58 | 2.73 | 1.09 | 0.3 | 0.47 | 0.63 | 0.63 | 0.68 | 11.82 | 0.3
| Analytical results for Sr in OREAS 84 (Certified Value 89 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [7 empty] | 0
1
| 1 | 1 | 81 | 89 | 67 | 91.0028 | 93 | 78 | 91.5 | 87 | 89.3 | 92.0988 | 89 | 89.7007 | 83 | 83 | [7 empty] | 1
| 1 | 2 | 86 | 89 | 68 | 94.2652 | 95 | 78 | 90.9 | 88 | 91 | 86.7091 | 92 | 92.2223 | 82 | 90 | [7 empty] | 16
| 1 | 3 | 86 | 91 | 73 | 90.8019 | 93 | 75 | 91.5 | 87 | 87.9 | 96.0938 | 92 | 87.1399 | 84 | 85 | [7 empty] | 16
| 1 | 4 | 81 | 90 | 68 | 92.1094 | 96 | 75 | 89.5 | 88 | 88.8 | 94.545 | 91 | 90.5877 | 83 | 81 | [7 empty] | 88.926
| 1 | 5 | 85 | 89 | 70 | 90.9255 | 95 | 76 | 91.3 | 86 | 89.1 | 87.4249 | 93 | 92.0183 | 84 | 85 | [7 empty] | 114
| 1 | 6 | 86 | 86 | 65 | 91.6572 | 91 | 76 | 88.4 | 87 | 88.4 | 90.3144 | 94 | 88.4228 | 83 | 91
| Mean |  | 84.1667 | 89 | 68.5 | 91.7937 | 93.8333 | 76.3333 | 90.5167 | 87.1667 | 89.0833 | 91.1977 | 91.8333 | 90.0153 | 83.1667 | 85.8333
| Median |  | 85.5 | 89 | 68 | 91.33 | 94 | 76 | 91.1 | 87 | 88.95 | 91.2066 | 92 | 90.1442 | 83 | 85
| Std Dev. |  | 2.4833 | 1.6733 | 2.7386 | 1.3107 | 1.8348 | 1.3663 | 1.2813 | 0.75277 | 1.0647 | 3.7723 | 1.7224 | 2.0057 | 0.75277 | 3.92
| Rel.Std.Dev. |  | 2.95% | 1.88% | 3.998% | 1.428% | 1.955% | 1.79% | 1.416% | 0.8636% | 1.195% | 4.136% | 1.876% | 2.228% | 0.9051% | 4.567%
| PDM3 |  | -5.352% | 0.08318% | -22.97% | 3.225% | 5.518% | -14.16% | 1.789% | -1.978% | 0.1769% | 2.555% | 3.269% | 1.225% | -6.477% | -3.478%
| Z-Score (Absolute) |  | 1.19 | 0.01 | 5 | 0.67 | 1.17 | 3.09 | 0.36 | 0.46 | 0.01 | 0.52 | 0.68 | 0.24 | 1.43 | 0.78
| Analytical results for Ta in OREAS 84 (Indicative Value 0.28 ppm) | [22 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.2 | 0.3 | < 10 | < 5 | < 10 | 2.2 | < 0.5 | 0.24 | < 0.2 | 0.04 | 0.61113 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.2 | 0.3 | < 10 | < 5 | < 10 | < 0.5 | < 0.5 | 0.21 | < 0.2 | < 0.01 | 0.58814 | < 0.5 | [9 empty] | 10
| 1 | 3 | < 0.2 | 0.3 | < 10 | < 5 | < 10 | 0.6 | < 0.5 | 0.22 | < 0.2 | 0.18 | 0.38842 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.2 | 0.4 | < 10 | < 5 | < 10 | < 0.5 | < 0.5 | 0.2 | < 0.2 | 0.11 | 0.32856 | < 0.5 | [9 empty] | 0.27894
| 1 | 5 | < 0.2 | 0.3 | < 10 | < 5 | < 10 | < 0.5 | < 0.5 | 0.22 | < 0.2 | < 0.01 | 0.60996 | < 0.5 | [9 empty] | 16
| 1 | 6 | < 0.2 | 0.2 | < 10 | < 5 | < 10 | < 0.5 | < 0.5 | 0.21 | < 0.2 | 0.02 | 0.46205 | < 0.5
| Mean |  |  | 0.3 |  |  |  | 1.4 |  | 0.21667 |  | 0.0875 | 0.49804
| Median |  |  | 0.3 |  |  |  | 1.4 |  | 0.215 |  | 0.075 | 0.5251
| Std Dev. |  |  | 0.063246 |  |  |  | 1.1314 |  | 0.013663 |  | 0.072744 | 0.12286
| Rel.Std.Dev. |  |  | 21.08% |  |  |  | 80.81% |  | 6.306% |  | 83.14% | 24.67%
| PDM3 |  |  | 7.549% |  |  |  | 401.9% |  | -22.33% |  | -68.63% | 78.55%
| Z-Score (Absolute) |  | 0.61 | 0.09 | 16.34 | 7.69 | 16.34 | 1.24 | 0.09 | 0.2 | 0.61 | 0.74 | 0.77 | 0.09
| Analytical results for Tb in OREAS 84 (Certified Value 0.16 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.1 | < 3 | 0.14339 | 0.1 | 0.15 | 0.18 | 0.16 | 0.17 | 0.17 | < 0.3 | 0.21 | 0.15 | [9 empty] | 1
| 1 | 2 | 0.2 | < 3 | 0.15087 | 0.14 | 0.14 | 0.18 | 0.19 | 0.16 | 0.16 | < 0.3 | 0.2 | 0.17 | [9 empty] | #N/A
| 1 | 3 | 0.2 | < 3 | 0.14017 | 0.09 | 0.16 | 0.17 | 0.18 | 0.17 | 0.15 | < 0.3 | 0.22 | 0.17 | [9 empty] | 16
| 1 | 4 | 0.1 | < 3 | 0.14053 | 0.13 | 0.14 | 0.17 | 0.17 | 0.16 | 0.16 | < 0.3 | 0.12 | 0.14 | [9 empty] | 0.16228
| 1 | 5 | 0.2 | < 3 | 0.14423 | 0.1 | 0.15 | 0.18 | 0.17 | 0.16 | 0.16 | < 0.3 | 0.12 | 0.17 | [9 empty] | 115
| 1 | 6 | 0.2 | < 3 | 0.15035 | 0.24 | 0.13 | 0.17 | 0.16 | 0.16 | 0.16 | < 0.3 | 0.19 | 0.17
| Mean |  | 0.16667 |  | 0.14492 | 0.13333 | 0.145 | 0.175 | 0.17167 | 0.16333 | 0.16 |  | 0.17667 | 0.16167
| Median |  | 0.2 |  | 0.14381 | 0.115 | 0.145 | 0.175 | 0.17 | 0.16 | 0.16 |  | 0.195 | 0.17
| Std Dev. |  | 0.05164 |  | 0.00468 | 0.055737 | 0.010488 | 0.0054772 | 0.01169 | 0.005164 | 0.0063246 |  | 0.045019 | 0.013292
| Rel.Std.Dev. |  | 30.98% |  | 3.229% | 41.8% | 7.233% | 3.13% | 6.81% | 3.162% | 3.953% |  | 25.48% | 8.222%
| PDM3 |  | 2.702% |  | -10.7% | -17.84% | -10.65% | 7.837% | 5.783% | 0.648% | -1.406% |  | 8.864% | -0.3791%
| Z-Score (Absolute) |  | NA | 67.68 | 0.85 | 1.43 | 0.84 | 0.67 | 0.51 | 0.08 | 0.08 | 0.59 | 0.76 | 0
Results from laboratory 1 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 84 (Indicative Value 1.05 ppm) | [22 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 07 | 12 | 15 | 16 | 18 | 19 | 20 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 6 | < 2 | < 5 | 1 | 1 | 1.18 | 0.3 | < 1 | < 5 | [12 empty] | 1
| 1 | 2 | < 6 | < 2 | < 5 | 1 | 2 | 1.2 | 0.8 | < 1 | < 5 | [12 empty] | 11
| 1 | 3 | < 6 | < 2 | < 5 | 1 | 1 | 1.19 | 0.6 | < 1 | < 5 | [12 empty] | 16
| 1 | 4 | < 6 | < 2 | < 5 | 2 | 2 | 1.21 | 0.4 | < 1 | < 5 | [12 empty] | 1.0475
| 1 | 5 | < 6 | < 2 | < 5 | 1 | 1 | 1.19 | 0.6 | < 1 | < 5 | [12 empty] | 17
| 1 | 6 | < 6 | < 2 | < 5 | 1 | 1 | 1.17 | 0.3 | < 1 | < 5
| Mean |  |  |  |  | 1.1667 | 1.3333 | 1.19 | 0.5
| Median |  |  |  |  | 1 | 1 | 1.19 | 0.5
| Std Dev. |  |  |  |  | 0.40825 | 0.5164 | 0.014142 | 0.2
| Rel.Std.Dev. |  |  |  |  | 34.99% | 38.73% | 1.188% | 40%
| PDM3 |  |  |  |  | 11.38% | 27.29% | 13.6% | -52.27%
| Z-Score (Absolute) |  | 1.77 | 0.19 | 1.28 | 0.02 | 0.14 | 0 | 0.67 | 0.67 | 1.28
| Analytical results for Th in OREAS 84 (Certified Value 3.04 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 2.7 | 3.3 | < 50 | 5.8 | 2.8 | 3.5 | 2.9 | 3.05 | 2.9195 | 3.71 | 3.1453 | 3 | [9 empty] | 1
| 1 | 2 | 2.9 | 3.2 | < 50 | 5.61 | 2.7 | 3.7 | 2.9 | 3.01 | 3.1224 | 3.19 | 2.7789 | 2.9 | [9 empty] | #N/A
| 1 | 3 | 2.8 | 3.2 | < 50 | 5.91 | 3 | 3.3 | 2.9 | 2.93 | 3.2741 | 4.01 | 2.8617 | 2.9 | [9 empty] | 16
| 1 | 4 | 2.6 | 3.4 | < 50 | 5.57 | 2.6 | 3.4 | 3 | 2.92 | 3.043 | 3.06 | 2.9257 | 2.8 | [9 empty] | 3.0386
| 1 | 5 | 2.8 | 3.1 | < 50 | 5.31 | 2.5 | 3.4 | 3 | 2.88 | 2.9468 | 3.67 | 3.1153 | 3 | [9 empty] | 116
| 1 | 6 | 2.8 | 3.2 | < 50 | 5.11 | 2.4 | 3.2 | 2.9 | 3.03 | 3.0167 | 3.74 | 3.0025 | 2.8
| Mean |  | 2.7667 | 3.2333 |  | 5.5517 | 2.6667 | 3.4167 | 2.9333 | 2.97 | 3.0538 | 3.5633 | 2.9716 | 2.9
| Median |  | 2.8 | 3.2 |  | 5.59 | 2.65 | 3.4 | 2.9 | 2.97 | 3.0298 | 3.69 | 2.9641 | 2.9
| Std Dev. |  | 0.10328 | 0.10328 |  | 0.29909 | 0.21602 | 0.17224 | 0.05164 | 0.068993 | 0.12978 | 0.36231 | 0.14362 | 0.089443
| Rel.Std.Dev. |  | 3.733% | 3.194% |  | 5.387% | 8.101% | 5.041% | 1.76% | 2.323% | 4.25% | 10.17% | 4.833% | 3.084%
| PDM3 |  | -8.949% | 6.409% |  | 82.7% | -12.24% | 12.44% | -3.464% | -2.258% | 0.4987% | 17.27% | -2.206% | -4.561%
| Z-Score (Absolute) |  | 0.71 | 0.64 | 63.54 | 7.34 | 1 | 1.17 | 0.23 | 0.12 | 0.12 | 1.59 | 0.12 | 0.33
| Analytical results for TiO2 in OREAS 84 (Certified Value 0.2 wt.%) | [22 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.183 | 0.25 | 0.2 | 0.2 | 0.20868 | 0.18225 | 0.2 | 0.2 | 0.17348 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2018 | 0.20256 | 0.19 | 0.20855 | 0.167 | 0.2 | 0.1752 |  | 1
| 1 | 2 | 0.2 | 0.28 | 0.183 | 0.2 | 0.20768 | 0.18231 | 0.2 | 0.21 | 0.16531 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2002 | 0.19982 | 0.2 | 0.19774 | 0.167 | 0.2 | 0.2002 |  | #N/A
| 1 | 3 | 0.2 | 0.13 | 0.2 | 0.2 | 0.20901 | 0.18247 | 0.2 | 0.2 | 0.16681 | 0.21 | 0.2 | 0.21 | 0.2 | 0.1985 | 0.19674 | 0.2 | 0.19564 | 0.167 | 0.2 | 0.1918 |  | 16
| 1 | 4 | 0.183 | 0.13 | 0.2 | 0.2 | 0.21602 | 0.18221 | 0.2 | 0.2 | 0.16848 | 0.2 | 0.2 | 0.21 | 0.2 | 0.2002 | 0.19569 | 0.2 | 0.20222 | 0.167 | 0.2 | 0.1818 |  | 0.20028
| 1 | 5 | 0.2 | 0.13 | 0.2 | 0.2 | 0.20768 | 0.18203 | 0.2 | 0.21 | 0.15997 | 0.2 | 0.2 | 0.21 | 0.2 | 0.1968 | 0.19846 | 0.2 | 0.20727 | 0.167 | 0.2 | 0.1902 |  | 117
| 1 | 6 | 0.2 | 0.13 | 0.2 | 0.2 | 0.21135 | 0.18183 | 0.183 | 0.2 | 0.17115 | 0.2 | 0.2 | 0.21 | 0.2 | 0.2002 | 0.19714 | 0.21 | 0.20618 | 0.15 | 0.2 | 0.2002
| Mean |  | 0.19433 | 0.175 | 0.19717 | 0.2 | 0.21007 | 0.18218 | 0.19717 | 0.20333 | 0.16753 | 0.20167 | 0.2 | 0.20667 | 0.2 | 0.19962 | 0.1984 | 0.2 | 0.20293 | 0.16417 | 0.2 | 0.1899
| Median |  | 0.2 | 0.13 | 0.2 | 0.2 | 0.20885 | 0.18223 | 0.2 | 0.2 | 0.16764 | 0.2 | 0.2 | 0.21 | 0.2 | 0.2002 | 0.1978 | 0.2 | 0.2042 | 0.167 | 0.2 | 0.191
| Std Dev. |  | 0.0087788 | 0.070356 | 0.0069402 | 3.0405e-17 | 0.00321 | 0.00022437 | 0.0069402 | 0.005164 | 0.0047331 | 0.0040825 | 3.0405e-17 | 0.005164 | 3.0405e-17 | 0.0017302 | 0.0024928 | 0.0063246 | 0.0053221 | 0.0069402 | 3.0405e-17 | 0.0099758
| Rel.Std.Dev. |  | 4.517% | 40.2% | 3.52% | 1.52e-14% | 1.528% | 0.1232% | 3.52% | 2.54% | 2.825% | 2.024% | 1.52e-14% | 2.499% | 1.52e-14% | 0.8668% | 1.256% | 3.162% | 2.623% | 4.228% | 1.52e-14% | 5.253%
| PDM3 |  | -2.97% | -12.62% | -1.556% | -0.1411% | 4.887% | -9.037% | -1.556% | 1.523% | -16.35% | 0.6911% | -0.1411% | 3.188% | -0.1411% | -0.3325% | -0.9395% | -0.1411% | 1.323% | -18.03% | -0.1411% | -5.184%
| Z-Score (Absolute) |  | 1.02 | 5 | 0.45 | NA | 2.13 | 3.54 | 0.45 | 0.76 | 6.52 | 0.42 | 0.11 | 1.43 | 0.11 | 0 | 0.25 | 0.08 | 0.67 | 7.24 | 0.11 | 1.98
Results from laboratory 4 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Tl in OREAS 84 (Certified Value 1.2 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 09 | 12 | 15 | 16 | 19 | 20 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [13 empty] | 2
3
| 1 | 1 | 1.1 | 1.2 | 1.1 | 1.3 | 1.2 | 1.15 | 1.2224 | 1 | [13 empty] | 1
| 1 | 2 | 1.2 | 1.2 | 1.1 | 1.4 | 1.2 | 1.17 | 1.3004 | 1.1 | [13 empty] | #N/A
| 1 | 3 | 1.2 | 1.3 | 1.2 | 1.3 | 1.2 | 1.16 | 1.2912 | 1.1 | [13 empty] | 16
| 1 | 4 | 1.1 | 1.3 | 1.1 | 1.4 | 1.2 | 1.15 | 1.2443 | 1 | [13 empty] | 1.1963
| 1 | 5 | 1.2 | 1.2 | 1.1 | 1.3 | 1.2 | 1.14 | 1.258 | 1.2 | [13 empty] | 118
| 1 | 6 | 1.2 | 1.2 | 1.1 | 1.4 | 1.2 | 1.16 | 1.2747 | 1.1
| Mean |  | 1.1667 | 1.2333 | 1.1167 | 1.35 | 1.2 | 1.155 | 1.2652 | 1.0833
| Median |  | 1.2 | 1.2 | 1.1 | 1.35 | 1.2 | 1.155 | 1.2664 | 1.1
| Std Dev. |  | 0.05164 | 0.05164 | 0.040825 | 0.054772 | 0 | 0.010488 | 0.029421 | 0.075277
| Rel.Std.Dev. |  | 4.426% | 4.187% | 3.656% | 4.057% | 0% | 0.9081% | 2.325% | 6.949%
| PDM3 |  | -2.475% | 3.098% | -6.654% | 12.85% | 0.3116% | -3.45% | 5.76% | -9.441%
| Z-Score (Absolute) |  | 0.19 | 0.58 | 0.77 | 1.93 | 0.19 | 0.33 | 0.95 | 1.16
| Analytical results for Tm in OREAS 84 (Certified Value 0.075 ppm) | [22 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 05 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [10 empty] | 3
3
| 1 | 1 | < 0.1 | < 2 | 0.1 | 0.06 | 0.08 | 0.08 | 0.07 | 0.08 | < 0.2 | 0.05 | 0.08 | [10 empty] | 1
| 1 | 2 | < 0.1 | < 2 | 0.07 | 0.08 | 0.09 | 0.08 | 0.07 | 0.08 | < 0.2 | 0.19 | 0.07 | [10 empty] | #N/A
| 1 | 3 | < 0.1 | < 2 | 0.16 | 0.06 | 0.09 | 0.09 | 0.07 | 0.08 | < 0.2 | 0.13 | 0.07 | [10 empty] | 16
| 1 | 4 | < 0.1 | < 2 | 0.09 | 0.07 | 0.09 | 0.09 | 0.07 | 0.07 | < 0.2 | 0.1 | 0.07 | [10 empty] | 0.075375
| 1 | 5 | < 0.1 | < 2 | 0.03 | 0.06 | 0.09 | 0.09 | 0.07 | 0.07 | < 0.2 | 0.03 | 0.07 | [10 empty] | 119
| 1 | 6 | < 0.1 | < 2 | 0.05 | 0.07 | 0.08 | 0.08 | 0.07 | 0.07 | < 0.2 | 0.09 | 0.07
| Mean |  |  |  | 0.083333 | 0.066667 | 0.086667 | 0.085 | 0.07 | 0.075 |  | 0.098333 | 0.071667
| Median |  |  |  | 0.08 | 0.065 | 0.09 | 0.085 | 0.07 | 0.075 |  | 0.095 | 0.07
| Std Dev. |  |  |  | 0.045461 | 0.008165 | 0.005164 | 0.0054772 | 0 | 0.0054772 |  | 0.057417 | 0.0040825
| Rel.Std.Dev. |  |  |  | 54.55% | 12.25% | 5.958% | 6.444% | 0% | 7.303% |  | 58.39% | 5.696%
| PDM3 |  |  |  | 10.56% | -11.55% | 14.98% | 12.77% | -7.131% | -0.4975% |  | 30.46% | -4.92%
| Z-Score (Absolute) |  | 1.69 | 46.36 | 0 | 0.84 | 0.17 | 0.08 | 0.67 | 0.42 | 0.84 | 0.76 | 0.59
| Analytical results for U in OREAS 84 (Certified Value 0.78 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.6 | 0.6 | < 50 | 0.78 | 0.69 | < 2 | 0.8 | 0.8 | < 5 | 1.21 | 0.85805 | 0.9 | [9 empty] | 1
| 1 | 2 | 0.7 | 0.8 | < 50 | 0.89 | 0.69 | < 2 | 0.76 | 0.7 | < 5 | 1.37 | 0.84581 | 0.9 | [9 empty] | #N/A
| 1 | 3 | 0.7 | 0.8 | < 50 | 0.9 | 0.67 | < 2 | 0.83 | 0.8 | < 5 | 0.94 | 0.76041 | 1 | [9 empty] | 16
| 1 | 4 | 0.6 | 0.7 | < 50 | 0.91 | 0.64 | < 2 | 0.82 | 0.8 | < 5 | 2.46 | 0.77082 | 0.9 | [9 empty] | 0.78191
| 1 | 5 | 0.7 | 0.8 | < 50 | 0.77 | 0.67 | < 2 | 0.76 | 0.8 | < 5 | 1.15 | 0.81335 | 1 | [9 empty] | 120
| 1 | 6 | 0.7 | 0.7 | < 50 | 0.81 | 0.69 | < 2 | 0.8 | 0.8 | < 5 | 6.33 | 0.90331 | 0.9
| Mean |  | 0.66667 | 0.73333 |  | 0.84333 | 0.675 |  | 0.795 | 0.78333 |  | 2.2433 | 0.82529 | 0.93333
| Median |  | 0.7 | 0.75 |  | 0.85 | 0.68 |  | 0.8 | 0.8 |  | 1.29 | 0.82958 | 0.9
| Std Dev. |  | 0.05164 | 0.08165 |  | 0.06377 | 0.019748 |  | 0.029496 | 0.040825 |  | 2.0723 | 0.054591 | 0.05164
| Rel.Std.Dev. |  | 7.746% | 11.13% |  | 7.562% | 2.926% |  | 3.71% | 5.212% |  | 92.38% | 6.615% | 5.533%
| PDM3 |  | -14.74% | -6.213% |  | 7.855% | -13.67% |  | 1.674% | 0.1818% |  | 186.9% | 5.548% | 19.37%
| Z-Score (Absolute) |  | 0.87 | 0.52 | 125.21 | 0.05 | 0.83 | 0.86 | 0.2 | 0.26 | 8.63 | 7.3 | 0.05 | 0.51
| Analytical results for V in OREAS 84 (Certified Value 100 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | [7 empty] | 0
0
| 1 | 1 | 93 | 109 | 102 | 114.044 | 107 | 93 | 89 | 108 | 98 | 104.9484 | 72 | 101.6068 | 104 | 82 | [7 empty] | 1
| 1 | 2 | 101 | 106 | 102 | 112.533 | 102 | 93 | 90 | 105 | 98 | 103.7746 | 101 | 99.7094 | 104 | 92 | [7 empty] | #N/A
| 1 | 3 | 101 | 109 | 105 | 115.013 | 101 | 88 | 92 | 106 | 99 | 105.5729 | 79 | 101.8335 | 105 | 91 | [7 empty] | 16
| 1 | 4 | 93 | 109 | 105 | 112.646 | 103 | 93 | 94 | 110 | 99 | 100.2762 | 76 | 104.4385 | 104 | 87 | [7 empty] | 100.1436
| 1 | 5 | 99 | 107 | 103 | 114.934 | 101 | 90 | 91 | 107 | 95 | 100.194 | 99 | 103.4266 | 108 | 89 | [7 empty] | 121
| 1 | 6 | 99 | 107 | 105 | 114.267 | 103 | 91 | 92 | 106 | 97 | 99.8585 | 95 | 105.9891 | 104 | 91
| Mean |  | 97.6667 | 107.8333 | 103.6667 | 113.9062 | 102.8333 | 91.3333 | 91.3333 | 107 | 97.6667 | 102.4374 | 87 | 102.834 | 104.8333 | 88.6667
| Median |  | 99 | 108 | 104 | 114.1555 | 102.5 | 92 | 91.5 | 106.5 | 98 | 102.0254 | 87 | 102.63 | 104 | 90
| Std Dev. |  | 3.7238 | 1.3292 | 1.5055 | 1.0867 | 2.2286 | 2.0656 | 1.7512 | 1.7889 | 1.5055 | 2.6183 | 12.7593 | 2.2428 | 1.6021 | 3.7238
| Rel.Std.Dev. |  | 3.813% | 1.233% | 1.452% | 0.954% | 2.167% | 2.262% | 1.917% | 1.672% | 1.542% | 2.556% | 14.67% | 2.181% | 1.528% | 4.2%
| PDM3 |  | -2.473% | 7.679% | 3.518% | 13.74% | 2.686% | -8.798% | -8.798% | 6.847% | -2.473% | 2.291% | -13.12% | 2.686% | 4.683% | -11.46%
| Z-Score (Absolute) |  | 0.67 | 0.71 | 0.14 | 1.53 | 0.03 | 1.53 | 1.53 | 0.59 | 0.67 | 0.03 | 2.12 | 0.03 | 0.3 | 1.9
| Analytical results for W in OREAS 84 (Indicative Value 1.21 ppm) | [22 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.4 | 1 | < 20 | 58 | < 100 | 2 | 1 | 1.4 | < 5 | < 1 | 1.2183 | < 5 | [9 empty] | 1
| 1 | 2 | 1.1 | 2 | < 20 | 53 | < 100 | < 2 | 1 | 1.6 | < 5 | < 1 | 1.1 | < 5 | [9 empty] | 12
| 1 | 3 | < 0.7 | 2 | < 20 | 69 | < 100 | < 2 | 1 | 1.4 | < 5 | < 1 | 1.0795 | < 5 | [9 empty] | 16
| 1 | 4 | 0.7 | 3 | < 20 | 54 | < 100 | 4 | 1 | 1.5 | < 5 | < 1 | 1.1865 | < 5 | [9 empty] | 1.2116
| 1 | 5 | < 0.7 | 2 | < 20 | 57 | < 100 | < 2 | 1 | 1.4 | < 5 | < 1 | 1.148 | < 5 | [9 empty] | 18
| 1 | 6 | 1.1 | 1 | < 20 | 54 | < 100 | < 2 | 1 | 1.6 | < 5 | < 1 | 1.1865 | < 5
| Mean |  | 1.075 | 1.8333 |  | 57.5 |  | 3 | 1 | 1.4833 |  |  | 1.1531
| Median |  | 1.1 | 2 |  | 55.5 |  | 3 | 1 | 1.45 |  |  | 1.1672
| Std Dev. |  | 0.28723 | 0.75277 |  | 5.9582 |  | 1.4142 | 0 | 0.098319 |  |  | 0.054293
| Rel.Std.Dev. |  | 26.72% | 41.06% |  | 10.36% |  | 47.14% | 0% | 6.628% |  |  | 4.708%
| PDM3 |  | -11.28% | 51.31% |  | 4646% |  | 147.6% | -17.47% | 22.42% |  |  | -4.83%
| Z-Score (Absolute) |  | 0.82 | 0.07 | 7.42 | 50.12 | 43.38 | 0.07 | 0.67 | 0.24 | 0.67 | 1.12 | 0.54 | 0.67
| Analytical results for Y in OREAS 84 (Certified Value 4.9 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [7 empty] | 2
3
| 1 | 1 | 4.5 | 5.2 | 6 | 3.0114 | < 5 | 4.21 | 5.2 | 5.1 | 4.8 | 5.3375 | 6.15 | 5.1015 | 4.6 | < 5 | [7 empty] | 1
| 1 | 2 | 5 | 5.1 | 6 | 2.9217 | 6 | 4.32 | 5.4 | 5 | 4.9 | 4.9786 | 6.74 | 4.9638 | 4.5 | < 5 | [7 empty] | 18
| 1 | 3 | 5 | 5.2 | 6 | 3.1164 | 5 | 4.42 | 5.3 | 5 | 4.9 | 5.1118 | 6.06 | 4.7989 | 4.6 | < 5 | [7 empty] | 16
| 1 | 4 | 4.6 | 5.3 | 6 | 3.0357 | 5 | 4.11 | 5.2 | 5.2 | 4.8 | 5.0329 | 6.25 | 4.8263 | 4.6 | < 5 | [7 empty] | 4.8956
| 1 | 5 | 4.9 | 5.2 | 7 | 2.978 | 5 | 4.17 | 5.2 | 5.1 | 4.8 | 5.1268 | 5.86 | 4.9852 | 4.8 | < 5 | [7 empty] | 122
| 1 | 6 | 5 | 5.1 | 6 | 3.0644 | 5 | 4.21 | 5.1 | 5 | 4.8 | 4.7326 | 5.54 | 5.1259 | 4.8 | < 5
| Mean |  | 4.8333 | 5.1833 | 6.1667 | 3.0213 | 5.2 | 4.24 | 5.2333 | 5.0667 | 4.8333 | 5.0534 | 6.1 | 4.9669 | 4.65
| Median |  | 4.95 | 5.2 | 6 | 3.0236 | 5 | 4.21 | 5.2 | 5.05 | 4.8 | 5.0724 | 6.105 | 4.9745 | 4.6
| Std Dev. |  | 0.22509 | 0.075277 | 0.40825 | 0.067836 | 0.44721 | 0.11171 | 0.10328 | 0.08165 | 0.05164 | 0.19922 | 0.40185 | 0.13546 | 0.12247
| Rel.Std.Dev. |  | 4.657% | 1.452% | 6.62% | 2.245% | 8.6% | 2.635% | 1.973% | 1.612% | 1.068% | 3.942% | 6.588% | 2.727% | 2.634%
| PDM3 |  | -1.272% | 5.878% | 25.96% | -38.29% | 6.218% | -13.39% | 6.899% | 3.495% | -1.272% | 3.223% | 24.6% | 1.457% | -5.017%
| Z-Score (Absolute) |  | 0.17 | 0.72 | NA | 4.75 | NA | 1.67 | 0.84 | 0.42 | 0.17 | 0.39 | 3.03 | 0.17 | 0.63 | 6.07
| Analytical results for Yb in OREAS 84 (Certified Value 0.46 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 05 | 07 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | 0.4 | < 2 | 0.54 | 0.4 | 0.4 | 0.5 | 0.5 | 0.45 | < 1 | 0.5 | 0.5 | [10 empty] | 1
| 1 | 2 | 0.5 | < 2 | 0.6 | 0.4 | 0.5 | 0.5 | 0.5 | 0.44 | < 1 | 0.42 | 0.4 | [10 empty] | #N/A
| 1 | 3 | 0.4 | < 2 | 0.46 | 0.5 | 0.5 | 0.5 | 0.4 | 0.44 | < 1 | 0.48 | 0.4 | [10 empty] | 16
| 1 | 4 | 0.4 | < 2 | 0.81 | 0.4 | 0.5 | 0.5 | 0.5 | 0.44 | < 1 | 0.58 | 0.5 | [10 empty] | 0.45833
| 1 | 5 | 0.5 | < 2 | 0.56 | 0.4 | 0.5 | 0.5 | 0.4 | 0.44 | < 1 | 0.42 | 0.5 | [10 empty] | 123
| 1 | 6 | 0.5 | < 2 | 0.38 | 0.4 | 0.5 | 0.4 | 0.5 | 0.45 | < 1 | 0.34 | 0.5
| Mean |  | 0.45 |  | 0.55833 | 0.41667 | 0.48333 | 0.48333 | 0.46667 | 0.44333 |  | 0.45667 | 0.46667
| Median |  | 0.45 |  | 0.55 | 0.4 | 0.5 | 0.5 | 0.5 | 0.44 |  | 0.45 | 0.5
| Std Dev. |  | 0.054772 |  | 0.14621 | 0.040825 | 0.040825 | 0.040825 | 0.05164 | 0.005164 |  | 0.082381 | 0.05164
| Rel.Std.Dev. |  | 12.17% |  | 26.19% | 9.798% | 8.447% | 8.447% | 11.07% | 1.165% |  | 18.04% | 11.07%
| PDM3 |  | -1.818% |  | 21.82% | -9.091% | 5.455% | 5.455% | 1.818% | -3.273% |  | -0.3636% | 1.818%
| Z-Score (Absolute) |  | 0.67 | 21.58 | 3.71 | 2.02 | 0.67 | 0.67 | 0 | 0.94 | 1.35 | 0.4 | 0
| Analytical results for Zn in OREAS 84 (Certified Value 820 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES/MS | BF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | 730 | 630 | 844 | 780 | 803 | 919 | 840 | 730 | 730 | 822 | 770 | 816 | 798 | 832.8967 | 895 | 872.9826 | 812 | 1500 | 852 |  |  | 1
| 1 | 2 | 780 | 670 | 824 | 790 | 797 | 915 | 840 | 712 | 730 | 828 | 780 | 816 | 803 | 806.0093 | 860 | 842.6382 | 794 | 1500 | 828 |  |  | #N/A
| 1 | 3 | 780 | 650 | 856 | 790 | 805 | 837 | 830 | 715 | 830 | 839 | 860 | 816 | 796 | 826.2768 | 885 | 845.8811 | 812 | 1600 | 829 |  |  | 16
| 1 | 4 | 730 | 680 | 848 | 790 | 799 | 843 | 830 | 695 | 760 | 821 | 810 | 817 | 793 | 804.0833 | 886 | 832.7286 | 808 | 1500 | 835 |  |  | 819.9599
| 1 | 5 | 770 | 650 | 844 | 770 | 802 | 894 | 840 | 693 | 780 | 807 | 820 | 810 | 799 | 804.651 | 930 | 865.1882 | 836 | 1600 | 833 |  |  | 124
| 1 | 6 | 770 | 650 | 852 | 790 | 793 | 813 | 840 | 729 | 790 | 829 | 810 | 814 | 791 | 798.5066 | 922 | 853.3078 | 809 | 1600 | 842
| Mean |  | 760 | 655 | 844.6667 | 785 | 799.8333 | 870.1667 | 836.6667 | 712.3333 | 770 | 824.3333 | 808.3333 | 814.8333 | 796.6667 | 812.0706 | 896.3333 | 852.1211 | 811.8333 | 1550 | 836.5
| Median |  | 770 | 650 | 846 | 790 | 800.5 | 868.5 | 840 | 713.5 | 770 | 825 | 810 | 816 | 797 | 805.3301 | 890.5 | 849.5944 | 810.5 | 1550 | 834
| Std Dev. |  | 23.6643 | 17.6068 | 11.1475 | 8.3666 | 4.4008 | 44.875 | 5.164 | 15.9457 | 38.4708 | 10.6521 | 31.8852 | 2.5626 | 4.3205 | 13.965 | 25.8818 | 14.915 | 13.6002 | 54.7723 | 9.094
| Rel.Std.Dev. |  | 3.114% | 2.688% | 1.32% | 1.066% | 0.5502% | 5.157% | 0.6172% | 2.239% | 4.996% | 1.292% | 3.945% | 0.3145% | 0.5423% | 1.72% | 2.888% | 1.75% | 1.675% | 3.534% | 1.087%
| PDM3 |  | -7.313% | -20.12% | 3.013% | -4.264% | -2.455% | 6.123% | 2.038% | -13.13% | -6.093% | 0.5334% | -1.418% | -0.6252% | -2.841% | -0.9622% | 9.314% | 3.922% | -0.9911% | 89.03% | 2.017%
| Z-Score (Absolute) |  | 1.36 | 4.1 | 0.85 | 0.7 | 0.32 | 1.52 | 0.65 | 2.6 | 1.1 | 0.32 | 0.09 | 0.08 | 0.4 | 0 | 2.2 | 1.05 | 0 | NA | 0.64
| Analytical results for Zr in OREAS 84 (Certified Value 34.7 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 07 | 09 | 15 | 16 | 17 | 18 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | BF*MS | [14 empty] | 1
1
| 1 | 1 | 32 | 27 | 34.3 | 31.8 | 39.1 | 34.6842 | 40 | [14 empty] | 1
| 1 | 2 | 37 | 22 | 32.3 | 32.7 | 37.4 | 33.3082 | 39 | [14 empty] | 19
| 1 | 3 | 40 | 24 | 27.4 | 32.5 | 38.6 | 33.3239 | 37 | [14 empty] | 16
| 1 | 4 | 32 | 23 | 41.3 | 32.8 | 37.9 | 33.0525 | 36 | [14 empty] | 34.7352
| 1 | 5 | 39 | 27 | 26.3 | 30.8 | 40.1 | 33.6702 | 37 | [14 empty] | 125
| 1 | 6 | 30 | 24 | 28.4 | 31.4 | 38.1 | 34.229 | 36
| Mean |  | 35 | 24.5 | 31.6667 | 32 | 38.5333 | 33.7113 | 37.5
| Median |  | 34.5 | 24 | 30.35 | 32.15 | 38.35 | 33.497 | 37
| Std Dev. |  | 4.1952 | 2.0736 | 5.6181 | 0.8025 | 0.96471 | 0.62655 | 1.6432
| Rel.Std.Dev. |  | 11.99% | 8.464% | 17.74% | 2.508% | 2.504% | 1.859% | 4.382%
| PDM3 |  | 0.7623% | -29.47% | -8.834% | -7.874% | 10.93% | -2.948% | 7.96%
| Z-Score (Absolute) |  | 0.42 | 3.04 | 0.67 | 0.56 | 1.59 | 0 | 1.25
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 84 (Certified Value 4.15 wt.%) | [17 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 22 | 23 | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 4.09 | 4.2 | 4.15 | 4.17 | 4.04 | 3.39 | 3.95 | 4.18 | 4.21 | NA | 4.3 | 4.18 | 4.2334 | 4.18 |  | 2.43 | 1
| 1 | 2 | 4.09 | 4.18 | 4.16 | 4.16 | 4.1 | 3.41 | 4.05 | 4.24 | 3.98 | NA | 4.2 | 4.2 | 4.2168 | 4.11 |  | 2.59 | #N/A
| 1 | 3 | 4.2 | 4.2 | 4.16 | 4.16 | 4 | 3.39 | 4.06 | 4.17 | 4.02 | NA | 4.2 | 4.17 | 4.1996 | 4.2 | 4.381 | 2.58 | 16
| 1 | 4 | 4.12 | 4.18 | 4.16 | 4.13 | 4.05 | 3.37 | 4.03 | 4.15 | 3.93 | NA | 4.3 | 4.19 | 4.216 | 4.14 | 4.338 | 2.58 | 4.1549
| 1 | 5 | 4 | 4.17 | 4.15 | 4.11 | 4.12 | 3.38 | 4.05 | 4.18 | 4.01 | NA | 4.2 | 4.2 | 4.1891 | 4.2 |  | 2.58 | 127
| 1 | 6 | 4.11 | 4.12 | 4.16 | 4.22 | 4.01 | 3.41 | 3.95 | 4.15 | 3.93 | NA | 4.2 | 4.2 | 4.2103 | 4.18 |  | 2.5
| Mean |  | 4.1017 | 4.175 | 4.1567 | 4.1583 | 4.0533 | 3.3917 | 4.015 | 4.1783 | 4.0133 |  | 4.2333 | 4.19 | 4.2109 | 4.1683 | 4.3595 | 2.5433
| Median |  | 4.1 | 4.18 | 4.16 | 4.16 | 4.045 | 3.39 | 4.04 | 4.175 | 3.995 |  | 4.2 | 4.195 | 4.2131 | 4.18 | 4.3595 | 2.58
| Std Dev. |  | 0.064317 | 0.029496 | 0.005164 | 0.037639 | 0.048028 | 0.016021 | 0.051284 | 0.033116 | 0.10367 |  | 0.05164 | 0.012649 | 0.015283 | 0.036009 | 0.030406 | 0.064704
| Rel.Std.Dev. |  | 1.568% | 0.7065% | 0.1242% | 0.9051% | 1.185% | 0.4724% | 1.277% | 0.7926% | 2.583% |  | 1.22% | 0.3019% | 0.3629% | 0.8639% | 0.6975% | 2.544%
| PDM3 |  | -1.281% | 0.4838% | 0.04254% | 0.08265% | -2.444% | -18.37% | -3.367% | 0.564% | -3.407% |  | 1.888% | 0.8448% | 1.347% | 0.3233% | 4.924% | -38.79%
| Z-Score (Absolute) |  | 0.67 | 0.2 | 0.02 | 0 | 1.25 | 9.12 | 1.71 | 0.24 | 1.73 | NA | 0.89 | 0.38 | 0.63 | 0.12 | 2.39 | 19.22
| Analytical results for As in OREAS 84 (Indicative Value 123 ppm) | [17 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 08 | 11 | 12 | 13 | 23 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 100 | 400 | < 80 | 200 | NA | 100 | 300 | 110 | [8 empty] | 1
| 1 | 2 | 100 | 500 | < 80 | 100 | NA | 100 | 400 | 110 | [8 empty] | 16
| 1 | 3 | 200 | 400 | 80 | 200 | NA | 200 | 300 | 110 | [8 empty] | 16
| 1 | 4 | 200 | 400 | < 80 | 100 | NA | 200 | 400 | 100 | [8 empty] | 122.8147
| 1 | 5 | 100 | 300 | < 80 | 200 | NA | 200 | 400 | 100 | [8 empty] | 22
| 1 | 6 | 100 | 500 | < 80 | 100 | NA | 100 | 500 | 100
| Mean |  | 133.3333 | 416.6667 | 80 | 150 |  | 150 | 383.3333 | 105
| Median |  | 100 | 400 | 80 | 150 |  | 150 | 400 | 105
| Std Dev. |  | 51.6398 | 75.2773 |  | 54.7723 |  | 54.7723 | 75.2773 | 5.4772
| Rel.Std.Dev. |  | 38.73% | 18.07% |  | 36.51% |  | 36.51% | 19.64% | 5.216%
| PDM3 |  | 8.565% | 239.3% | -34.86% | 22.14% |  | 22.14% | 212.1% | -14.51%
| Z-Score (Absolute) |  | 0.25 | 4 | 1.59 | 0 | NA | 0 | 3.5 | 0.67
| Analytical results for BaO in OREAS 84 (Indicative Value 249 ppm) | [17 empty] | Indicative
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 04 | 05 | 08 | 13 | 17 | 20 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | 300 | 400 | 100 | 400 | 200 | 158.0588 | 200 | [9 empty] | 1
| 1 | 2 | 300 | 300 | 200 | 300 | 300 | 168.7485 | 200 | [9 empty] | 17
| 1 | 3 | 300 | 400 | < 100 | 300 | 200 | 165.5448 | 200 | [9 empty] | 16
| 1 | 4 | 300 | 300 | 100 | 300 | 300 | 171.6199 | 200 | [9 empty] | 248.798
| 1 | 5 | 300 | 300 | 200 | 300 | 300 | 161.1472 | 200 | [9 empty] | 23
| 1 | 6 | 300 | 400 | 100 | 300 | 300 | 184.3954 | 200
| Mean |  | 300 | 350 | 140 | 316.6667 | 266.6667 | 168.2524 | 200
| Median |  | 300 | 350 | 100 | 300 | 300 | 167.1466 | 200
| Std Dev. |  | 0 | 54.7723 | 54.7723 | 40.8248 | 51.6398 | 9.3157 | 0
| Rel.Std.Dev. |  | 0% | 15.65% | 39.12% | 12.89% | 19.36% | 5.537% | 0%
| PDM3 |  | 20.58% | 40.68% | -43.73% | 27.28% | 7.182% | -32.37% | -19.61%
| Z-Score (Absolute) |  | 0.34 | 0.84 | 1.43 | 0.51 | 0 | 1 | 0.67
| Analytical results for Bi in OREAS 84 (Indicative Value < 100 ppm) | [17 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | 100 | < 100 | 100 | [12 empty] | 18
| 1 | 3 | < 100 | < 100 | < 100 | 100 | [12 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | 100 | [12 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | 100 | [12 empty] | 24
| 1 | 6 | < 100 | < 100 | < 100 | 100
| Mean |  |  | 100 |  | 100
| Median |  |  | 100 |  | 100
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 6.07
| Analytical results for CaO in OREAS 84 (Certified Value 2.91 wt.%) | [17 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 22 | 23 | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.83 | 2.92 | 2.89 | 2.93 | 2.9 | 2.27 | 3.14 | 2.97 | 2.86 | NA | 2.88 | 2.91 | 2.9583 | 2.9 | 2.979 | 1.88 | 1
| 1 | 2 | 2.77 | 2.92 | 2.89 | 2.93 | 2.93 | 2.27 | 3.17 | 2.99 | 2.77 | NA | 2.87 | 2.95 | 2.9553 | 2.97 | 2.99 | 1.91 | #N/A
| 1 | 3 | 2.78 | 2.94 | 2.89 | 2.9 | 2.89 | 2.3 | 3.13 | 2.96 | 2.78 | NA | 2.87 | 2.93 | 2.9524 | 2.97 |  | 1.96 | 16
| 1 | 4 | 2.74 | 2.93 | 2.89 | 2.86 | 2.88 | 2.28 | 3.15 | 2.97 | 2.76 | NA | 2.85 | 2.97 | 2.9728 | 2.91 |  | 1.86 | 2.9053
| 1 | 5 | 2.81 | 2.93 | 2.88 | 2.85 | 2.95 | 2.31 | 3.14 | 2.97 | 2.75 | NA | 2.89 | 2.94 | 2.951 | 2.93 |  | 1.88 | 128
| 1 | 6 | 2.8 | 2.89 | 2.91 | 2.93 | 2.93 | 2.29 | 3.15 | 2.96 | 2.71 | NA | 2.88 | 2.97 | 2.9642 | 2.99 |  | 1.83
| Mean |  | 2.7883 | 2.9217 | 2.8917 | 2.9 | 2.9133 | 2.2867 | 3.1467 | 2.97 | 2.7717 |  | 2.8733 | 2.945 | 2.959 | 2.945 | 2.9845 | 1.8867
| Median |  | 2.79 | 2.925 | 2.89 | 2.915 | 2.915 | 2.285 | 3.145 | 2.97 | 2.765 |  | 2.875 | 2.945 | 2.9568 | 2.95 | 2.9845 | 1.88
| Std Dev. |  | 0.031885 | 0.017224 | 0.0098319 | 0.036878 | 0.027325 | 0.01633 | 0.013663 | 0.010954 | 0.049565 |  | 0.013663 | 0.023452 | 0.0082335 | 0.036742 | 0.0077782 | 0.044572
| Rel.Std.Dev. |  | 1.144% | 0.5895% | 0.34% | 1.272% | 0.9379% | 0.7141% | 0.4342% | 0.3688% | 1.788% |  | 0.4755% | 0.7963% | 0.2783% | 1.248% | 0.2606% | 2.362%
| PDM3 |  | -4.026% | 0.5637% | -0.4689% | -0.1821% | 0.2768% | -21.29% | 8.308% | 2.227% | -4.599% |  | -1.1% | 1.367% | 1.848% | 1.367% | 2.726% | -35.06%
| Z-Score (Absolute) |  | 1.85 | 0.12 | 0.32 | 0.2 | 0 | 9.26 | 3.45 | 0.84 | 2.09 | NA | 0.59 | 0.47 | 0.67 | 0.47 | 1.05 | 15.16
| Analytical results for Ce in OREAS 84 (Indicative Value 112 ppm) | [17 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 80 | < 80 | < 80 | 80 | [12 empty] | 1
| 1 | 2 | 80 | 160 | 80 | 240 | [12 empty] | 19
| 1 | 3 | 160 | < 80 | 80 | 160 | [12 empty] | 16
| 1 | 4 | 160 | < 80 | 80 | 240 | [12 empty] | 111.9342
| 1 | 5 | 80 | < 80 | < 80 | 160 | [12 empty] | 25
| 1 | 6 | 80 | 160 | 80 | 240
| Mean |  | 106.6667 | 160 | 80 | 186.6667
| Median |  | 80 | 160 | 80 | 200
| Std Dev. |  | 41.3118 | 0 | 0 | 65.3197
| Rel.Std.Dev. |  | 38.73% | 0% | 0% | 34.99%
| PDM3 |  | -4.706% | 42.94% | -28.53% | 66.76%
| Z-Score (Absolute) |  | 0.45 | 0.45 | 0.9 | 3.15
| Analytical results for Co in OREAS 84 (Certified Value 1615 ppm) | [17 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 11 | 12 | 13 | 23 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 1700 | 1700 | 1670 | 1600 | 1570 | 1600 | 1700 | 1700 | NA | 1700 | 1600 | 970 |  |  |  |  | 1
| 1 | 2 | 1740 | 1700 | 1690 | 1700 | 1570 | 1500 | 1700 | 1600 | NA | 1700 | 1600 | 1020 |  |  |  |  | #N/A
| 1 | 3 | 1710 | 1700 | 1700 | 1700 | 1570 | 1500 | 1700 | 1500 | NA | 1700 | 1600 | 1080 |  |  |  |  | 16
| 1 | 4 | 1700 | 1700 | 1690 | 1600 | 1570 | 1500 | 1700 | 1600 | NA | 1600 | 1600 | 1030 |  |  |  |  | 1614.9264
| 1 | 5 | 1700 | 1700 | 1690 | 1600 | 1570 | 1500 | 1700 | 1600 | NA | 1600 | 1600 | 980 |  |  |  |  | 129
| 1 | 6 | 1740 | 1700 | 1680 | 1700 | 1570 | 1500 | 1700 | 1500 | NA | 1700 | 1600 | 1030
| Mean |  | 1715 | 1700 | 1686.6667 | 1650 | 1570 | 1516.6667 | 1700 | 1583.3333 |  | 1666.6667 | 1600 | 1018.3333
| Median |  | 1705 | 1700 | 1690 | 1650 | 1570 | 1500 | 1700 | 1600 |  | 1700 | 1600 | 1025
| Std Dev. |  | 19.7484 | 0 | 10.328 | 54.7723 | 0 | 40.8248 | 0 | 75.2773 |  | 51.6398 | 0 | 39.7073
| Rel.Std.Dev. |  | 1.152% | 0% | 0.6123% | 3.32% | 0% | 2.692% | 0% | 4.754% |  | 3.098% | 0% | 3.899%
| PDM3 |  | 6.197% | 5.268% | 4.442% | 2.172% | -2.782% | -6.084% | 5.268% | -1.956% |  | 3.204% | -0.9243% | -36.94%
| Z-Score (Absolute) |  | 1.09 | NA | 0.86 | NA | 0.04 | 0.49 | NA | 0.04 | NA | NA | NA | 4.44
Results from laboratories 2, 4, 8, 12 and 13 were removed due to their 100 ppm reading resolution.
| Analytical results for Cr2O3 in OREAS 84 (Certified Value 1.1 wt.%) | [17 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 1.052 | 1.111 | 1.105 | 1.125 | 1.089 | 0.9 | 1.03 | 1.125 | 1.09 | NA | 1.155 | 1.052 | 1.1802 | 1.08 | 0.90613 |  | 1
| 1 | 2 | 1.008 | 1.111 | 1.104 | 1.14 | 1.063 | 0.88 | 1.04 | 1.14 | 1.08 | NA | 1.155 | 1.082 | 1.1608 | 1.1 | 0.91636 |  | #N/A
| 1 | 3 | 1.023 | 1.125 | 1.097 | 1.14 | 1.091 | 0.89 | 1.03 | 1.125 | 1.06 | NA | 1.14 | 1.067 | 1.1535 | 1.12 | 0.94851 |  | 16
| 1 | 4 | 1.023 | 1.111 | 1.098 | 1.125 | 1.092 | 0.9 | 1.03 | 1.125 | 1.06 | NA | 1.14 | 1.096 | 1.1552 | 1.09 | 0.90467 |  | 1.0973
| 1 | 5 | 1.052 | 1.096 | 1.097 | 1.125 | 1.115 | 0.89 | 1.03 | 1.125 | 1.07 | NA | 1.155 | 1.082 | 1.1418 | 1.09 | 0.90321 |  | 130
| 1 | 6 | 1.052 | 1.096 | 1.102 | 1.096 | 1.087 | 0.89 | 1.03 | 1.125 | 1.03 | NA | 1.155 | 1.082 | 1.1769 | 1.1 | 0.87982
| Mean |  | 1.035 | 1.1083 | 1.1005 | 1.1252 | 1.0895 | 0.89167 | 1.0317 | 1.1275 | 1.065 |  | 1.15 | 1.0768 | 1.1614 | 1.0967 | 0.90978
| Median |  | 1.0375 | 1.111 | 1.1 | 1.125 | 1.09 | 0.89 | 1.03 | 1.125 | 1.065 |  | 1.155 | 1.082 | 1.158 | 1.095 | 0.9054
| Std Dev. |  | 0.019411 | 0.010985 | 0.0036194 | 0.016068 | 0.016538 | 0.0075277 | 0.0040825 | 0.0061237 | 0.020736 |  | 0.007746 | 0.015237 | 0.014674 | 0.013663 | 0.022466
| Rel.Std.Dev. |  | 1.875% | 0.9911% | 0.3289% | 1.428% | 1.518% | 0.8442% | 0.3957% | 0.5431% | 1.947% |  | 0.6736% | 1.415% | 1.263% | 1.246% | 2.469%
| PDM3 |  | -5.679% | 1.004% | 0.2904% | 2.538% | -0.7121% | -18.74% | -5.983% | 2.751% | -2.945% |  | 4.801% | -1.866% | 5.84% | -0.05895% | -17.09%
| Z-Score (Absolute) |  | 1.16 | 0.31 | 0.15 | 0.65 | 0.07 | 4.06 | 1.24 | 0.7 | 0.57 | NA | 1.14 | 0.33 | 1.38 | 0.07 | 3.69
| Analytical results for Cu in OREAS 84 (Certified Value 0.398 wt.%) | [17 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.406 | 0.358 | 0.44 | 0.386 | 0.4114 | 0.39 | 0.392 | 0.36 | NA | 0.43 | 0.367 | 0.41797 | 0.359 | 0.157 |  |  | 1
| 1 | 2 | 0.414 | 0.379 | 0.444 | 0.401 | 0.4074 | 0.4 | 0.406 | 0.35 | NA | 0.44 | 0.377 | 0.41601 | 0.367 | 0.153 |  |  | #N/A
| 1 | 3 | 0.411 | 0.372 | 0.452 | 0.392 | 0.3978 | 0.41 | 0.405 | 0.36 | NA | 0.42 | 0.377 | 0.43009 | 0.367 | 0.161 |  |  | 16
| 1 | 4 | 0.411 | 0.387 | 0.442 | 0.386 | 0.401 | 0.41 | 0.401 | 0.35 | NA | 0.43 | 0.379 | 0.41585 | 0.367 | 0.169 |  |  | 0.39752
| 1 | 5 | 0.41 | 0.374 | 0.44 | 0.382 | 0.4034 | 0.4 | 0.406 | 0.35 | NA | 0.43 | 0.372 | 0.41187 | 0.359 | 0.258 |  |  | 131
| 1 | 6 | 0.415 | 0.374 | 0.445 | 0.395 | 0.4082 | 0.4 | 0.402 | 0.33 | NA | 0.43 | 0.383 | 0.4304 | 0.375 | 0.236
| Mean |  | 0.41117 | 0.374 | 0.44383 | 0.39033 | 0.40487 | 0.40167 | 0.402 | 0.35 |  | 0.43 | 0.37583 | 0.42037 | 0.36567 | 0.189
| Median |  | 0.411 | 0.374 | 0.443 | 0.389 | 0.4054 | 0.4 | 0.4035 | 0.35 |  | 0.43 | 0.377 | 0.41699 | 0.367 | 0.165
| Std Dev. |  | 0.0031885 | 0.0095289 | 0.0044907 | 0.0070048 | 0.0050449 | 0.0075277 | 0.0053292 | 0.010954 |  | 0.0063246 | 0.0056006 | 0.0079057 | 0.0060222 | 0.045769
| Rel.Std.Dev. |  | 0.7755% | 2.548% | 1.012% | 1.795% | 1.246% | 1.874% | 1.326% | 3.13% |  | 1.471% | 1.49% | 1.881% | 1.647% | 24.22%
| PDM3 |  | 3.433% | -5.916% | 11.65% | -1.807% | 1.849% | 1.044% | 1.127% | -11.95% |  | 8.171% | -5.455% | 5.747% | -8.013% | -52.46%
| Z-Score (Absolute) |  | 0.25 | 0.72 | 1.1 | 0.3 | 0.08 | 0 | 0.01 | 1.35 | NA | 0.74 | 0.67 | 0.49 | 0.93 | 5.55
| Analytical results for Fe2O3 in OREAS 84 (Certified Value 44.93 wt.%) | [17 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 44.93 | 44.75 | 44.45 | 45.893 | 46.52 | 36.4 | 44.14 | 45.036 | 44.9 | NA | 45.3 | 43.749 | 45.8127 | 44.09 | 28.7 |  | 1
| 1 | 2 | 44.43 | 44.607 | 44.52 | 46.608 | 45.95 | 36.45 | 44.25 | 45.179 | 44.7 | NA | 45.43 | 44.321 | 45.9654 | 44.2 | 29 |  | #N/A
| 1 | 3 | 44.51 | 44.893 | 44.54 | 46.179 | 46.13 | 36.41 | 44.26 | 45.036 | 44 | NA | 45.15 | 43.892 | 45.8914 | 44.75 | 30 |  | 16
| 1 | 4 | 44.37 | 44.893 | 44.31 | 45.75 | 46.19 | 36.4 | 44.19 | 45.036 | 44 | NA | 45.29 | 44.321 | 46.1626 | 44.25 | 28.2 |  | 44.9255
| 1 | 5 | 44.96 | 44.893 | 44.35 | 45.607 | 47.45 | 36.48 | 44.19 | 45.036 | 44.6 | NA | 45.35 | 43.892 | 45.8386 | 43.98 | 28.7 |  | 132
| 1 | 6 | 44.46 | 44.464 | 44.46 | 45.607 | 46.25 | 36.43 | 44.22 | 44.893 | 43.4 | NA | 45.35 | 44.178 | 46.1101 | 44.62 | 27.8
| Mean |  | 44.61 | 44.75 | 44.4383 | 45.9407 | 46.415 | 36.4283 | 44.2083 | 45.036 | 44.2667 |  | 45.3117 | 44.0588 | 45.9635 | 44.315 | 28.7333
| Median |  | 44.485 | 44.8215 | 44.455 | 45.8215 | 46.22 | 36.42 | 44.205 | 45.036 | 44.3 |  | 45.325 | 44.035 | 45.9284 | 44.225 | 28.7
| Std Dev. |  | 0.26359 | 0.18088 | 0.091524 | 0.39075 | 0.53992 | 0.031885 | 0.04446 | 0.090441 | 0.56451 |  | 0.09347 | 0.2463 | 0.1447 | 0.30415 | 0.75277
| Rel.Std.Dev. |  | 0.5909% | 0.4042% | 0.206% | 0.8506% | 1.163% | 0.08753% | 0.1006% | 0.2008% | 1.275% |  | 0.2063% | 0.559% | 0.3148% | 0.6863% | 2.62%
| PDM3 |  | -0.7023% | -0.3907% | -1.084% | 2.26% | 3.315% | -18.91% | -1.596% | 0.2459% | -1.467% |  | 0.8595% | -1.929% | 2.31% | -1.359% | -36.04%
| Z-Score (Absolute) |  | 0.12 | 0.31 | 0.12 | 1.96 | 2.61 | 11.18 | 0.44 | 0.71 | 0.36 | NA | 1.09 | 0.64 | 1.99 | 0.29 | 21.8
| Analytical results for Hf in OREAS 84 (Certified Value < 80 ppm) | [17 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 04 | 08 | 13 | 17 | 20 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 80 | < 80 | 80 | < 80 | < 170 | < 80 | [10 empty] | 1
| 1 | 2 | < 80 | 80 | 80 | < 80 | < 170 | < 80 | [10 empty] | 14
| 1 | 3 | < 80 | < 80 | 80 | < 80 | < 170 | < 80 | [10 empty] | 16
| 1 | 4 | < 80 | < 80 | 80 | < 80 | < 170 | < 80 | [10 empty] | < 80
| 1 | 5 | < 80 | < 80 | 80 | < 80 | < 170 | < 80 | [10 empty] | 133
| 1 | 6 | < 80 | < 80 | 80 | 80 | < 170 | < 80
| Mean |  |  | 80 | 80 | 80
| Median |  |  | 80 | 80 | 80
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 3.37 | 0 | 3.37 | 0.67
| Analytical results for K2O in OREAS 84 (Certified Value 0.433 wt.%) | [17 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.43 | 0.43 | 0.43 | 0.44 | 0.418 | 0.34 | 0.41 | 0.42 | 0.35 | NA | 0.45 | 0.44 | 0.44136 | 0.43 | 0.29 |  | 1
| 1 | 2 | 0.41 | 0.43 | 0.43 | 0.45 | 0.428 | 0.35 | 0.41 | 0.45 | 0.32 | NA | 0.45 | 0.45 | 0.4368 | 0.43 | 0.29 |  | #N/A
| 1 | 3 | 0.42 | 0.43 | 0.43 | 0.45 | 0.413 | 0.34 | 0.4 | 0.44 | 0.32 | NA | 0.46 | 0.45 | 0.43723 | 0.44 | 0.3 |  | 16
| 1 | 4 | 0.41 | 0.43 | 0.43 | 0.44 | 0.412 | 0.34 | 0.41 | 0.44 | 0.32 | NA | 0.44 | 0.45 | 0.4399 | 0.44 | 0.28 |  | 0.43279
| 1 | 5 | 0.43 | 0.43 | 0.42 | 0.44 | 0.427 | 0.34 | 0.41 | 0.45 | 0.32 | NA | 0.44 | 0.44 | 0.43982 | 0.44 | 0.29 |  | 134
| 1 | 6 | 0.41 | 0.43 | 0.43 | 0.45 | 0.419 | 0.34 | 0.41 | 0.44 | 0.31 | NA | 0.45 | 0.45 | 0.44234 | 0.44 | 0.28
| Mean |  | 0.41833 | 0.43 | 0.42833 | 0.445 | 0.4195 | 0.34167 | 0.40833 | 0.44 | 0.32333 |  | 0.44833 | 0.44667 | 0.43958 | 0.43667 | 0.28833
| Median |  | 0.415 | 0.43 | 0.43 | 0.445 | 0.4185 | 0.34 | 0.41 | 0.44 | 0.32 |  | 0.45 | 0.45 | 0.43986 | 0.44 | 0.29
| Std Dev. |  | 0.0098319 | 0 | 0.0040825 | 0.0054772 | 0.006775 | 0.0040825 | 0.0040825 | 0.010954 | 0.013663 |  | 0.0075277 | 0.005164 | 0.0021986 | 0.005164 | 0.0075277
| Rel.Std.Dev. |  | 2.35% | 0% | 0.9531% | 1.231% | 1.615% | 1.195% | 0.9998% | 2.49% | 4.226% |  | 1.679% | 1.156% | 0.5002% | 1.183% | 2.611%
| PDM3 |  | -3.341% | -0.6457% | -1.031% | 2.82% | -3.072% | -21.06% | -5.652% | 1.665% | -25.29% |  | 3.59% | 3.205% | 1.567% | 0.8946% | -33.38%
| Z-Score (Absolute) |  | 0.55 | 0.04 | 0.04 | 0.8 | 0.49 | 4.43 | 1.05 | 0.55 | 5.35 | NA | 0.97 | 0.89 | 0.53 | 0.38 | 7.12
| Analytical results for La in OREAS 84 (Indicative Value 94 ppm) | [17 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 90 | 170 | 90 | 90 | [12 empty] | 1
| 1 | 2 | 170 | 90 | < 90 | 90 | [12 empty] | 20
| 1 | 3 | 90 | 170 | 90 | 90 | [12 empty] | 16
| 1 | 4 | 90 | 90 | < 90 | 90 | [12 empty] | 94.1479
| 1 | 5 | 90 | < 90 | 90 | 90 | [12 empty] | 26
| 1 | 6 | 90 | 170 | 90 | 90
| Mean |  | 103.3333 | 138 | 90 | 90
| Median |  | 90 | 170 | 90 | 90
| Std Dev. |  | 32.6599 | 43.8178 | 0 | 0
| Rel.Std.Dev. |  | 31.61% | 31.75% | 0% | 0%
| PDM3 |  | 9.756% | 46.58% | -4.406% | -4.406%
| Z-Score (Absolute) |  | 0.34 | 1.35 | 1.01 | 0.34
| Analytical results for MgO in OREAS 84 (Certified Value 2.7 wt.%) | [17 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 22 | 23 | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.61 | 2.82 | 2.69 | 2.73 | 2.54 | 1.95 | 2.72 | 2.84 | 2.66 | NA | 2.8 | 2.68 | 2.7602 | 2.58 | 2.78 | 1.76 | 1
| 1 | 2 | 2.5 | 2.82 | 2.7 | 2.79 | 2.53 | 1.94 | 2.77 | 2.85 | 2.58 | NA | 2.8 | 2.71 | 2.7582 | 2.6 | 2.79 | 1.79 | #N/A
| 1 | 3 | 2.62 | 2.84 | 2.69 | 2.74 | 2.48 | 1.96 | 2.72 | 2.84 | 2.55 | NA | 2.8 | 2.68 | 2.7538 | 2.62 | 2.69 | 1.86 | 16
| 1 | 4 | 2.49 | 2.84 | 2.68 | 2.76 | 2.47 | 1.95 | 2.72 | 2.83 | 2.52 | NA | 2.8 | 2.73 | 2.772 | 2.58 | 2.7 | 1.77 | 2.6988
| 1 | 5 | 2.62 | 2.85 | 2.69 | 2.73 | 2.57 | 1.96 | 2.69 | 2.81 | 2.59 | NA | 2.8 | 2.69 | 2.7536 | 2.6 |  | 1.77 | 135
| 1 | 6 | 2.63 | 2.79 | 2.69 | 2.7 | 2.56 | 1.95 | 2.74 | 2.82 | 2.5 | NA | 2.8 | 2.69 | 2.7602 | 2.63 |  | 1.72
| Mean |  | 2.5783 | 2.8267 | 2.69 | 2.7417 | 2.525 | 1.9517 | 2.7267 | 2.8317 | 2.5667 |  | 2.8 | 2.6967 | 2.7597 | 2.6017 | 2.74 | 1.7783
| Median |  | 2.615 | 2.83 | 2.69 | 2.735 | 2.535 | 1.95 | 2.72 | 2.835 | 2.565 |  | 2.8 | 2.69 | 2.7592 | 2.6 | 2.74 | 1.77
| Std Dev. |  | 0.064936 | 0.021602 | 0.0063246 | 0.030605 | 0.041352 | 0.0075277 | 0.026583 | 0.01472 | 0.057155 |  | 4.8648e-16 | 0.019664 | 0.0067251 | 0.020412 | 0.052281 | 0.046224
| Rel.Std.Dev. |  | 2.519% | 0.7642% | 0.2351% | 1.116% | 1.638% | 0.3857% | 0.9749% | 0.5198% | 2.227% |  | 1.737e-14% | 0.7292% | 0.2437% | 0.7846% | 1.908% | 2.599%
| PDM3 |  | -4.464% | 4.737% | -0.3268% | 1.588% | -6.441% | -27.68% | 1.032% | 4.922% | -4.897% |  | 3.749% | -0.07977% | 2.254% | -3.6% | 1.526% | -34.11%
| Z-Score (Absolute) |  | 0.77 | 0.85 | 0.04 | 0.29 | 1.12 | 4.86 | 0.2 | 0.88 | 0.85 | NA | 0.67 | 0 | 0.41 | 0.62 | 0.28 | 5.99
| Analytical results for MnO in OREAS 84 (Certified Value 0.087 wt.%) | [17 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.089 | 0.09 | 0.08 | 0.077 | 0.067 | 0.01 | 0.08 | 0.103 | 0.11 | NA | 0.09 | 0.09 | 0.084788 | 0.09 | 0.05 |  | 1
| 1 | 2 | 0.086 | 0.09 | 0.08 | 0.077 | 0.062 | 0.01 | 0.08 | 0.103 | 0.11 | NA | 0.09 | 0.09 | 0.084607 | 0.08 | 0.05 |  | #N/A
| 1 | 3 | 0.088 | 0.09 | 0.08 | 0.077 | 0.063 | 0.01 | 0.1 | 0.09 | 0.1 | NA | 0.09 | 0.09 | 0.083665 | 0.09 | 0.053 |  | 16
| 1 | 4 | 0.086 | 0.09 | 0.08 | 0.077 | 0.065 | 0.01 | 0.08 | 0.09 | 0.1 | NA | 0.09 | 0.09 | 0.085288 | 0.08 | 0.055 |  | 0.086806
| 1 | 5 | 0.089 | 0.09 | 0.08 | 0.077 | 0.074 | 0.01 | 0.08 | 0.103 | 0.11 | NA | 0.09 | 0.09 | 0.084631 | 0.09 | 0.054 |  | 136
| 1 | 6 | 0.091 | 0.09 | 0.08 | 0.077 | 0.078 | 0.01 | 0.08 | 0.103 | 0.1 | NA | 0.09 | 0.09 | 0.08561 | 0.08 | 0.053
| Mean |  | 0.088167 | 0.09 | 0.08 | 0.077 | 0.068167 | 0.01 | 0.083333 | 0.098667 | 0.105 |  | 0.09 | 0.09 | 0.084764 | 0.085 | 0.0525
| Median |  | 0.0885 | 0.09 | 0.08 | 0.077 | 0.066 | 0.01 | 0.08 | 0.103 | 0.105 |  | 0.09 | 0.09 | 0.084709 | 0.085 | 0.053
| Std Dev. |  | 0.0019408 | 1.5202e-17 | 0 | 0 | 0.0064317 | 0 | 0.008165 | 0.0067132 | 0.0054772 |  | 1.5202e-17 | 1.5202e-17 | 0.00066944 | 0.0054772 | 0.0020736
| Rel.Std.Dev. |  | 2.201% | 1.689e-14% | 0% | 0% | 9.435% | 0% | 9.798% | 6.804% | 5.216% |  | 1.689e-14% | 1.689e-14% | 0.7898% | 6.444% | 3.95%
| PDM3 |  | 1.568% | 3.68% | -7.84% | -11.3% | -21.47% | -88.48% | -4% | 13.66% | 20.96% |  | 3.68% | 3.68% | -2.352% | -2.08% | -39.52%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.6 | 0.91 | 2.05 | 9.17 | 0.19 | 1.73 | 2.46 | NA | 0.63 | 0.67 | 0.01 | 0.01 | 3.97
| Analytical results for Mo in OREAS 84 (Indicative Value < 50 ppm) | [17 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | 90 | [12 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | 110 | [12 empty] | 21
| 1 | 3 | < 50 | < 50 | < 50 | 100 | [12 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | 110 | [12 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | 110 | [12 empty] | 27
| 1 | 6 | < 50 | < 50 | < 50 | 130
| Mean |  |  |  |  | 108.3333
| Median |  |  |  |  | 110
| Std Dev. |  |  |  |  | 13.2916
| Rel.Std.Dev. |  |  |  |  | 12.27%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Na2O in OREAS 84 (Certified Value 0.461 wt.%) | [17 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 05 | 06 | 11 | 12 | 17 | 20 | 23 | [8 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 3
3
| 1 | 1 | 0.51 | 0.6 | 0.4 | NA | 0.7 | 0.53326 | 0.33 | 0.4 | [8 empty] | 1
| 1 | 2 | 0.5 | 0.71 | 0.39 | NA | 0.7 | 0.52701 | 0.32 | 0.4 | [8 empty] | 12
| 1 | 3 | 0.49 | 0.61 | 0.39 | NA | 0.7 | 0.52642 | 0.34 | 0.42 | [8 empty] | 16
| 1 | 4 | 0.49 | 0.61 | 0.4 | NA | 0.7 | 0.53233 | 0.31 | 0.41 | [8 empty] | 0.4613
| 1 | 5 | 0.51 | 0.61 | 0.39 | NA | 0.7 | 0.55693 | 0.32 | 0.39 | [8 empty] | 137
| 1 | 6 | 0.51 | 0.61 | 0.39 | NA | 0.7 | 0.543 | 0.33 | 0.4
| Mean |  | 0.50167 | 0.625 | 0.39333 |  | 0.7 | 0.53649 | 0.325 | 0.40333
| Median |  | 0.505 | 0.61 | 0.39 |  | 0.7 | 0.53279 | 0.325 | 0.4
| Std Dev. |  | 0.0098319 | 0.041833 | 0.005164 |  | 1.2162e-16 | 0.011655 | 0.010488 | 0.010328
| Rel.Std.Dev. |  | 1.96% | 6.693% | 1.313% |  | 1.737e-14% | 2.173% | 3.227% | 2.561%
| PDM3 |  | 8.75% | 35.49% | -14.73% |  | 51.74% | 16.3% | -29.55% | -12.57%
| Z-Score (Absolute) |  | 0.46 | 1.63 | 0.56 | NA | NA | 0.79 | 1.2 | 0.46
| Analytical results for Nb in OREAS 84 (Indicative Value 146 ppm) | [17 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 80 | 230 | 110 | 180 | [12 empty] | 1
| 1 | 2 | 70 | 250 | 90 | 150 | [12 empty] | 13
| 1 | 3 | 60 | 250 | 90 | 180 | [12 empty] | 16
| 1 | 4 | 70 | 240 | 100 | 170 | [12 empty] | 145.8333
| 1 | 5 | 80 | 250 | 100 | 170 | [12 empty] | 22
| 1 | 6 | 60 | 260 | 90 | 170
| Mean |  | 70 | 246.6667 | 96.6667 | 170
| Median |  | 70 | 250 | 95 | 170
| Std Dev. |  | 8.9443 | 10.328 | 8.165 | 10.9545
| Rel.Std.Dev. |  | 12.78% | 4.187% | 8.447% | 6.444%
| PDM3 |  | -52% | 69.14% | -33.71% | 16.57%
| Z-Score (Absolute) |  | 0.85 | 1.53 | 0.49 | 0.49
| Analytical results for Ni in OREAS 84 (Certified Value 10.37 wt.%) | [17 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 22 | 23 | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 10.5 | 9.88 | 10.103 | 10.75 | 11.261 | 7.4813 | 10.48 | 10.55 | 10.3 | NA | 10.87 | 9.67 | 10.2603 | 9.815 | 10.4381 | > 3.95 | 1
| 1 | 2 | 10.7 | 9.99 | 10.126 | 10.95 | 11.183 | 7.442 | 10.4 | 10.6 | 10.2 | NA | 10.9 | 9.85 | 10.2819 | 9.933 | 10.3831 | > 3.95 | #N/A
| 1 | 3 | 10.6 | 9.95 | 10.126 | 10.85 | 11.151 | 7.4656 | 10.36 | 10.55 | 10.1 | NA | 10.83 | 9.73 | 10.2612 | 9.823 | 10.361 | > 3.95 | 16
| 1 | 4 | 10.6 | 9.91 | 10.067 | 10.7 | 11.143 | 7.4578 | 10.38 | 10.55 | 10 | NA | 10.85 | 9.84 | 10.3494 | 9.776 | 10.417 | > 3.95 | 10.371
| 1 | 5 | 10.6 | 9.88 | 10.054 | 10.65 | 11.45 | 7.4499 | 10.38 | 10.55 | 10.2 | NA | 10.88 | 9.73 | 10.2381 | 9.729 |  | > 3.95 | 138
| 1 | 6 | 10.7 | 9.82 | 10.112 | 10.45 | 11.285 | 7.4578 | 10.46 | 10.55 | 9.87 | NA | 10.87 | 9.84 | 10.3411 | 9.847 |  | > 3.95
| Mean |  | 10.6167 | 9.905 | 10.098 | 10.725 | 11.2455 | 7.4591 | 10.41 | 10.5583 | 10.1117 |  | 10.8667 | 9.7767 | 10.2887 | 9.8205 | 10.3998
| Median |  | 10.6 | 9.895 | 10.1075 | 10.725 | 11.222 | 7.4578 | 10.39 | 10.55 | 10.15 |  | 10.87 | 9.785 | 10.2715 | 9.819 | 10.4001
| Std Dev. |  | 0.075277 | 0.059582 | 0.030614 | 0.17248 | 0.11573 | 0.013536 | 0.04858 | 0.020412 | 0.15626 |  | 0.024221 | 0.076333 | 0.046024 | 0.068917 | 0.034385
| Rel.Std.Dev. |  | 0.709% | 0.6015% | 0.3032% | 1.608% | 1.029% | 0.1815% | 0.4667% | 0.1933% | 1.545% |  | 0.2229% | 0.7808% | 0.4473% | 0.7018% | 0.3306%
| PDM3 |  | 2.369% | -4.493% | -2.632% | 3.414% | 8.432% | -28.08% | 0.3763% | 1.807% | -2.5% |  | 4.78% | -5.73% | -0.7937% | -5.308% | 0.278%
| Z-Score (Absolute) |  | 0.56 | 0.91 | 0.51 | 0.79 | 1.86 | 5.96 | 0.14 | 0.44 | 0.48 | NA | 1.08 | 1.17 | 0.11 | 1.08 | 0.11 | NA
| Analytical results for P2O5 in OREAS 84 (Certified Value 0.061 wt.%) | [17 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 17 | 20 | [5 empty] | wt.%
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.06 | 0.055 | 0.08 | 0.055 | 0.04 | 0.1 | 0.06 | 0.09 | 0.07 | 0.049599 | 0.05 | [5 empty] | 1
| 1 | 2 | 0.06 | 0.057 | 0.08 | 0.064 | 0.04 | 0.1 | 0.07 | 0.09 | 0.07 | 0.049045 | 0.05 | [5 empty] | #N/A
| 1 | 3 | 0.06 | 0.054 | 0.08 | 0.059 | 0.04 | 0.11 | 0.06 | 0.08 | 0.07 | 0.049673 | 0.05 | [5 empty] | 16
| 1 | 4 | 0.06 | 0.056 | 0.08 | 0.052 | 0.04 | 0.1 | 0.06 | 0.08 | 0.07 | 0.050093 | 0.05 | [5 empty] | 0.060518
| 1 | 5 | 0.06 | 0.057 | 0.07 | 0.058 | 0.04 | 0.1 | 0.06 | 0.08 | 0.07 | 0.049195 | 0.05 | [5 empty] | 139
| 1 | 6 | 0.06 | 0.056 | 0.07 | 0.05 | 0.04 | 0.09 | 0.06 | 0.08 | 0.07 | 0.05047 | 0.06
| Mean |  | 0.06 | 0.055833 | 0.076667 | 0.056333 | 0.04 | 0.1 | 0.061667 | 0.083333 | 0.07 | 0.049679 | 0.051667
| Median |  | 0.06 | 0.056 | 0.08 | 0.0565 | 0.04 | 0.1 | 0.06 | 0.08 | 0.07 | 0.049636 | 0.05
| Std Dev. |  | 0 | 0.001169 | 0.005164 | 0.0050859 | 0 | 0.0063246 | 0.0040825 | 0.005164 | 0 | 0.00053676 | 0.0040825
| Rel.Std.Dev. |  | 0% | 2.094% | 6.736% | 9.028% | 0% | 6.325% | 6.62% | 6.197% | 0% | 1.08% | 7.902%
| PDM3 |  | -0.8558% | -7.741% | 26.68% | -6.915% | -33.9% | 65.24% | 1.898% | 37.7% | 15.67% | -17.91% | -14.63%
| Z-Score (Absolute) |  | 0 | 0.28 | 1.12 | 0.25 | 1.35 | 2.7 | 0.11 | 1.57 | 0.67 | 0.7 | 0.56
| Analytical results for Pb in OREAS 84 (Indicative Value 177 ppm) | [17 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 06 | 08 | 09 | 13 | 17 | 20 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 120 | < 50 | 280 | 110 | 100 | 180 | 463.1761 | < 90 | [8 empty] | 1
| 1 | 2 | 120 | < 50 | 280 | 100 | 200 | 220 | 466.5804 | < 90 | [8 empty] | 17
| 1 | 3 | 130 | < 50 | 280 | 90 | 200 | 170 | 461.9749 | < 90 | [8 empty] | 16
| 1 | 4 | 120 | < 50 | 280 | 100 | 200 | 200 | 453.3161 | < 90 | [8 empty] | 177.0833
| 1 | 5 | 110 | < 50 | 280 | 100 | 200 | 190 | 461.4 | 190 | [8 empty] | 23
| 1 | 6 | 110 | < 50 | 280 | 100 | 200 | 220 | 475.2789 | < 90
| Mean |  | 118.3333 |  | 280 | 100 | 183.3333 | 196.6667 | 463.6211 | 190
| Median |  | 120 |  | 280 | 100 | 200 | 195 | 462.5755 | 190
| Std Dev. |  | 7.5277 |  | 0 | 6.3246 | 40.8248 | 20.6559 | 7.193
| Rel.Std.Dev. |  | 6.361% |  | 0% | 6.325% | 22.27% | 10.5% | 1.551%
| PDM3 |  | -33.18% |  | 58.12% | -43.53% | 3.529% | 11.06% | 161.8% | 7.294%
| Z-Score (Absolute) |  | 0.33 | 1.29 | 1.3 | 0.52 | 0.33 | 0.47 | 3.19 | 0.83
| Analytical results for Rb in OREAS 84 (Indicative Value 50 ppm) | [17 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 60 | 60 | 50 | < 50 | [12 empty] | 1
| 1 | 2 | 60 | 70 | < 50 | < 50 | [12 empty] | 18
| 1 | 3 | 60 | 70 | 50 | < 50 | [12 empty] | 16
| 1 | 4 | 60 | 80 | < 50 | < 50 | [12 empty] | 50
| 1 | 5 | 60 | 60 | 50 | < 50 | [12 empty] | 24
| 1 | 6 | 60 | 100 | 50 | 50
| Mean |  | 60 | 73.3333 | 50 | 50
| Median |  | 60 | 70 | 50 | 50
| Std Dev. |  | 0 | 15.0555 | 0
| Rel.Std.Dev. |  | 0% | 20.53% | 0%
| PDM3 |  | 20% | 46.67% | 0% | 0%
| Z-Score (Absolute) |  | 0.4 | 0.98 | 0.4 | 0.95
| Analytical results for S in OREAS 84 (Certified Value 22.63 wt.%) | [17 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 20 | 23 |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 23.4 | > 20 | 18.6466 | > 20 | 23.062 | 4.7573 | 24.688 | > 20 | 22.4 | NA | 24.9 | > 20 | 21.476 | N.A. |  |  | 1
| 1 | 2 | 23.2 | > 20 | 18.6666 | > 20 | 22.341 | 4.7614 | 24.784 | > 20 | 22.8 | NA | 25.1 | > 20 | 21.416 | N.A. |  |  | #N/A
| 1 | 3 | 23.5 | > 20 | 18.3014 | > 20 | 22.73 | 4.7533 | 24.732 | > 20 | 22.9 | NA | 24.9 | > 20 | 21.396 | N.A. |  |  | 16
| 1 | 4 | 23 | > 20 | 18.4291 | > 20 | 22.573 | 4.7493 | 24.668 | > 20 | 22.4 | NA | 25.2 | > 20 | 21.444 | N.A. |  |  | 22.6331
| 1 | 5 | 23.5 | > 20 | 18.3006 | > 20 | 23.45 | 4.7453 | 24.612 | > 20 | 22.6 | NA | 25.3 | > 20 | 21.44 | N.A. |  |  | 140
| 1 | 6 | 23.4 | > 20 | 18.6017 | > 20 | 23.278 | 4.7493 | 24.64 | > 20 | 21.8 | NA | 25.2 | > 20 | 21.416 | N.A.
| Mean |  | 23.3333 |  | 18.491 |  | 22.9057 | 4.7527 | 24.6873 |  | 22.4833 |  | 25.1 |  | 21.4313
| Median |  | 23.4 |  | 18.5154 |  | 22.896 | 4.7513 | 24.678 |  | 22.5 |  | 25.15 |  | 21.428
| Std Dev. |  | 0.19664 |  | 0.16931 |  | 0.42889 | 0.0058945 | 0.062657 |  | 0.392 |  | 0.16733 |  | 0.028105
| Rel.Std.Dev. |  | 0.8427% |  | 0.9156% |  | 1.872% | 0.124% | 0.2538% |  | 1.744% |  | 0.6667% |  | 0.1311%
| PDM3 |  | 3.094% |  | -18.3% |  | 1.204% | -79% | 9.076% |  | -0.6618% |  | 10.9% |  | -5.31%
| Z-Score (Absolute) |  | 0.26 | NA | 1.74 | NA | 0.09 | 7.43 | 0.83 | NA | 0.09 | NA | 1 | NA | 0.52 | NA
| Analytical results for Sb in OREAS 84 (Indicative Value 49.8 ppm) | [17 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | 50 | 70 | < 50 | < 50 | [12 empty] | 1
| 1 | 2 | 70 | 60 | < 50 | < 50 | [12 empty] | 19
| 1 | 3 | 60 | 70 | < 50 | 80 | [12 empty] | 16
| 1 | 4 | < 50 | 80 | 70 | 90 | [12 empty] | 49.75
| 1 | 5 | < 50 | 60 | 80 | < 50 | [12 empty] | 25
| 1 | 6 | < 50 | 80 | < 50 | 70
| Mean |  | 60 | 70 | 75 | 80
| Median |  | 60 | 70 | 75 | 80
| Std Dev. |  | 10 | 8.9443 | 7.0711 | 10
| Rel.Std.Dev. |  | 16.67% | 12.78% | 9.428% | 12.5%
| PDM3 |  | 20.6% | 40.7% | 50.75% | 60.8%
| Z-Score (Absolute) |  | 0.62 | 2.8 | 0.73 | 0.62
| Analytical results for Sc in OREAS 84 (Indicative Value < 1 ppm) | [17 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 03 | 07 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 2
2
| 1 | 1 | < 40 | < 1 | [14 empty] | 1
| 1 | 2 | < 40 | < 1 | [14 empty] | 20
| 1 | 3 | < 40 | < 1 | [14 empty] | 16
| 1 | 4 | < 40 | < 1 | [14 empty] | < 1
| 1 | 5 | < 40 | < 1 | [14 empty] | 26
| 1 | 6 | < 40 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for SiO2 in OREAS 84 (Certified Value 16.07 wt.%) | [17 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 15.68 | 16.2 | 15.99 | 16.15 | 14.62 | 12.27 | 16.11 | 16.2 | 15.3 | NA | 16.4 | 16 | 16.2488 | 15.78 | 10.2 |  | 1
| 1 | 2 | 15.38 | 16.1 | 16.02 | 16.35 | 14.75 | 12.23 | 16.48 | 16.2 | 14.8 | NA | 16.1 | 16.1 | 16.2152 | 15.98 | 10.7 |  | #N/A
| 1 | 3 | 15.33 | 16.1 | 16.02 | 16.15 | 14.39 | 12.28 | 16.15 | 16.15 | 14.7 | NA | 16.2 | 15.95 | 16.192 | 16.14 | 10.8 |  | 16
| 1 | 4 | 15.19 | 16.15 | 16.02 | 16.05 | 14.66 | 12.29 | 16.06 | 16.15 | 14.6 | NA | 16.3 | 16.1 | 16.2653 | 15.8 | 10.3 |  | 16.0654
| 1 | 5 | 15.96 | 16.2 | 15.96 | 16 | 15.06 | 12.26 | 16 | 16.25 | 14.8 | NA | 16.3 | 16 | 16.1518 | 15.97 | 10.3 |  | 141
| 1 | 6 | 15.29 | 16.05 | 15.96 | 16.15 | 14.39 | 12.25 | 16.02 | 16.05 | 14.5 | NA | 16.2 | 16.15 | 16.2154 | 16.2 | 10.1
| Mean |  | 15.4717 | 16.1333 | 15.995 | 16.1417 | 14.645 | 12.2633 | 16.1367 | 16.1667 | 14.7833 |  | 16.25 | 16.05 | 16.2148 | 15.9783 | 10.4
| Median |  | 15.355 | 16.125 | 16.005 | 16.15 | 14.64 | 12.265 | 16.085 | 16.175 | 14.75 |  | 16.25 | 16.05 | 16.2153 | 15.975 | 10.3
| Std Dev. |  | 0.29075 | 0.060553 | 0.029496 | 0.12007 | 0.25066 | 0.021602 | 0.17716 | 0.068313 | 0.27869 |  | 0.10488 | 0.07746 | 0.040481 | 0.17116 | 0.28284
| Rel.Std.Dev. |  | 1.879% | 0.3753% | 0.1844% | 0.7438% | 1.712% | 0.1762% | 1.098% | 0.4226% | 1.885% |  | 0.6454% | 0.4826% | 0.2497% | 1.071% | 2.72%
| PDM3 |  | -3.696% | 0.4229% | -0.4382% | 0.4748% | -8.841% | -23.67% | 0.4437% | 0.6304% | -7.98% |  | 1.149% | -0.09581% | 0.9297% | -0.5419% | -35.26%
| Z-Score (Absolute) |  | 2.21 | 0.44 | 0.11 | 0.48 | 5.52 | 15.07 | 0.46 | 0.58 | 4.97 | NA | 0.91 | 0.11 | 0.77 | 0.18 | 22.54
| Analytical results for Sn in OREAS 84 (Indicative Value < 50 ppm) | [17 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 02 | 04 | 08 | 13 | 17 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | < 50 | < 160 | [11 empty] | 1
| 1 | 2 | < 50 | 60 | < 50 | 60 | < 160 | [11 empty] | 21
| 1 | 3 | < 50 | < 50 | < 50 | < 50 | < 160 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | 50 | < 160 | [11 empty] | < 50
| 1 | 5 | < 50 | 60 | < 50 | 50 | < 160 | [11 empty] | 27
| 1 | 6 | < 50 | 50 | < 50 | 70 | < 160
| Mean |  |  | 56.6667 |  | 57.5
| Median |  |  | 60 |  | 55
| Std Dev. |  |  | 5.7735 |  | 9.5743
| Rel.Std.Dev. |  |  | 10.19% |  | 16.65%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 0.67 | 0.25 | 1.62
| Analytical results for Sr in OREAS 84 (Indicative Value 84 ppm) | [17 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 04 | 05 | 08 | 13 | 17 | 20 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | 100 | 100 | 51 | 100 | 200 | 54.1385 | 80 | [9 empty] | 1
| 1 | 2 | 100 | 100 | 59 | 100 | 200 | 44.9044 | 80 | [9 empty] | 16
| 1 | 3 | 100 | 100 | < 42 | 100 | 200 | 41.1464 | 170 | [9 empty] | 16
| 1 | 4 | 100 | 100 | 42 | 100 | 200 | 56.0262 | 170 | [9 empty] | 84.1532
| 1 | 5 | 100 | 100 | < 42 | 100 | 200 | 48.5631 | < 80 | [9 empty] | 22
| 1 | 6 | 100 | 100 | < 42 | 100 | 200 | 50.354 | 170
| Mean |  | 100 | 100 | 50.6667 | 100 | 200 | 49.1888 | 134
| Median |  | 100 | 100 | 51 | 100 | 200 | 49.4585 | 170
| Std Dev. |  | 0 | 0 | 8.5049 | 0 | 0 | 5.5863 | 49.295
| Rel.Std.Dev. |  | 0% | 0% | 16.79% | 0% | 0% | 11.36% | 36.79%
| PDM3 |  | 18.83% | 18.83% | -39.79% | 18.83% | 137.7% | -41.55% | 59.23%
| Z-Score (Absolute) |  | 0 | 0 | 2.18 | 0 | 3.41 | 1.73 | 0.67
| Analytical results for TiO2 in OREAS 84 (Certified Value 0.199 wt.%) | [17 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 17 | 20 | 23 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.19 | 0.21 | 0.2 | 0.21 | 0.16 | 0.17 | 0.21 | 0.23 | 0.21 | NA | 0.22 | 0.21 | 0.18999 | 0.19 | 0.13 |  | 1
| 1 | 2 | 0.2 | 0.22 | 0.2 | 0.2 | 0.16 | 0.17 | 0.21 | 0.22 | 0.22 | NA | 0.21 | 0.22 | 0.18226 | 0.19 | 0.14 |  | #N/A
| 1 | 3 | 0.2 | 0.23 | 0.2 | 0.2 | 0.16 | 0.17 | 0.23 | 0.22 | 0.2 | NA | 0.21 | 0.21 | 0.18544 | 0.19 | 0.13 |  | 16
| 1 | 4 | 0.19 | 0.21 | 0.2 | 0.2 | 0.16 | 0.17 | 0.21 | 0.22 | 0.19 | NA | 0.21 | 0.21 | 0.19386 | 0.2 | 0.13 |  | 0.19867
| 1 | 5 | 0.21 | 0.21 | 0.19 | 0.2 | 0.17 | 0.17 | 0.2 | 0.22 | 0.2 | NA | 0.21 | 0.21 | 0.19447 | 0.19 | 0.13 |  | 142
| 1 | 6 | 0.2 | 0.2 | 0.19 | 0.21 | 0.16 | 0.17 | 0.2 | 0.22 | 0.19 | NA | 0.21 | 0.21 | 0.18986 | 0.2 | 0.13
| Mean |  | 0.19833 | 0.21333 | 0.19667 | 0.20333 | 0.16167 | 0.17 | 0.21 | 0.22167 | 0.20167 |  | 0.21167 | 0.21167 | 0.18931 | 0.19333 | 0.13167
| Median |  | 0.2 | 0.21 | 0.2 | 0.2 | 0.16 | 0.17 | 0.21 | 0.22 | 0.2 |  | 0.21 | 0.21 | 0.18992 | 0.19 | 0.13
| Std Dev. |  | 0.0075277 | 0.010328 | 0.005164 | 0.005164 | 0.0040825 | 0 | 0.010954 | 0.0040825 | 0.01169 |  | 0.0040825 | 0.0040825 | 0.0047488 | 0.005164 | 0.0040825
| Rel.Std.Dev. |  | 3.795% | 4.841% | 2.626% | 2.54% | 2.525% | 0% | 5.216% | 1.842% | 5.797% |  | 1.929% | 1.929% | 2.508% | 2.671% | 3.101%
| PDM3 |  | -0.167% | 7.383% | -1.006% | 2.35% | -18.62% | -14.43% | 5.706% | 11.58% | 1.511% |  | 6.545% | 6.545% | -4.708% | -2.684% | -33.72%
| Z-Score (Absolute) |  | 0.1 | 0.8 | 0.2 | 0.2 | 2.31 | 1.81 | 0.6 | 1.31 | 0.1 | NA | 0.7 | 0.7 | 0.64 | 0.4 | 4.12
| Analytical results for V2O5 in OREAS 84 (Certified Value 250 ppm) | [17 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 04 | 05 | 08 | 13 | 17 | 20 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | 180 | 360 | 220 | 360 | 180 | 175.2844 | 100 | [9 empty] | 1
| 1 | 2 | 180 | 360 | 280 | 360 | 180 | 175.947 | 100 | [9 empty] | #N/A
| 1 | 3 | 180 | 360 | 70 | 360 | 180 | 175.6808 | 100 | [9 empty] | 16
| 1 | 4 | 180 | 360 | 210 | 360 | 360 | 177.5807 | 100 | [9 empty] | 249.5149
| 1 | 5 | 180 | 360 | 180 | 360 | 180 | 173.2466 | 100 | [9 empty] | 143
| 1 | 6 | 180 | 180 | 210 | 360 | 360 | 179.5561 | 100
| Mean |  | 180 | 330 | 195 | 360 | 240 | 176.2159 | 100
| Median |  | 180 | 360 | 210 | 360 | 180 | 175.8139 | 100
| Std Dev. |  | 0 | 73.4847 | 69.4982 | 0 | 92.9516 | 2.1477 | 0
| Rel.Std.Dev. |  | 0% | 22.27% | 35.64% | 0% | 38.73% | 1.219% | 0%
| PDM3 |  | -27.86% | 32.26% | -21.85% | 44.28% | -3.813% | -29.38% | -59.92%
| Z-Score (Absolute) |  | 0.65 | 1.91 | 0.37 | 2.42 | 0.37 | 0.69 | NA
Results from laboratory 20 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 84 (Indicative Value 22.5 ppm) | [17 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | 30 | 30 | 40 | < 10 | [12 empty] | 1
| 1 | 2 | 20 | 60 | 20 | < 10 | [12 empty] | 12
| 1 | 3 | 30 | 30 | 30 | < 10 | [12 empty] | 16
| 1 | 4 | 20 | 30 | 20 | < 10 | [12 empty] | 22.5
| 1 | 5 | 10 | 20 | 10 | < 10 | [12 empty] | 23
| 1 | 6 | < 10 | 70 | 20 | 20
| Mean |  | 22 | 40 | 23.3333 | 20
| Median |  | 20 | 30 | 20 | 20
| Std Dev. |  | 8.3666 | 20 | 10.328
| Rel.Std.Dev. |  | 38.03% | 50% | 44.26%
| PDM3 |  | -2.222% | 77.78% | 3.704% | -11.11%
| Z-Score (Absolute) |  | 0.18 | 1.6 | 0.18 | 1.17
| Analytical results for Y in OREAS 84 (Indicative Value 41.7 ppm) | [17 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 08 | 13 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | 47 | < 39 | 39 | < 39 | [12 empty] | 1
| 1 | 2 | 39 | < 39 | < 39 | 71 | [12 empty] | 18
| 1 | 3 | 39 | < 39 | < 39 | 55 | [12 empty] | 16
| 1 | 4 | 39 | < 39 | 39 | 63 | [12 empty] | 41.6918
| 1 | 5 | 39 | < 39 | < 39 | 63 | [12 empty] | 24
| 1 | 6 | < 39 | < 39 | < 39 | 55
| Mean |  | 40.6 |  | 39 | 61.4
| Median |  | 39 |  | 39 | 63
| Std Dev. |  | 3.5777 |  | 0 | 6.6933
| Rel.Std.Dev. |  | 8.812% |  | 0% | 10.9%
| PDM3 |  | -2.619% |  | -6.456% | 47.27%
| Z-Score (Absolute) |  | 0.42 | 0.92 | 0.42 | 1.72
| Analytical results for Zn in OREAS 84 (Certified Value 805 ppm) | [17 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 07 | 08 | 09 | 13 | 17 | 20 | [7 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 800 | 770 | 800 | 800 | 830 | 800 | 780 | 805.1649 | 800 | [7 empty] | 1
| 1 | 2 | 790 | 800 | 810 | 800 | 830 | 800 | 820 | 818.2019 | 720 | [7 empty] | #N/A
| 1 | 3 | 800 | 790 | 800 | 900 | 820 | 800 | 790 | 813.7787 | 720 | [7 empty] | 16
| 1 | 4 | 810 | 770 | 790 | 800 | 830 | 700 | 810 | 827.1693 | 720 | [7 empty] | 804.5695
| 1 | 5 | 790 | 780 | 790 | 800 | 870 | 800 | 790 | 814.5285 | 720 | [7 empty] | 144
| 1 | 6 | 780 | 800 | 850 | 800 | 830 | 700 | 830 | 829.6577 | 720
| Mean |  | 795 | 785 | 806.6667 | 816.6667 | 835 | 766.6667 | 803.3333 | 818.0835 | 733.3333
| Median |  | 795 | 785 | 800 | 800 | 830 | 800 | 800 | 816.3652 | 720
| Std Dev. |  | 10.4881 | 13.784 | 22.5093 | 40.8248 | 17.6068 | 51.6398 | 19.6638 | 9.1059 | 32.6599
| Rel.Std.Dev. |  | 1.319% | 1.756% | 2.79% | 4.999% | 2.109% | 6.736% | 2.448% | 1.113% | 4.454%
| PDM3 |  | -1.189% | -2.432% | 0.2607% | 1.504% | 3.782% | -4.711% | -0.1536% | 1.68% | -8.854%
| Z-Score (Absolute) |  | 0.38 | 0.84 | 0.15 | NA | 1.45 | NA | 0 | 0.67 | 3.06
Results from laboratories 7 and 9 were removed due to their 100 ppm reading resolution.
| Analytical results for Zr in OREAS 84 (Indicative Value 272 ppm) | [17 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 04 | 08 | 13 | 17 | 20 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 200 | 400 | < 100 | 200 | < 150 | < 70 | [10 empty] | 1
| 1 | 2 | 200 | 400 | < 100 | 200 | < 150 | < 70 | [10 empty] | 19
| 1 | 3 | 200 | 400 | < 100 | 200 | < 150 | < 70 | [10 empty] | 16
| 1 | 4 | 200 | 400 | < 100 | 200 | < 150 | < 70 | [10 empty] | 272.2222
| 1 | 5 | 300 | 400 | < 100 | 200 | < 150 | < 70 | [10 empty] | 25
| 1 | 6 | 200 | 400 | < 100 | 200 | < 150 | < 70
| Mean |  | 216.6667 | 400 |  | 200
| Median |  | 200 | 400 |  | 200
| Std Dev. |  | 40.8248 | 0 |  | 0
| Rel.Std.Dev. |  | 18.84% | 0% |  | 0%
| PDM3 |  | -20.41% | 46.94% |  | -26.53%
| Z-Score (Absolute) |  | 0.64 | 2.13 | 0.7 | 0.51 | 0.51 | 0.81
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 84 (Indicative Value 0.212 wt.%) | [17 empty] | Indicative
H2O- | Round | Replicate | Lab | [15 empty] | 1
| No | No | 03 | [15 empty] | wt.%
|  |  | LOI*TGA | [15 empty] | 3
3
| 1 | 1 | 0.17 | [15 empty] | 1
| 1 | 2 | 0.23 | [15 empty] | 23
| 1 | 3 | 0.15 | [15 empty] | 16
| 1 | 4 | 0.25 | [15 empty] | 0.21167
| 1 | 5 | 0.33 | [15 empty] | 29
| 1 | 6 | 0.14
| Mean |  | 0.21167
| Median |  | 0.2
| Std Dev. |  | 0.072778
| Rel.Std.Dev. |  | 34.38%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 84 (Certified Value 12.31 wt.%) | [17 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 11 | 12 | 16 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*XRF | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | PF*OES/MS | LOI*TGA | BF*MS | BF*XRF | LOI*Furnace | LOI*Furnace | 2
3
| 1 | 1 | 12.48 | 12.5 | 11.23 | 12.06 | 12.59 | 20.15 | 12.88 | 11.82 | NA | 11.59 | 12.42 | 11.6935 | 12.4 | 12.84 | 12.42 | 12.5 | 1
| 1 | 2 | 12.48 | 12.45 | 11.18 | 11.96 | 12.48 | 20.18 | 12.76 | 11.8 | NA | 11.54 | 12.54 | 11.5359 | 12.4 | 12.72 | 12.46 | 12.6 | #N/A
| 1 | 3 | 12.39 | 12.37 | 11.19 | 12.08 | 12.61 | 20.15 | 12.7 | 11.83 | NA | 11.58 | 12.49 | 11.5308 | 12.6 | 12.74 |  | 12.8 | 16
| 1 | 4 | 12.44 | 12.42 | 11.13 | 12.01 | 12.82 | 20.18 | 12.88 | 11.78 | NA | 11.62 | 12.53 | 11.6185 | 12.4 | 12.69 |  | 12.8 | 12.3141
| 1 | 5 | 12.43 | 12.07 | 11.29 | 12.06 | 12.61 | 20.14 | 12.85 | 11.85 | NA | 11.53 | 12.48 | 11.6398 | 12.4 | 12.65 |  | 12.5 | 146
| 1 | 6 | 12.4 | 12.2 | 11.21 | 12.26 | 12.74 | 20.18 | 12.72 | 11.97 | NA | 11.62 | 12.71 | 11.7308 | 12.4 | 12.77 |  | 12.5
| Mean |  | 12.4367 | 12.335 | 11.205 | 12.0717 | 12.6417 | 20.1633 | 12.7983 | 11.8417 |  | 11.58 | 12.5283 | 11.6249 | 12.4333 | 12.735 | 12.44 | 12.6167
| Median |  | 12.435 | 12.395 | 11.2 | 12.06 | 12.61 | 20.165 | 12.805 | 11.825 |  | 11.585 | 12.51 | 11.6292 | 12.4 | 12.73 | 12.44 | 12.55
| Std Dev. |  | 0.038297 | 0.16574 | 0.053572 | 0.10206 | 0.12024 | 0.018619 | 0.081588 | 0.067355 |  | 0.038471 | 0.098675 | 0.081214 | 0.08165 | 0.065955 | 0.028284 | 0.1472
| Rel.Std.Dev. |  | 0.3079% | 1.344% | 0.4781% | 0.8455% | 0.9511% | 0.09234% | 0.6375% | 0.5688% |  | 0.3322% | 0.7876% | 0.6986% | 0.6567% | 0.5179% | 0.2274% | 1.167%
| PDM3 |  | 0.9954% | 0.1698% | -9.007% | -1.969% | 2.66% | 63.74% | 3.932% | -3.836% |  | -5.961% | 1.74% | -5.597% | 0.9683% | 3.418% | 1.022% | 2.457%
| Z-Score (Absolute) |  | 0 | 0.23 | 2.78 | 0.82 | 0.46 | 17.46 | 0.82 | 1.34 | NA | 1.94 | 0.21 | 1.83 | 0.01 | 0.67 | 0.01 | 0.41
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.